3OQ222 | 3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_CoA_(mm) | Caliper |  |  | mm | FDA,2 | 0.4 | 0.48 | TM 102 | Extra_specs
3OQ222 | 3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD |  |  | g/in | FDA,0 | 600 | 1200 | TM 103 | Extra_specs
3OQ222 | 3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD |  |  | g/in | FDA,0 | 2200 | 3400 | TM 103 | Extra_specs
3OQ222 | 3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | [6 empty] | TM 208 | Extra_specs
3OQ222 | 3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 748 | Visual Inspection | Color | Color | [6 empty] | TM 208 | Extra_specs
3OQ222 | 3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 748 | Visual Inspection | Odor | Odor | [6 empty] | TM 208 | Extra_specs
3OQ222 | 3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Certificate of Analysis | Extra_specs
3OQ222 | 3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 44 |  | in | FDA,0 | 42 | 46 | Applicator Measurements | Extra_specs
3OQ222 | 3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 36 |  | in | FDA,3 | 35.875 | 36.125 | Applicator Measurements | Extra_specs
3OQ222 | 3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 39.61 |  | GSM | FDA,1 | 35.5 | 44.5 | TM 101 | Extra_specs
3OQ222 | 3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper | 0.46 |  | mm | FDA,2 | 0.4 | 0.48 | TM 102 | Extra_specs
3OQ222 | 3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 521.87 |  | g/in | FDA,0 | 600 | 1200 | TM 103 | Extra_specs
3OQ222 | 3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 1610.35 |  | g/in | FDA,0 | 2200 | 3400 | TM 103 | Extra_specs
3OQ222 | 3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | TM 208 | Extra_specs
3OQ222 | 3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 748 | Visual Inspection | Color | Color | White with no off colors or shading [Done] | [5 empty] | TM 208 | Extra_specs
3OQ222 | 3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) | [5 empty] | TM 208 | Extra_specs
3OQ222 | 3OQ222/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter |  |  | in | FDA,0 | 42 | 46 | Applicator Measurements | Extra_specs
3OQ444 | 3OQ444/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper | 0.41 |  | mm | FDA,2 | 0.4 | 0.48 | TM 102 | Extra_specs
3OQ850N | 3OQ850N/Applicator_01E | Applicator_01E | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [6 empty] | Applicator Measurements | Extra_specs
3OQ850N | 3OQ850N/Applicator_01E | Applicator_01E | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter |  |  | in | FDA,0 | 44 | 47 | Applicator Measurements | Extra_specs
3OQ850N | 3OQ850N/Applicator_01E | Applicator_01E | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width |  |  | in | FDA,3 | 35.875 | 36.125 | Applicator Measurements | Extra_specs
3OQ850N | 3OQ850N/Applicator_01E | Applicator_01E | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight |  |  | GSM | FDA,0 | 38 | 42 | TM 101 | Extra_specs
3OQ850N | 3OQ850N/Applicator_01E | Applicator_01E | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper |  |  | mm | FDA,2 | 0.35 | 0.45 | TM 102 | Extra_specs
3OQ850N | 3OQ850N/Applicator_01E | Applicator_01E | NY-Analytical | 758 | Tensile Strength | Dry_CD_(g/in) | Dry CD |  |  | g/in | FDA,0 | 513 | 1130 | TM 103 | Extra_specs
3OQ850N | 3OQ850N/Applicator_01E | Applicator_01E | NY-Analytical | 758 | Tensile Strength | Dry_MD_(g/in) | Dry CD |  |  | g/in | FDA,0 | 2054 | 4108 | TM 103 | Extra_specs
3OQ850N | 3OQ850N/Applicator_01E | Applicator_01E | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | [6 empty] | TM 208 | Extra_specs
3OQ850N | 3OQ850N/Applicator_01E | Applicator_01E | NY-Analytical | 748 | Visual Inspection | Color | Color | [6 empty] | TM 208 | Extra_specs
3OQ850N | 3OQ850N/Applicator_01E | Applicator_01E | NY-Analytical | 748 | Visual Inspection | Odor | Odor | [6 empty] | TM 208 | Extra_specs
3OS428 | 3OS428/REDUCED TESTING_NY_3OS428 | REDUCED TESTING_NY_3OS428 | NY-Analytical | 760 | Swab Measurements | AQL Sample Size | AQL Sample Size | 315 |  | in |  |  |  | TM 109 | Extra_specs
3OS428 | 3OS428/REDUCED TESTING_NY_3OS428 | REDUCED TESTING_NY_3OS428 | NY-Analytical | 760 | Swab Measurements | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | TM 109 | Extra_specs
3OS428 | 3OS428/REDUCED TESTING_NY_3OS428 | REDUCED TESTING_NY_3OS428 | NY-Analytical | 760 | Swab Measurements | Scrub_Test | Scrub Test | Pass (Done) | [5 empty] | TM 109 | Extra_specs
3OS428 | 3OS428/REDUCED TESTING_NY_3OS428 | REDUCED TESTING_NY_3OS428 | NY-Analytical | 760 | Swab Measurements | Tuft_Adhesion_(Lb.) | Tuft Adhesion | 2.4 |  | lb | FDA,1 | 1.4 |  | TM 109 | Extra_specs
3OS428 | 3OS428/REDUCED TESTING_NY_3OS428 | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 2.5 |  | g | FDA,2 |  |  | TM 610 | Extra_specs
3OS428 | 3OS428/REDUCED TESTING_NY_3OS428 | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA | [5 empty] | TM 610 | Extra_specs
3OS428 | 3OS428/REDUCED TESTING_NY_3OS428 | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA | [5 empty] | TM 610 | Extra_specs
3OS428 | 3OS428/REDUCED TESTING_NY_3OS428 | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Extra_specs
3OS428 | 3OS428/REDUCED TESTING_NY_3OS428 | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Extra_specs
3OS428 | 3OS428/REDUCED TESTING_NY_3OS428 | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Extra_specs
3OS428 | 3OS428/REDUCED TESTING_NY_3OS428 | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Extra_specs
3OS428 | 3OS428/REDUCED TESTING_NY_3OS428 | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Extra_specs
3OS428 | 3OS428/REDUCED TESTING_NY_3OS428 | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Extra_specs
3OU351 | 3OU351/Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | COA | Extra_specs
3OU351 | 3OU351/Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis (ELLI) | Certificate of Analysis | COA Meets Specification (Done) | [5 empty] | COA | Extra_specs
3OU351 | 3OU351/Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 5.968 |  | in | FDA,3 | 5.875 | 6.125 | TM 143 | Extra_specs
3OU351 | 3OU351/Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | NY-Analytical | 397 | Appearance | Appearance/Matches Embossing Pattern | Appearance/Matches Embossing Pattern | Free of Hard wrinkles. Formation uniform (Done) | [5 empty] | TM 208 | Extra_specs
3OU351 | 3OU351/Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | NY-Analytical | 395 | Color | Color | Color | White with no off colors or shading (Done) | [5 empty] | TM 208 | Extra_specs
3OU351 | 3OU351/Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | NY-Analytical | 1038 | KC Part Number | Verify KC Part Number | Verify KC Part Number | SPEC-08468 (Done) | [6 empty] | Extra_specs
3OX465H | 3OX465H/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 765 | Canister & Lid Inspection | Handle Internal Diameter | Handle Internal Diameter | 0.9545 |  | in | FDA,3 | 0.951 | 0.957 | Canister & Lid Inspection | Extra_specs
3XP540 | 3XP540/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (COA) | Roll Diameter (COA) |  |  | in | FDA,0 | 41 | 45 | Applicator Measurements | Extra_specs
3XP540 | 3XP540/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 759 | Applicator Measurements | Roll width (COA) | Roll width (COA) |  |  | in | FDA,2 | 41.75 | 42.25 | Applicator Measurements | Extra_specs
3XP540 | 3XP540/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 49.551 |  | GSM | FDA,0 | 45 | 55 | TM 101 | Extra_specs
3XP540 | 3XP540/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper | 0.46 |  | mm | FDA,2 | 0.51 | 0.63 | TM 102 | Extra_specs
3XP540 | 3XP540/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 758 | Tensile Strength | Dry_CD_(g/in) | Dry CD |  |  | g/in | FDA,0 | 1801 | 2821 | TM 103 | Extra_specs
3XP540 | 3XP540/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 758 | Tensile Strength | Dry_MD_(g/in) | Dry CD |  |  | g/in | FDA,0 | 6675 | 9447 | TM 103 | Extra_specs
3XP540 | 3XP540/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | TM 208 | Extra_specs
3XP540 | 3XP540/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) | [5 empty] | TM 208 | Extra_specs
3XP540 | 3XP540/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 748 | Visual Inspection | Odor | Odor | No objectionable odor Wet or Dry [Done] | [5 empty] | TM 208 | Extra_specs
3XP667 | 3XP667/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 758 | Tensile Strength | Dry_CD_(g/in) | Dry CD | 1805 |  | g/in | FDA,0 | 1801 | 2821 | TM 103 | Extra_specs
3XP667 | 3XP667/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 758 | Tensile Strength | Dry_MD_(g/in) | Dry CD | 8039 |  | g/in | FDA,0 | 6675 | 9447 | TM 103 | Extra_specs
4YBBY12 | 4YBBY12/4YBBY12 | 4YBBY12 | JB-Analytical | 398 | pH | pH | pH |  |  |  | FDA,1 | 10.8 | 11.3 | TM200 | Extra_specs
4YBBY12 | 4YBBY12/4YBBY12 | 4YBBY12 | JB-Analytical | 902 | Bulk Liquid Inspection | Appearance (TM 208) | Appearance | [6 empty] | See Analyte for Method # | Extra_specs
4YBBY12 | 4YBBY12/4YBBY12 | 4YBBY12 | JB-Analytical | 902 | Bulk Liquid Inspection | Color (TM 208) | Color | [6 empty] | See Analyte for Method # | Extra_specs
4YBBY12 | 4YBBY12/4YBBY12 | 4YBBY12 | JB-Analytical | 902 | Bulk Liquid Inspection | Contamination (TM 208) | Contamination | [6 empty] | See Analyte for Method # | Extra_specs
4YBBY12 | 4YBBY12/4YBBY12 | 4YBBY12 | JB-Analytical | 902 | Bulk Liquid Inspection | Odor (TM 208) | Odor | [6 empty] | See Analyte for Method # | Extra_specs
4YBBY12 | 4YBBY12/4YBBY12 | 4YBBY12 | JB-Analytical | 425 | % Sodium Benzoate | Sodium Benzoate | Sodium Benzoate |  |  | % | FDA,2 | 8.46 | 10.34 | TM 395 | Extra_specs
4YBBY12 | 4YBBY12/4YBBY12 | 4YBBY12 | JB-Analytical | 425 | % Sodium Benzoate | Sodium Benzoate Final | Sodium Benzoate Final |  |  | % | FDA,2 | 8.46 | 10.34 | TM 395 | Extra_specs
4YBBY12 | 4YBBY12/BL_JB 20X Sodium Benz Base Blend | BL_JB 20X Sodium Benz Base Blend | JB-Analytical | 447 | pH as is | pH | pH | 11.86 |  |  | FDA,2 | 11.2 | 12.6 | TM 200 | Extra_specs
4YBBY12 | 4YBBY12/BL_JB 20X Sodium Benz Base Blend | BL_JB 20X Sodium Benz Base Blend | JB-Analytical | 748 | Visual Inspection | Appearance | Appearance | Slighlty Hazy (Done) | [5 empty] | TM 208 | Extra_specs
4YBBY12 | 4YBBY12/BL_JB 20X Sodium Benz Base Blend | BL_JB 20X Sodium Benz Base Blend | JB-Analytical | 748 | Visual Inspection | Color | Color | Colorless to Slight Yellow (Done) | [5 empty] | TM 208 | Extra_specs
4YBBY12 | 4YBBY12/BL_JB 20X Sodium Benz Base Blend | BL_JB 20X Sodium Benz Base Blend | JB-Analytical | 748 | Visual Inspection | Contamination | Contamination | None (Done) | [5 empty] | TM 208 | Extra_specs
4YBBY12 | 4YBBY12/BL_JB 20X Sodium Benz Base Blend | BL_JB 20X Sodium Benz Base Blend | JB-Analytical | 748 | Visual Inspection | Odor | Odor | Unscented (Done) | [5 empty] | TM 208 | Extra_specs
4YBBY12 | 4YBBY12/BL_JB 20X Sodium Benz Base Blend | BL_JB 20X Sodium Benz Base Blend | JB-Analytical | 425 | % Sodium Benzoate | Sodium Benzoate | Sodium Benzoate | 20.47 |  | % | FDA,2 | 17.86 | 20.68 | TM 395 | Extra_specs
4YBBY12 | 4YBBY12/BL_JB 20X Sodium Benz Base Blend | BL_JB 20X Sodium Benz Base Blend | JB-Analytical | 425 | % Sodium Benzoate | Sodium Benzoate Final | Sodium Benzoate Final | 20.41 |  | % | FDA,2 | 17.86 | 20.68 | TM 395 | Extra_specs
4YBBY23 | 4YBBY23/4YBBY23 | 4YBBY23 | JB-Analytical | 902 | Bulk Liquid Inspection | Appearance (TM 208) | Appearance | [6 empty] | See Analyte for Method # | Extra_specs
4YBBY23 | 4YBBY23/4YBBY23 | 4YBBY23 | JB-Analytical | 902 | Bulk Liquid Inspection | Color (TM 208) | Color | [6 empty] | See Analyte for Method # | Extra_specs
4YBBY23 | 4YBBY23/4YBBY23 | 4YBBY23 | JB-Analytical | 902 | Bulk Liquid Inspection | Contamination (TM 208) | Contamination | [6 empty] | See Analyte for Method # | Extra_specs
4YBBY23 | 4YBBY23/4YBBY23 | 4YBBY23 | JB-Analytical | 902 | Bulk Liquid Inspection | Odor (TM 208) | Odor | [6 empty] | See Analyte for Method # | Extra_specs
4YBBY23 | 4YBBY23/4YBBY23 | 4YBBY23 | JB-Analytical | 398 | pH | pH | pH |  |  |  | FDA,1 | 10.8 | 11.3 | TM 200 | Extra_specs
4YBBY23 | 4YBBY23/4YBBY23 | 4YBBY23 | JB-Analytical | 425 | % Sodium Benzoate | Sodium Benzoate | Sodium Benzoate |  |  | % | FDA,2 | 8.46 | 10.34 | TM 395 | Extra_specs
4YBBY23 | 4YBBY23/4YBBY23 | 4YBBY23 | JB-Analytical | 425 | % Sodium Benzoate | Sodium Benzoate Final | Sodium Benzoate Final | 0 |  | % | FDA,2 | 8.46 | 10.34 | TM 395 | Extra_specs
4YBBY23 | 4YBBY23/BL_JB 40X Natural Baby W/Vitamin E | BL_JB 40X Natural Baby W/Vitamin E | JB-Analytical | 902 | Bulk Liquid Inspection | Appearance (TM 208) | Appearance | Slightly Hazy Liquid (Done) | [5 empty] | See Analyte for Method # | Extra_specs
4YBBY23 | 4YBBY23/BL_JB 40X Natural Baby W/Vitamin E | BL_JB 40X Natural Baby W/Vitamin E | JB-Analytical | 902 | Bulk Liquid Inspection | Color (TM 208) | Color | Slightly Yellow Liquid (Done) | [5 empty] | See Analyte for Method # | Extra_specs
4YBBY23 | 4YBBY23/BL_JB 40X Natural Baby W/Vitamin E | BL_JB 40X Natural Baby W/Vitamin E | JB-Analytical | 902 | Bulk Liquid Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | See Analyte for Method # | Extra_specs
4YBBY23 | 4YBBY23/BL_JB 40X Natural Baby W/Vitamin E | BL_JB 40X Natural Baby W/Vitamin E | JB-Analytical | 902 | Bulk Liquid Inspection | Odor (TM 208) | Odor | Unscented (Done) | [5 empty] | See Analyte for Method # | Extra_specs
4YBBY23 | 4YBBY23/BL_JB 40X Natural Baby W/Vitamin E | BL_JB 40X Natural Baby W/Vitamin E | JB-Analytical | 902 | Bulk Liquid Inspection | pH (TM 200) | pH | 11.88 |  |  | FDA,2 | 11.2 | 12.6 | See Analyte for Method # | Extra_specs
4YBBY23 | 4YBBY23/BL_JB 40X Natural Baby W/Vitamin E | BL_JB 40X Natural Baby W/Vitamin E | JB-Analytical | 425 | % Sodium Benzoate | Sodium Benzoate | Sodium Benzoate | 20.11 |  | % | FDA,2 | 17.86 | 20.68 | TM 395 | Extra_specs
4YBBY23 | 4YBBY23/BL_JB 40X Natural Baby W/Vitamin E | BL_JB 40X Natural Baby W/Vitamin E | JB-Analytical | 425 | % Sodium Benzoate | Sodium Benzoate Final | Sodium Benzoate Final | 20.13 |  | % | FDA,2 | 17.86 | 20.68 | TM 395 | Extra_specs
5AA347BT | 5AA347BT/Box_NY_Skip lot testing | Box_NY_Skip lot testing | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Box Inspection | Extra_specs
5AA404S48 | 5AA404S48/Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Extra_specs
5AA404S48 | 5AA404S48/Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.018 |  | in | FDA,3 | 0.017 | 0.019 | Box Inspection | Extra_specs
5AA404S48 | 5AA404S48/Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) | [5 empty] | Box Inspection | Extra_specs
5AA404S48 | 5AA404S48/Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) | [5 empty] | Box Inspection | Extra_specs
5AA404S48 | 5AA404S48/Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | NY-Analytical | 772 | Box Inspection | Height | Height | 6.0625 |  | in | FDA,5 | 6.0312 | 6.0938 | Box Inspection | Extra_specs
5AA404S48 | 5AA404S48/Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | NY-Analytical | 772 | Box Inspection | Length | Length | 7.1875 |  | in | FDA,5 | 7.1562 | 7.2188 | Box Inspection | Extra_specs
5AA404S48 | 5AA404S48/Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.6848 |  | in | FDA,5 | 3.6562 | 3.7188 | Box Inspection | Extra_specs
5AB123CP | 5AB123CP/Label_NY_Skip lot testing | Label_NY_Skip lot testing | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Certificate of Conformance | Extra_specs
5AU66522 | 5AU66522/Box_NY_0.016BMPB_122_3.625-3-3.8125 | Box_NY_0.016BMPB_122_3.625-3-3.8125 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Extra_specs
5AU66522 | 5AU66522/Box_NY_0.016BMPB_122_3.625-3-3.8125 | Box_NY_0.016BMPB_122_3.625-3-3.8125 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.017 |  | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Extra_specs
5AU66522 | 5AU66522/Box_NY_0.016BMPB_122_3.625-3-3.8125 | Box_NY_0.016BMPB_122_3.625-3-3.8125 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | Bright Metalized Poly Board (Done) | [5 empty] | Box Inspection | Extra_specs
5AU66522 | 5AU66522/Box_NY_0.016BMPB_122_3.625-3-3.8125 | Box_NY_0.016BMPB_122_3.625-3-3.8125 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Straight Tuck End (Done) | [5 empty] | Box Inspection | Extra_specs
5AU66522 | 5AU66522/Box_NY_0.016BMPB_122_3.625-3-3.8125 | Box_NY_0.016BMPB_122_3.625-3-3.8125 | NY-Analytical | 772 | Box Inspection | Height | Height | 3.812 |  | in | FDA,5 | 3.7812 | 3.8438 | Box Inspection | Extra_specs
5AU66522 | 5AU66522/Box_NY_0.016BMPB_122_3.625-3-3.8125 | Box_NY_0.016BMPB_122_3.625-3-3.8125 | NY-Analytical | 772 | Box Inspection | Length | Length | 3.625 |  | in | FDA,5 | 3.5938 | 3.6562 | Box Inspection | Extra_specs
5AU66522 | 5AU66522/Box_NY_0.016BMPB_122_3.625-3-3.8125 | Box_NY_0.016BMPB_122_3.625-3-3.8125 | NY-Analytical | 772 | Box Inspection | Width | Width | 3 |  | in | FDA,5 | 2.9688 | 3.0312 | Box Inspection | Extra_specs
5HB59800 | 5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 125 |  |  | FDA,0 |  |  | Box Inspection | Extra_specs
5HB59800 | 5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Extra_specs
5HB59800 | 5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | Back Flip | Back Flip | 5.158 |  | in | FDA,5 | 4.6875 | 5.3125 | Box Inspection | Extra_specs
5HB59800 | 5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.001 |  | in | FDA,4 | 0.0008 | 0.0028 | Box Inspection | Extra_specs
5HB59800 | 5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | LDPE (Done) | [5 empty] | Box Inspection | Extra_specs
5HB59800 | 5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Poly Bag w/fold over top on 2” wicketts (Done) | [5 empty] | Box Inspection | Extra_specs
5HB59800 | 5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | Gussett | Gussett | 3.8 |  | in | FDA,5 | 3.6875 | 4.3125 | Box Inspection | Extra_specs
5HB59800 | 5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | Height | Height | 5.76 |  | in | FDA,5 | 5.1875 | 5.8125 | Box Inspection | Extra_specs
5HB59800 | 5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | Width | Width | 11.812 |  | in | FDA,5 | 11.6875 | 12.3125 | Box Inspection | Extra_specs
5MS48050 | 5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | UCSBS | [5 empty] | Box Inspection | Extra_specs
5MS48050 | 5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | Box Inspection | Extra_specs
5MS48050 | 5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Extra_specs
5MS48050 | 5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0155 |  | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Extra_specs
5MS48050 | 5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) | [5 empty] | Box Inspection | Extra_specs
5MS48050 | 5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | NY-Analytical | 772 | Box Inspection | Height | Height | 5.8125 |  | in | FDA,5 | 5.7812 | 5.8438 | Box Inspection | Extra_specs
5MS48050 | 5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | NY-Analytical | 772 | Box Inspection | Length | Length | 6.5 |  | in | FDA,5 | 6.4688 | 6.5312 | Box Inspection | Extra_specs
5MS48050 | 5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.3125 |  | in | FDA,5 | 3.2812 | 3.3438 | Box Inspection | Extra_specs
6AA34724 | 6AA34724/Shipper_NY_Skip lot testing | Shipper_NY_Skip lot testing | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Shipper Inspection | Extra_specs
6AP07207 | 6AP07207/Simple Kind To Skin facial 7ct Shipper Canada | Simple Kind To Skin facial 7ct Shipper Canada | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Extra_specs
6AP07207 | 6AP07207/Simple Kind To Skin facial 7ct Shipper Canada | Simple Kind To Skin facial 7ct Shipper Canada | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Extra_specs
6AP07207 | 6AP07207/Simple Kind To Skin facial 7ct Shipper Canada | Simple Kind To Skin facial 7ct Shipper Canada | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Extra_specs
6AP07207 | 6AP07207/Simple Kind To Skin facial 7ct Shipper Canada | Simple Kind To Skin facial 7ct Shipper Canada | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 4 |  | in | FDA,4 | 3.9375 | 4.0625 | TM 141 / Visual | Extra_specs
6AP07207 | 6AP07207/Simple Kind To Skin facial 7ct Shipper Canada | Simple Kind To Skin facial 7ct Shipper Canada | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 13 |  | in | FDA,4 | 12.9375 | 13.0625 | TM 141 / Visual | Extra_specs
6AP07207 | 6AP07207/Simple Kind To Skin facial 7ct Shipper Canada | Simple Kind To Skin facial 7ct Shipper Canada | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Extra_specs
6AP07207 | 6AP07207/Simple Kind To Skin facial 7ct Shipper Canada | Simple Kind To Skin facial 7ct Shipper Canada | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Extra_specs
6AP07207 | 6AP07207/Simple Kind To Skin facial 7ct Shipper Canada | Simple Kind To Skin facial 7ct Shipper Canada | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.375 |  | in | FDA,4 | 7.25 | 7.375 | TM 141 / Visual | Extra_specs
6AP09384 | 6AP09384/Shipper_NY_Skip lot testing | Shipper_NY_Skip lot testing | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | TM 141 / Visual | Extra_specs
6AP09384 | 6AP09384/Shipper_NY_Skip lot testing | Shipper_NY_Skip lot testing | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Extra_specs
6AP09384 | 6AP09384/Shipper_NY_Skip lot testing | Shipper_NY_Skip lot testing | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Extra_specs
6AP09384 | 6AP09384/Shipper_NY_Skip lot testing | Shipper_NY_Skip lot testing | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Extra_specs
6AP09384 | 6AP09384/Shipper_NY_Skip lot testing | Shipper_NY_Skip lot testing | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 3.75 |  | in | FDA,4 | 3.5625 | 3.8125 | TM 141 / Visual | Extra_specs
6AP09384 | 6AP09384/Shipper_NY_Skip lot testing | Shipper_NY_Skip lot testing | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 8.5 |  | in | FDA,4 | 8.375 | 8.625 | TM 141 / Visual | Extra_specs
6AP09384 | 6AP09384/Shipper_NY_Skip lot testing | Shipper_NY_Skip lot testing | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Extra_specs
6AP09384 | 6AP09384/Shipper_NY_Skip lot testing | Shipper_NY_Skip lot testing | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Extra_specs
6AP09384 | 6AP09384/Shipper_NY_Skip lot testing | Shipper_NY_Skip lot testing | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 5.75 |  | in | FDA,4 | 5.625 | 5.875 | TM 141 / Visual | Extra_specs
6AP71310 | 6AP71310/SHIP_NY_44K_100_18.875-11.625-9 | SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | TM 141 / Visual | Extra_specs
6AP71310 | 6AP71310/SHIP_NY_44K_100_18.875-11.625-9 | SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Extra_specs
6AP71310 | 6AP71310/SHIP_NY_44K_100_18.875-11.625-9 | SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | Extra_specs
6AP71310 | 6AP71310/SHIP_NY_44K_100_18.875-11.625-9 | SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Extra_specs
6AP71310 | 6AP71310/SHIP_NY_44K_100_18.875-11.625-9 | SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 9 |  | in | FDA,4 | 8.8125 | 9.1875 | TM 141 / Visual | Extra_specs
6AP71310 | 6AP71310/SHIP_NY_44K_100_18.875-11.625-9 | SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 18.75 |  | in | FDA,4 | 18.6875 | 19.0625 | TM 141 / Visual | Extra_specs
6AP71310 | 6AP71310/SHIP_NY_44K_100_18.875-11.625-9 | SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Extra_specs
6AP71310 | 6AP71310/SHIP_NY_44K_100_18.875-11.625-9 | SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Extra_specs
6AP71310 | 6AP71310/SHIP_NY_44K_100_18.875-11.625-9 | SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 11.625 |  | in | FDA,4 | 11.4375 | 11.8125 | TM 141 / Visual | Extra_specs
6JA580FW | 6JA580FW/Shipper_NY_Skip lot testing | Shipper_NY_Skip lot testing | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Shipper Inspection | Extra_specs
7A4OB314 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Certificate of Analysis | Extra_specs
7A4OB314 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.7837 |  |  | FDA,3 | 0.783 | 0.787 | TM 202 | Extra_specs
7A4OB314 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C |  |  |  | FDA,3 | 1.376 | 1.378 | TM 204 | Extra_specs
7A4OB314 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 797 | Nonvolatile Residue | End Weight (g) | End Weight |  |  | g | FDA,4 |  |  | TM 226 | Extra_specs
7A4OB314 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 797 | Nonvolatile Residue | Nonvolatile Residue (calc) | Nonvolatile Residue |  |  | % | FDA,3 |  | 0.005 | TM 226 | Extra_specs
7A4OB314 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 797 | Nonvolatile Residue | Nonvolatile Residue Final (calc) | Nonvolatile Residue Final |  |  | % | FDA,3 |  | 0.005 | TM 226 | Extra_specs
7A4OB314 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 797 | Nonvolatile Residue | Start Weight (g) | Start Weight |  |  | g | FDA,4 |  |  | TM 226 | Extra_specs
7A4OB314 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 797 | Nonvolatile Residue | Weight of Residue (g) | Weight of Residue |  |  | g | FDA,4 |  |  | TM 226 | Extra_specs
7A4OB314 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 795 | Acidity | Acidity (calc) | Acidity |  |  | mL | FDA,2 |  | 0.7 | TM 375 | Extra_specs
7A4OB314 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 795 | Acidity | Acidity Final (Calc) | Acidity Final (Calc) |  |  | mL | FDA,2 |  | 0.7 | TM 375 | Extra_specs
7A4OB314 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 795 | Acidity | Titrant End Volume (mL) | Titrant End Volume |  |  | mL | FDA,2 |  |  | TM 375 | Extra_specs
7A4OB314 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 795 | Acidity | Titrant Start Volume (mL) | Titrant Start Volume |  |  | mL | FDA,2 |  |  | TM 375 | Extra_specs
7A4OB314 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 796 | Assay | Assay | Assay |  |  | % | FDA,1 | 99 |  | TM 376 | Extra_specs
7A4OB314 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 796 | Assay | Assay Final | Assay Final |  |  | % | FDA,1 | 99 |  | TM 376 | Extra_specs
7A4OB314 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 390 | Identification | IR Identification | IR Identification | [6 empty] | TM 411 | Extra_specs
7A4OB314 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | [6 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7A4OB314 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | [6 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7A4OB314 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | [6 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7A4OB314 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | [6 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7A4OB314 | 7A4OB314/Chem_125A | Chem_125A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [6 empty] | Certificate of Analysis | Extra_specs
7A4OB314 | 7A4OB314/Chem_125A | Chem_125A | NY-Analytical | 398 | pH | pH (10% Solution) | pH (10% Solution) |  |  |  | FDA,1 | 6 | 8 | TM 200 | Extra_specs
7A4OB314 | 7A4OB314/Chem_125A | Chem_125A | NY-Analytical | 796 | Assay | Assay | Assay |  |  | % | FDA,1 | 76 | 84 | TM 301/434 | Extra_specs
7A4OB314 | 7A4OB314/Chem_125A | Chem_125A | NY-Analytical | 796 | Assay | Assay Final | Assay Final |  |  | % | FDA,1 | 76 | 84 | TM 301/434 | Extra_specs
7A4OB314 | 7A4OB314/Chem_125A | Chem_125A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | [6 empty] | TM 411 | Extra_specs
7A4OB314 | 7A4OB314/Chem_125A | Chem_125A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear viscous liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7A4OB314 | 7A4OB314/Chem_125A | Chem_125A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | [6 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7A4OB314 | 7A4OB314/Chem_125A | Chem_125A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7A4OB314 | 7A4OB314/Chem_125A | Chem_125A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | [6 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7A4OB314 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | pH | Extra_specs
7A4OB314 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.7839 |  |  | FDA,3 | 0.783 | 0.787 | TM 202 | Extra_specs
7A4OB314 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, mobile liquid (Done) | [5 empty] | TM 208 | Extra_specs
7A4OB314 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) | [5 empty] | TM 208 | Extra_specs
7A4OB314 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Extra_specs
7A4OB314 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | IPA - Characteristic (Done) | [5 empty] | TM 208 | Extra_specs
7A4OB314 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 796 | Assay | Assay | Assay | 99.96 |  | % | FDA,1 | 99 |  | TM 376 | Extra_specs
7A4OB314 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 796 | Assay | Assay Final | Assay Final |  |  | % | FDA,1 | 99 |  | TM 376 | Extra_specs
7A4OB314 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Match standard, per USP (Done) | [5 empty] | TM 411 | Extra_specs
7A4OP027 | 7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Extra_specs
7A4OP027 | 7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | A clear colorless to pale straw colored liquid. (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Extra_specs
7A4OP027 | 7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 792 | Chemical Inspection | Odor (J&J TM-003188) | Odor | Passes Test (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Extra_specs
7A4OP027 | 7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Extra_specs
7A4OP027 | 7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) | [5 empty] | J&J TMWW78013 | Extra_specs
7A4OP027 | 7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Microbiology | 524 | J&J Fungal | Fungal | Fungal | 0 |  | CFU/g | FDA,0 | 0 | 200 | TM 869 | Extra_specs
7A4OP027 | 7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Microbiology | 700 | J&J SDA Controls | Negative Control | Negative Control | [6 empty] | TM 869 | Extra_specs
7A4OP027 | 7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | NegativeControl | No Growth TSB | [5 empty] | TM 869 | Extra_specs
7A4OP027 | 7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | TSA JJ | 0 |  | cfu/mL | FDA,0 | 0 | 500 | TM 869 | Extra_specs
7A4OP448 | 7A4OP448/Chem_197A | Chem_197A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | pH | Extra_specs
7A4OP448 | 7A4OP448/Chem_197A | Chem_197A | NY-Analytical | 398 | pH | pH (10% in water) | pH (10% in water) | 5.9 |  |  | FDA,1 | 9 | 11 | TM 200 | Extra_specs
7A4OP448 | 7A4OP448/Chem_197A | Chem_197A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Extra_specs
7A4OP448 | 7A4OP448/Chem_197A | Chem_197A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7A4OP448 | 7A4OP448/Chem_197A | Chem_197A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Yellow (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7A4OP448 | 7A4OP448/Chem_197A | Chem_197A | NY-Analytical | 792 | Chemical Inspection | Color (TM 230) | Color |  |  |  | FDA,1 | 0 | 3 | TM208/220/230/USP/Kao LM | Extra_specs
7A4OP448 | 7A4OP448/Chem_197A | Chem_197A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7A4OP448 | 7A4OP448/Chem_197A | Chem_197A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Fatty odor (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7A4OP806 | 7A4OP806/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Extra_specs
7A4OP806 | 7A4OP806/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C | 1.1155 |  |  | FDA,4 | 1.11 | 1.12 | J&J TMWW78012 | Extra_specs
7A4OP806 | 7A4OP806/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.391 |  |  | FDA,4 | 1.386 | 1.396 | J&J TMWW78015 | Extra_specs
7A4OP806 | 7A4OP806/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 825 | Limit of Ethylene Glycol and Diethylene Glycol | Diethylene Glycol | Diethylene Glycol | 0 |  | % | FDA,2 |  | 0.1 | USP | Extra_specs
7A4OP806 | 7A4OP806/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 825 | Limit of Ethylene Glycol and Diethylene Glycol | Ethylene Glycol | Ethylene Glycol | 0 |  | % | FDA,2 |  | 0.1 | USP | Extra_specs
7A4OQ203 | 7A4OQ203/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 802 | Loss on Drying | Loss on Drying | Loss on Drying | 0.75 |  | % | FDA,2 |  | 2 | J&J TM7221 | Extra_specs
7A4OQ203 | 7A4OQ203/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 802 | Loss on Drying | Loss on Drying Final | Loss on Drying Final | 0.75 |  | % | FDA,2 |  | 2 | J&J TM7221 | Extra_specs
7A4OQ203 | 7A4OQ203/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 747 | Viscosity | Viscosity (Spindle 7@20RPM@30 sec 0.5% solution) | Viscosity (Spindle 7@20RPM@30 sec 0.5% solution) | 60667 |  | cps | FDA,0 | 45000 | 65000 | J&J TM7423 | Extra_specs
7A4OQ229 | 7A4OQ229/Chem_NY_Witch Hazel | Chem_NY_Witch Hazel | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Specific Gravity at 25°C | Extra_specs
7A4OQ229 | 7A4OQ229/Chem_NY_Witch Hazel | Chem_NY_Witch Hazel | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.981 |  |  | FDA,3 | 0.979 | 0.983 | TM 202 | Extra_specs
7A4OQ229 | 7A4OQ229/Chem_NY_Witch Hazel | Chem_NY_Witch Hazel | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7A4OQ229 | 7A4OQ229/Chem_NY_Witch Hazel | Chem_NY_Witch Hazel | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7A4OQ229 | 7A4OQ229/Chem_NY_Witch Hazel | Chem_NY_Witch Hazel | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7A4OQ229 | 7A4OQ229/Chem_NY_Witch Hazel | Chem_NY_Witch Hazel | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic witch hazel, free from acetous odor (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7A4OQ229 | 7A4OQ229/Chem_NY_Witch Hazel | Chem_NY_Witch Hazel | NY-Analytical | 797 | Nonvolatile Residue | End Weight (g) | End Weight | 240.3127 |  | g | FDA,0 |  |  | USP | Extra_specs
7A4OQ229 | 7A4OQ229/Chem_NY_Witch Hazel | Chem_NY_Witch Hazel | NY-Analytical | 797 | Nonvolatile Residue | Start Weight (g) | Start Weight | 240.3107 |  | g | FDA,0 |  |  | USP | Extra_specs
7A4OQ229 | 7A4OQ229/Chem_NY_Witch Hazel | Chem_NY_Witch Hazel | NY-Analytical | 797 | Nonvolatile Residue | Total non-volatile reside (calc) | Total non-volatile reside | 2 |  | mg/100 mL | FDA,0 |  | 25 | USP | Extra_specs
7A4OQ229 | 7A4OQ229/Chem_NY_Witch Hazel | Chem_NY_Witch Hazel | NY-Analytical | 447 | pH as is | pH | pH | 3.64 |  |  | FDA,1 | 3 | 5 | USP <791> | Extra_specs
7A4OU555 | 7A4OU555/Chem_NY_Phenylephrine Reduced | Chem_NY_Phenylephrine Reduced | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Certificate of Analysis | Extra_specs
7A4OU555 | 7A4OU555/Chem_NY_Phenylephrine Reduced | Chem_NY_Phenylephrine Reduced | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Crystalline Powder or Crystals (Done) | [5 empty] | TM 208 | Extra_specs
7A4OU555 | 7A4OU555/Chem_NY_Phenylephrine Reduced | Chem_NY_Phenylephrine Reduced | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White to almost white (Done) | [5 empty] | TM 208 | Extra_specs
7A4OU555 | 7A4OU555/Chem_NY_Phenylephrine Reduced | Chem_NY_Phenylephrine Reduced | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Extra_specs
7A4OU555 | 7A4OU555/Chem_NY_Phenylephrine Reduced | Chem_NY_Phenylephrine Reduced | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) | [5 empty] | TM 208 | Extra_specs
7A4OU555 | 7A4OU555/Chem_NY_Phenylephrine Reduced | Chem_NY_Phenylephrine Reduced | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Extra_specs
7A4OU555 | 7A4OU555/Chem_NY_Phenylephrine Reduced | Chem_NY_Phenylephrine Reduced | NY-Analytical | 796 | Assay | Assay | Assay | 99.9 |  | % | FDA,1 | 98 | 102 | USP | Extra_specs
7B4OC109 | 7B4OC109/CHE ANTIFOAM AF EMULSION | CHE ANTIFOAM AF EMULSION | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Identification | Extra_specs
7B4OC109 | 7B4OC109/CHE ANTIFOAM AF EMULSION | CHE ANTIFOAM AF EMULSION | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Extra_specs
7B4OC109 | 7B4OC109/CHE ANTIFOAM AF EMULSION | CHE ANTIFOAM AF EMULSION | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per 10mL) | TAMC (Average per 10mL) | 0 |  | cfu/10mL | FDA,1 |  | 0 | TM 610 | Extra_specs
7B4OC109 | 7B4OC109/CHE ANTIFOAM AF EMULSION | CHE ANTIFOAM AF EMULSION | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per mL - Undiluted) | TAMC (Average per mL - Undiluted) | 0 |  | cfu/mL |  |  | 0 | TM 610 | Extra_specs
7B4OC109 | 7B4OC109/CHE ANTIFOAM AF EMULSION | CHE ANTIFOAM AF EMULSION | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  | 0 | TM 610 | Extra_specs
7B4OC109 | 7B4OC109/CHE ANTIFOAM AF EMULSION | CHE ANTIFOAM AF EMULSION | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per 10mL) | TMYC (Average per 10mL) | 0 |  | cfu/10mL | FDA,1 |  | 0 | TM 610 | Extra_specs
7B4OC109 | 7B4OC109/CHE ANTIFOAM AF EMULSION | CHE ANTIFOAM AF EMULSION | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per mL - Undiluted) | TMYC (Average per mL - Undiluted) | 0 |  | cfu/mL |  |  | 0 | TM 610 | Extra_specs
7B4OC109 | 7B4OC109/CHE ANTIFOAM AF EMULSION | CHE ANTIFOAM AF EMULSION | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  | 0 | TM 610 | Extra_specs
7B4OC109 | 7B4OC109/CHE ANTIFOAM AF EMULSION | CHE ANTIFOAM AF EMULSION | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Homogeneous Liquid, Free of Dark Particles (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7B4OC109 | 7B4OC109/CHE ANTIFOAM AF EMULSION | CHE ANTIFOAM AF EMULSION | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White to off White (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7B4OC109 | 7B4OC109/CHE ANTIFOAM AF EMULSION | CHE ANTIFOAM AF EMULSION | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7B4OC109 | 7B4OC109/CHE ANTIFOAM AF EMULSION | CHE ANTIFOAM AF EMULSION | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7B4OC109 | 7B4OC109/CHE ANTIFOAM AF EMULSION | CHE ANTIFOAM AF EMULSION | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | [6 empty] | WI QAS 076 | Extra_specs
7B4OM799 | 7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | Chem_NY_Lipotec_Cucumber Extract in Glycerin | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Specific Gravity at 25°C | Extra_specs
7B4OM799 | 7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | Chem_NY_Lipotec_Cucumber Extract in Glycerin | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.2627 |  |  | FDA,3 | 1.25 | 1.29 | TM 202 | Extra_specs
7B4OM799 | 7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | Chem_NY_Lipotec_Cucumber Extract in Glycerin | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.473 |  |  | FDA,3 | 1.45 | 1.49 | TM 204 | Extra_specs
7B4OM799 | 7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | Chem_NY_Lipotec_Cucumber Extract in Glycerin | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear viscous liquid (Done) | [5 empty] | TM 208 | Extra_specs
7B4OM799 | 7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | Chem_NY_Lipotec_Cucumber Extract in Glycerin | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Light to Medium Yellow (Done) | [5 empty] | TM 208 | Extra_specs
7B4OM799 | 7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | Chem_NY_Lipotec_Cucumber Extract in Glycerin | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Extra_specs
7B4OM799 | 7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | Chem_NY_Lipotec_Cucumber Extract in Glycerin | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) | [5 empty] | TM 208 | Extra_specs
7B4OM799 | 7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | Chem_NY_Lipotec_Cucumber Extract in Glycerin | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Extra_specs
7B4OP209 | 7B4OP209/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 808 | Absorption | Absorption (1%/1cm) | Absorption (1%/1cm) | 404.16 |  |  | FDA,1 | 401.8 |  | J&J TM-002747 | Extra_specs
7B4OP209 | 7B4OP209/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 398 | pH | pH (1% Solution) | pH (1% Solution) | 6.2 |  |  | FDA,1 | 4.5 | 7 | J&J TMWW78014 | Extra_specs
7C4OA709 | 7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 961 | Peroxide Value | Peroxide Value | Peroxide Value |  |  |  | FDA,1 |  | 1 | AOCS Cd 8-53 | Extra_specs
7C4OA709 | 7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 944 | color (Gardner) | color (Gardner) | color (Gardner) |  |  |  | FDA,0 |  | 3 | AOCS Td 1a-64) | Extra_specs
7C4OA709 | 7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | COA | Extra_specs
7C4OA709 | 7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 826 | Acid Value | Acid Value | Acid Value | 1.7 |  | mg KOH/g | FDA,2 |  | 2 | IUPAC 2.201(M) | Extra_specs
7C4OA709 | 7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 826 | Acid Value | Acid Value Final | Acid Value Final | 1.7 |  | mg KOH/g | FDA,2 |  | 2 | IUPAC 2.201(M) | Extra_specs
7C4OA709 | 7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 810 | Saponification Value | Saponification Value | Saponification Value | 22.1 |  | mg/g | FDA,1 | 17 | 25 | IUPAC 2.202 | Extra_specs
7C4OA709 | 7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 810 | Saponification Value | Saponification Value Final | Saponification Value Final | 22.1 |  | mg/g | FDA,1 | 17 | 25 | IUPAC 2.202 | Extra_specs
7C4OA709 | 7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 1000 | Cloudpoint | Cloudpoint | Cloudpoint | 67.4 |  | C | FDA,1 | 66 | 71 | TC276/WG2N30 930809C | Extra_specs
7C4OA709 | 7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 959 | Water Content | Water Content | Water Content |  |  | % | FDA,2 |  | 1.5 | Vatten LB | Extra_specs
7C4OD232 | 7C4OD232/Chem_NY_Propylparaben USP | Chem_NY_Propylparaben USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Identification | Extra_specs
7C4OD232 | 7C4OD232/Chem_NY_Propylparaben USP | Chem_NY_Propylparaben USP | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Small, colorless crystals or white crystalline powder (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7C4OD232 | 7C4OD232/Chem_NY_Propylparaben USP | Chem_NY_Propylparaben USP | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7C4OD232 | 7C4OD232/Chem_NY_Propylparaben USP | Chem_NY_Propylparaben USP | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7C4OD232 | 7C4OD232/Chem_NY_Propylparaben USP | Chem_NY_Propylparaben USP | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | USP <197M> | Extra_specs
7C4OD232 | 7C4OD232/Chem_NY_Propylparaben USP | Chem_NY_Propylparaben USP | NY-Analytical | 796 | Assay | Assay | Assay | 99.4 |  | % | FDA,1 | 98 | 102 | USP/NF | Extra_specs
7C4OD232 | 7C4OD232/Chem_NY_Propylparaben USP | Chem_NY_Propylparaben USP | NY-Analytical | 796 | Assay | Assay Final | Assay Final | 99.4 |  | % | FDA,1 | 98 | 102 | USP/NF | Extra_specs
7C4OP230 | 7C4OP230/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 826 | Acid Value | Acid Value | Acid Value | 0.01 |  |  | FDA,2 | 0 | 0.2 | J&J TMWW78011 | Extra_specs
7C4OP230 | 7C4OP230/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 831 | Density @ 20° C | Density | Density | 0.8916 |  |  | FDA,4 | 0.89 | 0.893 | J&J TMWW78012 | Extra_specs
7C4OP601 | 7C4OP601/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 748 | Visual Inspection | Appearance of Packaging | Appearance of Packaging | Meets specifications (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Extra_specs
7F4OP601 | 7F4OP601/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 396 | Odor | Odor | Odor | Passes Test (Done) | [5 empty] | J&J TM-003188 | Extra_specs
7F4OP601 | 7F4OP601/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear colorless liquid with oily consistency (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Extra_specs
7G4OP027 | 7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [6 empty] | Identity of Denatonium Benzoate | Extra_specs
7G4OP027 | 7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | NY-Analytical | 824 | Identity of Denatonium Benzoate | Identification | Identification | [6 empty] | J&J TM4028 | Extra_specs
7G4OP027 | 7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | [6 empty] | J&J TMWW78010/J&J TM7252 | Extra_specs
7G4OP027 | 7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | NY-Analytical | 792 | Chemical Inspection | Odor (J&J TM-003188) | Odor | [6 empty] | J&J TMWW78010/J&J TM7252 | Extra_specs
7G4OP027 | 7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | [6 empty] | J&J TMWW78010/J&J TM7252 | Extra_specs
7G4OP027 | 7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C |  |  |  | FDA,4 | 0.7905 | 0.792 | J&J TMWW78012 | Extra_specs
7G4OP027 | 7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | NY-Analytical | 822 | Specific Gravity @ 15.56°C/ 60° F | Specific Gravity | Specific Gravity at 15.56°C | 0.7942 |  |  | FDA,4 |  | 0.7952 | J&J TMWW78012 | Extra_specs
7G4OP027 | 7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | NY-Analytical | 390 | Identification | IR Identification | IR Identification | [6 empty] | J&J TMWW78013 | Extra_specs
7G4OP601 | 7G4OP601/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 828 | Iodine Value | Iodine Value | Iodine Value | 0.6 |  | g I2/100g | FDA,1 |  | 2 | J&J TM-003029 | Extra_specs
7G4OP601 | 7G4OP601/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 828 | Iodine Value | Iodine Value Final | Iodine Value Final | 0.6 |  | g I2/100g | FDA,1 |  | 2 | J&J TM-003029 | Extra_specs
7G4OP601 | 7G4OP601/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 792 | Chemical Inspection | Color (J&J TM-000839) | Color | 1 |  |  | FDA,1 |  | 1 | J&J TMWW78010/J&J TM7252 | Extra_specs
7G4OP601 | 7G4OP601/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 826 | Acid Value | Acid Value Final | Acid Value Final | 0.1 |  | mg KOH/g | FDA,1 |  | 2 | J&J TMWW78011 | Extra_specs
7G4OP601 | 7G4OP601/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.451 |  |  | FDA,4 | 1.451 | 1.453 | J&J TMWW78015 | Extra_specs
7G4OP601 | 7G4OP601/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 810 | Saponification Value | Saponification Value | Saponification Value | 108.1 |  | mg KOH/g | FDA,1 | 100 | 115 | J&J TMWW78016 | Extra_specs
7G4OP601 | 7G4OP601/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 810 | Saponification Value | Saponification Value Final | Saponification Value Final | 108.1 |  | mg KOH/g | FDA,1 | 100 | 115 | J&J TMWW78016 | Extra_specs
7G4OQ028 | 7G4OQ028/Chem_Aminomethylpropanol | Chem_Aminomethylpropanol | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Identification | Extra_specs
7G4OQ028 | 7G4OQ028/Chem_Aminomethylpropanol | Chem_Aminomethylpropanol | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Extra_specs
7G4OQ028 | 7G4OQ028/Chem_Aminomethylpropanol | Chem_Aminomethylpropanol | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Match standard (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7G4OQ028 | 7G4OQ028/Chem_Aminomethylpropanol | Chem_Aminomethylpropanol | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Match standard (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7G4OQ028 | 7G4OQ028/Chem_Aminomethylpropanol | Chem_Aminomethylpropanol | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7G4OQ028 | 7G4OQ028/Chem_Aminomethylpropanol | Chem_Aminomethylpropanol | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7H4OP601 | 7H4OP601/J&J Chemical CHE DUB PTO | J&J Chemical CHE DUB PTO | NY-Microbiology | 524 | J&J Fungal | Fungal | Fungal |  |  | CFU/g | FDA,0 | 0 | 200 | TM 869 | Extra_specs
7H4OP601 | 7H4OP601/J&J Chemical CHE DUB PTO | J&J Chemical CHE DUB PTO | NY-Microbiology | 701 | J&J TSA Controls | Negative Control | Negative Control | [6 empty] | TM 869 | Extra_specs
7H4OP601 | 7H4OP601/J&J Chemical CHE DUB PTO | J&J Chemical CHE DUB PTO | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | NegativeControl | [6 empty] | TM 869 | Extra_specs
7H4OP601 | 7H4OP601/J&J Chemical CHE DUB PTO | J&J Chemical CHE DUB PTO | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | TSA JJ |  |  | cfu/mL | FDA,0 | 0 | 2000 | TM 869 | Extra_specs
7J4OP027 | 7J4OP027/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 792 | Chemical Inspection | Solubility | Solubility | Passes Test Soluble in any proportion of water (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Extra_specs
7J4OP027 | 7J4OP027/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 447 | pH as is | pH | pH | 5.59 |  |  | FDA,1 | 4 | 6.5 | TM WW78014 | Extra_specs
7L4OP027 | 7L4OP027/Chem_NY_Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Extra_specs
7L4OP027 | 7L4OP027/Chem_NY_Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Colorless to white needle crystal or granule (Done) | [5 empty] | JJTMWW78010 / JJTM7252 | Extra_specs
7L4OP027 | 7L4OP027/Chem_NY_Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) | [5 empty] | JJTMWW78010 / JJTM7252 | Extra_specs
7L4OP027 | 7L4OP027/Chem_NY_Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 809 | Melting Range | Melting Range - End | Melting Range - End | 42.5 |  | C | FDA,0 | 41 | 44 | USP | Extra_specs
7L4OP027 | 7L4OP027/Chem_NY_Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 809 | Melting Range | Melting Range - Start | Melting Range - Start | 41.2 |  | C | FDA,0 | 41 | 44 | USP | Extra_specs
7L4OP027 | 7L4OP027/Chem_NY_Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 797 | Nonvolatile Residue | Nonvolatile Residue | Nonvolatile Residue | 0.02 |  | % w/v | FDA,2 |  | 0.05 | USP | Extra_specs
7L4OP027 | 7L4OP027/Chem_NY_Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 863 | Specific Rotation | Specific Rotation | Specific Rotation | -49.9 |  | ° | FDA,0 | -51 | -45 | USP | Extra_specs
7L4OP027 | 7L4OP027/Chem_NY_Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 796 | Assay | Assay | Assay | 100.4 |  | % | FDA,1 | 98 | 102 | USP/NF | Extra_specs
7L4OP027 | 7L4OP027/Chem_NY_Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 796 | Assay | Assay Final | Assay Final | 100.4 |  | % | FDA,1 | 98 | 102 | USP/NF | Extra_specs
7L4OP027 | 7L4OP027/Chem_NY_Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 455 | Identification A | Identification A | Identification A | Passes Test (Done) | [5 empty] | USP/NF | Extra_specs
7L4OP027 | 7L4OP027/Chem_NY_Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 456 | Identification B | Identification B | Identification B | Passes Test (Done) | [5 empty] | USP/NF | Extra_specs
7L4OP027 | 7L4OP027/Chem_NY_Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 964 | Individual Impurity | Iso-Menthone | Iso-Menthone | 0.3 |  | % | FDA,1 |  | 0.3 | USP/NF | Extra_specs
7L4OP027 | 7L4OP027/Chem_NY_Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 964 | Individual Impurity | Iso-Pulegol | Iso-Pulegol | 0.14 |  | % | FDA,1 |  | 0.3 | USP/NF | Extra_specs
7L4OP027 | 7L4OP027/Chem_NY_Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 964 | Individual Impurity | Menthone | Menthone | 0.04 |  | % | FDA,1 |  | 0.3 | USP/NF | Extra_specs
7L4OP027 | 7L4OP027/Chem_NY_Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 964 | Individual Impurity | Neo-Menthol | Neo-Menthol | 0.18 |  | % | FDA,1 |  | 0.3 | USP/NF | Extra_specs
7L4OP027 | 7L4OP027/Chem_NY_Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 964 | Individual Impurity | Pulegone | Pulegone | 0.08 |  | % | FDA,1 |  | 0.3 | USP/NF | Extra_specs
7L4OP027 | 7L4OP027/Chem_NY_Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 965 | Total Impurities | Total Impurities | Total Impurities | 0.53 |  | % | FDA,1 |  | 2 | USP/NF | Extra_specs
7L4OP027 | 7L4OP027/Chem_NY_Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 964 | Individual Impurity | Unknown impurities | Unknown impurities | 0.05 |  | % | FDA,1 |  | 0.1 | USP/NF | Extra_specs
7O4AQ084 | 7O4AQ084/Chem_162A | Chem_162A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Extra_specs
7O4AQ084 | 7O4AQ084/Chem_162A | Chem_162A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7O4AQ084 | 7O4AQ084/Chem_162A | Chem_162A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Clear/Colorless (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7O4AQ084 | 7O4AQ084/Chem_162A | Chem_162A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7O4AQ084 | 7O4AQ084/Chem_162A | Chem_162A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Light Amine (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7O4OH460 | 7O4OH460/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 801 | % Water | Moisture Content | Moisture Content | 0.6 |  | % | FDA,1 |  | 1.5 | J&J TMWW78017 | Extra_specs
7O4OP027 | 7O4OP027/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 827 | Hydroxyl Value | Hydroxyl Value | Hydroxyl Value | 38.5 |  | mg KOH/g | FDA,1 | 35 | 50 | J&J TMWW7671 | Extra_specs
7O4OP027 | 7O4OP027/MASTER - Microbiology - J&J Chemical Testing) | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 827 | Hydroxyl Value | Hydroxyl Value Final | Hydroxyl Value Final | 38.5 |  | mg KOH/g | FDA,1 | 35 | 50 | J&J TMWW7671 | Extra_specs
7O4OQ028 | 7O4OQ028/Chem_NY_Euxyl K703 | Chem_NY_Euxyl K703 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Density @ 20° C | Extra_specs
7O4OQ028 | 7O4OQ028/Chem_NY_Euxyl K703 | Chem_NY_Euxyl K703 | NY-Analytical | 831 | Density @ 20° C | Density | Density | 1.1282 |  |  | FDA,3 | 1.124 | 1.128 | TM 202 | Extra_specs
7O4OQ028 | 7O4OQ028/Chem_NY_Euxyl K703 | Chem_NY_Euxyl K703 | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.539 |  |  | FDA,3 | 1.539 | 1.543 | TM 204 | Extra_specs
7O4OQ028 | 7O4OQ028/Chem_NY_Euxyl K703 | Chem_NY_Euxyl K703 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Extra_specs
7O4OQ028 | 7O4OQ028/Chem_NY_Euxyl K703 | Chem_NY_Euxyl K703 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7O4OQ028 | 7O4OQ028/Chem_NY_Euxyl K703 | Chem_NY_Euxyl K703 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Yellow (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7O4OQ028 | 7O4OQ028/Chem_NY_Euxyl K703 | Chem_NY_Euxyl K703 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7O4OQ028 | 7O4OQ028/Chem_NY_Euxyl K703 | Chem_NY_Euxyl K703 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Phenoxyethanol-like (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Extra_specs
7O4OU351 | 7O4OU351/Chem_NY_Glycopyrronium Tosylate Monohydrate | Chem_NY_Glycopyrronium Tosylate Monohydrate | NY-Analytical | 1051 | Temptale Reading | Shipment Temperatutre | Shipment Temperatutre | Send Temperature data to Dermira (Done) | [5 empty] | Shipment Temperatutre | Extra_specs
7O4OU351 | 7O4OU351/Chem_NY_Glycopyrronium Tosylate Monohydrate | Chem_NY_Glycopyrronium Tosylate Monohydrate | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | COA | Extra_specs
7O4OU351 | 7O4OU351/Chem_NY_Glycopyrronium Tosylate Monohydrate | Chem_NY_Glycopyrronium Tosylate Monohydrate | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis (ELLI) | Certificate of Analysis | COA Meets Specification (Done) | [5 empty] | COA | Extra_specs
7O4OU555 | 7O4OU555/Chem_NY_Lidocaine Reduced testing | Chem_NY_Lidocaine Reduced testing | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Certificate of Analysis | Extra_specs
7O4OU555 | 7O4OU555/Chem_NY_Lidocaine Reduced testing | Chem_NY_Lidocaine Reduced testing | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Match standard (Done) | [5 empty] | TM 208 | Extra_specs
7O4OU555 | 7O4OU555/Chem_NY_Lidocaine Reduced testing | Chem_NY_Lidocaine Reduced testing | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Match standard (Done) | [5 empty] | TM 208 | Extra_specs
7O4OU555 | 7O4OU555/Chem_NY_Lidocaine Reduced testing | Chem_NY_Lidocaine Reduced testing | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Extra_specs
7O4OU555 | 7O4OU555/Chem_NY_Lidocaine Reduced testing | Chem_NY_Lidocaine Reduced testing | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) | [5 empty] | TM 208 | Extra_specs
7O4OU555 | 7O4OU555/Chem_NY_Lidocaine Reduced testing | Chem_NY_Lidocaine Reduced testing | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Extra_specs
7O4OU555 | 7O4OU555/Chem_NY_Lidocaine Reduced testing | Chem_NY_Lidocaine Reduced testing | NY-Analytical | 796 | Assay | Assay | Assay | 100.8 |  | % | FDA,1 | 97.5 | 102.5 | USP | Extra_specs
8AA654SI | 8AA654SI/Label_A10 | Label_A10 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Label Inspection | Extra_specs
8AA654SI | 8AA654SI/Label_A10 | Label_A10 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Extra_specs
8AA654SI | 8AA654SI/Label_A10 | Label_A10 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Extra_specs
8AA654SI | 8AA654SI/Label_A10 | Label_A10 | NY-Analytical | 770 | Label Inspection | Length | Length | 1.375 |  | in | FDA,4 | 1.3125 | 1.4375 | TM 143 / Visual | Extra_specs
8AA654SI | 8AA654SI/Label_A10 | Label_A10 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Extra_specs
8AA654SI | 8AA654SI/Label_A10 | Label_A10 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Extra_specs
8AA654SI | 8AA654SI/Label_A10 | Label_A10 | NY-Analytical | 770 | Label Inspection | Width | Width | 2.5625 |  | in | FDA,3 | 2.5 | 2.625 | TM 143 / Visual | Extra_specs
8AA654SI | 8AA654SI/Label_NY_Skip lot testing | Label_NY_Skip lot testing | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | [6 empty] | TM 143 / Visual | Extra_specs
8AA654SI | 8AA654SI/Label_NY_Skip lot testing | Label_NY_Skip lot testing | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Extra_specs
8AA654SI | 8AA654SI/Label_NY_Skip lot testing | Label_NY_Skip lot testing | NY-Analytical | 770 | Label Inspection | Length | Length | 1.375 |  | in | FDA,4 | 1.3125 | 1.4375 | TM 143 / Visual | Extra_specs
8AA654SI | 8AA654SI/Label_NY_Skip lot testing | Label_NY_Skip lot testing | NY-Analytical | 770 | Label Inspection | Odor | Odor | [6 empty] | TM 143 / Visual | Extra_specs
8AA654SI | 8AA654SI/Label_NY_Skip lot testing | Label_NY_Skip lot testing | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Extra_specs
8AA654SI | 8AA654SI/Label_NY_Skip lot testing | Label_NY_Skip lot testing | NY-Analytical | 770 | Label Inspection | Width | Width | 2.562 |  | in | FDA,3 | 2.5 | 2.625 | TM 143 / Visual | Extra_specs
8AH12880 | 8AH12880/Label_NY_NP7030_9-2.625 | Label_NY_NP7030_9-2.625 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Label Inspection | Extra_specs
8AH12880 | 8AH12880/Label_NY_NP7030_9-2.625 | Label_NY_NP7030_9-2.625 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Extra_specs
8AH12880 | 8AH12880/Label_NY_NP7030_9-2.625 | Label_NY_NP7030_9-2.625 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Extra_specs
8AH12880 | 8AH12880/Label_NY_NP7030_9-2.625 | Label_NY_NP7030_9-2.625 | NY-Analytical | 770 | Label Inspection | Length | Length | 9 |  | in | FDA,4 | 8.9375 | 9.0625 | TM 143 / Visual | Extra_specs
8AH12880 | 8AH12880/Label_NY_NP7030_9-2.625 | Label_NY_NP7030_9-2.625 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Extra_specs
8AH12880 | 8AH12880/Label_NY_NP7030_9-2.625 | Label_NY_NP7030_9-2.625 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Extra_specs
8AH12880 | 8AH12880/Label_NY_NP7030_9-2.625 | Label_NY_NP7030_9-2.625 | NY-Analytical | 770 | Label Inspection | Width | Width | 2.625 |  | in | FDA,4 | 2.5625 | 2.6875 | TM 143 / Visual | Extra_specs
8CP63801 | 8CP63801/Label_A20 | Label_A20 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Label Inspection | Extra_specs
8CP63801 | 8CP63801/Label_A20 | Label_A20 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | [6 empty] | TM 143 / Visual | Extra_specs
8CP63801 | 8CP63801/Label_A20 | Label_A20 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | [6 empty] | TM 143 / Visual | Extra_specs
8CP63801 | 8CP63801/Label_A20 | Label_A20 | NY-Analytical | 770 | Label Inspection | Length | Length |  |  | in | FDA,4 | 5.9375 | 6.0625 | TM 143 / Visual | Extra_specs
8CP63801 | 8CP63801/Label_A20 | Label_A20 | NY-Analytical | 770 | Label Inspection | Odor | Odor | [6 empty] | TM 143 / Visual | Extra_specs
8CP63801 | 8CP63801/Label_A20 | Label_A20 | NY-Analytical | 770 | Label Inspection | Print | Print | [6 empty] | TM 143 / Visual | Extra_specs
8CP63801 | 8CP63801/Label_A20 | Label_A20 | NY-Analytical | 770 | Label Inspection | Width | Width |  |  | in | FDA,4 | 3.6875 | 3.8125 | TM 143 / Visual | Extra_specs
8OX351CK | 8OX351CK/Label_NY_50#W_5_4 | Label_NY_50#W_5_4 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Label Inspection | Extra_specs
8OX351CK | 8OX351CK/Label_NY_50#W_5_4 | Label_NY_50#W_5_4 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Extra_specs
8OX351CK | 8OX351CK/Label_NY_50#W_5_4 | Label_NY_50#W_5_4 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Extra_specs
8OX351CK | 8OX351CK/Label_NY_50#W_5_4 | Label_NY_50#W_5_4 | NY-Analytical | 770 | Label Inspection | Length | Length | 5.01 |  | in | FDA,4 | 4.9375 | 5.0625 | TM 143 / Visual | Extra_specs
8OX351CK | 8OX351CK/Label_NY_50#W_5_4 | Label_NY_50#W_5_4 | NY-Analytical | 770 | Label Inspection | NP Structure | NP Structure | 50# White Paper Stock (Done) |  | in |  |  |  | TM 143 / Visual | Extra_specs
8OX351CK | 8OX351CK/Label_NY_50#W_5_4 | Label_NY_50#W_5_4 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Extra_specs
8OX351CK | 8OX351CK/Label_NY_50#W_5_4 | Label_NY_50#W_5_4 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Extra_specs
8OX351CK | 8OX351CK/Label_NY_50#W_5_4 | Label_NY_50#W_5_4 | NY-Analytical | 770 | Label Inspection | Width | Width | 3.99 |  | in | FDA,4 | 3.9375 | 4.0625 | TM 143 / Visual | Extra_specs
9OX465GL | 9OX465GL/NY_Glue_01 | NY_Glue_01 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Visual | Extra_specs
9OX465GL | 9OX465GL/NY_Glue_01 | NY_Glue_01 | NY-Analytical | 395 | Color | Color | Color | Very Pale Amber (Done) | [5 empty] | Visual | Extra_specs
| / | [13 empty] | Extra_specs
| / | [13 empty] | Extra_specs

## Count
Plan.C | Count
2AA284SLI/Lid_01B | 1
2AA347/Web_NP8550_10.25 | 1
2AA592/Web_NP8504_11.0 | 3
2AA592SLI/Lid_NY_NP7600_22R1 | 8
2AA632/Web_NP8504_11.75 | 5
2AA709/Web_NP8504_11.0 | 1
2AA922SLI/Lid_01V | 2
2AA922SLI/Lid_NY_Skip lot testing | 1
2AD173/Web_NP5115_4.5 | 4
2AJ123/Web_NY_Skip lot testing | 1
2AJ123SLI/Lid_NY_Skip lot testing | 2
2AJ143SLI/Lid_NY_Skip lot testing | 1
2AJ170/Web_NP8530_11.0 | 1
2AJ170SLI/Lid_NY_NP7600_22FA | 1
2AJ190SLI/Lid_NY_NP7600_22V | 1
2AJ324/Web_NY_Skip lot testing | 1
2AJ324SLI/Lid_NY_NP7600_22X | 1
2AP034/Web_NY_Skip lot testing | 2
2AP075/Web_NP8542_12.205 | 3
2AP087/Web_NY_Skip lot testing | 4
2AP091/Web_NP8550_12.205 | 1
2AP159LI/Lid_NY_Skip lot testing | 2
2AP210/Web_NP8542_11.0 | 1
2AP211/Web_NP8512_12.205 | 1
2AP227/Web_NP8542_11.0 | 3
2AP328/Web_NP8512_12.205 | 1
2AP345/Web_NY_Skip lot testing | 1
2AP713/Web_NY_Skip lot testing | 2
2AQ471LI/Lid_04B_560_1 | 4
2AU191/Web_NP5115_7.5 | 1
2AU252/Web_3459_8 | 1
2AU252/Web_NP3459_8.0 | 1
2AU265/Web_NY_Skip lot testing | 1
2AU372W/Web_NY_Skip lot testing | 2
2AU444/Web_NP5115_7.5 | 7
2AU555/Web_NP3460_7.5 | 6
2BC801/Web_NP3870_6.0 | 2
2BH592/Web_NP3895_6.0 | 1
2BJ123/Web_NP8550_11 | 1
2BJ143/Web_NP8530_16.25 | 1
2BJ213/Web_NP8530_18 | 1
2BJ227/Web_NY_Skip lot testing | 2
2BP308SLI/Lid_Aveeno Ultra-Calming MUR | 2
2BQ084LI/Lid_11A | 1
2BQ084LI/Lid_NY_Skip lot testing | 6
2BQ471LI/Lid_04F_560 | 2
2BQ471LI/Lid_NY_Skip lot testing | 3
2BQ551LI/Applicator_NY_Skip lot testing | 1
2BQ551LI/Lid_13A | 10
2BQ551LI/Lid_NY_Skip lot testing | 80
2BS420/Web_NP3870_11.5 | 1
2BU275/Web_NP3895_8.5 | 1
2CA580SLI/Lid_NY_Skip lot testing | 5
2CA735/Web_NP8810_12.125 | 1
2CA911/Web_NP8810_12.125 | 2
2CB196/Web_NP5114_41.0 | 2
2CB512/Web_NP3870_4.5 | 2
2CD333/Web_NP5115_4.5 | 2
2DH128/Web_NP5115_7.5 | 2
2DH211/Web_NP5116_6.0 | 1
2ED110/Web_NP5410_4.5 | 1
2ED110/Web_NY_Skip lot testing | 2
2ED419/Web_NY_Skip lot testing | 7
2ED436/Web_NY_Skip lot testing | 6
2EU351/Web_NP5211_7.0 | 1
2FB6038/Web_NP5410_8.0 | 7
2FH070/Web_NP5115_6.0 | 1
2GB470/Web_NP5115_4.5 | 2
2HA580/Web_NP8540_17.0 | 1
2HA580/Web_NY_Skip lot testing | 4
2HD748/Web_NP5115_4.5 | 1
2HD748/Web_NY_Skip lot testing | 4
2HU120/Web_NY_Skip lot testing | 2
2JC223/Web_NP3870_6.0 | 2
2JC699/Web_NP5410_6.0 | 6
2KH040/Web_NY_Skip lot testing | 7
2LB406/Web_NP3870_10.0 | 1
2MB164/Web_NP5410_8.0 | 5
2MB712/Web_NP5115_4.5 | 1
2MS413/Web_NP3870_11.5 | 2
2OA347/Web_NY_Skip lot testing | 2
2OA347SLI/Lid_NY_Skip lot testing | 3
2OA632/Web_NP8504_11.75 | 1
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 5
2OB123/Web_NP5114_41.0 | 2
2OB963/Web_NP5115_4.5 | 1
2OD024/Web_NP5115_7.5 | 1
2OD029/Web_NP5115_7.5 | 2
2OH450/Web_NY_Skip lot testing | 1
2OH581/Web_NY_Skip lot testing | 2
2OJ213SLI/Lid_NY_Skip lot testing | 2
2OM304/Web_NP8540_16.0 | 1
2OM304/Web_NY_Skip lot testing | 6
2OM304SLI/Lid_NY_Skip lot testing | 5
2OM701/Web_NP8540_16.0 | 1
2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | 1
2OP036LI/Lid_NY_Skip lot testing | 4
2OP037/Web_NP8511_245 | 3
2OP063/Web_NP8511_320 | 1
2OP065/Web_NP8536_11.875 | 1
2OP067/Web_NP8536_11.875 | 1
2OP067/Web_NY_Skip lot testing | 1
2OP072/Web_NP8511_245 | 1
2OP073/Web_NP8536_11.875 | 1
2OP079/Web_NP8541_12.205 | 1
2OP081/Web_NP8536_11.875 | 2
2OP092/Web_NP8536_10 | 1
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 12
2OP093LI/Lid_NY_Threaded Blue Cap | 3
2OP134/Reduced Testing_NY_Can 2OP134 | 8
2OP134/Skip Lot Release_NY_2OP134 | 8
2OP138/Reduced Testing_NY_Can 2OP138 | 9
2OP138/Skip Lot Relrase_NY_2OP138 | 220
2OP138/Skip Lot Relrase_NY_2OP138C | 1
2OP138/Skip Lot Relrase_NY_2OP638 | 2
2OP138C/Skip Lot Relrase_NY_2OP138 | 1
2OP138C/Skip Lot Relrase_NY_2OP138C | 11
2OP159LI/Lid_03B | 1
2OP159LI/Lid_NY_Skip lot testing | 4
2OP195/Web_NP8511_320 | 1
2OP195/Web_NY_Skip lot testing | 1
2OP213/Web_NP8541_11 | 1
2OP223/Web_NP8511_320 | 4
2OP242LI/Lid_17C | 4
2OP280LI/Lid_17A | 1
2OP280LI/Lid_NY_Skip lot testing | 17
2OP300LI/Lid_NY_Skip lot testing | 2
2OP520LI/Lid_NY_Skip lot testing | 4
2OP587/Web_NP8550_12.205 | 1
2OP638/Lid_NY_Skip lot testing | 1
2OP638/Reduced Testing_NY_Can 2OP638 | 8
2OP638/Skip Lot Relrase_NY_2OP138 | 2
2OP638/Skip Lot Relrase_NY_2OP638 | 242
2OP638LI/Lid_13D | 3
2OP638LI/Lid_NY_Skip lot testing | 27
2OP638LI/Skip Lot Relrase_NY_2OP638 | 1
2OQ753/Reduced Testing_NY_Can 2OQ753 | 18
2OQ753/Skip Lot Release_NY_2OQ753 | 103
2OQ753/Skip Lot Release_NY_2OQ850 | 1
2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 1
2OQ753X/Skip Lot Relrase_NY_2OQ753X | 11
2OQ850/Reduced Testing_NY_Can 2OQ850 | 16
2OQ850/Skip Lot Release_NY_2OQ753 | 1
2OQ850/Skip Lot Release_NY_2OQ850 | 98
2OQ850LI/Lid_11B | 1
2OQ850LI/Lid_NY_Skip lot testing | 16
2OQ850X/Reduced Testing_NY_Can 2OQ850X | 2
2OQ850X/Skip Lot Release_NY_2OQ850X | 29
2OQ882LI/Lid_12A | 1
2OS123/Web_NP3870_11.5 | 1
2OS456/Web_NP3870_11.5 | 1
2OU500/Web_NY_Skip lot testing | 3
2OU573/Web_NY_Skip lot testing | 1
2OU611/Web_NP5211_7.0 | 1
2OX465B/Bottle, 26ml | 1
2OX510B/Bottle, 5.1ml | 1
2OX510B/Microbiologu - Bioburden_Applicator_NY | 1
2PD351/Web_NP5116_4.5 | 5
2PU872/Lid_NY_Skip lot testing | 8
2PU872/Web_NP5115_8.5 | 1
2QB114X/Web_NP3860_5.0 | 1
2QS411/Web_NP3870_11.5 | 4
2RB108/Web_NP5114_41.0 | 3
2RB114X/Web_NP3860_5.0 | 1
2RQ890/Web_NP9508_18.125 | 2
2RQ890/Web_NY_Skip lot testing | 8
2RQ890P/Web_NP9508_18.125 | 2
2RS324/Web_NP3860_12.5 | 1
2RS407/Web_NP3885_13.0 | 4
2RS419/Web_NP3885_16 | 2
2SB108/Web_NP5114_41.0 | 4
2SS273/Web_NP3860_15.5 | 1
2SS428/Web_NP3860_8.0 | 2
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 1
3AU228/Applicator_NY_Skip lot testing | 2
3BJ141/Applicator_NY_Skip lot testing | 5
3DQ595/Applicator_10D | 1
3DQ595N/Applicator_10I | 1
3DQ595N/Applicator_NY_Skip lot testing | 14
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 19
3OA347/Microbiologu - Bioburden_Applicator_NY | 17
3OA426/Applicator_NY_Skip lot testing | 4
3OA580/Applicator_38A | 2
3OA580/Applicator_NY_Skip lot testing | 25
3OA580/Lid_NY_Skip lot testing | 1
3OB108/REDUCED TESTING_NY_3OB108 | 1
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 1
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 1
3OD438/REDUCED TESTING_NY_3OD438 | 1
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 1
3OH581/Applicator_NY_Skip lot testing | 1
3OH592/Applicator_NY_Skip lot testing | 4
3OJ213/Applicator_08A | 4
3OJ213/Applicator_NY_Skip lot testing | 6
3OM304/Applicator_NY_Skip lot testing | 28
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 2
3OP065/Applicator_NY_Skip lot testing | 3
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 1
3OP093PD/Applicator_NY_Die Cut Pads | 3
3OP113/Applicator_NY_Skip lot testing | 3
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3
3OP138/Applicator_36D | 2
3OP138/Applicator_NY_Skip lot testing | 19
3OP195/Applicator_NY_Skip lot testing | 1
3OP197A-SU/Label_NY_NP7200_6.25-4.5 | 1
3OP197A/Applicator_NY_Skip lot testing | 2
3OP257/Applicator_NY_Skip lot testing | 7
3OP705/Applicator_NY_Skip lot testing | 3
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 1
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 19
3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 1
3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 1
3OQ222/Applicator_NY_Skip lot testing | 46
3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 31
3OQ333/Applicator_NY_36" Embossed Spunlaced Nonwoven | 1
3OQ333/Applicator_NY_Skip lot testing | 58
3OQ444/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 1
3OQ444/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 1
3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 19
3OQ444/Applicator_NY_Skip lot testing | 43
3OQ551/Applicator_15E | 4
3OQ551/Applicator_NY_Skip lot testing | 33
3OQ850N/Applicator_NY_Skip lot testing | 7
3OQ890N/Applicator_NY_Skip lot testing | 8
3OS407/Applicator_25B | 6
3OS419/Applicator_25C | 3
3OU252/Applicator_NY_Skip lot testing | 1
3OU265/Applicator_NY_Skip lot testing | 2
3OU275/Applicator_NY_Skip lot testing | 4
3OU351/Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | 3
3OX105F/Applicatpr_NY_Foam 10.5ml | 1
3OX465F/Microbiologu - Bioburden_Applicator_NY | 2
3OX465H/Handle, 26 ml | 1
3OX465H/Microbiologu - Bioburden_Applicator_NY | 1
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 1
3OX510H/Handle, 5.1 ml | 1
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 4
3XP138/Applicator_36B | 1
3XP138/Applicator_NY_Skip lot testing | 11
3XP540/Applicator_NY_Skip lot testing | 18
3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 12
3XP667/Applicator_NY_Skip lot testing | 19
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 1
4ABBY12/Skip Lot Relrase_NY_2OP638 | 1
4FQ51801/BL_Sani_07C | 1
5AA347BT/Box_NY_Skip lot testing | 1
5AA540S48/Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | 2
5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | 6
5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | 2
5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | 1
5AB123CP/Label_NY_Skip lot testing | 1
5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | 3
5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | 1
5AD24301/Box_NY_0.018SBS_107_4-2.03125-4.5 | 1
5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | 2
5AH58195/Box_NY_Skip lot testing | 3
5AP027TY/Box_57A | 7
5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | 1
5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | 3
5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | 2
5AP308TY/Box_63A | 3
5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | 2
5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | 2
5AU50010/Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | 1
5AU50010/Box_NY_Skip lot testing | 2
5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | 1
5AU57314/Box_NY_Skip lot testing | 1
5BA347BT/Box_NY_Skip lot testing | 1
5BA617S96/Box_NY_0.018SBS_47B_3.625-3.875-6 | 2
5BA735TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | 1
5BH59200/Box_NY_0.020SBS_127_6.25-2.5625-4.625 | 1
5BJ123BT/Box_NY_Skip lot testing | 3
5BJ170BT/Box_NY_Skip lot testing | 1
5BJ190BT/Box_NY_Skip lot testing | 1
5BJ324BT/Box_NY_Skip lot testing | 1
5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | 1
5BU26595/Box_NY_Skip lot testing | 4
5BU27500/Box_NY_Skip lot testing | 2
5BU37230/Box_NY_Skip lot testing | 1
5BU44410/Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | 4
5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | 2
5BU62620/Box_NY_0.016BMPB_122_3.625-3-3.8125 | 5
5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | 2
5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | 1
5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 5
5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | 2
5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | 3
5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 4
5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 1
5CU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | 2
5DA941TY/Box_NY_Skip lot testing | 1
5DD43600/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 6
5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | 2
5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | 1
5FD41900/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 1
5FD41900/Box_NY_Skip lot testing | 6
5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | 1
5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 1
5HB59800/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | 3
5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | 1
5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | 5
5HD74800/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 1
5HD74800/Box_NY_Skip lot testing | 3
5HS48450/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 1
5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | 4
5KB71200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | 1
5KC69900/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | 6
5KH04095/Box_NY_0.020SBS_127_6.25-2.5625-4.625 | 1
5KH04095/Box_NY_Skip lot testing | 7
5KS41125/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 6
5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | 7
5LS41350/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 4
5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | 1
5MS82125/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 2
5NB10800/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 4
5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | 3
5NS23125/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 4
5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | 1
5NU87250/Box_NY_Skip lot testing | 8
5OH382PD/Box_NY_Skip lot testing | 1
5OJ124BT/Box_NY_Skip lot testing | 1
5OP037TY/Box_NY_Skip lot testing | 3
5OP065BT/Box_NY_Skip lot testing | 2
5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | 1
5OP073BT/Box_NY_Skip lot testing | 1
5OP081BT/Box_NY_Skip lot testing | 3
5OP081PB/Box_NY_Skip lot testing | 3
5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | 1
5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | 1
5OP195TY/Box_NY_Skip lot testing | 4
5OP308PD/Box_50C | 1
5OP624PB/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 1
5OP624PB/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | 1
5OP624PB/Box_NY_Skip lot testing | 1
5OP705PB/Box_NY_Skip lot testing | 2
5PB10800/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 3
5PB114IS/Box_NY_NA50#_402_2.5-3.25 | 1
5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | 2
5QB114IS/Box_NY_NA50#_402_2.5-3.25 | 1
5QS273IS/Box_NY_NA50#_402_5-3.25 | 1
5QS324IS/Box_NY_NA50#_402_5-3.25 | 1
5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | 2
5QS428IS/Box_NY_NA50#_402_5-3.25 | 2
6AA34724/Shipper_NY_Skip lot testing | 2
6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | 2
6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | 7
6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 5
6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 1
6AA76930/SHIP_NY_44K_100_9-9-4.75 | 1
6AA88230/SHIP_NY_44K_100_9-9-4.75 | 1
6AA91130/SHIP_NY_32K_100_10-8.375-8.25 | 2
6AA92230/SHIP_NY_32K_100_10-8.375-8.25 | 1
6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 4
6AD11055/Shipper_NY_Skip lot testing | 3
6AD41900/Shipper_NY_Skip lot testing | 6
6AD43600/SHIP_NY_32K_100_13.25-12-6.125 | 5
6AH58195/Shipper_NY_Skip lot testing | 3
6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 1
6AJ19025/Shipper_NY_Skip lot testing | 6
6AM30472/SHIP_NY_44K_100_15.5-9-12 | 3
6AM30472/Shipper_NY_Skip lot testing | 10
6AM70172/Shipper_NY_Skip lot testing | 1
6AP03672/Shipper_NY_Skip lot testing | 5
6AP03707/Shipper_NY_Skip lot testing | 5
6AP06325/Shipper_NY_Skip lot testing | 2
6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | 2
6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 1
6AP07207/Simple Kind To Skin facial 7ct Shipper Canada | 1
6AP07328/Shipper_NY_Skip lot testing | 1
6AP08128/SHIP_NY_32K_100_13.25-7.75-5.75 | 6
6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 1
6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 9
6AP09384/Shipper_NY_Skip lot testing | 1
6AP11325/Shipper_NY_Skip lot testing | 3
6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | 5
6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | 2
6AP19525/Shipper_NY_Skip lot testing | 6
6AP22325/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | 1
6AP22325/Shipper_NY_Skip lot testing | 6
6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | 1
6AP24284/Shipper_NY_Skip lot testing | 4
6AP24684/SHIP_NY_40K_100_15.375-10.5625-9.625 | 1
6AP25784/Shipper_NY_Skip lot testing | 5
6AP27284/Shipper_NY_Skip lot testing | 3
6AP32825/SHIP_32_1P | 1
6AP56784/SHIP_NY_44K_100_15.375-10.5625-9.625 | 1
6AP56784/Shipper_NY_Skip lot testing | 10
6AP66784/Shipper_NY_Skip lot testing | 3
6AP71310/Shipper_NY_Skip lot testing | 3
6AQ20372/Shipper_NY_Skip lot testing | 2
6AQ76584/SHIP_NY_44K_100_15.25-10.375-7.6875 | 1
6AS27350/SHIP_44_2O | 8
6AS27350/Shipper_NY_Skip lot testing | 1
6AS40750/SHIP_32_1AP | 3
6AS41950/SHIP_40_1O | 2
6AU12310/SHIP_32_1U | 6
6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 1
6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 2
6AU57314/Shipper_NY_Skip lot testing | 1
6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 1
6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 3
6BB10800/SHIP_32_1S | 6
6BD17385/SHIP_44_2AB | 3
6BH59200/SHIP_NY_32K_100_13.4375-6.25-9.6875 | 1
6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 3
6BP28425/SHIP_32_1P | 6
6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | 2
6BP54072/SHIP_44_2L | 1
6BP54072/Shipper_NY_Skip lot testing | 7
6BU26595/Shipper_NY_Skip lot testing | 5
6BU27500/Shipper_NY_Skip lot testing | 2
6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | 1
6CJ22750/Shipper_NY_Skip lot testing | 1
6CP02724/SHIP_NY_32K_100_11-5.8125-4.0625 | 2
6CP13872/Shipper_NY_Skip lot testing | 1
6CP22884/SHIP_NY_44K_100_15.375-10.5625-9.625 | 2
6CP22884/Shipper_NY_Skip lot testing | 2
6CP33724/SHIP_NY_32K_100_11-5.8125-4.0625 | 1
6CP60125/SHIP_32_1P | 2
6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 3
6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | 1
6DD74800/Shipper_NY_Skip lot testing | 5
6DP13872/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 3
6DP13872/Shipper_NY_Skip lot testing | 3
6DP44584/Shipper_NY_Skip lot testing | 4
6DP63884/SHIP_NY_44K_100_15.375-10.5625-9.625 | 1
6DP63884/Shipper_NY_Skip lot testing | 14
6ED12555/Shipper_NY_Skip lot testing | 2
6EQ85484/Shipper_NY_Skip lot testing | 2
6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 1
6FB71200/Shipper_NY_Skip lot testing | 1
6FJ14108/Shipper_NY_Skip lot testing | 1
6FJ14143/Shipper_NY_Skip lot testing | 1
6FP92084/SHIP_NY_40K_100_15.375-10.5625-9.625 | 4
6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 3
6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | 1
6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | 4
6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 5
6HH04082/SHIP_NY_32K_100_13.4375-6.25-9.6875 | 1
6HH04082/Shipper_NY_Skip lot testing | 6
6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | 1
6JA580FW/Shipper_NY_Skip lot testing | 9
6JQ08472/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 3
6KS41125/SHIP_NY_29K_100_17.125-6.6875-11.75 | 3
6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | 2
6LQ50772/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 1
6LQ85084/Shipper_NY_Skip lot testing | 3
6LS41350/SHIP_NY_29K_100_17.125-6.6875-11.75 | 4
6LS82125/SHIP_NY_29K_100_17.125-6.6875-11.75 | 1
6MS48050/SHIP_NY_29K_100_17.125-6.6875-11.75 | 1
6MU87295/SHIP_NY_32K_100_13.625-4.5-12.25 | 1
6MU87295/Shipper_NY_Skip lot testing | 4
6NQ86984/SHIP_NY_44K_100_15.375-10.5625-9.625 | 5
6NQ86984/Shipper_NY_Skip lot testing | 41
6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | 2
6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 1
6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 4
6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 3
6OB31400/SHIP_32_1AL | 2
6OC22324/Shipper_NY_Skip lot testing | 2
6OD01524/Shipper_NY_Skip lot testing | 1
6OD13500A/SHIP_32_1AD | 1
6OD438SS/Shipper_NY_Skip lot testing | 9
6OH14324/SHIP_32_1Y | 2
6OH21112/SHIP_32_1F | 1
6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | 1
6OK75595/SHIP_32_1R | 1
6OM960CB/SHIP_NY_NA_200_2.5-2.5-33.75 | 1
6OP027LP/SHIP_32_13A | 9
6OP07530/Shipper_NY_Skip lot testing | 3
6OP08725/Shipper_NY_Skip lot testing | 4
6OP34524/SHIP_44_2G | 1
6OP705PD/Shipper_NY_Skip lot testing | 3
6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 1
6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | 3
6OQ22272/Shipper_NY_Skip lot testing | 3
6OQ85984/SHIP_NY_44K_100_15.375-10.5625-9.625 | 1
6OQ85984/Shipper_NY_Skip lot testing | 2
6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 1
6OS42850/SHIP_44_2N | 2
6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 1
6OS46250/SHIP_29_1J | 1
6OS46250/Shipper_NY_Skip lot testing | 1
6OU12066/Shipper_NY_Skip lot testing | 1
6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 1
6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 1
6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | 2
6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 2
6OU37230/Shipper_NY_Skip lot testing | 1
6OU38510/SHIP_32_1BG | 1
6OU38510/Shipper_NY_Skip lot testing | 2
6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 3
6OU62620/SHIP_NY_32K_100_15-9.5-8 | 3
6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 1
6PQ86984/SHIP_NY_32K_100_15.1875-10.1875-9.625 | 1
6PQ89072/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 3
6PQ89072/Shipper_NY_Skip lot testing | 4
6PS23125/SHIP_NY_29K_100_17.125-6.6875-11.75 | 2
6RQ55172/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 6
6RQ55172/Shipper_NY_Skip lot testing | 31
6ZS505PD/Shipper_NY_Skip lot testing | 3
7A4AL930/Chem_08A | 6
7A4AS162/Chem_09A | 4
7A4BL625/Chem_11A | 1
7A4CD419/Chem_13A | 2
7A4FD110/Chem_Sodium Chloride USP/NF | 4
7A4OA709/Chem_NY_Chamomile Extract # 11-19-2 | 2
7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | 1
7A4OB223/Chem_189A | 1
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 125
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | 1
7A4OB339/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 1
7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 138
7A4OB397/Chem_04A | 2
7A4OB690/Chem_01A | 5
7A4OF602/Chem_19A | 1
7A4OG169/Chem_20A | 1
7A4OH247/Chem_21A | 5
7A4OK459/Chem_NY_Salicylic Acid USP | 3
7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 1
7A4OP073/Chem_NY_Fragrance Bali Leia 5 | 1
7A4OP081/Chem_NY_Fragrance SC 2705 LF | 2
7A4OP093/Chem_NY_Essential Oil 27A Fragrance | 1
7A4OP197/Chem_NY_D-Panthenol 75W_01 | 12
7A4OP448/Chem_197A | 1
7A4OP448/Chem_208A | 2
7A4OP567/Chem_NY_Surfonic L12-6_Surfactant | 5
7A4OQ028/Chem_NY_Emulgin VL-75 | 3
7A4OQ084/Chem_23A | 14
7A4OQ132/Chem_25A | 5
7A4OQ229/Chem_NY_Witch Hazel | 1
7A4OQ229/Chem_Trisodium Citrate Dihydrate USP | 16
7A4OQ518/Chem_31A | 24
7A4OQ890/Chem_34A | 6
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 14
7A4OS411/Chem_36A | 4
7A4OU555/Chem_NY_Phenylephrine HCL | 2
7A4OU555/Chem_NY_Phenylephrine Reduced | 2
7A4OX110/Chem_39A | 3
7A4OX231/Chem_NY_D&C Yellow 10 | 2
7ACHE61P/Chem_NY_Sodium Benzoate | 3
7B4AA580/Chem_NY_Acetic Acid - USP | 1
7B4AL930/Chem_42A | 2
7B4OA242/Chem_48A | 9
7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | 2
7B4OA709/Chem_NY_Bell fragrance | 2
7B4OD232/Chem_52A | 11
7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | 1
7B4OP065/Chem_NY_Chamomile Flower Extract | 2
7B4OP093/Chem_NY_PEG-4 | 10
7B4OP113/Chem_NY_Mentha Piperita Extract_01 | 1
7B4OQ028/Chem_NY_Sorbitan Monolaurate | 3
7B4OQ084/Chem_60A | 1
7B4OQ209/Chem_61A | 3
7B4OQ214/Chem_62A | 4
7B4OQ595/Chem_63A | 2
7B4OQ794/Chem_NY_Disodium EDTA-USP | 2
7B4OX232/Chem_40A | 34
7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | 3
7C4FD110/Chem_69A | 2
7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 2
7C4OD232/Chem_NY_Methylparaben NF/USP | 7
7C4OG169/Chem_75A | 2
7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | 1
7C4OP601/MASTER - Microbiology - J&J Chemical Testing) | 1
7C4OQ028/Chem_NY_Isohexadecane | 6
7C4OQ595/Chem_81A | 3
7C4OX110/Chem_84A | 2
7D4AA580/Chem_NY_Monoethanolamine | 2
7D4OD232/Chem_NY_Propylparaben USP | 7
7D4OG169/Chem_88A | 1
7D4OP065/Chem_NY_ Glydant Plus Liquid | 2
7D4OP230/Chem_NY_Zemea_01 | 1
7D4OQ028/Chem_NY_Dicaprylyl Ether | 6
7D4OQ214/Chem_93A | 14
7D4OQ967/Chem_94A | 1
7E4OD232/Chem_97A | 5
7E4OP065/Chem_NY_Benzyl Alcohol | 6
7E4OP806/Chem_58B | 9
7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | 2
7F4BP869/Chem_NY_Citric Acid Anhydrous - USP | 13
7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | 3
7F4OD110/Chem_106A | 6
7F4OP601/MASTER - Microbiology - J&J Chemical Testing) | 4
7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | 2
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | 2
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | 7
7G4OP308/Chem_111A | 1
7L4OP027/Chem_NY_Menthol USP | 2
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 1
7O4AA580/Chem_NY_Ethylene Glycol Monohexyl Ether | 3
7O4AQ084/Chem_162A | 2
7O4BJ131/Chem_NY_Acticide CBM-2 | 4
7O4BM301/Chem_126A | 3
7O4BU125/Chem_09A | 19
7O4DH247/Chem_201A | 3
7O4DX220/Chem_NY_Neutralizing Agent | 3
7O4OA234/Chem_131A | 1
7O4OA347/Chem_NY_Lauryl Glucoside | 1
7O4OA401/Chem_132A | 1
7O4OA490/Chem_133A | 13
7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | 1
7O4OA709/Chem_NY_Cucumber Extract # 11-19-3 | 3
7O4OA911/Chem_NY_Zinc Sulfate Heptahydrate | 3
7O4OD351/Chem_NY_Benzalkonium Chloride 50% | 1
7O4OD363/Chem_137A | 4
7O4OD419/Chem_138A | 6
7O4OD748/Chem_139A | 2
7O4OF575/Che_NY_Plantaren 2000N - UP | 1
7O4OF602/Chem_141A | 3
7O4OG169/Chem_142A | 1
7O4OH247/Chem_143A | 2
7O4OJ123/Chem_NY_NET-KOB | 4
7O4OM165/Chem_146A | 6
7O4OM666/Chem_NY_Sodium Benzoate EDF NF/FCC | 8
7O4OP027/MASTER - Microbiology - J&J Chemical Testing) | 1
7O4OP065/Chem_NY_Actipone White Tea | 2
7O4OP075/Chem_NY_Magique Lift 02LF | 3
7O4OP093/Chem_NY_PEG-11 | 3
7O4OP197/Chem_NY_Emulgade Smart/MB_01 | 12
7O4OP209/MASTER - Microbiology - J&J Chemical Testing) | 2
7O4OP253/Chem_NY_BTC 1010 | 1
7O4OP345/Chem_157A | 2
7O4OP448/Chem_197A | 5
7O4OP448/Chem_208A | 1
7O4OP567/Chem_NY_BTC 1210-80% | 8
7O4OP638/Chem_192A | 4
7O4OP705/Chem_NY_Cetaphil Concentrate | 8
7O4OQ028/Chem_NY_Euxyl K703 | 1
7O4OQ028/Chem_NY_Hexylene Glycol | 4
7O4OQ084/Chem_125A | 2
7O4OQ229/Chem_NY_Witch Hazel | 24
7O4OQ229/Chem_Trisodium Citrate Dihydrate USP | 1
7O4OQ243/Chem_165A | 8
7O4OQ268/Chem_NY_Euxyl K703 | 7
7O4OQ268/Chem_NY_Hexylene Glycol | 1
7O4OQ342/Chem_166A | 1
7O4OQ437/Chem_168A | 5
7O4OQ890/Chem_170A | 10
7O4OS407/Chem_173A | 14
7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 11
7O4OU351/Chem_NY_Glycopyrronium Tosylate Monohydrate | 3
7O4OU555/Chem_NY_Lidocaine Base USP | 5
7O4OU555/Chem_NY_Lidocaine Reduced testing | 3
7O4OX231/Chem_NY_FD&C Red 40 | 2
8AA654SI/Label_NY_Skip lot testing | 7
8AA70900/Label_NY_NP9990_2.5-4.625 | 1
8AP13800/Label_NY_Skip lot testing | 3
8AP22800/Label_NY_NP7200_5.8125-6.5 | 1
8AP22800/Label_NY_Skip lot testing | 1
8AP24200/Label_NY_NP7200_5.8125-6.5 | 2
8AP24200/Label_NY_Skip lot testing | 1
8AP25300/Label_NY_NP7200_6.25-4.5 | 1
8AP25301/Label_NY_NP7200_6.25-4.5 | 2
8AP25301/Label_NY_Skip lot testing | 1
8AP32800/Label_NY_NP98372_73-85 | 1
8AP34500A/Label_NY_Skip lot testing | 1
8AP520CK/Label_NY_Skip lot testing | 1
8AP56701/Label_NY_Skip lot testing | 5
8AP63800/Label_NY_NP7200_5.8125-6.5 | 1
8AP63800/Label_NY_Skip lot testing | 8
8AP66701/Label_NY_NP7200_5.8125-6.5 | 1
8AP66701/Label_NY_Skip lot testing | 1
8AP713RSL/Label_NY_Skip lot testing | 2
8AU120CK/Label_NY_NP7030_6-4 | 1
8BA70900/Label_NY_NP9990_2.5-4.625 | 1
8BB196CK/Label_NY_NP7030_6-5 | 2
8BP13800/Label_NY_NP7200_6.25-4.5 | 3
8BP13800/Label_NY_Skip lot testing | 1
8BP13801/Label_NY_Skip lot testing | 3
8BP22801/Label_NY_NP7200_5.8125-6.5 | 1
8BP22801/Label_NY_Skip lot testing | 1
8BP25300/Label_NY_NP7200_6.25-4.5 | 2
8BP25300/Label_NY_Skip lot testing | 1
8BP54001/Label_NY_Skip lot testing | 6
8BQ76500/Label_NY_NP7200_6-3.75 | 1
8BQ76501/Label_NY_NP7200_3.75-6 | 1
8CH128CK/Label_NY_NP7030_9-2.625 | 2
8CH211CK/Label_A12 | 1
8CP13400/Label_NY_NP7100_4.5-4 | 4
8CP13400/Label_NY_Skip lot testing | 1
8CP13801/Label_NY_NP7200_6.25-4.5 | 3
8CP13801/Label_NY_Skip lot testing | 1
8CP15900/Label_NY_NP7100/NP7200_6-3.75 | 2
8CP44500/Label_NY_NP7200_5.8125-6.5 | 2
8CP63801/Label_A20 | 1
8CP63801/Label_NY_NP7200_5.8125-6.5 | 1
8CP63801/Label_NY_Skip lot testing | 6
8CP68301/Label_A20 | 1
8CP68301/Label_NY_NP7200_5.8125-6.5 | 1
8DP13401/Label_NY_NP7200_4.5-4 | 3
8DP13401/Label_NY_Skip lot testing | 1
8DP54000/Label_NY_Skip lot testing | 8
8DP68300/Label_NY_Skip lot testing | 1
8DP68301/Label_NY_Skip lot testing | 1
8DP683CK/Label_NY_Skip lot testing | 4
8DP92000/Label_NY_NP7200_6.5-5.8125 | 2
8DQ91401/Label_NY_NP7200_6.25-4.5 | 1
8DQ914SL/Label_NY_NP7030_9-7 | 1
8EP15901/Label_NY_NP7200_6.25-4 | 2
8EP44501/Label_NY_NP7200_5.8125-6.5 | 2
8EP68300/Label_NY_NP7200_6.25-4.5 | 3
8EP68301/Label_NY_NP7200_6.25-4.5 | 3
8EP92001/Label_NY_NP7200_6.5-5.8125 | 2
8EQ08400/Label_NY_Skip lot testing | 1
8EQ08401/Label_NY_Skip lot testing | 1
8EQ90100/Label_NY_NP7200_6.25-4.5 | 1
8EQ901CK/Label_NY_NP7030_9-7 | 1
8FQ08400/Label_NY_NP7200_6.25-4.5 | 3
8FQ08401/Label_NY_NP7200_6.25-4.5 | 2
8FQ85000/Label_NY_Skip lot testing | 4
8FQ859CK/Label_NY_NP7030_9-7 | 1
8FQ91400/Label_NY_NP7200_6.25-4.5 | 1
8GQ772SK/Label_B03 | 6
8GQ90101/Label_NY_NP7200_6.25-4.5 | 2
8HQ50701/Label_NY_NP7200_6.25-4.5 | 1
8HQ85001/Label_NY_Skip lot testing | 3
8HQ86900/Label_NY_NP7200_6.5-5.8125 | 1
8HQ86900/Label_NY_Skip lot testing | 11
8HQ89000/Label_NY_Skip lot testing | 2
8IQ50700/Label_NY_NP7200_6.25-4.5 | 1
8JQ86901/Label_NY_NP7200_6.5-5.8125 | 1
8JQ86901/Label_NY_Skip lot testing | 10
8JQ89000/Label_NY_NP7200_6.25-4.5 | 3
8JQ89000/Label_NY_Skip lot testing | 1
8JQ89001/Label_NY_Skip lot testing | 3
8KQ89001/Label_NY_NP7200_6.25-4.5 | 3
8MB10800/Label_NY_NP7030_6-5 | 2
8MQ55100/Label_NY_NP7200_6.25-4.5 | 3
8MQ55100/Label_NY_Skip lot testing | 7
8MQ55101/Label_NY_NP7200_6.25-4.5 | 3
8MQ55101/Label_NY_Skip lot testing | 7
8NB10800/Label_NY_NP7030_6-5 | 1
8NS40750/Label_NY_NP7030_6-5 | 2
8NS41950/Label_NY_NP7030_6-5 | 1
8OA407SI/Label_A10 | 10
8OA664CK/Label_NY_Skip lot testing | 1
8OB123CK/Label_NY_NP7030_6-5 | 2
8OD024CK/Label_NY_NP7030_9-2.625 | 1
8OD243CK/Label_NY_NP7030_9-2.625 | 2
8OH450CK/Label_NY_NP7030_9-2.625 | 1
8OH450CK/Label_NY_Skip lot testing | 1
8OJ22750/Label_NY_Skip lot testing | 1
8OP03600/Label_NY_Skip lot testing | 4
8OP03601/Label_NY_Skip lot testing | 4
8OP03700/Label_NY_NP9840_87-75 | 2
8OP06300/Label_NY_NP9840_87-75 | 2
8OP065RSL/Label_NY_NP9995_2.625-3.375 | 1
8OP067RSL/Label_NY_NP9995_2.625-3.375 | 1
8OP07200/Label_NY_NP9840_87-75 | 1
8OP073RSL/Label_NY_NP9995_2.625-3.375 | 1
8OP081RSL/Label_NY_NP9995_2.625-3.375 | 3
8OP08700/Label_NY_Skip lot testing | 2
8OP09100/Label_NY_NP9991_3.375-2.625 | 1
8OP09300/Label_NY_NP7110_2.55-3 | 3
8OP09301/Label_NY_NP7110_2.55-5 | 3
8OP19500/Label_NY_NP9840_87-75 | 1
8OP214RSL/Label_NY_NP9991_3.375-2.625 | 1
8OP223RSL/Label_NY_NP9840_87-75 | 4
8OP227RSL/Label_NY_NP9997_2.625-3.375 | 1
8OP24200/Label_NY_Skip lot testing | 1
8OP24201/Label_NY_Skip lot testing | 3
8OP26300/Label_NY_NP7200_6.5-5.8125 | 1
8OP26301/Label_NY_NP7200_6.5-5.8125 | 1
8OP263CK/Label_NY_NP7030_8.5-6.5 | 1
8OP27200/Label_NY_NP7200_5.8125-6.5 | 1
8OP27200/Label_NY_Skip lot testing | 4
8OP27201/Label_NY_NP7200_5.8125-6.5 | 1
8OP27201/Label_NY_Skip lot testing | 4
8OP29400/Label_NY_NP7200_6.5-5.8125 | 1
8OP29401/Label_NY_NP7200_6.5-5.8125 | 1
8OP294CK/Label_NY_NP7030_8.5-6.5 | 2
8OP47101/Label_NY_NP7200_6.5-5.8125 | 1
8OP56700/Label_NY_NP7200_5.8125-6.5 | 1
8OP56700/Label_NY_Skip lot testing | 5
8OP66700/Label_NY_Skip lot testing | 2
8OU372CK/Label_NY_Skip lot testing | 1
8OU500CK/Label_NY_Skip lot testing | 2
8OX351CK/Label_NY_50#W_5_4 | 1
8OX461CK/Label_NY_50#W_5_4 | 1
8OX465CK/Label_NY_50#W_5_4 | 1
8OX510CK/Label_NY_50#W_5_4 | 1
8QB11400/Label_NY_NP7030_6-5 | 1
8RB11400/Label_NY_NP7030_6-5 | 1
8RS32450/Label_NY_NP7030_6-5 | 2
8RS428CK/Label_NY_NP7030_6-5 | 2
8SS27350/Label_NY_NP7030_6-5 | 3

## Planscount
plan.m | Mat_MatCode | SP_TestPlan | Test_Lab | Res_TestCode | TestName | Analyte | Analyte.C | Numeric_Result | Res_Units | Rounding_Rule | Low_Release_Limit | High_Release_Limit | METHOD | Plan.C | Count | _merge
2AA284SLI/Lid_01B | 2AA284SLI | Lid_01B | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection | 2AA284SLI/Lid_01B | 1 | both
2AA284SLI/Lid_01B | 2AA284SLI | Lid_01B | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | 2AA284SLI/Lid_01B | 1 | both
2AA284SLI/Lid_01B | 2AA284SLI | Lid_01B | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | 2AA284SLI/Lid_01B | 1 | both
2AA284SLI/Lid_01B | 2AA284SLI | Lid_01B | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual | 2AA284SLI/Lid_01B | 1 | both
2AA284SLI/Lid_01B | 2AA284SLI | Lid_01B | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.6875 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | 2AA284SLI/Lid_01B | 1 | both
2AA284SLI/Lid_01B | 2AA284SLI | Lid_01B | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.875 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | 2AA284SLI/Lid_01B | 1 | both
2AA284SLI/Lid_01B | 2AA284SLI | Lid_01B | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual | 2AA284SLI/Lid_01B | 1 | both
2AA347/Web_NP8550_10.25 | 2AA347 | Web_NP8550_10.25 | NY-Analytical | 885 | Web Inspection | AQL Sample Size | AQL Sample Size | 315 |  | FDA,0 | 0 |  | TM101/TM107/TM132/Measure | 2AA347/Web_NP8550_10.25 | 1 | both
2AA347/Web_NP8550_10.25 | 2AA347 | Web_NP8550_10.25 | NY-Analytical | 885 | Web Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Sample(s) all meet specifications (Done) |  |  |  |  | TM101/TM107/TM132/Measure | 2AA347/Web_NP8550_10.25 | 1 | both
2AA347/Web_NP8550_10.25 | 2AA347 | Web_NP8550_10.25 | NY-Analytical | 885 | Web Inspection | Basis_Weight | Basis Weight | 41.3 | #/Ream | FDA,1 | 37.8 | 48.2 | TM101/TM107/TM132/Measure | 2AA347/Web_NP8550_10.25 | 1 | both
2AA347/Web_NP8550_10.25 | 2AA347 | Web_NP8550_10.25 | NY-Analytical | 885 | Web Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | TM101/TM107/TM132/Measure | 2AA347/Web_NP8550_10.25 | 1 | both
2AA347/Web_NP8550_10.25 | 2AA347 | Web_NP8550_10.25 | NY-Analytical | 885 | Web Inspection | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 5762.95 | g/in | FDA,0 | 3500 |  | TM101/TM107/TM132/Measure | 2AA347/Web_NP8550_10.25 | 1 | both
2AA347/Web_NP8550_10.25 | 2AA347 | Web_NP8550_10.25 | NY-Analytical | 885 | Web Inspection | Web Width (Inches) | Web Width | 10.25 | in | FDA,4 | 10.1875 | 10.3125 | TM101/TM107/TM132/Measure | 2AA347/Web_NP8550_10.25 | 1 | both
2AA592/Web_NP8504_11.0 | 2AA592 | Web_NP8504_11.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Dyne Level | 2AA592/Web_NP8504_11.0 | 3 | both
2AA592/Web_NP8504_11.0 | 2AA592 | Web_NP8504_11.0 | NY-Analytical | 836 | Dyne Level | Dyne Level (outside film) | Dyne Level (outside film) | 36 |  | FDA,0 | 36 | 38 | Dyne Test Kit | 2AA592/Web_NP8504_11.0 | 3 | both
2AA592/Web_NP8504_11.0 | 2AA592 | Web_NP8504_11.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 10.968 | in | FDA,4 | 10.9375 | 11.0625 | Measurement | 2AA592/Web_NP8504_11.0 | 3 | both
2AA592/Web_NP8504_11.0 | 2AA592 | Web_NP8504_11.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 57.54 | #/Ream | FDA,1 | 53.3 | 65.1 | TM 101 | 2AA592/Web_NP8504_11.0 | 3 | both
2AA592/Web_NP8504_11.0 | 2AA592 | Web_NP8504_11.0 | NY-Analytical | 762 | Seal Strength | Seal Strength(g/in@350F/30psi/0.5s dwell Top Jaw) | Seal Strength (@350°F/30psi/0.5s dwell Top Jaw) | 5339.63 | g/in | FDA,0 | 2700 |  | TM 107 | 2AA592/Web_NP8504_11.0 | 3 | both
2AA592/Web_NP8504_11.0 | 2AA592 | Web_NP8504_11.0 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 3.65 | mil | FDA,1 | 3.2 | 3.9 | TM 132 | 2AA592/Web_NP8504_11.0 | 3 | both
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Lid_NY_NP7600_22R1 | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection | 2AA592SLI/Lid_NY_NP7600_22R1 | 8 | both
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Lid_NY_NP7600_22R1 | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Solo Lid Inspection | 2AA592SLI/Lid_NY_NP7600_22R1 | 8 | both
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Lid_NY_NP7600_22R1 | NY-Analytical | 818 | Solo Lid Inspection | Color | Color | Meets the specification (Done) |  |  |  |  | Solo Lid Inspection | 2AA592SLI/Lid_NY_NP7600_22R1 | 8 | both
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Lid_NY_NP7600_22R1 | NY-Analytical | 818 | Solo Lid Inspection | Functionality | Functionality | Meet specification (Done) |  |  |  |  | Solo Lid Inspection | 2AA592SLI/Lid_NY_NP7600_22R1 | 8 | both
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Lid_NY_NP7600_22R1 | NY-Analytical | 818 | Solo Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.028 | in | FDA,3 | 0.025 | 0.035 | Solo Lid Inspection | 2AA592SLI/Lid_NY_NP7600_22R1 | 8 | both
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Lid_NY_NP7600_22R1 | NY-Analytical | 818 | Solo Lid Inspection | Open Lid Overall Width | Open Lid Overall Width | 4.36 | in | FDA,3 | 4.351 | 4.401 | Solo Lid Inspection | 2AA592SLI/Lid_NY_NP7600_22R1 | 8 | both
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Lid_NY_NP7600_22R1 | NY-Analytical | 818 | Solo Lid Inspection | Overall closed lid height | Overall closed lid height | 0.37 | in | FDA,3 | 0.355 | 0.385 | Solo Lid Inspection | 2AA592SLI/Lid_NY_NP7600_22R1 | 8 | both
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Lid_NY_NP7600_22R1 | NY-Analytical | 818 | Solo Lid Inspection | Weight | Weight | 11.241 | g | FDA,1 | 10.5 | 12.5 | Measurement | 2AA592SLI/Lid_NY_NP7600_22R1 | 8 | both
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Lid_NY_NP7600_22R1 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | 2AA592SLI/Lid_NY_NP7600_22R1 | 8 | both
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Lid_NY_NP7600_22R1 | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual | 2AA592SLI/Lid_NY_NP7600_22R1 | 8 | both
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Lid_NY_NP7600_22R1 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.7025 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | 2AA592SLI/Lid_NY_NP7600_22R1 | 8 | both
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Lid_NY_NP7600_22R1 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.8705 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | 2AA592SLI/Lid_NY_NP7600_22R1 | 8 | both
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Lid_NY_NP7600_22R1 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual | 2AA592SLI/Lid_NY_NP7600_22R1 | 8 | both
2AA632/Web_NP8504_11.75 | 2AA632 | Web_NP8504_11.75 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Dyne Level | 2AA632/Web_NP8504_11.75 | 5 | both
2AA632/Web_NP8504_11.75 | 2AA632 | Web_NP8504_11.75 | NY-Analytical | 836 | Dyne Level | Dyne Level (outside film) | Dyne Level (outside film) | 38 |  | FDA,0 | 36 | 38 | Dyne Test Kit | 2AA632/Web_NP8504_11.75 | 5 | both
2AA632/Web_NP8504_11.75 | 2AA632 | Web_NP8504_11.75 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 11.75 | in | FDA,4 | 11.6875 | 11.8125 | Measurement | 2AA632/Web_NP8504_11.75 | 5 | both
2AA632/Web_NP8504_11.75 | 2AA632 | Web_NP8504_11.75 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 59.05 | #/Ream | FDA,1 | 53.3 | 65.1 | TM 101 | 2AA632/Web_NP8504_11.75 | 5 | both
2AA632/Web_NP8504_11.75 | 2AA632 | Web_NP8504_11.75 | NY-Analytical | 762 | Seal Strength | Seal Strength(g/in@350F/30psi/0.5s dwell Top Jaw) | Seal Strength (@350°F/30psi/0.5s dwell Top Jaw) | 5664 | g/in | FDA,0 | 2700 |  | TM 107 | 2AA632/Web_NP8504_11.75 | 5 | both
2AA632/Web_NP8504_11.75 | 2AA632 | Web_NP8504_11.75 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 3.7 | mil | FDA,1 | 3.2 | 3.9 | TM 132 | 2AA632/Web_NP8504_11.75 | 5 | both
2AA714SLI/Lid_01V | 2AA714SLI | Lid_01V | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width |  | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual |  |  | left_only
2AA714SLI/Lid_01V | 2AA714SLI | Lid_01V | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length |  | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual |  |  | left_only
2AA714SLI/Lid_01V | 2AA714SLI | Lid_01V | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2AA714SLI/Lid_01V | 2AA714SLI | Lid_01V | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2AA714SLI/Lid_01V | 2AA714SLI | Lid_01V | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection |  |  | left_only
2AA714SLI/Lid_01V | 2AA714SLI | Lid_01V | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2AA714SLI/Lid_01V | 2AA714SLI | Lid_01V | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection |  |  | left_only
2AA922SLI/Lid_NY_Skip lot testing | 2AA922SLI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection | 2AA922SLI/Lid_NY_Skip lot testing | 1 | both
2AA922SLI/Lid_01V | 2AA922SLI | Lid_01V | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis | 2AA922SLI/Lid_01V | 2 | both
2AA922SLI/Lid_NY_Skip lot testing | 2AA922SLI | Lid_NY_Skip lot testing | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Compliance | 2AA922SLI/Lid_NY_Skip lot testing | 1 | both
2AA922SLI/Lid_NY_Skip lot testing | 2AA922SLI | Lid_NY_Skip lot testing | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | 2AA922SLI/Lid_NY_Skip lot testing | 1 | both
2AA922SLI/Lid_NY_Skip lot testing | 2AA922SLI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | 2AA922SLI/Lid_NY_Skip lot testing | 1 | both
2AA922SLI/Lid_NY_Skip lot testing | 2AA922SLI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual | 2AA922SLI/Lid_NY_Skip lot testing | 1 | both
2AA922SLI/Lid_NY_Skip lot testing | 2AA922SLI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.6875 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | 2AA922SLI/Lid_NY_Skip lot testing | 1 | both
2AA922SLI/Lid_NY_Skip lot testing | 2AA922SLI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.8438 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | 2AA922SLI/Lid_NY_Skip lot testing | 1 | both
2AA922SLI/Lid_NY_Skip lot testing | 2AA922SLI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual | 2AA922SLI/Lid_NY_Skip lot testing | 1 | both
2AD173/Web_NP5115_4.5 | 2AD173 | Web_NP5115_4.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2AD173/Web_NP5115_4.5 | 4 | both
2AD173/Web_NP5115_4.5 | 2AD173 | Web_NP5115_4.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 4.5 | in | FDA,4 | 4.4375 | 4.5625 | Measurement | 2AD173/Web_NP5115_4.5 | 4 | both
2AD173/Web_NP5115_4.5 | 2AD173 | Web_NP5115_4.5 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Printed/Varnished) | Basis Weight | 59.5 | #/Ream | FDA,1 | 55.6 | 68 | TM 101 | 2AD173/Web_NP5115_4.5 | 4 | both
2AD173/Web_NP5115_4.5 | 2AD173 | Web_NP5115_4.5 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@325DegF/40PSI/0.5Sec_Dwell) | Seal Strength (@325°F/40PSI/0.5Sec Dwell) | 2057.76 | g/in | FDA,0 | 1500 |  | TM 107 | 2AD173/Web_NP5115_4.5 | 4 | both
2AJ123/Web_NY_Skip lot testing | 2AJ123 | Web_NY_Skip lot testing | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2AJ123/Web_NY_Skip lot testing | 1 | both
2AJ123/Web_NY_Skip lot testing | 2AJ123 | Web_NY_Skip lot testing | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 11 | in | FDA,4 | 10.9375 | 11.0625 | Measurement | 2AJ123/Web_NY_Skip lot testing | 1 | both
2AJ123/Web_NY_Skip lot testing | 2AJ123 | Web_NY_Skip lot testing | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 42.33 | #/Ream | FDA,1 | 37.8 | 48.2 | TM 101 | 2AJ123/Web_NY_Skip lot testing | 1 | both
2AJ123/Web_NY_Skip lot testing | 2AJ123 | Web_NY_Skip lot testing | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 7486 | g/in | FDA,0 | 3500 |  | TM 107 | 2AJ123/Web_NY_Skip lot testing | 1 | both
2AJ143SLI/Lid_NY_NP7600_01C | 2AJ143SLI | Lid_NY_NP7600_01C | IN-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width |  | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual |  |  | left_only
2AJ143SLI/Lid_NY_NP7600_01C | 2AJ143SLI | Lid_NY_NP7600_01C | IN-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length |  | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual |  |  | left_only
2AJ143SLI/Lid_NY_NP7600_01C | 2AJ143SLI | Lid_NY_NP7600_01C | IN-Analytical | 765 | Canister & Lid Inspection | AQL Sample Size | AQL Sample Size |  |  |  | 0 |  | TM 143 / Visual |  |  | left_only
2AJ143SLI/Lid_NY_NP7600_01C | 2AJ143SLI | Lid_NY_NP7600_01C | IN-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2AJ143SLI/Lid_NY_NP7600_01C | 2AJ143SLI | Lid_NY_NP7600_01C | IN-Analytical | 765 | Canister & Lid Inspection | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2AJ143SLI/Lid_NY_NP7600_01C | 2AJ143SLI | Lid_NY_NP7600_01C | IN-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2AJ143SLI/Lid_NY_NP7600_01C | 2AJ143SLI | Lid_NY_NP7600_01C | IN-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection |  |  | left_only
2AJ143SLI/Lid_NY_NP7600_01C | 2AJ143SLI | Lid_NY_NP7600_01C | IN-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2AJ170/Web_NP8530_11.0 | 2AJ170 | Web_NP8530_11.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2AJ170/Web_NP8530_11.0 | 1 | both
2AJ170/Web_NP8530_11.0 | 2AJ170 | Web_NP8530_11.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 11 | in | FDA,4 | 10.9375 | 11.0625 | Measurement | 2AJ170/Web_NP8530_11.0 | 1 | both
2AJ170/Web_NP8530_11.0 | 2AJ170 | Web_NP8530_11.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 34.39 | #/Ream | FDA,2 | 29.4 | 37.42 | TM 101 | 2AJ170/Web_NP8530_11.0 | 1 | both
2AJ170/Web_NP8530_11.0 | 2AJ170 | Web_NP8530_11.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 5442 | g/in | FDA,0 | 3500 |  | TM 107 | 2AJ170/Web_NP8530_11.0 | 1 | both
2AJ170/Web_NP8530_11.0 | 2AJ170 | Web_NP8530_11.0 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 1.9 | mil | FDA,1 | 1.8 | 2.2 | TM 132 | 2AJ170/Web_NP8530_11.0 | 1 | both
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Lid_NY_NP7600_22FA | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection | 2AJ170SLI/Lid_NY_NP7600_22FA | 1 | both
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Lid_NY_NP7600_22FA | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Solo Lid Inspection | 2AJ170SLI/Lid_NY_NP7600_22FA | 1 | both
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Lid_NY_NP7600_22FA | NY-Analytical | 818 | Solo Lid Inspection | Color | Color | Meets the specification (Done) |  |  |  |  | Solo Lid Inspection | 2AJ170SLI/Lid_NY_NP7600_22FA | 1 | both
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Lid_NY_NP7600_22FA | NY-Analytical | 818 | Solo Lid Inspection | Functionality | Functionality | Meet specification (Done) |  |  |  |  | Solo Lid Inspection | 2AJ170SLI/Lid_NY_NP7600_22FA | 1 | both
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Lid_NY_NP7600_22FA | NY-Analytical | 818 | Solo Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.028 | in | FDA,3 | 0.025 | 0.035 | Solo Lid Inspection | 2AJ170SLI/Lid_NY_NP7600_22FA | 1 | both
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Lid_NY_NP7600_22FA | NY-Analytical | 818 | Solo Lid Inspection | Open Lid Overall Width | Open Lid Overall Width | 4.36 | in | FDA,3 | 4.351 | 4.401 | Solo Lid Inspection | 2AJ170SLI/Lid_NY_NP7600_22FA | 1 | both
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Lid_NY_NP7600_22FA | NY-Analytical | 818 | Solo Lid Inspection | Overall closed lid height | Overall closed lid height | 0.37 | in | FDA,3 | 0.355 | 0.385 | Solo Lid Inspection | 2AJ170SLI/Lid_NY_NP7600_22FA | 1 | both
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Lid_NY_NP7600_22FA | NY-Analytical | 818 | Solo Lid Inspection | Weight | Weight | 11.25 | g | FDA,1 | 10.5 | 12.5 | Measurement | 2AJ170SLI/Lid_NY_NP7600_22FA | 1 | both
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Lid_NY_NP7600_22FA | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | 2AJ170SLI/Lid_NY_NP7600_22FA | 1 | both
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Lid_NY_NP7600_22FA | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual | 2AJ170SLI/Lid_NY_NP7600_22FA | 1 | both
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Lid_NY_NP7600_22FA | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.67 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | 2AJ170SLI/Lid_NY_NP7600_22FA | 1 | both
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Lid_NY_NP7600_22FA | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.84 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | 2AJ170SLI/Lid_NY_NP7600_22FA | 1 | both
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Lid_NY_NP7600_22FA | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual | 2AJ170SLI/Lid_NY_NP7600_22FA | 1 | both
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Lid_NY_NP7600_22V | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection | 2AJ190SLI/Lid_NY_NP7600_22V | 1 | both
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Lid_NY_NP7600_22V | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Solo Lid Inspection | 2AJ190SLI/Lid_NY_NP7600_22V | 1 | both
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Lid_NY_NP7600_22V | NY-Analytical | 818 | Solo Lid Inspection | Appearance | Appearance | Meet specification (Done) |  |  |  |  | Solo Lid Inspection | 2AJ190SLI/Lid_NY_NP7600_22V | 1 | both
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Lid_NY_NP7600_22V | NY-Analytical | 818 | Solo Lid Inspection | Color | Color | Meets the specification (Done) |  |  |  |  | Solo Lid Inspection | 2AJ190SLI/Lid_NY_NP7600_22V | 1 | both
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Lid_NY_NP7600_22V | NY-Analytical | 818 | Solo Lid Inspection | Functionality | Functionality | Meet specification (Done) |  |  |  |  | Solo Lid Inspection | 2AJ190SLI/Lid_NY_NP7600_22V | 1 | both
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Lid_NY_NP7600_22V | NY-Analytical | 818 | Solo Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.0295 | in | FDA,3 | 0.025 | 0.035 | Solo Lid Inspection | 2AJ190SLI/Lid_NY_NP7600_22V | 1 | both
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Lid_NY_NP7600_22V | NY-Analytical | 818 | Solo Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | Solo Lid Inspection | 2AJ190SLI/Lid_NY_NP7600_22V | 1 | both
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Lid_NY_NP7600_22V | NY-Analytical | 818 | Solo Lid Inspection | Open Lid Overall Width | Open Lid Overall Width | 4.374 | in | FDA,3 | 4.351 | 4.401 | Solo Lid Inspection | 2AJ190SLI/Lid_NY_NP7600_22V | 1 | both
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Lid_NY_NP7600_22V | NY-Analytical | 818 | Solo Lid Inspection | Overall closed lid height | Overall closed lid height | 0.373 | in | FDA,3 | 0.355 | 0.385 | Solo Lid Inspection | 2AJ190SLI/Lid_NY_NP7600_22V | 1 | both
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Lid_NY_NP7600_22V | NY-Analytical | 818 | Solo Lid Inspection | Weight | Weight | 11.19 | g | FDA,1 | 10.5 | 12.5 | Measurement | 2AJ190SLI/Lid_NY_NP7600_22V | 1 | both
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Lid_NY_NP7600_22V | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual | 2AJ190SLI/Lid_NY_NP7600_22V | 1 | both
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Lid_NY_NP7600_22V | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.7188 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | 2AJ190SLI/Lid_NY_NP7600_22V | 1 | both
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Lid_NY_NP7600_22V | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.875 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | 2AJ190SLI/Lid_NY_NP7600_22V | 1 | both
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Lid_NY_NP7600_22X | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection | 2AJ324SLI/Lid_NY_NP7600_22X | 1 | both
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Lid_NY_NP7600_22X | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Solo Lid Inspection | 2AJ324SLI/Lid_NY_NP7600_22X | 1 | both
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Lid_NY_NP7600_22X | NY-Analytical | 818 | Solo Lid Inspection | Color | Color | Meets the specification (Done) |  |  |  |  | Solo Lid Inspection | 2AJ324SLI/Lid_NY_NP7600_22X | 1 | both
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Lid_NY_NP7600_22X | NY-Analytical | 818 | Solo Lid Inspection | Functionality | Functionality | Meet specification (Done) |  |  |  |  | Solo Lid Inspection | 2AJ324SLI/Lid_NY_NP7600_22X | 1 | both
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Lid_NY_NP7600_22X | NY-Analytical | 818 | Solo Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.0297 | in | FDA,3 | 0.025 | 0.035 | Solo Lid Inspection | 2AJ324SLI/Lid_NY_NP7600_22X | 1 | both
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Lid_NY_NP7600_22X | NY-Analytical | 818 | Solo Lid Inspection | Open Lid Overall Width | Open Lid Overall Width | 4.371 | in | FDA,3 | 4.351 | 4.401 | Solo Lid Inspection | 2AJ324SLI/Lid_NY_NP7600_22X | 1 | both
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Lid_NY_NP7600_22X | NY-Analytical | 818 | Solo Lid Inspection | Overall closed lid height | Overall closed lid height | 0.3705 | in | FDA,3 | 0.355 | 0.385 | Solo Lid Inspection | 2AJ324SLI/Lid_NY_NP7600_22X | 1 | both
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Lid_NY_NP7600_22X | NY-Analytical | 818 | Solo Lid Inspection | Weight | Weight | 11.385 | g | FDA,1 | 10.5 | 12.5 | Measurement | 2AJ324SLI/Lid_NY_NP7600_22X | 1 | both
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Lid_NY_NP7600_22X | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | 2AJ324SLI/Lid_NY_NP7600_22X | 1 | both
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Lid_NY_NP7600_22X | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual | 2AJ324SLI/Lid_NY_NP7600_22X | 1 | both
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Lid_NY_NP7600_22X | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.7088 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | 2AJ324SLI/Lid_NY_NP7600_22X | 1 | both
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Lid_NY_NP7600_22X | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.875 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | 2AJ324SLI/Lid_NY_NP7600_22X | 1 | both
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Lid_NY_NP7600_22X | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual | 2AJ324SLI/Lid_NY_NP7600_22X | 1 | both
2AP027/Web_NP8512_310 | 2AP027 | Web_NP8512_310 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  |  | left_only
2AP027/Web_NP8512_310 | 2AP027 | Web_NP8512_310 | NY-Analytical | 748 | Visual Inspection | Color/Copy | Color/Copy | Meets specifications (Done) |  |  |  |  | TM 208 |  |  | left_only
2AP027/Web_NP8512_310 | 2AP027 | Web_NP8512_310 | NY-Analytical | 761 | Web Measurement | Web Width (mm) | Web Width |  | mm | FDA,0 | 308 | 312 | Measurement |  |  | left_only
2AP027/Web_NP8512_310 | 2AP027 | Web_NP8512_310 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/inch@130DegC/14psi/1.0 sec dwell) | Seal Strength (@130°C/14psi/1.0 sec dwell) |  | g/in | FDA,0 | 2000 |  | TM 107 |  |  | left_only
2AP027/Web_NP8512_310 | 2AP027 | Web_NP8512_310 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness |  | mil | FDA,2 | 3.06 | 3.74 | TM 132 |  |  | left_only
2AP075/Web_NP8542_12.205 | 2AP075 | Web_NP8542_12.205 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2AP075/Web_NP8542_12.205 | 3 | both
2AP075/Web_NP8542_12.205 | 2AP075 | Web_NP8542_12.205 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 12.25 | in | FDA,4 | 12.1425 | 12.2675 | Measurement | 2AP075/Web_NP8542_12.205 | 3 | both
2AP075/Web_NP8542_12.205 | 2AP075 | Web_NP8542_12.205 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 60.85 | #/Ream | FDA,2 | 52.53 | 64.21 | TM 101 | 2AP075/Web_NP8542_12.205 | 3 | both
2AP075/Web_NP8542_12.205 | 2AP075 | Web_NP8542_12.205 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/mm@130DegC/1bar/0.5sec dwell) | Seal Strength (@130°C/1bar/0.5sec dwell) | 7718 | g/mm | FDA,0 | 25 |  | TM 107 | 2AP075/Web_NP8542_12.205 | 3 | both
2AP087/Web_NP8550_12.205 | 2AP087 | Web_NP8550_12.205 | NY-Analytical | 885 | Web Inspection | Web Width (Inches) | Web Width |  | in | FDA,4 | 12.1425 | 12.2675 | TM101/TM107/TM132/Measure |  |  | left_only
2AP087/Web_NP8550_12.205 | 2AP087 | Web_NP8550_12.205 | NY-Analytical | 885 | Web Inspection | Basis_Weight | Basis Weight |  | #/Ream | FDA,1 | 37.8 | 48.2 | TM101/TM107/TM132/Measure |  |  | left_only
2AP087/Web_NP8550_12.205 | 2AP087 | Web_NP8550_12.205 | NY-Analytical | 885 | Web Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 | 0 |  | TM101/TM107/TM132/Measure |  |  | left_only
2AP087/Web_NP8550_12.205 | 2AP087 | Web_NP8550_12.205 | NY-Analytical | 885 | Web Inspection | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) |  | g/in | FDA,0 | 3500 |  | TM101/TM107/TM132/Measure |  |  | left_only
2AP087/Web_NP8550_12.205 | 2AP087 | Web_NP8550_12.205 | NY-Analytical | 885 | Web Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | TM101/TM107/TM132/Measure |  |  | left_only
2AP087/Web_NP8550_12.205 | 2AP087 | Web_NP8550_12.205 | NY-Analytical | 885 | Web Inspection | Wind Direction | Wind Direction | Meets specification (Indicate numeric value - Done) |  |  |  |  | TM101/TM107/TM132/Measure |  |  | left_only
2AP087/Web_NP8550_12.205 | 2AP087 | Web_NP8550_12.205 | NY-Analytical | 885 | Web Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Sample(s) all meet specifications (Done) |  |  |  |  | TM101/TM107/TM132/Measure |  |  | left_only
2AP087/Web_NP8550_12.205 | 2AP087 | Web_NP8550_12.205 | NY-Analytical | 885 | Web Inspection | UPC Verification | UPC Verification | UPC Scan is acceptable (Scan into comment section below - Done) |  |  |  |  | TM101/TM107/TM132/Measure |  |  | left_only
2AP113/Web_NP8542_11.0 | 2AP113 | Web_NP8542_11.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  |  | left_only
2AP113/Web_NP8542_11.0 | 2AP113 | Web_NP8542_11.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight |  | #/Ream | FDA,2 | 52.53 | 64.21 | TM 101 |  |  | left_only
2AP113/Web_NP8542_11.0 | 2AP113 | Web_NP8542_11.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width |  | in | FDA,4 | 10.9375 | 11.0625 | Measurement |  |  | left_only
2AP113/Web_NP8542_11.0 | 2AP113 | Web_NP8542_11.0 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/mm@130DegC/1bar/0.5sec dwell) | Seal Strength (@130°C/1bar/0.5sec dwell) |  | g/mm | FDA,0 | 25 |  | TM 107 |  |  | left_only
2AP211/Web_NP8512_12.205 | 2AP211 | Web_NP8512_12.205 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2AP211/Web_NP8512_12.205 | 1 | both
2AP211/Web_NP8512_12.205 | 2AP211 | Web_NP8512_12.205 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 12.25 | in | FDA,4 | 12.1425 | 12.2675 | Measurement | 2AP211/Web_NP8512_12.205 | 1 | both
2AP211/Web_NP8512_12.205 | 2AP211 | Web_NP8512_12.205 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/inch@130DegC/14psi/1.0 sec dwell) | Seal Strength (@130°C/14psi/1.0 sec dwell) | 7698.56 | g/in | FDA,0 | 2000 |  | TM 107 | 2AP211/Web_NP8512_12.205 | 1 | both
2AP211/Web_NP8512_12.205 | 2AP211 | Web_NP8512_12.205 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 3.5 | mil | FDA,2 | 3.06 | 3.74 | TM 132 | 2AP211/Web_NP8512_12.205 | 1 | both
2AP211/Web_NP8512_12.205 | 2AP211 | Web_NP8512_12.205 | NY-Analytical | 748 | Visual Inspection | Color/Copy | Color/Copy | Meets specifications (Done) |  |  |  |  | TM 208 | 2AP211/Web_NP8512_12.205 | 1 | both
2AP229/Web_NP8512_9.5 | 2AP229 | Web_NP8512_9.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  |  | left_only
2AP229/Web_NP8512_9.5 | 2AP229 | Web_NP8512_9.5 | NY-Analytical | 748 | Visual Inspection | Color/Copy | Color/Copy | Meets specifications (Done) |  |  |  |  | TM 208 |  |  | left_only
2AP229/Web_NP8512_9.5 | 2AP229 | Web_NP8512_9.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width |  | in | FDA,4 | 9.4375 | 9.5625 | Measurement |  |  | left_only
2AP229/Web_NP8512_9.5 | 2AP229 | Web_NP8512_9.5 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/inch@130DegC/14psi/1.0 sec dwell) | Seal Strength (@130°C/14psi/1.0 sec dwell) |  | g/in | FDA,0 | 2000 |  | TM 107 |  |  | left_only
2AP229/Web_NP8512_9.5 | 2AP229 | Web_NP8512_9.5 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness |  | mil | FDA,2 | 3.06 | 3.74 | TM 132 |  |  | left_only
2AP260/Web_NP8535_12.75 | 2AP260 | Web_NP8535_12.75 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  |  | left_only
2AP260/Web_NP8535_12.75 | 2AP260 | Web_NP8535_12.75 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight |  | #/Ream | FDA,2 | 35.82 | 45.6 | TM 101 |  |  | left_only
2AP260/Web_NP8535_12.75 | 2AP260 | Web_NP8535_12.75 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width |  | in | FDA,4 | 12.6875 | 12.8125 | Measurement |  |  | left_only
2AP260/Web_NP8535_12.75 | 2AP260 | Web_NP8535_12.75 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) |  | g/in | FDA,0 | 3500 |  | TM 107 |  |  | left_only
2AP260/Web_NP8535_12.75 | 2AP260 | Web_NP8535_12.75 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness |  | mil | FDA,1 | 2.2 | 2.8 | TM 132 |  |  | left_only
2AP806/Web_NP8430_310 | 2AP806 | Web_NP8430_310 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  |  | left_only
2AP806/Web_NP8430_310 | 2AP806 | Web_NP8430_310 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight |  | #/Ream | FDA,2 | 44.88 | 57.12 | TM 101 |  |  | left_only
2AP806/Web_NP8430_310 | 2AP806 | Web_NP8430_310 | NY-Analytical | 761 | Web Measurement | Web Width (mm) | Web Width |  | mm | FDA,0 | 308 | 312 | Measurement |  |  | left_only
2AP806/Web_NP8430_310 | 2AP806 | Web_NP8430_310 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) |  | g/in | FDA,0 | 3500 |  | TM 107 |  |  | left_only
2AP806/Web_NP8430_310 | 2AP806 | Web_NP8430_310 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness |  | mil | FDA,2 | 2.64 | 3.36 | TM 132 |  |  | left_only
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Lid_04B_560_1 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | 2AQ471LI/Lid_04B_560_1 | 4 | both
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | 2AQ471LI/Lid_04B_560_1 | 4 | both
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Cap Thickness | Cap Thickness | 0.154 | in | FDA,3 | 0.147 | 0.157 | TM 143 / Visual | 2AQ471LI/Lid_04B_560_1 | 4 | both
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Blue (Done) |  |  |  |  | TM 143 / Visual | 2AQ471LI/Lid_04B_560_1 | 4 | both
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) |  |  |  |  | TM 143 / Visual | 2AQ471LI/Lid_04B_560_1 | 4 | both
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter (as per specification) | Lid Inside Diameter | 3.7695 | in | FDA,3 | 3.76 | 3.77 | TM 143 / Visual | 2AQ471LI/Lid_04B_560_1 | 4 | both
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.05 | in | FDA,3 | 0.04 | 0.05 | TM 143 / Visual | 2AQ471LI/Lid_04B_560_1 | 4 | both
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual | 2AQ471LI/Lid_04B_560_1 | 4 | both
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.5845 | in | FDA,3 | 1.58 | 1.6 | TM 143 / Visual | 2AQ471LI/Lid_04B_560_1 | 4 | both
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Outer Diameter (per specification) | Pop-Up Hole Outer Diameter | 0.565 | in | FDA,3 | 0.555 | 0.565 | TM 143 / Visual | 2AQ471LI/Lid_04B_560_1 | 4 | both
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Thicknes | Pop-Up Hole Thicknes | 0.02 | in | FDA,3 | 0.015 | 0.025 | TM 143 / Visual | 2AQ471LI/Lid_04B_560_1 | 4 | both
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 22.773 | g | FDA,1 | 20 | 24 | TM 143 / Visual | 2AQ471LI/Lid_04B_560_1 | 4 | both
2AU191/Web_NP5115_7.5 | 2AU191 | Web_NP5115_7.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2AU191/Web_NP5115_7.5 | 1 | both
2AU191/Web_NP5115_7.5 | 2AU191 | Web_NP5115_7.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 7.4375 | in | FDA,4 | 7.4375 | 7.5625 | Measurement | 2AU191/Web_NP5115_7.5 | 1 | both
2AU191/Web_NP5115_7.5 | 2AU191 | Web_NP5115_7.5 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Printed/Varnished) | Basis Weight | 60.5 | #/Ream | FDA,1 | 55.6 | 68 | TM 101 | 2AU191/Web_NP5115_7.5 | 1 | both
2AU191/Web_NP5115_7.5 | 2AU191 | Web_NP5115_7.5 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@325DegF/40PSI/0.5Sec_Dwell) | Seal Strength (@325°F/40PSI/0.5Sec Dwell) | 1823.34 | g/in | FDA,0 | 1500 |  | TM 107 | 2AU191/Web_NP5115_7.5 | 1 | both
2AU252/Web_NP3459_8.0 | 2AU252 | Web_NP3459_8.0 |  | 1044 | Basis Weight (lbs/ream) | Basis Weight | Basis Weight | 63.4 | Lb/Ream | FDA,2 | 57.56 | 70.4 | ASTM D 646 | 2AU252/Web_NP3459_8.0 | 1 | both
2AU252/Web_NP3459_8.0 | 2AU252 | Web_NP3459_8.0 |  | 757 | Caliper | caliper (COA) | Caliper | 3.8 | mil | FDA,3 | 3.515 | 3.515 | Caliper | 2AU252/Web_NP3459_8.0 | 1 | both
2AU252/Web_3459_8 | 2AU252 | Web_3459_8 | NY-Analytical | 1044 | Basis Weight (lbs/ream) | Basis Weight | Basis Weight |  | Lb/Ream | FDA,2 | 57.56 | 70.4 | TM 101 | 2AU252/Web_3459_8 | 1 | both
2AU252/Web_NP3459_8.0 | 2AU252 | Web_NP3459_8.0 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/inch@290DegF/40psi/0.5 sec dwell) | Seal Strength (@290°F/40psi/0.5 sec dwell) |  | g/in | FDA,0 | 7500 |  | TM 107 | 2AU252/Web_NP3459_8.0 | 1 | both
2AU252/Web_3459_8 | 2AU252 | Web_3459_8 | NY-Analytical | 761 | Web Measurement | Core ID (Inches) | Core ID |  | in | FDA,3 | 2.9375 | 3.0625 | TM 143 | 2AU252/Web_3459_8 | 1 | both
2AU252/Web_3459_8 | 2AU252 | Web_3459_8 | NY-Analytical | 761 | Web Measurement | Roll OD | Roll OD | 15 | in | FDA,3 | 14.9375 | 15.0625 | TM 143 | 2AU252/Web_3459_8 | 1 | both
2AU252/Web_NP3459_8.0 | 2AU252 | Web_NP3459_8.0 |  | 761 | Web Measurement | Web Width (Inches) | Web Width | 8 | in | FDA,4 | 7.9375 | 8.0625 | TM 143 | 2AU252/Web_NP3459_8.0 | 1 | both
2AU252/Web_3459_8 | 2AU252 | Web_3459_8 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 8 | in | FDA,4 | 7.9375 | 8.0625 | TM 143 | 2AU252/Web_3459_8 | 1 | both
2AU252/Web_NP3459_8.0 | 2AU252 | Web_NP3459_8.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Visual | 2AU252/Web_NP3459_8.0 | 1 | both
2AU252/Web_3459_8 | 2AU252 | Web_3459_8 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Visual | 2AU252/Web_3459_8 | 1 | both
2AU252/Web_NP3459_8.0 | 2AU252 | Web_NP3459_8.0 |  | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | Visual Observation | 2AU252/Web_NP3459_8.0 | 1 | both
2AU252/Web_NP3459_8.0 | 2AU252 | Web_NP3459_8.0 |  | 748 | Visual Inspection | Color/Copy | Color/Copy | Meets specifications (Done) |  |  |  |  | Visual Observation | 2AU252/Web_NP3459_8.0 | 1 | both
2AU555/Web_NP3460_7.5 | 2AU555 | Web_NP3460_7.5 | NY-Analytical | 1044 | Basis Weight (lbs/ream) | Basis Weight | Basis Weight | 72.89 | Lb/Ream | FDA,1 | 66.6 | 81.4 | ASTM D 646 | 2AU555/Web_NP3460_7.5 | 6 | both
2AU555/Web_NP3460_7.5 | 2AU555 | Web_NP3460_7.5 | NY-Analytical | 757 | Caliper | caliper (COA) | Caliper | 4.22 | mil | FDA,4 | 3.995 | 3.995 | Caliper | 2AU555/Web_NP3460_7.5 | 6 | both
2AU555/Web_NP3460_7.5 | 2AU555 | Web_NP3460_7.5 | NY-Analytical | 836 | Dyne Level | Dyne Level (outside film) | Dyne Level (outside film) | 36 |  | FDA,0 |  |  | Dyne Test Kit | 2AU555/Web_NP3460_7.5 | 6 | both
2AU555/Web_NP3460_7.5 | 2AU555 | Web_NP3460_7.5 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/inch@290DegF/40psi/0.5 sec dwell) | Seal Strength (@290°F/40psi/0.5 sec dwell) | 8720.24 | g/in | FDA,0 | 5850 |  | TM 107 | 2AU555/Web_NP3460_7.5 | 6 | both
2AU555/Web_NP3460_7.5 | 2AU555 | Web_NP3460_7.5 | NY-Analytical | 761 | Web Measurement | Core ID (Inches) | Core ID | 3 | in | FDA,4 | 2.9375 | 3.0625 | TM 143 | 2AU555/Web_NP3460_7.5 | 6 | both
2AU555/Web_NP3460_7.5 | 2AU555 | Web_NP3460_7.5 | NY-Analytical | 761 | Web Measurement | Roll OD | Roll OD | 15 | in | FDA,4 | 14.9375 | 15.0625 | TM 143 | 2AU555/Web_NP3460_7.5 | 6 | both
2AU555/Web_NP3460_7.5 | 2AU555 | Web_NP3460_7.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 7.5 | in | FDA,4 | 7.4375 | 7.5625 | TM 143 | 2AU555/Web_NP3460_7.5 | 6 | both
2AU555/Web_NP3460_7.5 | 2AU555 | Web_NP3460_7.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Visual | 2AU555/Web_NP3460_7.5 | 6 | both
2AU555/Web_NP3460_7.5 | 2AU555 | Web_NP3460_7.5 | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | Visual Observation | 2AU555/Web_NP3460_7.5 | 6 | both
2AU555/Web_NP3460_7.5 | 2AU555 | Web_NP3460_7.5 | NY-Analytical | 748 | Visual Inspection | Color/Copy | Color/Copy | Meets specifications (Done) |  |  |  |  | Visual Observation | 2AU555/Web_NP3460_7.5 | 6 | both
2BA585SLI/Lid_19AB | 2BA585SLI | Lid_19AB | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width |  | mm | FDA,0 | 52 | 56 | TM 143 / Visual |  |  | left_only
2BA585SLI/Lid_19AB | 2BA585SLI | Lid_19AB | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length |  | mm | FDA,0 | 71 | 75 | TM 143 / Visual |  |  | left_only
2BA585SLI/Lid_19AB | 2BA585SLI | Lid_19AB | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2BA585SLI/Lid_19AB | 2BA585SLI | Lid_19AB | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2BA585SLI/Lid_19AB | 2BA585SLI | Lid_19AB | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection |  |  | left_only
2BA585SLI/Lid_19AB | 2BA585SLI | Lid_19AB | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2BA585SLI/Lid_19AB | 2BA585SLI | Lid_19AB | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection |  |  | left_only
2BA972/Web_NP8540_13.5 | 2BA972 | Web_NP8540_13.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Web Measurement |  |  | left_only
2BA972/Web_NP8540_13.5 | 2BA972 | Web_NP8540_13.5 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight |  | #/Ream | FDA,1 | 45.4 | 57.8 | TM 101 |  |  | left_only
2BA972/Web_NP8540_13.5 | 2BA972 | Web_NP8540_13.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width |  | in | FDA,4 | 13.4375 | 13.5625 | Measurement |  |  | left_only
2BA972/Web_NP8540_13.5 | 2BA972 | Web_NP8540_13.5 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) |  | g/in | FDA,0 | 3500 |  | TM 107 |  |  | left_only
2BA972/Web_NP8540_13.5 | 2BA972 | Web_NP8540_13.5 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness |  | mil | FDA,2 | 2.86 | 3.64 | TM 132 |  |  | left_only
2BC801/Web_NP3870_6.0 | 2BC801 | Web_NP3870_6.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2BC801/Web_NP3870_6.0 | 2 | both
2BC801/Web_NP3870_6.0 | 2BC801 | Web_NP3870_6.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 6 | in | FDA,4 | 5.9375 | 6.0625 | Measurement | 2BC801/Web_NP3870_6.0 | 2 | both
2BC801/Web_NP3870_6.0 | 2BC801 | Web_NP3870_6.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 92.46 | #/Ream | FDA,1 | 81 | 99 | TM 101 | 2BC801/Web_NP3870_6.0 | 2 | both
2BC801/Web_NP3870_6.0 | 2BC801 | Web_NP3870_6.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@350DegF/40psi/1.0 sec dwell) | Seal Strength (@350°F/40psi/1.0 sec dwell) | 10097 | g/in | FDA,0 | 5000 |  | TM 107 | 2BC801/Web_NP3870_6.0 | 2 | both
2BH592/Web_NP3895_6.0 | 2BH592 | Web_NP3895_6.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2BH592/Web_NP3895_6.0 | 1 | both
2BH592/Web_NP3895_6.0 | 2BH592 | Web_NP3895_6.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 6 | in | FDA,4 | 5.9375 | 6.0625 | Measurement | 2BH592/Web_NP3895_6.0 | 1 | both
2BH592/Web_NP3895_6.0 | 2BH592 | Web_NP3895_6.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/3000_sq_ft) | Basis Weight | 69.71 | #/3000sqft | FDA,1 | 65.3 | 79.8 | TM 101 | 2BH592/Web_NP3895_6.0 | 1 | both
2BH592/Web_NP3895_6.0 | 2BH592 | Web_NP3895_6.0 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/inch@250DegF/30psi/0.5 sec dwell) | Seal Strength (@250°F/30psi/0.5 sec dwell) | 5740.1 | g/in | FDA,0 | 2270 |  | TM 107 | 2BH592/Web_NP3895_6.0 | 1 | both
2BH592/Web_NP3895_6.0 | 2BH592 | Web_NP3895_6.0 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 4.5 | mil | FDA,1 | 3.9 | 4.7 | TM 132 | 2BH592/Web_NP3895_6.0 | 1 | both
2BJ123/Web_NP8550_11 | 2BJ123 | Web_NP8550_11 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2BJ123/Web_NP8550_11 | 1 | both
2BJ123/Web_NP8550_11 | 2BJ123 | Web_NP8550_11 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 11 | in | FDA,4 | 10.9375 | 11.0625 | Measurement | 2BJ123/Web_NP8550_11 | 1 | both
2BJ123/Web_NP8550_11 | 2BJ123 | Web_NP8550_11 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 43.12 | #/Ream | FDA,1 | 37.8 | 48.2 | TM 101 | 2BJ123/Web_NP8550_11 | 1 | both
2BJ123/Web_NP8550_11 | 2BJ123 | Web_NP8550_11 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 6855 | g/in | FDA,0 | 3500 |  | TM 107 | 2BJ123/Web_NP8550_11 | 1 | both
2BJ143/Web_NP8530_16.25 | 2BJ143 | Web_NP8530_16.25 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2BJ143/Web_NP8530_16.25 | 1 | both
2BJ143/Web_NP8530_16.25 | 2BJ143 | Web_NP8530_16.25 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 16.25 | in | FDA,4 | 16.1875 | 16.3125 | Measurement | 2BJ143/Web_NP8530_16.25 | 1 | both
2BJ143/Web_NP8530_16.25 | 2BJ143 | Web_NP8530_16.25 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 34.26 | #/Ream | FDA,2 | 29.4 | 37.42 | TM 101 | 2BJ143/Web_NP8530_16.25 | 1 | both
2BJ143/Web_NP8530_16.25 | 2BJ143 | Web_NP8530_16.25 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 7619 | g/in | FDA,0 | 3500 |  | TM 107 | 2BJ143/Web_NP8530_16.25 | 1 | both
2BJ143/Web_NP8530_16.25 | 2BJ143 | Web_NP8530_16.25 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 2.1 | mil | FDA,1 | 1.8 | 2.2 | TM 132 | 2BJ143/Web_NP8530_16.25 | 1 | both
2BJ213/Web_NP8530_18 | 2BJ213 | Web_NP8530_18 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2BJ213/Web_NP8530_18 | 1 | both
2BJ213/Web_NP8530_18 | 2BJ213 | Web_NP8530_18 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 18 | in | FDA,4 | 17.9375 | 18.0625 | Measurement | 2BJ213/Web_NP8530_18 | 1 | both
2BJ213/Web_NP8530_18 | 2BJ213 | Web_NP8530_18 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 33.33 | #/Ream | FDA,2 | 29.4 | 37.42 | TM 101 | 2BJ213/Web_NP8530_18 | 1 | both
2BJ213/Web_NP8530_18 | 2BJ213 | Web_NP8530_18 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 6908 | g/in | FDA,0 | 3500 |  | TM 107 | 2BJ213/Web_NP8530_18 | 1 | both
2BJ213/Web_NP8530_18 | 2BJ213 | Web_NP8530_18 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 1.85 | mil | FDA,1 | 1.8 | 2.2 | TM 132 | 2BJ213/Web_NP8530_18 | 1 | both
2BJ227/Web_NP8530_15.75 | 2BJ227 | Web_NP8530_15.75 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  |  | left_only
2BJ227/Web_NP8530_15.75 | 2BJ227 | Web_NP8530_15.75 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight |  | #/Ream | FDA,2 | 29.4 | 37.42 | TM 101 |  |  | left_only
2BJ227/Web_NP8530_15.75 | 2BJ227 | Web_NP8530_15.75 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width |  | in | FDA,4 | 15.6875 | 15.8125 | Measurement |  |  | left_only
2BJ227/Web_NP8530_15.75 | 2BJ227 | Web_NP8530_15.75 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) |  | g/in | FDA,0 | 3500 |  | TM 107 |  |  | left_only
2BJ227/Web_NP8530_15.75 | 2BJ227 | Web_NP8530_15.75 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness |  | mil | FDA,1 | 1.8 | 2.2 | TM 132 |  |  | left_only
2BP308SLI/Lid_Aveeno Ultra-Calming MUR | 2BP308SLI | Lid_Aveeno Ultra-Calming MUR | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection | 2BP308SLI/Lid_Aveeno Ultra-Calming MUR | 2 | both
2BP308SLI/Lid_Aveeno Ultra-Calming MUR | 2BP308SLI | Lid_Aveeno Ultra-Calming MUR | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | 2BP308SLI/Lid_Aveeno Ultra-Calming MUR | 2 | both
2BP308SLI/Lid_Aveeno Ultra-Calming MUR | 2BP308SLI | Lid_Aveeno Ultra-Calming MUR | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | 2BP308SLI/Lid_Aveeno Ultra-Calming MUR | 2 | both
2BP308SLI/Lid_Aveeno Ultra-Calming MUR | 2BP308SLI | Lid_Aveeno Ultra-Calming MUR | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual | 2BP308SLI/Lid_Aveeno Ultra-Calming MUR | 2 | both
2BP308SLI/Lid_Aveeno Ultra-Calming MUR | 2BP308SLI | Lid_Aveeno Ultra-Calming MUR | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 72 | mm | FDA,0 | 71 | 75 | TM 143 / Visual | 2BP308SLI/Lid_Aveeno Ultra-Calming MUR | 2 | both
2BP308SLI/Lid_Aveeno Ultra-Calming MUR | 2BP308SLI | Lid_Aveeno Ultra-Calming MUR | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 53 | mm | FDA,0 | 52 | 56 | TM 143 / Visual | 2BP308SLI/Lid_Aveeno Ultra-Calming MUR | 2 | both
2BP308SLI/Lid_Aveeno Ultra-Calming MUR | 2BP308SLI | Lid_Aveeno Ultra-Calming MUR | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual | 2BP308SLI/Lid_Aveeno Ultra-Calming MUR | 2 | both
2BQ084LI/Lid_11A | 2BQ084LI | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid fits with a tight seal on |  |  |  |  | Canister & Lid Inspection | 2BQ084LI/Lid_11A | 1 | both
2BQ084LI/Lid_11A | 2BQ084LI | Lid_11A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Flip Lid Force | 2BQ084LI/Lid_11A | 1 | both
2BQ084LI/Lid_11A | 2BQ084LI | Lid_11A | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | Flip Lid Force | 2.32 | lb | FDA,2 | 0.75 |  | TM 143 | 2BQ084LI/Lid_11A | 1 | both
2BQ084LI/Lid_11A | 2BQ084LI | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | TM 143 / Visual | 2BQ084LI/Lid_11A | 1 | both
2BQ084LI/Lid_11A | 2BQ084LI | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | Bead Diameter | 4.85 | in | FDA,3 | 4.847 | 4.867 | TM 143 / Visual | 2BQ084LI/Lid_11A | 1 | both
2BQ084LI/Lid_11A | 2BQ084LI | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Translucent Teal (Done) |  |  |  |  | TM 143 / Visual | 2BQ084LI/Lid_11A | 1 | both
2BQ084LI/Lid_11A | 2BQ084LI | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Insert Fin Gap (per specification) | Insert Fin Gap | 0.071 | in | FDA,3 | 0.065 | 0.075 | TM 143 / Visual | 2BQ084LI/Lid_11A | 1 | both
2BQ084LI/Lid_11A | 2BQ084LI | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter (as per specification) | Lid Inside Diameter | 4.92 | in | FDA,3 | 4.917 | 4.937 | TM 143 / Visual | 2BQ084LI/Lid_11A | 1 | both
2BQ084LI/Lid_11A | 2BQ084LI | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.044 | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual | 2BQ084LI/Lid_11A | 1 | both
2BQ084LI/Lid_11A | 2BQ084LI | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual | 2BQ084LI/Lid_11A | 1 | both
2BQ084LI/Lid_11A | 2BQ084LI | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness (per specification) | Orifice Flap Thickness | 0.023 | in | FDA,3 | 0.018 | 0.028 | TM 143 / Visual | 2BQ084LI/Lid_11A | 1 | both
2BQ084LI/Lid_11A | 2BQ084LI | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.48 | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual | 2BQ084LI/Lid_11A | 1 | both
2BQ084LI/Lid_11A | 2BQ084LI | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Outer Diameter (per specification) | Pop-Up Hole Outer Diameter | 0.166 | in | FDA,3 | 0.16 | 0.17 | TM 143 / Visual | 2BQ084LI/Lid_11A | 1 | both
2BQ084LI/Lid_11A | 2BQ084LI | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 34.11 | g | FDA,1 | 32 | 42 | TM 143 / Visual | 2BQ084LI/Lid_11A | 1 | both
2BQ471LI/Lid_04F_560 | 2BQ471LI | Lid_04F_560 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | 2BQ471LI/Lid_04F_560 | 2 | both
2BQ471LI/Lid_04F_560 | 2BQ471LI | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | 2BQ471LI/Lid_04F_560 | 2 | both
2BQ471LI/Lid_04F_560 | 2BQ471LI | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Cap Thickness | Cap Thickness | 0.151 | in | FDA,3 | 0.147 | 0.157 | TM 143 / Visual | 2BQ471LI/Lid_04F_560 | 2 | both
2BQ471LI/Lid_04F_560 | 2BQ471LI | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Translucent Blue (Done) |  |  |  |  | TM 143 / Visual | 2BQ471LI/Lid_04F_560 | 2 | both
2BQ471LI/Lid_04F_560 | 2BQ471LI | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) |  |  |  |  | TM 143 / Visual | 2BQ471LI/Lid_04F_560 | 2 | both
2BQ471LI/Lid_04F_560 | 2BQ471LI | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter (as per specification) | Lid Inside Diameter | 3.7695 | in | FDA,3 | 3.76 | 3.77 | TM 143 / Visual | 2BQ471LI/Lid_04F_560 | 2 | both
2BQ471LI/Lid_04F_560 | 2BQ471LI | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.046 | in | FDA,3 | 0.04 | 0.05 | TM 143 / Visual | 2BQ471LI/Lid_04F_560 | 2 | both
2BQ471LI/Lid_04F_560 | 2BQ471LI | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual | 2BQ471LI/Lid_04F_560 | 2 | both
2BQ471LI/Lid_04F_560 | 2BQ471LI | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.486 | in | FDA,3 | 1.473 | 1.493 | TM 143 / Visual | 2BQ471LI/Lid_04F_560 | 2 | both
2BQ471LI/Lid_04F_560 | 2BQ471LI | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Outer Diameter (per specification) | Pop-Up Hole Outer Diameter | 0.555 | in | FDA,3 | 0.555 | 0.565 | TM 143 / Visual | 2BQ471LI/Lid_04F_560 | 2 | both
2BQ471LI/Lid_04F_560 | 2BQ471LI | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Thicknes | Pop-Up Hole Thicknes | 0.015 | in | FDA,3 | 0.015 | 0.025 | TM 143 / Visual | 2BQ471LI/Lid_04F_560 | 2 | both
2BQ471LI/Lid_04F_560 | 2BQ471LI | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 22.425 | g | FDA,1 | 20 | 24 | TM 143 / Visual | 2BQ471LI/Lid_04F_560 | 2 | both
2BQ551LI/Skip Lot Release_NY_2OQ850 | 2BQ551LI | Skip Lot Release_NY_2OQ850 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA |  |  | left_only
2BQ551LI/Applicator_NY_Skip lot testing | 2BQ551LI | Applicator_NY_Skip lot testing | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Certificate of Analysis | 2BQ551LI/Applicator_NY_Skip lot testing | 1 | both
2BQ551LI/Lid_13A | 2BQ551LI | Lid_13A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Flip Lid Force | 2BQ551LI/Lid_13A | 10 | both
2BQ551LI/Lid_13A | 2BQ551LI | Lid_13A | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | Flip Lid Force | 1.61 | lb | FDA,2 | 0.75 |  | TM 143 | 2BQ551LI/Lid_13A | 10 | both
2BQ551LI/Lid_NY_Skip lot testing | 2BQ551LI | Lid_NY_Skip lot testing | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | Flip Lid Force | 1.07 | lb | FDA,2 | 0.75 |  | TM 143 | 2BQ551LI/Lid_NY_Skip lot testing | 80 | both
2BQ551LI/Lid_13A | 2BQ551LI | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | TM 143 / Visual | 2BQ551LI/Lid_13A | 10 | both
2BQ551LI/Lid_13A | 2BQ551LI | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | Bead Diameter | 4.8545 | in | FDA,3 | 4.847 | 4.867 | TM 143 / Visual | 2BQ551LI/Lid_13A | 10 | both
2BQ551LI/Lid_NY_Skip lot testing | 2BQ551LI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | Bead Diameter | 4.85 | in | FDA,3 | 4.847 | 4.867 | TM 143 / Visual | 2BQ551LI/Lid_NY_Skip lot testing | 80 | both
2BQ551LI/Lid_13A | 2BQ551LI | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Center to Center Distance | Center to Center Distance | 0.45 | in | FDA,3 | 0.445 | 0.455 | TM 143 / Visual | 2BQ551LI/Lid_13A | 10 | both
2BQ551LI/Lid_NY_Skip lot testing | 2BQ551LI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Center to Center Distance | Center to Center Distance | 0.45 | in | FDA,3 | 0.445 | 0.455 | TM 143 / Visual | 2BQ551LI/Lid_NY_Skip lot testing | 80 | both
2BQ551LI/Lid_13A | 2BQ551LI | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Translucent Purple (Done) |  |  |  |  | TM 143 / Visual | 2BQ551LI/Lid_13A | 10 | both
2BQ551LI/Lid_NY_Skip lot testing | 2BQ551LI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Translucent Purple (Done) |  |  |  |  | TM 143 / Visual | 2BQ551LI/Lid_NY_Skip lot testing | 80 | both
2BQ551LI/Lid_13A | 2BQ551LI | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | End Circle Diameter | End Circle Diameter | 0.123 | in | FDA,3 | 0.12 | 0.13 | TM 143 / Visual | 2BQ551LI/Lid_13A | 10 | both
2BQ551LI/Lid_NY_Skip lot testing | 2BQ551LI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | End Circle Diameter | End Circle Diameter | 0.13 | in | FDA,3 | 0.12 | 0.13 | TM 143 / Visual | 2BQ551LI/Lid_NY_Skip lot testing | 80 | both
2BQ551LI/Lid_13A | 2BQ551LI | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) |  |  |  |  | TM 143 / Visual | 2BQ551LI/Lid_13A | 10 | both
2BQ551LI/Lid_NY_Skip lot testing | 2BQ551LI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) |  |  |  |  | TM 143 / Visual | 2BQ551LI/Lid_NY_Skip lot testing | 80 | both
2BQ551LI/Lid_13A | 2BQ551LI | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Insert Fin Gap (per specification) | Insert Fin Gap | 0.076 | in | FDA,3 | 0.07 | 0.08 | TM 143 / Visual | 2BQ551LI/Lid_13A | 10 | both
2BQ551LI/Lid_NY_Skip lot testing | 2BQ551LI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Insert Fin Gap (per specification) | Insert Fin Gap | 0.076 | in | FDA,3 | 0.07 | 0.08 | TM 143 / Visual | 2BQ551LI/Lid_NY_Skip lot testing | 80 | both
2BQ551LI/Lid_13A | 2BQ551LI | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter (as per specification) | Lid Inside Diameter | 4.923 | in | FDA,3 | 4.917 | 4.937 | TM 143 / Visual | 2BQ551LI/Lid_13A | 10 | both
2BQ551LI/Lid_NY_Skip lot testing | 2BQ551LI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter (as per specification) | Lid Inside Diameter | 4.92 | in | FDA,3 | 4.917 | 4.937 | TM 143 / Visual | 2BQ551LI/Lid_NY_Skip lot testing | 80 | both
2BQ551LI/Lid_13A | 2BQ551LI | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.04 | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual | 2BQ551LI/Lid_13A | 10 | both
2BQ551LI/Lid_NY_Skip lot testing | 2BQ551LI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.042 | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual | 2BQ551LI/Lid_NY_Skip lot testing | 80 | both
2BQ551LI/Lid_13A | 2BQ551LI | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual | 2BQ551LI/Lid_13A | 10 | both
2BQ551LI/Lid_13A | 2BQ551LI | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness (per specification) | Orifice Flap Thickness | 0.04 | in | FDA,3 | 0.03 | 0.04 | TM 143 / Visual | 2BQ551LI/Lid_13A | 10 | both
2BQ551LI/Lid_NY_Skip lot testing | 2BQ551LI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness (per specification) | Orifice Flap Thickness | 0.031 | in | FDA,3 | 0.03 | 0.04 | TM 143 / Visual | 2BQ551LI/Lid_NY_Skip lot testing | 80 | both
2BQ551LI/Lid_13A | 2BQ551LI | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.483 | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual | 2BQ551LI/Lid_13A | 10 | both
2BQ551LI/Lid_NY_Skip lot testing | 2BQ551LI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.49 | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual | 2BQ551LI/Lid_NY_Skip lot testing | 80 | both
2BQ551LI/Lid_13A | 2BQ551LI | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 34.73 | g | FDA,1 | 32 | 42 | TM 143 / Visual | 2BQ551LI/Lid_13A | 10 | both
2BQ551LI/Lid_NY_Skip lot testing | 2BQ551LI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 34.65 | g | FDA,1 | 32 | 42 | TM 143 / Visual | 2BQ551LI/Lid_NY_Skip lot testing | 80 | both
2BS420/Web_NP3870_11.5 | 2BS420 | Web_NP3870_11.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2BS420/Web_NP3870_11.5 | 1 | both
2BS420/Web_NP3870_11.5 | 2BS420 | Web_NP3870_11.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 11.5 | in | FDA,4 | 11.4375 | 11.5625 | Measurement | 2BS420/Web_NP3870_11.5 | 1 | both
2BS420/Web_NP3870_11.5 | 2BS420 | Web_NP3870_11.5 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 91.8 | #/Ream | FDA,1 | 81 | 99 | TM 101 | 2BS420/Web_NP3870_11.5 | 1 | both
2BS420/Web_NP3870_11.5 | 2BS420 | Web_NP3870_11.5 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@350DegF/40psi/1.0 sec dwell) | Seal Strength (@350°F/40psi/1.0 sec dwell) | 10192.62 | g/in | FDA,0 | 5000 |  | TM 107 | 2BS420/Web_NP3870_11.5 | 1 | both
2BU275/Web_NP3895_8.5 | 2BU275 | Web_NP3895_8.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2BU275/Web_NP3895_8.5 | 1 | both
2BU275/Web_NP3895_8.5 | 2BU275 | Web_NP3895_8.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 8.5 | in | FDA,4 | 8.4375 | 8.5625 | Measurement | 2BU275/Web_NP3895_8.5 | 1 | both
2BU275/Web_NP3895_8.5 | 2BU275 | Web_NP3895_8.5 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/3000_sq_ft) | Basis Weight | 71.69 | #/3000sqft | FDA,1 | 65.3 | 79.8 | TM 101 | 2BU275/Web_NP3895_8.5 | 1 | both
2BU275/Web_NP3895_8.5 | 2BU275 | Web_NP3895_8.5 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/inch@250DegF/30psi/0.5 sec dwell) | Seal Strength (@250°F/30psi/0.5 sec dwell) | 5589 | g/in | FDA,0 | 2270 |  | TM 107 | 2BU275/Web_NP3895_8.5 | 1 | both
2BU275/Web_NP3895_8.5 | 2BU275 | Web_NP3895_8.5 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 4.05 | mil | FDA,1 | 3.9 | 4.7 | TM 132 | 2BU275/Web_NP3895_8.5 | 1 | both
2CA580SLI/Lid_01R | 2CA580SLI | Lid_01R | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width |  | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual |  |  | left_only
2CA580SLI/Lid_01R | 2CA580SLI | Lid_01R | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length |  | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual |  |  | left_only
2CA580SLI/Lid_01R | 2CA580SLI | Lid_01R | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2CA580SLI/Lid_01R | 2CA580SLI | Lid_01R | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2CA580SLI/Lid_01R | 2CA580SLI | Lid_01R | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection |  |  | left_only
2CA580SLI/Lid_01R | 2CA580SLI | Lid_01R | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2CA580SLI/Lid_01R | 2CA580SLI | Lid_01R | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection |  |  | left_only
2CA735/Web_NP8810_12.125 | 2CA735 | Web_NP8810_12.125 | NY-Analytical | 1019 | Moisture Vapor Transmission rate | Report result from COA | Report result from COA | 5 | GSM | FDA,0 |  | 10 | Report result from COA | 2CA735/Web_NP8810_12.125 | 1 | both
2CA735/Web_NP8810_12.125 | 2CA735 | Web_NP8810_12.125 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Dyne Level | 2CA735/Web_NP8810_12.125 | 1 | both
2CA735/Web_NP8810_12.125 | 2CA735 | Web_NP8810_12.125 | NY-Analytical | 836 | Dyne Level | Dyne Level (outside film) | Dyne Level (outside film) | 36 |  | FDA,0 | 36 | 38 | Dyne Test Kit | 2CA735/Web_NP8810_12.125 | 1 | both
2CA735/Web_NP8810_12.125 | 2CA735 | Web_NP8810_12.125 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 12.125 | in | FDA,4 | 12.0625 | 12.1875 | Measurement | 2CA735/Web_NP8810_12.125 | 1 | both
2CA735/Web_NP8810_12.125 | 2CA735 | Web_NP8810_12.125 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 42.33 | #/Ream | FDA,1 | 39.4 | 48.2 | TM 101 | 2CA735/Web_NP8810_12.125 | 1 | both
2CA735/Web_NP8810_12.125 | 2CA735 | Web_NP8810_12.125 | NY-Analytical | 762 | Seal Strength | Seal Strength(g/in@350F/30psi/0.5s dwell Top Jaw) | Seal Strength (@350°F/30psi/0.5s dwell Top Jaw) | 5303 | gms/f | FDA,0 | 500 |  | TM 107 | 2CA735/Web_NP8810_12.125 | 1 | both
2CA735/Web_NP8810_12.125 | 2CA735 | Web_NP8810_12.125 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 2.7 | mil | FDA,2 | 2.48 | 3.03 | TM 132 | 2CA735/Web_NP8810_12.125 | 1 | both
2CA735/Web_NP8810_12.125 | 2CA735 | Web_NP8810_12.125 | NY-Analytical | 397 | Appearance | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | 2CA735/Web_NP8810_12.125 | 1 | both
2CB196/Web_NP5114_41.0 | 2CB196 | Web_NP5114_41.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2CB196/Web_NP5114_41.0 | 2 | both
2CB196/Web_NP5114_41.0 | 2CB196 | Web_NP5114_41.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 40.9375 | in | FDA,4 | 40.9375 | 41.0625 | Measurement | 2CB196/Web_NP5114_41.0 | 2 | both
2CB196/Web_NP5114_41.0 | 2CB196 | Web_NP5114_41.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 59.79 | #/Ream | FDA,1 | 53.1 | 64.9 | TM 101 | 2CB196/Web_NP5114_41.0 | 2 | both
2CB196/Web_NP5114_41.0 | 2CB196 | Web_NP5114_41.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/40psi/0.5 sec dewll) | Seal Strength (@300°F/40psi/0.5 sec dewll) | 2891.81 | g/in | FDA,0 | 1300 |  | TM 107 | 2CB196/Web_NP5114_41.0 | 2 | both
2CB196/Web_NP5114_41.0 | 2CB196 | Web_NP5114_41.0 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 3.2 | mil | FDA,2 | 3.15 | 3.85 | TM 132 | 2CB196/Web_NP5114_41.0 | 2 | both
2CB512/Web_NP3870_4.5 | 2CB512 | Web_NP3870_4.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2CB512/Web_NP3870_4.5 | 2 | both
2CB512/Web_NP3870_4.5 | 2CB512 | Web_NP3870_4.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 4.5 | in | FDA,4 | 4.4375 | 4.5625 | Measurement | 2CB512/Web_NP3870_4.5 | 2 | both
2CB512/Web_NP3870_4.5 | 2CB512 | Web_NP3870_4.5 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 92.6 | #/Ream | FDA,1 | 81 | 99 | TM 101 | 2CB512/Web_NP3870_4.5 | 2 | both
2CB512/Web_NP3870_4.5 | 2CB512 | Web_NP3870_4.5 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@350DegF/40psi/1.0 sec dwell) | Seal Strength (@350°F/40psi/1.0 sec dwell) | 9864.87 | g/in | FDA,0 | 5000 |  | TM 107 | 2CB512/Web_NP3870_4.5 | 2 | both
2DH211/Web_NP5116_6.0 | 2DH211 | Web_NP5116_6.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2DH211/Web_NP5116_6.0 | 1 | both
2DH211/Web_NP5116_6.0 | 2DH211 | Web_NP5116_6.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 6 | in | FDA,4 | 5.9375 | 6.0625 | Measurement | 2DH211/Web_NP5116_6.0 | 1 | both
2DH211/Web_NP5116_6.0 | 2DH211 | Web_NP5116_6.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Printed/Varnished) | Basis Weight | 57.67 | #/Ream | FDA,1 | 55.8 | 68.2 | TM 101 | 2DH211/Web_NP5116_6.0 | 1 | both
2DH211/Web_NP5116_6.0 | 2DH211 | Web_NP5116_6.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/40psi/0.5 sec dewll) | Seal Strength (@300°F/40psi/0.5 sec dewll) | 1800.28 | g/in | FDA,0 | 1500 |  | TM 107 | 2DH211/Web_NP5116_6.0 | 1 | both
2ED110/Web_NP5410_4.5 | 2ED110 | Web_NP5410_4.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2ED110/Web_NP5410_4.5 | 1 | both
2ED110/Web_NP5410_4.5 | 2ED110 | Web_NP5410_4.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 4.5 | in | FDA,4 | 4.4375 | 4.5625 | Measurement | 2ED110/Web_NP5410_4.5 | 1 | both
2ED110/Web_NP5410_4.5 | 2ED110 | Web_NP5410_4.5 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Printed) | Basis Weight | 54.5 | #/Ream | FDA,1 | 49.8 | 60.8 | TM 101 | 2ED110/Web_NP5410_4.5 | 1 | both
2ED110/Web_NP5410_4.5 | 2ED110 | Web_NP5410_4.5 | NY-Analytical | 762 | Seal Strength | Destruct Bond | Destruct Bond | Break occurs off the seal (Done) | g/in |  |  |  | TM 107 | 2ED110/Web_NP5410_4.5 | 1 | both
2ED110/Web_NP5410_4.5 | 2ED110 | Web_NP5410_4.5 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/40psi/0.5 sec dewll) | Seal Strength (@300°F/40psi/0.5 sec dewll) | 1814.1 | g/in | FDA,0 | 1500 |  | TM 107 | 2ED110/Web_NP5410_4.5 | 1 | both
2EU351/Web_NP5211_7.0 | 2EU351 | Web_NP5211_7.0 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Eyespot Repeat | 2EU351/Web_NP5211_7.0 | 1 | both
2EU351/Web_NP5211_7.0 | 2EU351 | Web_NP5211_7.0 | NY-Analytical | 999 | Eyespot Repeat | Eye spot repeat over 7 impressions | Eye spot repeat over 7 impressions | 488 | mm | FDA,0 | 488 | 490 | TM 208 | 2EU351/Web_NP5211_7.0 | 1 | both
2EU351/Web_NP5211_7.0 | 2EU351 | Web_NP5211_7.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 7 | in | FDA,4 | 6.9375 | 7.0625 | TM 143 | 2EU351/Web_NP5211_7.0 | 1 | both
2EU351/Web_NP5211_7.0 | 2EU351 | Web_NP5211_7.0 | NY-Analytical | 1031 | 2D Matrix (Barcode) | 2D Matrix (Barcode) | 2D Matrix (Barcode) | Meet Specifications (Done) |  |  |  |  | Visual | 2EU351/Web_NP5211_7.0 | 1 | both
2EU351/Web_NP5211_7.0 | 2EU351 | Web_NP5211_7.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Visual | 2EU351/Web_NP5211_7.0 | 1 | both
2EU351/Web_NP5211_7.0 | 2EU351 | Web_NP5211_7.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis (ELLI) | Certificate of Analysis | COA Meets Specification (Done) |  |  |  |  | Visual | 2EU351/Web_NP5211_7.0 | 1 | both
2EU351/Web_NP5211_7.0 | 2EU351 | Web_NP5211_7.0 | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | Visual Observation | 2EU351/Web_NP5211_7.0 | 1 | both
2EU351/Web_NP5211_7.0 | 2EU351 | Web_NP5211_7.0 | NY-Analytical | 748 | Visual Inspection | Color/Copy | Color/Copy | Meets specifications (Done) |  |  |  |  | Visual Observation | 2EU351/Web_NP5211_7.0 | 1 | both
2FB6038/Web_NP5410_8.0 | 2FB6038 | Web_NP5410_8.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2FB6038/Web_NP5410_8.0 | 7 | both
2FB6038/Web_NP5410_8.0 | 2FB6038 | Web_NP5410_8.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width |  | in | FDA,4 | 7.9375 | 8.0625 | Measurement | 2FB6038/Web_NP5410_8.0 | 7 | both
2FB6038/Web_NP5410_8.0 | 2FB6038 | Web_NP5410_8.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Printed) | Basis Weight | 52.91 | #/Ream | FDA,1 | 49.8 | 60.8 | TM 101 | 2FB6038/Web_NP5410_8.0 | 7 | both
2FB6038/Web_NP5410_8.0 | 2FB6038 | Web_NP5410_8.0 | NY-Analytical | 762 | Seal Strength | Destruct Bond | Destruct Bond |  | g/in |  |  |  | TM 107 | 2FB6038/Web_NP5410_8.0 | 7 | both
2FB6038/Web_NP5410_8.0 | 2FB6038 | Web_NP5410_8.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/40psi/0.5 sec dewll) | Seal Strength (@300°F/40psi/0.5 sec dewll) |  | g/in | FDA,0 | 1500 |  | TM 107 | 2FB6038/Web_NP5410_8.0 | 7 | both
2FH070/Web_NP5115_6.0 | 2FH070 | Web_NP5115_6.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2FH070/Web_NP5115_6.0 | 1 | both
2FH070/Web_NP5115_6.0 | 2FH070 | Web_NP5115_6.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 6 | in | FDA,4 | 5.9375 | 6.0625 | Measurement | 2FH070/Web_NP5115_6.0 | 1 | both
2FH070/Web_NP5115_6.0 | 2FH070 | Web_NP5115_6.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Printed/Varnished) | Basis Weight | 59.66 | #/Ream | FDA,1 | 55.6 | 68 | TM 101 | 2FH070/Web_NP5115_6.0 | 1 | both
2FH070/Web_NP5115_6.0 | 2FH070 | Web_NP5115_6.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@325DegF/40PSI/0.5Sec_Dwell) | Seal Strength (@325°F/40PSI/0.5Sec Dwell) | 2304.16 | g/in | FDA,0 | 1500 |  | TM 107 | 2FH070/Web_NP5115_6.0 | 1 | both
2FJ141R/Web_NP8535_16 | 2FJ141R | Web_NP8535_16 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  |  | left_only
2FJ141R/Web_NP8535_16 | 2FJ141R | Web_NP8535_16 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight |  | #/Ream | FDA,2 | 35.82 | 45.6 | TM 101 |  |  | left_only
2FJ141R/Web_NP8535_16 | 2FJ141R | Web_NP8535_16 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width |  | in | FDA,4 | 15.9375 | 16.0625 | Measurement |  |  | left_only
2FJ141R/Web_NP8535_16 | 2FJ141R | Web_NP8535_16 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) |  | g/in | FDA,0 | 3500 |  | TM 107 |  |  | left_only
2FJ141R/Web_NP8535_16 | 2FJ141R | Web_NP8535_16 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness |  | mil | FDA,1 | 2.2 | 2.8 | TM 132 |  |  | left_only
2HA580/Web_NP8540_17.0 | 2HA580 | Web_NP8540_17.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2HA580/Web_NP8540_17.0 | 1 | both
2HA580/Web_NP8540_17.0 | 2HA580 | Web_NP8540_17.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 17 | in | FDA,4 | 16.9375 | 17.0625 | Measurement | 2HA580/Web_NP8540_17.0 | 1 | both
2HA580/Web_NP8540_17.0 | 2HA580 | Web_NP8540_17.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 52.91 | #/Ream | FDA,1 | 45.4 | 57.8 | TM 101 | 2HA580/Web_NP8540_17.0 | 1 | both
2HA580/Web_NP8540_17.0 | 2HA580 | Web_NP8540_17.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 7254 | g/in | FDA,0 | 3500 |  | TM 107 | 2HA580/Web_NP8540_17.0 | 1 | both
2HA580/Web_NP8540_17.0 | 2HA580 | Web_NP8540_17.0 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 3 | mil | FDA,2 | 2.86 | 3.64 | TM 132 | 2HA580/Web_NP8540_17.0 | 1 | both
2HU120/Web_NP5115_8.5 | 2HU120 | Web_NP5115_8.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  |  | left_only
2HU120/Web_NP5115_8.5 | 2HU120 | Web_NP5115_8.5 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Printed/Varnished) | Basis Weight |  | #/Ream | FDA,1 | 55.6 | 68 | TM 101 |  |  | left_only
2HU120/Web_NP5115_8.5 | 2HU120 | Web_NP5115_8.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width |  | in | FDA,4 | 8.4375 | 8.5625 | Measurement |  |  | left_only
2HU120/Web_NP5115_8.5 | 2HU120 | Web_NP5115_8.5 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@325DegF/40PSI/0.5Sec_Dwell) | Seal Strength (@325°F/40PSI/0.5Sec Dwell) |  | g/in | FDA,0 | 1500 |  | TM 107 |  |  | left_only
2JC699/Web_NP5410_6.0 | 2JC699 | Web_NP5410_6.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2JC699/Web_NP5410_6.0 | 6 | both
2JC699/Web_NP5410_6.0 | 2JC699 | Web_NP5410_6.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 6 | in | FDA,4 | 5.9375 | 6.0625 | Measurement | 2JC699/Web_NP5410_6.0 | 6 | both
2JC699/Web_NP5410_6.0 | 2JC699 | Web_NP5410_6.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Printed) | Basis Weight | 55.6 | #/Ream | FDA,1 | 49.8 | 60.8 | TM 101 | 2JC699/Web_NP5410_6.0 | 6 | both
2JC699/Web_NP5410_6.0 | 2JC699 | Web_NP5410_6.0 | NY-Analytical | 762 | Seal Strength | Destruct Bond | Destruct Bond | Break occurs off the seal (Done) | g/in |  |  |  | TM 107 | 2JC699/Web_NP5410_6.0 | 6 | both
2JC699/Web_NP5410_6.0 | 2JC699 | Web_NP5410_6.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/40psi/0.5 sec dewll) | Seal Strength (@300°F/40psi/0.5 sec dewll) | 1686.57 | g/in | FDA,0 | 1500 |  | TM 107 | 2JC699/Web_NP5410_6.0 | 6 | both
2LB406/Web_NP3870_10.0 | 2LB406 | Web_NP3870_10.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2LB406/Web_NP3870_10.0 | 1 | both
2LB406/Web_NP3870_10.0 | 2LB406 | Web_NP3870_10.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 10 | in | FDA,4 | 9.9375 | 10.0625 | Measurement | 2LB406/Web_NP3870_10.0 | 1 | both
2LB406/Web_NP3870_10.0 | 2LB406 | Web_NP3870_10.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 93.53 | #/Ream | FDA,1 | 83.7 | 96.3 | TM 101 | 2LB406/Web_NP3870_10.0 | 1 | both
2LB406/Web_NP3870_10.0 | 2LB406 | Web_NP3870_10.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@350DegF/40psi/1.0 sec dwell) | Seal Strength (@350°F/40psi/1.0 sec dwell) | 10196.54 | g/in | FDA,0 | 5000 |  | TM 107 | 2LB406/Web_NP3870_10.0 | 1 | both
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width |  | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual |  |  | left_only
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length |  | in | FDA,4 | 3.6875 | 3.8125 | TM 143 / Visual |  |  | left_only
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection |  |  | left_only
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Solo Lid Inspection |  |  | left_only
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 818 | Solo Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness |  | in | FDA,3 | 0.025 | 0.035 | Solo Lid Inspection |  |  | left_only
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 818 | Solo Lid Inspection | Overall closed lid height | Overall closed lid height |  | in | FDA,3 | 0.355 | 0.385 | Solo Lid Inspection |  |  | left_only
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 818 | Solo Lid Inspection | Open Lid Overall Width | Open Lid Overall Width |  | in | FDA,3 | 4.351 | 4.401 | Solo Lid Inspection |  |  | left_only
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 818 | Solo Lid Inspection | Weight | Weight |  | g | FDA,1 | 10.5 | 12.5 | Measurement |  |  | left_only
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 818 | Solo Lid Inspection | Appearance | Appearance | Meet specification (Done) |  |  |  |  | Solo Lid Inspection |  |  | left_only
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 818 | Solo Lid Inspection | Functionality | Functionality | Meet specification (Done) |  |  |  |  | Solo Lid Inspection |  |  | left_only
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 818 | Solo Lid Inspection | Color | Color | Meets the specification (Done) |  |  |  |  | Solo Lid Inspection |  |  | left_only
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 818 | Solo Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | Solo Lid Inspection |  |  | left_only
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Lid_NY_NP7600_22S | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width |  | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual |  |  | left_only
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Lid_NY_NP7600_22S | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length |  | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual |  |  | left_only
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Lid_NY_NP7600_22S | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Lid_NY_NP7600_22S | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection |  |  | left_only
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Lid_NY_NP7600_22S | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Solo Lid Inspection |  |  | left_only
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Lid_NY_NP7600_22S | NY-Analytical | 818 | Solo Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness |  | in | FDA,3 | 0.025 | 0.035 | Solo Lid Inspection |  |  | left_only
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Lid_NY_NP7600_22S | NY-Analytical | 818 | Solo Lid Inspection | Overall closed lid height | Overall closed lid height |  | in | FDA,3 | 0.355 | 0.385 | Solo Lid Inspection |  |  | left_only
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Lid_NY_NP7600_22S | NY-Analytical | 818 | Solo Lid Inspection | Open Lid Overall Width | Open Lid Overall Width |  | in | FDA,3 | 4.351 | 4.401 | Solo Lid Inspection |  |  | left_only
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Lid_NY_NP7600_22S | NY-Analytical | 818 | Solo Lid Inspection | Weight | Weight |  | g | FDA,1 | 10.5 | 12.5 | Measurement |  |  | left_only
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Lid_NY_NP7600_22S | NY-Analytical | 818 | Solo Lid Inspection | Appearance | Appearance | Meet specification (Done) |  |  |  |  | Solo Lid Inspection |  |  | left_only
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Lid_NY_NP7600_22S | NY-Analytical | 818 | Solo Lid Inspection | Functionality | Functionality | Meet specification (Done) |  |  |  |  | Solo Lid Inspection |  |  | left_only
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Lid_NY_NP7600_22S | NY-Analytical | 818 | Solo Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | Solo Lid Inspection |  |  | left_only
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Lid_NY_NP7600_22S | NY-Analytical | 818 | Solo Lid Inspection | Color | Color | Teal Pantone 3145C |  |  |  |  | Solo Lid Inspection |  |  | left_only
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection | 2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 5 | both
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Solo Lid Inspection | 2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 5 | both
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 818 | Solo Lid Inspection | Color | Color | Meets the specification (Done) |  |  |  |  | Solo Lid Inspection | 2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 5 | both
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 818 | Solo Lid Inspection | Functionality | Functionality | Meet specification (Done) |  |  |  |  | Solo Lid Inspection | 2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 5 | both
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 818 | Solo Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.027 | in | FDA,3 | 0.025 | 0.035 | Solo Lid Inspection | 2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 5 | both
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 818 | Solo Lid Inspection | Open Lid Overall Width | Open Lid Overall Width | 4.358 | in | FDA,3 | 4.351 | 4.401 | Solo Lid Inspection | 2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 5 | both
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 818 | Solo Lid Inspection | Overall closed lid height | Overall closed lid height | 0.3785 | in | FDA,3 | 0.355 | 0.385 | Solo Lid Inspection | 2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 5 | both
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 818 | Solo Lid Inspection | Weight | Weight | 11.235 | g | FDA,1 | 10.5 | 12.5 | Measurement | 2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 5 | both
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | 2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 5 | both
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual | 2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 5 | both
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.69 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | 2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 5 | both
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.83 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | 2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 5 | both
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual | 2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 5 | both
2OD024/Web_NY_Skip lot testing | 2OD024 | Web_NY_Skip lot testing | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Printed/Varnished) | Basis Weight |  | #/Ream | FDA,1 | 55.6 | 68 | TM 101 |  |  | left_only
2OD024/Web_NY_Skip lot testing | 2OD024 | Web_NY_Skip lot testing | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@325DegF/40PSI/0.5Sec_Dwell) | Seal Strength (@325°F/40PSI/0.5Sec Dwell) |  | g/in | FDA,0 | 1500 |  | TM 107 |  |  | left_only
2OJ124SLI/Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | 2OJ124SLI | Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width |  | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual |  |  | left_only
2OJ124SLI/Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | 2OJ124SLI | Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length |  | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual |  |  | left_only
2OJ124SLI/Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | 2OJ124SLI | Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2OJ124SLI/Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | 2OJ124SLI | Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2OJ124SLI/Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | 2OJ124SLI | Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection |  |  | left_only
2OJ124SLI/Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | 2OJ124SLI | Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2OJ124SLI/Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | 2OJ124SLI | Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection |  |  | left_only
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Lid_NY_NP7600_22Y | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width |  | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual |  |  | left_only
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Lid_NY_NP7600_22Y | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length |  | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual |  |  | left_only
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Lid_NY_NP7600_22Y | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Lid_NY_NP7600_22Y | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Lid_NY_NP7600_22Y | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection |  |  | left_only
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Lid_NY_NP7600_22Y | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Lid_NY_NP7600_22Y | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Solo Lid Inspection |  |  | left_only
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Lid_NY_NP7600_22Y | NY-Analytical | 818 | Solo Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness |  | in | FDA,3 | 0.025 | 0.035 | Solo Lid Inspection |  |  | left_only
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Lid_NY_NP7600_22Y | NY-Analytical | 818 | Solo Lid Inspection | Overall closed lid height | Overall closed lid height |  | in | FDA,3 | 0.355 | 0.385 | Solo Lid Inspection |  |  | left_only
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Lid_NY_NP7600_22Y | NY-Analytical | 818 | Solo Lid Inspection | Open Lid Overall Width | Open Lid Overall Width |  | in | FDA,3 | 4.351 | 4.401 | Solo Lid Inspection |  |  | left_only
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Lid_NY_NP7600_22Y | NY-Analytical | 818 | Solo Lid Inspection | Weight | Weight |  | g | FDA,1 | 10.5 | 12.5 | Measurement |  |  | left_only
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Lid_NY_NP7600_22Y | NY-Analytical | 818 | Solo Lid Inspection | Functionality | Functionality | Meet specification (Done) |  |  |  |  | Solo Lid Inspection |  |  | left_only
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Lid_NY_NP7600_22Y | NY-Analytical | 818 | Solo Lid Inspection | Color | Color | Meets the specification (Done) |  |  |  |  | Solo Lid Inspection |  |  | left_only
2OJ213SLI/Lid_01A | 2OJ213SLI | Lid_01A | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length |  | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual |  |  | left_only
2OJ213SLI/Lid_01A | 2OJ213SLI | Lid_01A | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width |  | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual |  |  | left_only
2OJ213SLI/Lid_01A | 2OJ213SLI | Lid_01A | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2OJ213SLI/Lid_01A | 2OJ213SLI | Lid_01A | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2OJ213SLI/Lid_01A | 2OJ213SLI | Lid_01A | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection |  |  | left_only
2OJ213SLI/Lid_01A | 2OJ213SLI | Lid_01A | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2OJ213SLI/Lid_01A | 2OJ213SLI | Lid_01A | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection |  |  | left_only
2OM304/Web_NP8540_16.0 | 2OM304 | Web_NP8540_16.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2OM304/Web_NP8540_16.0 | 1 | both
2OM304/Web_NP8540_16.0 | 2OM304 | Web_NP8540_16.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 16 | in | FDA,4 | 15.9375 | 16.0625 | Measurement | 2OM304/Web_NP8540_16.0 | 1 | both
2OM304/Web_NP8540_16.0 | 2OM304 | Web_NP8540_16.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 53.176 | #/Ream | FDA,1 | 45.4 | 57.8 | TM 101 | 2OM304/Web_NP8540_16.0 | 1 | both
2OM304/Web_NP8540_16.0 | 2OM304 | Web_NP8540_16.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 7038.28 | g/in | FDA,0 | 3500 |  | TM 107 | 2OM304/Web_NP8540_16.0 | 1 | both
2OM304/Web_NP8540_16.0 | 2OM304 | Web_NP8540_16.0 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 3.3 | mil | FDA,2 | 2.86 | 3.64 | TM 132 | 2OM304/Web_NP8540_16.0 | 1 | both
2OM304SLI/Lid_No-Rinse sanitizing Solo | 2OM304SLI | Lid_No-Rinse sanitizing Solo | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width |  | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual |  |  | left_only
2OM304SLI/Lid_No-Rinse sanitizing Solo | 2OM304SLI | Lid_No-Rinse sanitizing Solo | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length |  | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual |  |  | left_only
2OM304SLI/Lid_No-Rinse sanitizing Solo | 2OM304SLI | Lid_No-Rinse sanitizing Solo | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2OM304SLI/Lid_No-Rinse sanitizing Solo | 2OM304SLI | Lid_No-Rinse sanitizing Solo | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2OM304SLI/Lid_No-Rinse sanitizing Solo | 2OM304SLI | Lid_No-Rinse sanitizing Solo | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection |  |  | left_only
2OM304SLI/Lid_No-Rinse sanitizing Solo | 2OM304SLI | Lid_No-Rinse sanitizing Solo | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2OM304SLI/Lid_No-Rinse sanitizing Solo | 2OM304SLI | Lid_No-Rinse sanitizing Solo | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection |  |  | left_only
2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | 2OM701SLI | Lid_No-Rinse sanitizing Solo_01 | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection | 2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | 1 | both
2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | 2OM701SLI | Lid_No-Rinse sanitizing Solo_01 | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | 2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | 1 | both
2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | 2OM701SLI | Lid_No-Rinse sanitizing Solo_01 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | 2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | 1 | both
2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | 2OM701SLI | Lid_No-Rinse sanitizing Solo_01 | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual | 2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | 1 | both
2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | 2OM701SLI | Lid_No-Rinse sanitizing Solo_01 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.687 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | 2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | 1 | both
2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | 2OM701SLI | Lid_No-Rinse sanitizing Solo_01 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.8125 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | 2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | 1 | both
2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | 2OM701SLI | Lid_No-Rinse sanitizing Solo_01 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual | 2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | 1 | both
2OP037/Web_NP8511_245 | 2OP037 | Web_NP8511_245 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2OP037/Web_NP8511_245 | 3 | both
2OP037/Web_NP8511_245 | 2OP037 | Web_NP8511_245 | NY-Analytical | 761 | Web Measurement | Web Width (mm) | Web Width | 245 | mm | FDA,0 | 243 | 247 | Measurement | 2OP037/Web_NP8511_245 | 3 | both
2OP037/Web_NP8511_245 | 2OP037 | Web_NP8511_245 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 47.88 | Lbs/Ream | FDA,1 | 43.1 | 52.7 | TM 101 | 2OP037/Web_NP8511_245 | 3 | both
2OP037/Web_NP8511_245 | 2OP037 | Web_NP8511_245 | NY-Analytical | 1048 | Seal strength (NP8511) | Seal strength 2500g/in@140DegC/1sec/14psi | Seal Strength (@140°C/1sec/14psi | 6296.05 | g/in | FDA,0 | 2500 |  | TM 107 | 2OP037/Web_NP8511_245 | 3 | both
2OP037/Web_NP8511_245 | 2OP037 | Web_NP8511_245 | NY-Analytical | 763 | Web Thickness | Thickness | Thickness | 71 | micron | FDA,1 | 64.8 | 79.2 | TM 132 | 2OP037/Web_NP8511_245 | 3 | both
2OP037/Web_NP8511_245 | 2OP037 | Web_NP8511_245 | NY-Analytical | 885 | Web Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Sample(s) all meet specifications (Done) |  |  |  |  | Visual Observation | 2OP037/Web_NP8511_245 | 3 | both
2OP063/Web_NP8511_320 | 2OP063 | Web_NP8511_320 | NY-Analytical | 836 | Dyne Level | Dyne Level (outside film) | Dyne Level (outside film) | 36 |  | FDA,0 |  |  | Dyne Level | 2OP063/Web_NP8511_320 | 1 | both
2OP063/Web_NP8511_320 | 2OP063 | Web_NP8511_320 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2OP063/Web_NP8511_320 | 1 | both
2OP063/Web_NP8511_320 | 2OP063 | Web_NP8511_320 | NY-Analytical | 761 | Web Measurement | Web Width (mm) | Web Width | 320 | mm | FDA,0 | 318 | 322 | Measurement | 2OP063/Web_NP8511_320 | 1 | both
2OP063/Web_NP8511_320 | 2OP063 | Web_NP8511_320 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 47.8 | Lbs/Ream | FDA,1 | 43.1 | 52.7 | TM 101 | 2OP063/Web_NP8511_320 | 1 | both
2OP063/Web_NP8511_320 | 2OP063 | Web_NP8511_320 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/inch@140C/14psi/1 s dwellTopJaw) | Seal Strength (@140°C/14psi/1 s dwellTopJaw) | 7207.05 | g/in | FDA,0 | 2500 |  | TM 107 | 2OP063/Web_NP8511_320 | 1 | both
2OP063/Web_NP8511_320 | 2OP063 | Web_NP8511_320 | NY-Analytical | 1048 | Seal strength (NP8511) | Seal strength 2500g/in@140DegC/1sec/14psi | Seal Strength (@140°C/1sec/14psi | 5948 | g/in | FDA,0 | 2500 |  | TM 107 | 2OP063/Web_NP8511_320 | 1 | both
2OP063/Web_NP8511_320 | 2OP063 | Web_NP8511_320 | NY-Analytical | 763 | Web Thickness | Thickness | Thickness | 74 | micron | FDA,1 | 64.8 | 79.2 | TM 132 | 2OP063/Web_NP8511_320 | 1 | both
2OP065/Web_NP8536_11.875 | 2OP065 | Web_NP8536_11.875 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2OP065/Web_NP8536_11.875 | 1 | both
2OP065/Web_NP8536_11.875 | 2OP065 | Web_NP8536_11.875 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 11.875 | in | FDA,4 | 11.8125 | 11.9375 | Measurement | 2OP065/Web_NP8536_11.875 | 1 | both
2OP065/Web_NP8536_11.875 | 2OP065 | Web_NP8536_11.875 | NY-Analytical | 1044 | Basis Weight (lbs/ream) | Basis Weight | Basis Weight | 41.4 | Lb/Ream | FDA,2 | 40.1 | 49.02 | TM 101 | 2OP065/Web_NP8536_11.875 | 1 | both
2OP065/Web_NP8536_11.875 | 2OP065 | Web_NP8536_11.875 | NY-Analytical | 1049 | Seal strength (NP8536) | Seal Strength (g/inch@320DegF/40psi/0.5 sec dwell) | Seal Strength (@320°F/40psi/0.5 sec dwell) | 8624 | g/in | FDA,0 | 2000 |  | TM 107 | 2OP065/Web_NP8536_11.875 | 1 | both
2OP065/Web_NP8536_11.875 | 2OP065 | Web_NP8536_11.875 | NY-Analytical | 763 | Web Thickness | Thickness | Thickness | 2.4 | mil | FDA,3 | 2.367 | 2.893 | TM 132 | 2OP065/Web_NP8536_11.875 | 1 | both
2OP065/Web_NP8536_11.875 | 2OP065 | Web_NP8536_11.875 | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | Visual | 2OP065/Web_NP8536_11.875 | 1 | both
2OP065/Web_NP8536_11.875 | 2OP065 | Web_NP8536_11.875 | NY-Analytical | 748 | Visual Inspection | Barcode Legibility/Quality | Barcode Legibility/Quality | Barcode should scan and code should be matched under the barcode (Done) |  |  |  |  | Visual | 2OP065/Web_NP8536_11.875 | 1 | both
2OP065/Web_NP8536_11.875 | 2OP065 | Web_NP8536_11.875 | NY-Analytical | 748 | Visual Inspection | Color/Copy | Color/Copy | Meets specifications (Done) |  |  |  |  | Visual | 2OP065/Web_NP8536_11.875 | 1 | both
2OP067/Web_NY_Skip lot testing | 2OP067 | Web_NY_Skip lot testing | NY-Analytical | 1044 | Basis Weight (lbs/ream) | Basis Weight | Basis Weight | 42.86 | Lb/Ream | FDA,2 | 40.1 | 49.02 | TM 101 | 2OP067/Web_NY_Skip lot testing | 1 | both
2OP067/Web_NY_Skip lot testing | 2OP067 | Web_NY_Skip lot testing | NY-Analytical | 1049 | Seal strength (NP8536) | Seal Strength (g/inch@320DegF/40psi/0.5 sec dwell) | Seal Strength (@320°F/40psi/0.5 sec dwell) | 9089 | g/in | FDA,0 | 2000 |  | TM 107 | 2OP067/Web_NY_Skip lot testing | 1 | both
2OP067/Web_NY_Skip lot testing | 2OP067 | Web_NY_Skip lot testing | NY-Analytical | 763 | Web Thickness | Thickness | Thickness | 2.6 | mil | FDA,3 | 2.367 | 2.893 | TM 132 | 2OP067/Web_NY_Skip lot testing | 1 | both
2OP067/Web_NY_Skip lot testing | 2OP067 | Web_NY_Skip lot testing | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | Visual | 2OP067/Web_NY_Skip lot testing | 1 | both
2OP067/Web_NY_Skip lot testing | 2OP067 | Web_NY_Skip lot testing | NY-Analytical | 748 | Visual Inspection | Barcode Legibility/Quality | Barcode Legibility/Quality | Barcode should scan and code should be matched under the barcode (Done) |  |  |  |  | Visual | 2OP067/Web_NY_Skip lot testing | 1 | both
2OP067/Web_NY_Skip lot testing | 2OP067 | Web_NY_Skip lot testing | NY-Analytical | 748 | Visual Inspection | Color/Copy | Color/Copy | Meets specifications (Done) |  |  |  |  | Visual | 2OP067/Web_NY_Skip lot testing | 1 | both
2OP072/Web_NP8511_245 | 2OP072 | Web_NP8511_245 | NY-Analytical | 836 | Dyne Level | Dyne Level (outside film) | Dyne Level (outside film) | 34 |  | FDA,0 |  |  | Dyne Level | 2OP072/Web_NP8511_245 | 1 | both
2OP073/Web_NP8536_11.875 | 2OP073 | Web_NP8536_11.875 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight |  | lbs/rm | FDA,2 |  |  | TM 101 | 2OP073/Web_NP8536_11.875 | 1 | both
2OP073/Web_NP8536_11.875 | 2OP073 | Web_NP8536_11.875 | NY-Analytical | 762 | Seal Strength | Seal Strength(g/in@350DegF/40psi/0.5sec dwell) | Seal Strength (@350°F/40psi/0.5sec dwell) |  | g/in | FDA,0 | 2500 |  | TM 107 | 2OP073/Web_NP8536_11.875 | 1 | both
2OP075/Web_NP8541_12.205 | 2OP075 | Web_NP8541_12.205 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  |  | left_only
2OP075/Web_NP8541_12.205 | 2OP075 | Web_NP8541_12.205 | NY-Analytical | 748 | Visual Inspection | Barcode Legibility/Quality | Barcode Legibility/Quality | Barcode should scan and code should be matched under the barcode (Done) |  |  |  |  | Visual |  |  | left_only
2OP075/Web_NP8541_12.205 | 2OP075 | Web_NP8541_12.205 | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | Visual |  |  | left_only
2OP075/Web_NP8541_12.205 | 2OP075 | Web_NP8541_12.205 | NY-Analytical | 748 | Visual Inspection | Color/Copy | Color/Copy | Meets specifications (Done) |  |  |  |  | Visual |  |  | left_only
2OP075/Web_NP8541_12.205 | 2OP075 | Web_NP8541_12.205 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight |  | GSM | FDA,1 | 74.3 | 90.9 | TM 101 |  |  | left_only
2OP075/Web_NP8541_12.205 | 2OP075 | Web_NP8541_12.205 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width |  | in | FDA,4 | 12.1425 | 12.2675 | Measurement |  |  | left_only
2OP075/Web_NP8541_12.205 | 2OP075 | Web_NP8541_12.205 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/inch@130DegC/0.5 sec/1 bar) | Seal Strength (@130°C/0.5 sec/1 bar) |  | g/in | FDA,0 | 1000 |  | TM 107 |  |  | left_only
2OP075/Web_NP8541_12.205 | 2OP075 | Web_NP8541_12.205 | NY-Analytical | 763 | Web Thickness | Thickness | Thickness |  | mm | FDA,3 | 0.074 | 0.094 | TM 132 |  |  | left_only
2OP092/Web_NP8536_10 | 2OP092 | Web_NP8536_10 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2OP092/Web_NP8536_10 | 1 | both
2OP092/Web_NP8536_10 | 2OP092 | Web_NP8536_10 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 10 | in | FDA,4 | 9.9375 | 10.0625 | Measurement | 2OP092/Web_NP8536_10 | 1 | both
2OP092/Web_NP8536_10 | 2OP092 | Web_NP8536_10 | NY-Analytical | 1044 | Basis Weight (lbs/ream) | Basis Weight | Basis Weight | 42.3 | Lb/Ream | FDA,2 | 40.1 | 49.02 | TM 101 | 2OP092/Web_NP8536_10 | 1 | both
2OP092/Web_NP8536_10 | 2OP092 | Web_NP8536_10 | NY-Analytical | 1049 | Seal strength (NP8536) | Seal Strength (g/inch@320DegF/40psi/0.5 sec dwell) | Seal Strength (@320°F/40psi/0.5 sec dwell) | 8116.37 | g/in | FDA,0 | 2000 |  | TM 107 | 2OP092/Web_NP8536_10 | 1 | both
2OP092/Web_NP8536_10 | 2OP092 | Web_NP8536_10 | NY-Analytical | 763 | Web Thickness | Thickness | Thickness | 2.8 | mil | FDA,3 | 2.367 | 2.893 | TM 132 | 2OP092/Web_NP8536_10 | 1 | both
2OP092/Web_NP8536_10 | 2OP092 | Web_NP8536_10 | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | Visual | 2OP092/Web_NP8536_10 | 1 | both
2OP092/Web_NP8536_10 | 2OP092 | Web_NP8536_10 | NY-Analytical | 748 | Visual Inspection | Barcode Legibility/Quality | Barcode Legibility/Quality | Barcode should scan and code should be matched under the barcode (Done) |  |  |  |  | Visual | 2OP092/Web_NP8536_10 | 1 | both
2OP092/Web_NP8536_10 | 2OP092 | Web_NP8536_10 | NY-Analytical | 748 | Visual Inspection | Color/Copy | Color/Copy | Meets specifications (Done) |  |  |  |  | Visual | 2OP092/Web_NP8536_10 | 1 | both
2OP093/Skip Lot Release_NY_2OP093 | 2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection |  |  | left_only
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis | 2OP093/ Jar _10 oz Blue HDPE Grade Jar | 12 | both
2OP093/Skip Lot Release_NY_2OP093 | 2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 765 | Canister & Lid Inspection | S Dimension | S Dimension |  | in | FDA,3 | 0.035 | 0.065 | Canister & Lid Inspection |  |  | left_only
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection | 2OP093/ Jar _10 oz Blue HDPE Grade Jar | 12 | both
2OP093/Skip Lot Release_NY_2OP093 | 2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 765 | Canister & Lid Inspection | H Dimension | H Dimension |  | in | FDA,3 | 0.457 | 0.487 | Canister & Lid Inspection |  |  | left_only
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Matches Specification (Done) |  |  |  |  | Canister & Lid Inspection | 2OP093/ Jar _10 oz Blue HDPE Grade Jar | 12 | both
2OP093/Skip Lot Release_NY_2OP093 | 2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 765 | Canister & Lid Inspection | E Diameter | E Diameter |  | in | FDA,3 | 2.509 | 2.539 | Canister & Lid Inspection |  |  | left_only
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | Diameter | Diameter | 2.746 | in | FDA,3 | 2.709 | 2.803 | Canister & Lid Inspection | 2OP093/ Jar _10 oz Blue HDPE Grade Jar | 12 | both
2OP093/Skip Lot Release_NY_2OP093 | 2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 765 | Canister & Lid Inspection | T Diameter | T Diameter |  | in | FDA,3 | 2.593 | 2.623 | Canister & Lid Inspection |  |  | left_only
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | E Diameter | E Diameter | 2.51 | in | FDA,3 | 2.509 | 2.539 | Canister & Lid Inspection | 2OP093/ Jar _10 oz Blue HDPE Grade Jar | 12 | both
2OP093/Skip Lot Release_NY_2OP093 | 2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 765 | Canister & Lid Inspection | Diameter | Diameter |  | in | FDA,3 | 2.709 | 2.803 | Canister & Lid Inspection |  |  | left_only
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) |  |  |  |  | Canister & Lid Inspection | 2OP093/ Jar _10 oz Blue HDPE Grade Jar | 12 | both
2OP093/Skip Lot Release_NY_2OP093 | 2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 765 | Canister & Lid Inspection | Height | Height |  | in | FDA,3 | 3.45 | 3.544 | Canister & Lid Inspection |  |  | left_only
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | H Dimension | H Dimension | 0.458 | in | FDA,3 | 0.457 | 0.487 | Canister & Lid Inspection | 2OP093/ Jar _10 oz Blue HDPE Grade Jar | 12 | both
2OP093/Skip Lot Release_NY_2OP093 | 2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight |  | g | FDA,2 | 24.5 | 27.5 | Measurement |  |  | left_only
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | Height | Height | 3.52 | in | FDA,3 | 3.45 | 3.544 | Canister & Lid Inspection | 2OP093/ Jar _10 oz Blue HDPE Grade Jar | 12 | both
2OP093/Skip Lot Release_NY_2OP093 | 2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Matches Specification (Done) |  |  |  |  | Canister & Lid Inspection |  |  | left_only
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | Canister & Lid Inspection | 2OP093/ Jar _10 oz Blue HDPE Grade Jar | 12 | both
2OP093/Skip Lot Release_NY_2OP093 | 2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection |  |  | left_only
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | S Dimension | S Dimension | 0.053 | in | FDA,3 | 0.035 | 0.065 | Canister & Lid Inspection | 2OP093/ Jar _10 oz Blue HDPE Grade Jar | 12 | both
2OP093/Skip Lot Release_NY_2OP093 | 2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | Canister & Lid Inspection |  |  | left_only
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | T Diameter | T Diameter | 2.6 | in | FDA,3 | 2.593 | 2.623 | Canister & Lid Inspection | 2OP093/ Jar _10 oz Blue HDPE Grade Jar | 12 | both
2OP093/Skip Lot Release_NY_2OP093 | 2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) |  |  |  |  | Canister & Lid Inspection |  |  | left_only
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight | 26.19 | g | FDA,2 | 24.5 | 27.5 | Measurement | 2OP093/ Jar _10 oz Blue HDPE Grade Jar | 12 | both
2OP093/Skip Lot Release_NY_2OP093 | 2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA |  |  | left_only
2OP093LI/Lid_NY_Threaded Blue Cap | 2OP093LI | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid fits with a tight seal on |  |  |  |  | Canister & Lid Inspection | 2OP093LI/Lid_NY_Threaded Blue Cap | 3 | both
2OP093LI/Lid_NY_Threaded Blue Cap | 2OP093LI | Lid_NY_Threaded Blue Cap | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | 2OP093LI/Lid_NY_Threaded Blue Cap | 3 | both
2OP093LI/Lid_NY_Threaded Blue Cap | 2OP093LI | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | TM 143 / Visual | 2OP093LI/Lid_NY_Threaded Blue Cap | 3 | both
2OP093LI/Lid_NY_Threaded Blue Cap | 2OP093LI | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Matches Specification (Done) |  |  |  |  | TM 143 / Visual | 2OP093LI/Lid_NY_Threaded Blue Cap | 3 | both
2OP093LI/Lid_NY_Threaded Blue Cap | 2OP093LI | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | Diameter | Diameter | 2.7515 | in | FDA,3 | 2.748 | 2.764 | TM 143 / Visual | 2OP093LI/Lid_NY_Threaded Blue Cap | 3 | both
2OP093LI/Lid_NY_Threaded Blue Cap | 2OP093LI | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | E Diameter | E Diameter | 2.5495 | in | FDA,3 | 2.548 | 2.562 | TM 143 / Visual | 2OP093LI/Lid_NY_Threaded Blue Cap | 3 | both
2OP093LI/Lid_NY_Threaded Blue Cap | 2OP093LI | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | Height | Height | 0.5625 | in | FDA,3 | 0.56 | 0.58 | TM 143 / Visual | 2OP093LI/Lid_NY_Threaded Blue Cap | 3 | both
2OP093LI/Lid_NY_Threaded Blue Cap | 2OP093LI | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual | 2OP093LI/Lid_NY_Threaded Blue Cap | 3 | both
2OP093LI/Lid_NY_Threaded Blue Cap | 2OP093LI | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | T Diameter | T Diameter | 2.619 | in | FDA,3 | 2.619 | 2.639 | TM 143 / Visual | 2OP093LI/Lid_NY_Threaded Blue Cap | 3 | both
2OP093LI/Lid_NY_Threaded Blue Cap | 2OP093LI | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 10.465 | g | FDA,2 |  |  | TM 143 / Visual | 2OP093LI/Lid_NY_Threaded Blue Cap | 3 | both
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Cap Closure | Functionality - Cap Closure | Cap of lid |  |  |  |  | Canister & Lid Inspection | 2OP134/Reduced Testing_NY_Can 2OP134 | 8 | both
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid fits with a tight seal on |  |  |  |  | Canister & Lid Inspection | 2OP134/Reduced Testing_NY_Can 2OP134 | 8 | both
2OP134/Skip Lot Release_NY_2OP134 | 2OP134 | Skip Lot Release_NY_2OP134 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA | 2OP134/Skip Lot Release_NY_2OP134 | 8 | both
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA | 2OP134/Reduced Testing_NY_Can 2OP134 | 8 | both
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | 2OP134/Reduced Testing_NY_Can 2OP134 | 8 | both
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection | 2OP134/Reduced Testing_NY_Can 2OP134 | 8 | both
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | White (Done) |  |  |  |  | Canister & Lid Inspection | 2OP134/Reduced Testing_NY_Can 2OP134 | 8 | both
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Bead OD | Bead OD | 3.77 | in | FDA,3 | 3.768 | 3.792 | Canister & Lid Inspection | 2OP134/Reduced Testing_NY_Can 2OP134 | 8 | both
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Canister Height OD | Canister Height OD | 6.66 | in | FDA,3 | 6.608 | 6.732 | Canister & Lid Inspection | 2OP134/Reduced Testing_NY_Can 2OP134 | 8 | both
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Lip (width) | Lip Width | 0.138 | in | FDA,3 | 0.127 | 0.143 | Canister & Lid Inspection | 2OP134/Reduced Testing_NY_Can 2OP134 | 8 | both
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Neck OD | Neck OD | 3.68 | in | FDA,3 | 3.673 | 3.697 | Canister & Lid Inspection | 2OP134/Reduced Testing_NY_Can 2OP134 | 8 | both
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | HDPE Resin Identification Marker | Meets the Specification (Done) |  |  |  |  | Canister & Lid Inspection | 2OP134/Reduced Testing_NY_Can 2OP134 | 8 | both
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | Canister & Lid Inspection | 2OP134/Reduced Testing_NY_Can 2OP134 | 8 | both
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight | 53.94 | g | FDA,1 | 52 | 56 | Measurement | 2OP134/Reduced Testing_NY_Can 2OP134 | 8 | both
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 766 | Lid Pull Force | Lid Pull-Off Force | Lid Pull-Off Force | 2997 | g | FDA,0 | 2000 | 3500 | TM 148 | 2OP134/Reduced Testing_NY_Can 2OP134 | 8 | both
2OP138/Skip Lot Relrase_NY_2OP138C | 2OP138 | Skip Lot Relrase_NY_2OP138C | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA | 2OP138/Skip Lot Relrase_NY_2OP138C | 1 | both
2OP138/Skip Lot Relrase_NY_2OP638 | 2OP138 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA | 2OP138/Skip Lot Relrase_NY_2OP638 | 2 | both
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA | 2OP138/Reduced Testing_NY_Can 2OP138 | 9 | both
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | 2OP138/Reduced Testing_NY_Can 2OP138 | 9 | both
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | [5 empty] | Canister & Lid Inspection | 2OP138/Reduced Testing_NY_Can 2OP138 | 9 | both
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | [5 empty] | Canister & Lid Inspection | 2OP138/Reduced Testing_NY_Can 2OP138 | 9 | both
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Bead OD | Bead OD |  | in | FDA,3 | 4.922 | 4.958 | Canister & Lid Inspection | 2OP138/Reduced Testing_NY_Can 2OP138 | 9 | both
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Canister OD | Canister OD |  | in | FDA,3 | 4.956 | 5.076 | Canister & Lid Inspection | 2OP138/Reduced Testing_NY_Can 2OP138 | 9 | both
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Neck OD | Neck OD |  | in | FDA,3 | 4.824 | 4.86 | Canister & Lid Inspection | 2OP138/Reduced Testing_NY_Can 2OP138 | 9 | both
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Total Height OD | Total Height OD |  | in | FDA,3 | 6.745 | 6.865 | Canister & Lid Inspection | 2OP138/Reduced Testing_NY_Can 2OP138 | 9 | both
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | [5 empty] | Canister & Lid Inspection | 2OP138/Reduced Testing_NY_Can 2OP138 | 9 | both
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | HDPE Resin Identification Marker | [5 empty] | Canister & Lid Inspection | 2OP138/Reduced Testing_NY_Can 2OP138 | 9 | both
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Lip Width (for reference only) | Lip Width (for reference only) |  | in | FDA,3 | 0.14 |  | Canister & Lid Inspection | 2OP138/Reduced Testing_NY_Can 2OP138 | 9 | both
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | [5 empty] | Canister & Lid Inspection | 2OP138/Reduced Testing_NY_Can 2OP138 | 9 | both
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Wall Thickness (Bottom base area) | Wall Thickness (Bottom base area) |  | in | FDA,3 | 0.01 |  | Canister & Lid Inspection | 2OP138/Reduced Testing_NY_Can 2OP138 | 9 | both
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Wall Thickness (Top area) | Wall Thickness (Top area) |  | in | FDA,3 | 0.02 |  | Canister & Lid Inspection | 2OP138/Reduced Testing_NY_Can 2OP138 | 9 | both
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight |  | g | FDA,1 | 66.5 | 73.5 | Measurement | 2OP138/Reduced Testing_NY_Can 2OP138 | 9 | both
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 1052 | Top Load Compression Strength | Top load Compression Strength (0.5"/mSR,0.450"TD) | Top load Compression Strength (0.5"/mSR,0.450"TD) | 150.3 | lb | FDA,0 | 84 |  | NPTM 1019 | 2OP138/Reduced Testing_NY_Can 2OP138 | 9 | both
2OP138/Skip Lot Relrase_NY_2OP638 | 2OP138 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | [5 empty] | TM 143 / Visual | 2OP138/Skip Lot Relrase_NY_2OP638 | 2 | both
2OP138/Skip Lot Relrase_NY_2OP638 | 2OP138 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | White (Done) |  |  |  |  | TM 143 / Visual | 2OP138/Skip Lot Relrase_NY_2OP638 | 2 | both
2OP138/Skip Lot Relrase_NY_2OP638 | 2OP138 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | [5 empty] | TM 143 / Visual | 2OP138/Skip Lot Relrase_NY_2OP638 | 2 | both
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 1053 | Cavity | cavity Inserts | cavity Inserts | [5 empty] | Visual | 2OP138/Reduced Testing_NY_Can 2OP138 | 9 | both
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection |  |  | left_only
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Total Height OD | Total Height OD |  | in | FDA,3 | 6.745 | 6.865 | Canister & Lid Inspection |  |  | left_only
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | [5 empty] | Canister & Lid Inspection |  |  | left_only
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | [5 empty] | Canister & Lid Inspection |  |  | left_only
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Bead OD | Bead OD |  | in | FDA,3 | 4.922 | 4.958 | Canister & Lid Inspection |  |  | left_only
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Canister OD | Canister OD |  | in | FDA,3 | 4.956 | 5.076 | Canister & Lid Inspection |  |  | left_only
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Neck OD | Neck OD |  | in | FDA,3 | 4.824 | 4.86 | Canister & Lid Inspection |  |  | left_only
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | [5 empty] | Canister & Lid Inspection |  |  | left_only
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | HDPE Resin Identification Marker | [5 empty] | Canister & Lid Inspection |  |  | left_only
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | [5 empty] | Canister & Lid Inspection |  |  | left_only
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Wall Thickness (Bottom base area) | Wall Thickness (Bottom base area) |  | in | FDA,3 | 0.01 |  | Canister & Lid Inspection |  |  | left_only
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Wall Thickness (Top area) | Wall Thickness (Top area) |  | in | FDA,3 | 0.02 |  | Canister & Lid Inspection |  |  | left_only
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight |  | g | FDA,1 | 66.5 | 73.5 | Measurement |  |  | left_only
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 1052 | Top Load Compression Strength | Top load Compression Strength (0.5"/mSR,0.450"TD) | Top load Compression Strength (0.5"/mSR,0.450"TD) |  | lb | FDA,0 | 84 |  | NPTM 1019 |  |  | left_only
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 1053 | Cavity | cavity Inserts | cavity Inserts | [5 empty] | Visual |  |  | left_only
2OP159LI/Lid_03B | 2OP159LI | Lid_03B | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | 2OP159LI/Lid_03B | 1 | both
2OP159LI/Lid_03B | 2OP159LI | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | 2OP159LI/Lid_03B | 1 | both
2OP159LI/Lid_03B | 2OP159LI | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | Bead Diameter | 4.856 | in | FDA,3 | 4.847 | 4.867 | TM 143 / Visual | 2OP159LI/Lid_03B | 1 | both
2OP159LI/Lid_03B | 2OP159LI | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Cap Thickness | Cap Thickness | 0.049 | in | FDA,3 | 0.045 | 0.055 | TM 143 / Visual | 2OP159LI/Lid_03B | 1 | both
2OP159LI/Lid_03B | 2OP159LI | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Blue (Done) |  |  |  |  | TM 143 / Visual | 2OP159LI/Lid_03B | 1 | both
2OP159LI/Lid_03B | 2OP159LI | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) |  |  |  |  | TM 143 / Visual | 2OP159LI/Lid_03B | 1 | both
2OP159LI/Lid_03B | 2OP159LI | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Insert Fin Gap (per specification) | Insert Fin Gap | 0.07 | in | FDA,3 | 0.065 | 0.075 | TM 143 / Visual | 2OP159LI/Lid_03B | 1 | both
2OP159LI/Lid_03B | 2OP159LI | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter (as per specification) | Lid Inside Diameter | 4.929 | in | FDA,3 | 4.922 | 4.932 | TM 143 / Visual | 2OP159LI/Lid_03B | 1 | both
2OP159LI/Lid_03B | 2OP159LI | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.042 | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual | 2OP159LI/Lid_03B | 1 | both
2OP159LI/Lid_03B | 2OP159LI | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual | 2OP159LI/Lid_03B | 1 | both
2OP159LI/Lid_03B | 2OP159LI | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.486 | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual | 2OP159LI/Lid_03B | 1 | both
2OP159LI/Lid_03B | 2OP159LI | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Outer Diameter (per specification) | Pop-Up Hole Outer Diameter | 0.163 | in | FDA,3 | 0.16 | 0.17 | TM 143 / Visual | 2OP159LI/Lid_03B | 1 | both
2OP159LI/Lid_03B | 2OP159LI | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Thicknes | Pop-Up Hole Thicknes | 0.02 | in | FDA,3 | 0.012 | 0.022 | TM 143 / Visual | 2OP159LI/Lid_03B | 1 | both
2OP159LI/Lid_03B | 2OP159LI | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 34.335 | g | FDA,1 | 32 | 42 | TM 143 / Visual | 2OP159LI/Lid_03B | 1 | both
2OP213/Web_NP8541_11 | 2OP213 | Web_NP8541_11 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2OP213/Web_NP8541_11 | 1 | both
2OP213/Web_NP8541_11 | 2OP213 | Web_NP8541_11 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 11 | in | FDA,4 | 10.9375 | 11.0625 | Measurement | 2OP213/Web_NP8541_11 | 1 | both
2OP213/Web_NP8541_11 | 2OP213 | Web_NP8541_11 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 86.76 | GSM | FDA,1 | 74.3 | 90.9 | TM 101 | 2OP213/Web_NP8541_11 | 1 | both
2OP213/Web_NP8541_11 | 2OP213 | Web_NP8541_11 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/inch@130DegC/0.5 sec/1 bar) | Seal Strength (@130°C/0.5 sec/1 bar) | 8561 | g/in | FDA,0 | 1000 |  | TM 107 | 2OP213/Web_NP8541_11 | 1 | both
2OP213/Web_NP8541_11 | 2OP213 | Web_NP8541_11 | NY-Analytical | 763 | Web Thickness | Thickness | Thickness | 0.088 | mm | FDA,3 | 0.074 | 0.094 | TM 132 | 2OP213/Web_NP8541_11 | 1 | both
2OP213/Web_NP8541_11 | 2OP213 | Web_NP8541_11 | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | Visual | 2OP213/Web_NP8541_11 | 1 | both
2OP213/Web_NP8541_11 | 2OP213 | Web_NP8541_11 | NY-Analytical | 748 | Visual Inspection | Barcode Legibility/Quality | Barcode Legibility/Quality | Barcode should scan and code should be matched under the barcode (Done) |  |  |  |  | Visual | 2OP213/Web_NP8541_11 | 1 | both
2OP213/Web_NP8541_11 | 2OP213 | Web_NP8541_11 | NY-Analytical | 748 | Visual Inspection | Color/Copy | Color/Copy | Meets specifications (Done) |  |  |  |  | Visual | 2OP213/Web_NP8541_11 | 1 | both
2OP223/Web_NP8511_320 | 2OP223 | Web_NP8511_320 | NY-Analytical | 885 | Web Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Sample(s) all meet specifications (Done) |  |  |  |  | Visual Observation | 2OP223/Web_NP8511_320 | 4 | both
2OP242LI/Lid_17C | 2OP242LI | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Cap Closure | Functionality - Cap Closure | Cap of lid |  |  |  |  | Canister & Lid Inspection | 2OP242LI/Lid_17C | 4 | both
2OP242LI/Lid_17C | 2OP242LI | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid fits with a tight seal on |  |  |  |  | Canister & Lid Inspection | 2OP242LI/Lid_17C | 4 | both
2OP242LI/Lid_17C | 2OP242LI | Lid_17C | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Flip Lid Force | 2OP242LI/Lid_17C | 4 | both
2OP242LI/Lid_17C | 2OP242LI | Lid_17C | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | Flip Lid Force | 1.11 | lb | FDA,2 | 0.75 |  | TM 143 | 2OP242LI/Lid_17C | 4 | both
2OP242LI/Lid_17C | 2OP242LI | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | TM 143 / Visual | 2OP242LI/Lid_17C | 4 | both
2OP242LI/Lid_17C | 2OP242LI | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | Bead Diameter | 4.835 | in | FDA,3 | 4.831 | 4.851 | TM 143 / Visual | 2OP242LI/Lid_17C | 4 | both
2OP242LI/Lid_17C | 2OP242LI | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Translucent Plum (Done) |  |  |  |  | TM 143 / Visual | 2OP242LI/Lid_17C | 4 | both
2OP242LI/Lid_17C | 2OP242LI | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.045 | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual | 2OP242LI/Lid_17C | 4 | both
2OP242LI/Lid_17C | 2OP242LI | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual | 2OP242LI/Lid_17C | 4 | both
2OP242LI/Lid_17C | 2OP242LI | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Diameter | Orifice Diameter | 0.62 | in | FDA,3 | 0.62 | 0.63 | TM 143 / Visual | 2OP242LI/Lid_17C | 4 | both
2OP242LI/Lid_17C | 2OP242LI | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness (per specification) | Orifice Flap Thickness | 0.021 | in | FDA,3 | 0.017 | 0.023 | TM 143 / Visual | 2OP242LI/Lid_17C | 4 | both
2OP242LI/Lid_17C | 2OP242LI | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.483 | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual | 2OP242LI/Lid_17C | 4 | both
2OP242LI/Lid_17C | 2OP242LI | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 38.41 | g | FDA,1 | 32.5 | 41.5 | TM 143 / Visual | 2OP242LI/Lid_17C | 4 | both
2OP280LI/Lid_17A | 2OP280LI | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Cap Closure | Functionality - Cap Closure | Cap of lid |  |  |  |  | Canister & Lid Inspection | 2OP280LI/Lid_17A | 1 | both
2OP280LI/Lid_17A | 2OP280LI | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid fits with a tight seal on |  |  |  |  | Canister & Lid Inspection | 2OP280LI/Lid_17A | 1 | both
2OP280LI/Lid_17A | 2OP280LI | Lid_17A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Flip Lid Force | 2OP280LI/Lid_17A | 1 | both
2OP280LI/Lid_17A | 2OP280LI | Lid_17A | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | Flip Lid Force | 2.88 | lb | FDA,2 | 0.75 |  | TM 143 | 2OP280LI/Lid_17A | 1 | both
2OP280LI/Lid_17A | 2OP280LI | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | TM 143 / Visual | 2OP280LI/Lid_17A | 1 | both
2OP280LI/Lid_17A | 2OP280LI | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | Bead Diameter | 4.851 | in | FDA,3 | 4.831 | 4.851 | TM 143 / Visual | 2OP280LI/Lid_17A | 1 | both
2OP280LI/Lid_17A | 2OP280LI | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Translucent Orange (Done) |  |  |  |  | TM 143 / Visual | 2OP280LI/Lid_17A | 1 | both
2OP280LI/Lid_NY_Skip lot testing | 2OP280LI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Cap Closure | Functionality - Cap Closure | [5 empty] | TM 143 / Visual | 2OP280LI/Lid_NY_Skip lot testing | 17 | both
2OP280LI/Lid_NY_Skip lot testing | 2OP280LI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | [5 empty] | TM 143 / Visual | 2OP280LI/Lid_NY_Skip lot testing | 17 | both
2OP280LI/Lid_17A | 2OP280LI | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.035 | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual | 2OP280LI/Lid_17A | 1 | both
2OP280LI/Lid_17A | 2OP280LI | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual | 2OP280LI/Lid_17A | 1 | both
2OP280LI/Lid_17A | 2OP280LI | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Diameter | Orifice Diameter | 1.468 | in | FDA,3 | 1.468 | 1.478 | TM 143 / Visual | 2OP280LI/Lid_17A | 1 | both
2OP280LI/Lid_NY_Skip lot testing | 2OP280LI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Diameter | Orifice Diameter |  | in | FDA,3 | 1.468 | 1.478 | TM 143 / Visual | 2OP280LI/Lid_NY_Skip lot testing | 17 | both
2OP280LI/Lid_17A | 2OP280LI | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness (per specification) | Orifice Flap Thickness | 0.035 | in | FDA,3 | 0.03 | 0.04 | TM 143 / Visual | 2OP280LI/Lid_17A | 1 | both
2OP280LI/Lid_17A | 2OP280LI | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.484 | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual | 2OP280LI/Lid_17A | 1 | both
2OP280LI/Lid_17A | 2OP280LI | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 37.555 | g | FDA,1 | 36 | 42 | TM 143 / Visual | 2OP280LI/Lid_17A | 1 | both
2OP300LI/Lid_13E | 2OP300LI | Lid_13E | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection |  |  | left_only
2OP300LI/Lid_13E | 2OP300LI | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness (per specification) | Orifice Flap Thickness |  | in | FDA,3 | 0.03 | 0.04 | TM 143 / Visual |  |  | left_only
2OP300LI/Lid_13E | 2OP300LI | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness |  | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual |  |  | left_only
2OP300LI/Lid_13E | 2OP300LI | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Insert Fin Gap (per specification) | Insert Fin Gap |  | in | FDA,3 | 0.07 | 0.08 | TM 143 / Visual |  |  | left_only
2OP300LI/Lid_13E | 2OP300LI | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | End Circle Diameter | End Circle Diameter |  | in | FDA,3 | 0.12 | 0.13 | TM 143 / Visual |  |  | left_only
2OP300LI/Lid_13E | 2OP300LI | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Center to Center Distance | Center to Center Distance |  | in | FDA,3 | 0.445 | 0.455 | TM 143 / Visual |  |  | left_only
2OP300LI/Lid_13E | 2OP300LI | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height |  | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual |  |  | left_only
2OP300LI/Lid_13E | 2OP300LI | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | Bead Diameter |  | in | FDA,3 | 4.847 | 4.867 | TM 143 / Visual |  |  | left_only
2OP300LI/Lid_13E | 2OP300LI | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight |  | g | FDA,1 | 32 | 42 | TM 143 / Visual |  |  | left_only
2OP300LI/Lid_13E | 2OP300LI | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2OP300LI/Lid_13E | 2OP300LI | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2OP300LI/Lid_13E | 2OP300LI | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2OP300LI/Lid_13E | 2OP300LI | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Translucent Blue (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2OP300LI/Lid_13E | 2OP300LI | Lid_13E | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | Flip Lid Force |  | lb | FDA,2 | 0.75 |  | TM 148 |  |  | left_only
2OP435LI/Lid_04C_560 | 2OP435LI | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Thicknes | Pop-Up Hole Thicknes |  | in | FDA,3 | 0.015 | 0.025 | TM 143 / Visual |  |  | left_only
2OP435LI/Lid_04C_560 | 2OP435LI | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness |  | in | FDA,3 | 0.04 | 0.05 | TM 143 / Visual |  |  | left_only
2OP435LI/Lid_04C_560 | 2OP435LI | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Cap Thickness | Cap Thickness |  | in | FDA,3 | 0.147 | 0.157 | TM 143 / Visual |  |  | left_only
2OP435LI/Lid_04C_560 | 2OP435LI | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Outer Diameter (per specification) | Pop-Up Hole Outer Diameter |  | in | FDA,3 | 0.555 | 0.565 | TM 143 / Visual |  |  | left_only
2OP435LI/Lid_04C_560 | 2OP435LI | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height |  | in | FDA,3 | 1.473 | 1.493 | TM 143 / Visual |  |  | left_only
2OP435LI/Lid_04C_560 | 2OP435LI | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter (as per specification) | Lid Inside Diameter |  | in | FDA,3 | 3.76 | 3.77 | TM 143 / Visual |  |  | left_only
2OP435LI/Lid_04C_560 | 2OP435LI | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight |  | g | FDA,1 | 20 | 24 | TM 143 / Visual |  |  | left_only
2OP435LI/Lid_04C_560 | 2OP435LI | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2OP435LI/Lid_04C_560 | 2OP435LI | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2OP435LI/Lid_04C_560 | 2OP435LI | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2OP435LI/Lid_04C_560 | 2OP435LI | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Translucent Purple (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
2OP435LI/Lid_04C_560 | 2OP435LI | Lid_04C_560 | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection |  |  | left_only
2OP630/Web_NP8440_310 | 2OP630 | Web_NP8440_310 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  |  | left_only
2OP630/Web_NP8440_310 | 2OP630 | Web_NP8440_310 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight |  | #/Ream | FDA,2 | 44.88 | 57.12 | TM 101 |  |  | left_only
2OP630/Web_NP8440_310 | 2OP630 | Web_NP8440_310 | NY-Analytical | 761 | Web Measurement | Web Width (mm) | Web Width |  | mm | FDA,0 | 308 | 312 | Measurement |  |  | left_only
2OP630/Web_NP8440_310 | 2OP630 | Web_NP8440_310 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) |  | g/in | FDA,0 | 3500 |  | TM 107 |  |  | left_only
2OP630/Web_NP8440_310 | 2OP630 | Web_NP8440_310 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness |  | mil | FDA,0 | 2.64 | 3.36 | TM 132 |  |  | left_only
2OP638/Skip Lot Relrase_NY_2OP638 | 2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid has a tight friction fit |  |  |  |  | Canister & Lid Inspection | 2OP638/Skip Lot Relrase_NY_2OP638 | 242 | both
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA | 2OP638/Reduced Testing_NY_Can 2OP638 | 8 | both
2OP638/Skip Lot Relrase_NY_2OP638 | 2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis | 2OP638/Skip Lot Relrase_NY_2OP638 | 242 | both
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | 2OP638/Reduced Testing_NY_Can 2OP638 | 8 | both
2OP638/Skip Lot Relrase_NY_2OP638 | 2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Bead OD | Bead OD | 4.9245 | in | FDA,3 | 4.922 | 4.958 | Canister & Lid Inspection | 2OP638/Skip Lot Relrase_NY_2OP638 | 242 | both
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Bead OD | Bead OD | 4.9545 | in | FDA,3 | 4.922 | 4.958 | Canister & Lid Inspection | 2OP638/Reduced Testing_NY_Can 2OP638 | 8 | both
2OP638/Skip Lot Relrase_NY_2OP638 | 2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Canister OD | Canister OD | 5.03 | in | FDA,3 | 4.956 | 5.076 | Canister & Lid Inspection | 2OP638/Skip Lot Relrase_NY_2OP638 | 242 | both
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Canister OD | Canister OD | 5.031 | in | FDA,3 | 4.956 | 5.076 | Canister & Lid Inspection | 2OP638/Reduced Testing_NY_Can 2OP638 | 8 | both
2OP638/Skip Lot Relrase_NY_2OP638 | 2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Neck OD | Neck OD | 4.8545 | in | FDA,3 | 4.824 | 4.86 | Canister & Lid Inspection | 2OP638/Skip Lot Relrase_NY_2OP638 | 242 | both
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Neck OD | Neck OD | 4.8425 | in | FDA,3 | 4.824 | 4.86 | Canister & Lid Inspection | 2OP638/Reduced Testing_NY_Can 2OP638 | 8 | both
2OP638/Skip Lot Relrase_NY_2OP638 | 2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Total Height OD | Total Height OD | 8.75 | in | FDA,3 | 8.688 | 8.808 | Canister & Lid Inspection | 2OP638/Skip Lot Relrase_NY_2OP638 | 242 | both
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Total Height OD | Total Height OD | 8.75 | in | FDA,3 | 8.688 | 8.808 | Canister & Lid Inspection | 2OP638/Reduced Testing_NY_Can 2OP638 | 8 | both
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | [5 empty] | Canister & Lid Inspection | 2OP638/Reduced Testing_NY_Can 2OP638 | 8 | both
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | HDPE Resin Identification Marker | [5 empty] | Canister & Lid Inspection | 2OP638/Reduced Testing_NY_Can 2OP638 | 8 | both
2OP638/Skip Lot Relrase_NY_2OP638 | 2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | HDPE Resin Identification Marker | Meets the Specification (Done) |  |  |  |  | Canister & Lid Inspection | 2OP638/Skip Lot Relrase_NY_2OP638 | 242 | both
2OP638/Skip Lot Relrase_NY_2OP638 | 2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Lip Width (for reference only) | Lip Width (for reference only) | 0.154 | in | FDA,3 | 0.14 |  | Canister & Lid Inspection | 2OP638/Skip Lot Relrase_NY_2OP638 | 242 | both
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Lip Width (for reference only) | Lip Width (for reference only) | 0.15 | in | FDA,3 | 0.14 |  | Canister & Lid Inspection | 2OP638/Reduced Testing_NY_Can 2OP638 | 8 | both
2OP638/Skip Lot Relrase_NY_2OP638 | 2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Wall Thickness (Bottom base area) | Wall Thickness (Bottom base area) | 0.0325 | in | FDA,3 | 0.01 |  | Canister & Lid Inspection | 2OP638/Skip Lot Relrase_NY_2OP638 | 242 | both
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Wall Thickness (Bottom base area) | Wall Thickness (Bottom base area) |  | in | FDA,3 | 0.01 |  | Canister & Lid Inspection | 2OP638/Reduced Testing_NY_Can 2OP638 | 8 | both
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Wall Thickness (Top area) | Wall Thickness (Top area) |  | in | FDA,3 | 0.02 |  | Canister & Lid Inspection | 2OP638/Reduced Testing_NY_Can 2OP638 | 8 | both
2OP638/Skip Lot Relrase_NY_2OP638 | 2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Wall Thickness (Top area) | Wall Thickness (Top area) | 0.043 | in | FDA,3 | 0.02 |  | Canister & Lid Inspection | 2OP638/Skip Lot Relrase_NY_2OP638 | 242 | both
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight | 90.2 | g | FDA,1 | 85.5 | 94.5 | Measurement | 2OP638/Reduced Testing_NY_Can 2OP638 | 8 | both
2OP638/Skip Lot Relrase_NY_2OP638 | 2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight | 89.75 | g | FDA,1 | 85.5 | 94.5 | Measurement | 2OP638/Skip Lot Relrase_NY_2OP638 | 242 | both
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 1052 | Top Load Compression Strength | Top load Compression Strength (0.5"/mSR,0.450"TD) | Top load Compression Strength (0.5"/mSR,0.450"TD) |  | lb | FDA,0 | 97 |  | NPTM 1019 | 2OP638/Reduced Testing_NY_Can 2OP638 | 8 | both
2OP638/Skip Lot Relrase_NY_2OP638 | 2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 1052 | Top Load Compression Strength | Top load Compression Strength (0.5"/mSR,0.450"TD) | Top load Compression Strength (0.5"/mSR,0.450"TD) | 138.8 | lb | FDA,0 | 97 |  | NPTM 1019 | 2OP638/Skip Lot Relrase_NY_2OP638 | 242 | both
2OP638/Bulk_IN_4ABBY12 - 20190218 | 2OP638 | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 902 | Bulk Liquid Inspection | Appearance (TM 208) | Appearance | [5 empty] | See Analyte for Method # |  |  | left_only
2OP638/Bulk_IN_4ABBY12 - 20190218 | 2OP638 | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 902 | Bulk Liquid Inspection | Color (TM 208) | Color | [5 empty] | See Analyte for Method # |  |  | left_only
2OP638/Bulk_IN_4ABBY12 - 20190218 | 2OP638 | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 902 | Bulk Liquid Inspection | Contamination (TM 208) | Contamination | [5 empty] | See Analyte for Method # |  |  | left_only
2OP638/Bulk_IN_4ABBY12 - 20190218 | 2OP638 | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 902 | Bulk Liquid Inspection | Odor (TM 208) | Odor | [5 empty] | See Analyte for Method # |  |  | left_only
2OP638/Bulk_IN_4ABBY12 - 20190218 | 2OP638 | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 902 | Bulk Liquid Inspection | pH (TM 200) | pH |  |  | FDA,2 | 4.5 | 4.9 | See Analyte for Method # |  |  | left_only
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | TM 143 / Visual | 2OP638/Reduced Testing_NY_Can 2OP638 | 8 | both
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | White (Done) |  |  |  |  | TM 143 / Visual | 2OP638/Reduced Testing_NY_Can 2OP638 | 8 | both
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual | 2OP638/Reduced Testing_NY_Can 2OP638 | 8 | both
2OP638/Bulk_IN_4ABBY12 - 20190218 | 2OP638 | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 943 | %Phenoxyethanol & Sodium Benzoate | Phenoxyethanol | Phenoxyethanol |  | % | FDA,3 | 0.63 | 0.77 | TM 395 |  |  | left_only
2OP638/Bulk_IN_4ABBY12 - 20190218 | 2OP638 | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 943 | %Phenoxyethanol & Sodium Benzoate | Phenoxyethanol_Final | Phenoxyethanol Final |  | % | FDA,3 | 0.63 | 0.77 | TM 395 |  |  | left_only
2OP638/Bulk_IN_4ABBY12 - 20190218 | 2OP638 | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 943 | %Phenoxyethanol & Sodium Benzoate | Sodium_Benzoate | Sodium Benzoate |  | % | FDA,3 | 0.423 | 0.517 | TM 395 |  |  | left_only
2OP638/Bulk_IN_4ABBY12 - 20190218 | 2OP638 | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 943 | %Phenoxyethanol & Sodium Benzoate | Sodium_Benzoate_Final | Sodium Benzoate Final |  | % | FDA,3 | 0.423 | 0.517 | TM 395 |  |  | left_only
2OP638/Skip Lot Relrase_NY_2OP638 | 2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 1053 | Cavity | cavity Inserts | cavity Inserts | Meets specification (Done) |  |  |  |  | Visual | 2OP638/Skip Lot Relrase_NY_2OP638 | 242 | both
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 1053 | Cavity | cavity Inserts | cavity Inserts | [5 empty] | Visual | 2OP638/Reduced Testing_NY_Can 2OP638 | 8 | both
2OP638LI/Lid_13D | 2OP638LI | Lid_13D | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | 2OP638LI/Lid_13D | 3 | both
2OP638LI/Lid_13D | 2OP638LI | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | 2OP638LI/Lid_13D | 3 | both
2OP638LI/Lid_13D | 2OP638LI | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | Bead Diameter | 4.8615 | in | FDA,3 | 4.847 | 4.867 | TM 143 / Visual | 2OP638LI/Lid_13D | 3 | both
2OP638LI/Lid_13D | 2OP638LI | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Center to Center Distance | Center to Center Distance | 0.4535 | in | FDA,3 | 0.445 | 0.455 | TM 143 / Visual | 2OP638LI/Lid_13D | 3 | both
2OP638LI/Lid_13D | 2OP638LI | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Translucent Gray (Done) |  |  |  |  | TM 143 / Visual | 2OP638LI/Lid_13D | 3 | both
2OP638LI/Lid_13D | 2OP638LI | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | End Circle Diameter | End Circle Diameter | 0.122 | in | FDA,3 | 0.12 | 0.13 | TM 143 / Visual | 2OP638LI/Lid_13D | 3 | both
2OP638LI/Lid_13D | 2OP638LI | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) |  |  |  |  | TM 143 / Visual | 2OP638LI/Lid_13D | 3 | both
2OP638LI/Lid_13D | 2OP638LI | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Insert Fin Gap (per specification) | Insert Fin Gap | 0.0745 | in | FDA,3 | 0.07 | 0.08 | TM 143 / Visual | 2OP638LI/Lid_13D | 3 | both
2OP638LI/Lid_13D | 2OP638LI | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter (as per specification) | Lid Inside Diameter | 4.9218 | in | FDA,3 | 4.917 | 4.937 | TM 143 / Visual | 2OP638LI/Lid_13D | 3 | both
2OP638LI/Lid_13D | 2OP638LI | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.041 | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual | 2OP638LI/Lid_13D | 3 | both
2OP638LI/Lid_13D | 2OP638LI | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual | 2OP638LI/Lid_13D | 3 | both
2OP638LI/Lid_13D | 2OP638LI | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness (per specification) | Orifice Flap Thickness | 0.0345 | in | FDA,3 | 0.03 | 0.04 | TM 143 / Visual | 2OP638LI/Lid_13D | 3 | both
2OP638LI/Lid_13D | 2OP638LI | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.487 | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual | 2OP638LI/Lid_13D | 3 | both
2OP638LI/Lid_13D | 2OP638LI | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 33.92 | g | FDA,1 | 32 | 42 | TM 143 / Visual | 2OP638LI/Lid_13D | 3 | both
2OP638LI/Lid_13D | 2OP638LI | Lid_13D | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | Flip Lid Force | 1.79 | lb | FDA,2 | 0.75 |  | TM 148 | 2OP638LI/Lid_13D | 3 | both
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid has a tight friction fit |  |  |  |  | Canister & Lid Inspection | 2OQ753/Reduced Testing_NY_Can 2OQ753 | 18 | both
2OQ753/Skip Lot Release_NY_2OQ753 | 2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid has a tight friction fit |  |  |  |  | Canister & Lid Inspection | 2OQ753/Skip Lot Release_NY_2OQ753 | 103 | both
2OQ753/Skip Lot Release_NY_2OQ753 | 2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA | 2OQ753/Skip Lot Release_NY_2OQ753 | 103 | both
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA | 2OQ753/Reduced Testing_NY_Can 2OQ753 | 18 | both
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis | 2OQ753/Reduced Testing_NY_Can 2OQ753 | 18 | both
2OQ753/Skip Lot Release_NY_2OQ753 | 2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | 2OQ753/Skip Lot Release_NY_2OQ753 | 103 | both
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection | 2OQ753/Reduced Testing_NY_Can 2OQ753 | 18 | both
2OQ753/Skip Lot Release_NY_2OQ753 | 2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection | 2OQ753/Skip Lot Release_NY_2OQ753 | 103 | both
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | White (Done) |  |  |  |  | Canister & Lid Inspection | 2OQ753/Reduced Testing_NY_Can 2OQ753 | 18 | both
2OQ753/Skip Lot Release_NY_2OQ753 | 2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | White (Done) |  |  |  |  | Canister & Lid Inspection | 2OQ753/Skip Lot Release_NY_2OQ753 | 103 | both
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Bead OD | Bead OD | 4.9355 | in | FDA,3 | 4.908 | 4.944 | Canister & Lid Inspection | 2OQ753/Reduced Testing_NY_Can 2OQ753 | 18 | both
2OQ753/Skip Lot Release_NY_2OQ753 | 2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Bead OD | Bead OD | 4.911 | in | FDA,3 | 4.908 | 4.944 | Canister & Lid Inspection | 2OQ753/Skip Lot Release_NY_2OQ753 | 103 | both
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Canister OD | Canister OD | 5.0145 | in | FDA,3 | 4.998 | 5.034 | Canister & Lid Inspection | 2OQ753/Reduced Testing_NY_Can 2OQ753 | 18 | both
2OQ753/Skip Lot Release_NY_2OQ753 | 2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Canister OD | Canister OD | 4.998 | in | FDA,3 | 4.998 | 5.034 | Canister & Lid Inspection | 2OQ753/Skip Lot Release_NY_2OQ753 | 103 | both
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Lip (width) | Lip Width | 0.17 | in | FDA,3 | 0.142 | 0.178 | Canister & Lid Inspection | 2OQ753/Reduced Testing_NY_Can 2OQ753 | 18 | both
2OQ753/Skip Lot Release_NY_2OQ753 | 2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Lip (width) | Lip Width | 0.15 | in | FDA,3 | 0.142 | 0.178 | Canister & Lid Inspection | 2OQ753/Skip Lot Release_NY_2OQ753 | 103 | both
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Total Height OD | Total Height OD | 6.748 | in | FDA,3 | 6.706 | 6.826 | Canister & Lid Inspection | 2OQ753/Reduced Testing_NY_Can 2OQ753 | 18 | both
2OQ753/Skip Lot Release_NY_2OQ753 | 2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Total Height OD | Total Height OD | 6.825 | in | FDA,3 | 6.706 | 6.826 | Canister & Lid Inspection | 2OQ753/Skip Lot Release_NY_2OQ753 | 103 | both
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | HDPE Resin Identification Marker | Meets the Specification (Done) |  |  |  |  | Canister & Lid Inspection | 2OQ753/Reduced Testing_NY_Can 2OQ753 | 18 | both
2OQ753/Skip Lot Release_NY_2OQ753 | 2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | HDPE Resin Identification Marker | Meets the Specification (Done) |  |  |  |  | Canister & Lid Inspection | 2OQ753/Skip Lot Release_NY_2OQ753 | 103 | both
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | Canister & Lid Inspection | 2OQ753/Reduced Testing_NY_Can 2OQ753 | 18 | both
2OQ753/Skip Lot Release_NY_2OQ753 | 2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | Canister & Lid Inspection | 2OQ753/Skip Lot Release_NY_2OQ753 | 103 | both
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight | 69.21 | g | FDA,1 | 67.5 | 72.5 | Measurement | 2OQ753/Reduced Testing_NY_Can 2OQ753 | 18 | both
2OQ753/Skip Lot Release_NY_2OQ753 | 2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight | 68.86 | g | FDA,1 | 67.5 | 72.5 | Measurement | 2OQ753/Skip Lot Release_NY_2OQ753 | 103 | both
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 766 | Lid Pull Force | Lid Pull-Off Force | Lid Pull-Off Force | 2859.39 | g | FDA,0 | 2000 | 3500 | TM 148 | 2OQ753/Reduced Testing_NY_Can 2OQ753 | 18 | both
2OQ753/Skip Lot Release_NY_2OQ753 | 2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 766 | Lid Pull Force | Lid Pull-Off Force | Lid Pull-Off Force | 2979.67 | g | FDA,0 | 2000 | 3500 | TM 148 | 2OQ753/Skip Lot Release_NY_2OQ753 | 103 | both
2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 2OQ753X | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid has a tight friction fit |  |  |  |  | Canister & Lid Inspection | 2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 1 | both
2OQ753X/Skip Lot Relrase_NY_2OQ753X | 2OQ753X | Skip Lot Relrase_NY_2OQ753X | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA | 2OQ753X/Skip Lot Relrase_NY_2OQ753X | 11 | both
2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 2OQ753X | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA | 2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 1 | both
2OQ753X/Skip Lot Release_NY_2OQ850X | 2OQ753X | Skip Lot Release_NY_2OQ850X | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA |  |  | left_only
2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 2OQ753X | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Canister & Lid Inspection | 2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 1 | both
2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 2OQ753X | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection | 2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 1 | both
2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 2OQ753X | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | White (Done) |  |  |  |  | Canister & Lid Inspection | 2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 1 | both
2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 2OQ753X | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | HDPE Resin Identification Marker | Meets the Specification (Done) |  |  |  |  | Canister & Lid Inspection | 2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 1 | both
2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 2OQ753X | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | Canister & Lid Inspection | 2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 1 | both
2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 2OQ753X | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight | 70.655 | g | FDA,1 | 66.5 | 73.5 | Measurement | 2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 1 | both
2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 2OQ753X | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 1052 | Top Load Compression Strength | Top load Compression Strength (0.5"/mSR,0.450"TD) | Top load Compression Strength (0.5"/mSR,0.450"TD) |  | lb | FDA,0 | 84 |  | NPTM 1019 | 2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 1 | both
2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 2OQ753X | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 1053 | Cavity | cavity Inserts | cavity Inserts | Meets specification (Done) |  |  |  |  | Visual | 2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 1 | both
2OQ850/Reduced Testing_NY_Can 2OQ850X | 2OQ850 | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection |  |  | left_only
2OQ850/Reduced Testing_NY_Can 2OQ850 | 2OQ850 | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid has a tight friction fit |  |  |  |  | Canister & Lid Inspection | 2OQ850/Reduced Testing_NY_Can 2OQ850 | 16 | both
2OQ850/Skip Lot Release_NY_2OQ850 | 2OQ850 | Skip Lot Release_NY_2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid has a tight friction fit |  |  |  |  | Canister & Lid Inspection | 2OQ850/Skip Lot Release_NY_2OQ850 | 98 | both
2OQ850/Reduced Testing_NY_Can 2OQ850 | 2OQ850 | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA | 2OQ850/Reduced Testing_NY_Can 2OQ850 | 16 | both
2OQ850/Skip Lot Release_NY_2OQ850 | 2OQ850 | Skip Lot Release_NY_2OQ850 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis | 2OQ850/Skip Lot Release_NY_2OQ850 | 98 | both
2OQ850/Reduced Testing_NY_Can 2OQ850 | 2OQ850 | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis | 2OQ850/Reduced Testing_NY_Can 2OQ850 | 16 | both
2OQ850/Reduced Testing_NY_Can 2OQ850X | 2OQ850 | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight |  | g | FDA,1 | 85.5 | 94.5 | Measurement |  |  | left_only
2OQ850/Reduced Testing_NY_Can 2OQ850X | 2OQ850 | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid has a tight friction fit |  |  |  |  | Canister & Lid Inspection |  |  | left_only
2OQ850/Reduced Testing_NY_Can 2OQ850 | 2OQ850 | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection | 2OQ850/Reduced Testing_NY_Can 2OQ850 | 16 | both
2OQ850/Skip Lot Release_NY_2OQ850 | 2OQ850 | Skip Lot Release_NY_2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection | 2OQ850/Skip Lot Release_NY_2OQ850 | 98 | both
2OQ850/Reduced Testing_NY_Can 2OQ850X | 2OQ850 | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection |  |  | left_only
2OQ850/Reduced Testing_NY_Can 2OQ850 | 2OQ850 | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | White (Done) |  |  |  |  | Canister & Lid Inspection | 2OQ850/Reduced Testing_NY_Can 2OQ850 | 16 | both
2OQ850/Skip Lot Release_NY_2OQ850 | 2OQ850 | Skip Lot Release_NY_2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | White (Done) |  |  |  |  | Canister & Lid Inspection | 2OQ850/Skip Lot Release_NY_2OQ850 | 98 | both
2OQ850/Reduced Testing_NY_Can 2OQ850X | 2OQ850 | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | HDPE Resin Identification Marker | Meets the Specification (Done) |  |  |  |  | Canister & Lid Inspection |  |  | left_only
2OQ850/Reduced Testing_NY_Can 2OQ850 | 2OQ850 | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | HDPE Resin Identification Marker | Meets the Specification (Done) |  |  |  |  | Canister & Lid Inspection | 2OQ850/Reduced Testing_NY_Can 2OQ850 | 16 | both
2OQ850/Skip Lot Release_NY_2OQ850 | 2OQ850 | Skip Lot Release_NY_2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | HDPE Resin Identification Marker | Meets the Specification (Done) |  |  |  |  | Canister & Lid Inspection | 2OQ850/Skip Lot Release_NY_2OQ850 | 98 | both
2OQ850/Reduced Testing_NY_Can 2OQ850X | 2OQ850 | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | Canister & Lid Inspection |  |  | left_only
2OQ850/Reduced Testing_NY_Can 2OQ850 | 2OQ850 | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | Canister & Lid Inspection | 2OQ850/Reduced Testing_NY_Can 2OQ850 | 16 | both
2OQ850/Skip Lot Release_NY_2OQ850 | 2OQ850 | Skip Lot Release_NY_2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | Canister & Lid Inspection | 2OQ850/Skip Lot Release_NY_2OQ850 | 98 | both
2OQ850/Reduced Testing_NY_Can 2OQ850X | 2OQ850 | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | White (Done) |  |  |  |  | Canister & Lid Inspection |  |  | left_only
2OQ850/Reduced Testing_NY_Can 2OQ850 | 2OQ850 | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight | 88.41 | g | FDA,1 | 87 | 93 | Measurement | 2OQ850/Reduced Testing_NY_Can 2OQ850 | 16 | both
2OQ850/Skip Lot Release_NY_2OQ850 | 2OQ850 | Skip Lot Release_NY_2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight | 89.99 | g | FDA,1 | 87 | 93 | Measurement | 2OQ850/Skip Lot Release_NY_2OQ850 | 98 | both
2OQ850/Reduced Testing_NY_Can 2OQ850X | 2OQ850 | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 1052 | Top Load Compression Strength | Top load Compression Strength (0.5"/mSR,0.450"TD) | Top load Compression Strength (0.5"/mSR,0.450"TD) |  | lb | FDA,0 | 97 |  | NPTM 1019 |  |  | left_only
2OQ850/Reduced Testing_NY_Can 2OQ850 | 2OQ850 | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 766 | Lid Pull Force | Lid Pull-Off Force | Lid Pull-Off Force |  | g | FDA,0 | 2000 | 3500 | TM 148 | 2OQ850/Reduced Testing_NY_Can 2OQ850 | 16 | both
2OQ850/Skip Lot Release_NY_2OQ850 | 2OQ850 | Skip Lot Release_NY_2OQ850 | NY-Analytical | 766 | Lid Pull Force | Lid Pull-Off Force | Lid Pull-Off Force | 2719.98 | g | FDA,0 | 2000 | 3500 | TM 148 | 2OQ850/Skip Lot Release_NY_2OQ850 | 98 | both
2OQ850/Reduced Testing_NY_Can 2OQ850X | 2OQ850 | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 1053 | Cavity | cavity Inserts | cavity Inserts | [5 empty] | Visual |  |  | left_only
2OQ850LI/Lid_11B | 2OQ850LI | Lid_11B | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Flip Lid Force | 2OQ850LI/Lid_11B | 1 | both
2OQ850LI/Lid_11B | 2OQ850LI | Lid_11B | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | Flip Lid Force | 2.41 | lb | FDA,2 | 0.75 |  | TM 143 | 2OQ850LI/Lid_11B | 1 | both
2OQ850LI/Lid_11B | 2OQ850LI | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | 2OQ850LI/Lid_11B | 1 | both
2OQ850LI/Lid_11B | 2OQ850LI | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | Bead Diameter | 4.861 | in | FDA,3 | 4.847 | 4.867 | TM 143 / Visual | 2OQ850LI/Lid_11B | 1 | both
2OQ850LI/Lid_11B | 2OQ850LI | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Translucent Red (Done) |  |  |  |  | TM 143 / Visual | 2OQ850LI/Lid_11B | 1 | both
2OQ850LI/Lid_11B | 2OQ850LI | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) |  |  |  |  | TM 143 / Visual | 2OQ850LI/Lid_11B | 1 | both
2OQ850LI/Lid_11B | 2OQ850LI | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Insert Fin Gap (per specification) | Insert Fin Gap | 0.073 | in | FDA,3 | 0.065 | 0.075 | TM 143 / Visual | 2OQ850LI/Lid_11B | 1 | both
2OQ850LI/Lid_11B | 2OQ850LI | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter (as per specification) | Lid Inside Diameter | 4.919 | in | FDA,3 | 4.917 | 4.937 | TM 143 / Visual | 2OQ850LI/Lid_11B | 1 | both
2OQ850LI/Lid_11B | 2OQ850LI | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.0425 | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual | 2OQ850LI/Lid_11B | 1 | both
2OQ850LI/Lid_11B | 2OQ850LI | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual | 2OQ850LI/Lid_11B | 1 | both
2OQ850LI/Lid_11B | 2OQ850LI | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness (per specification) | Orifice Flap Thickness | 0.024 | in | FDA,3 | 0.018 | 0.028 | TM 143 / Visual | 2OQ850LI/Lid_11B | 1 | both
2OQ850LI/Lid_11B | 2OQ850LI | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.488 | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual | 2OQ850LI/Lid_11B | 1 | both
2OQ850LI/Lid_11B | 2OQ850LI | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Outer Diameter (per specification) | Pop-Up Hole Outer Diameter | 0.165 | in | FDA,3 | 0.16 | 0.17 | TM 143 / Visual | 2OQ850LI/Lid_11B | 1 | both
2OQ850LI/Lid_NY_Skip lot testing | 2OQ850LI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Outer Diameter (per specification) | Pop-Up Hole Outer Diameter | 0.163 | in | FDA,3 | 0.16 | 0.17 | TM 143 / Visual | 2OQ850LI/Lid_NY_Skip lot testing | 16 | both
2OQ850LI/Lid_11B | 2OQ850LI | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 34.04 | g | FDA,1 | 33 | 42 | TM 143 / Visual | 2OQ850LI/Lid_11B | 1 | both
2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | 2OQ850X | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 |  |  | TM 141 / Visual |  |  | left_only
2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | 2OQ850X | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | [5 empty] | TM 141 / Visual |  |  | left_only
2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | 2OQ850X | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual |  |  | left_only
2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | 2OQ850X | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | [5 empty] | TM 141 / Visual |  |  | left_only
2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | 2OQ850X | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height |  | in | FDA,4 | 4.0625 | 4.3125 | TM 141 / Visual |  |  | left_only
2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | 2OQ850X | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 12.3125 | 12.5625 | TM 141 / Visual |  |  | left_only
2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | 2OQ850X | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | [5 empty] | TM 141 / Visual |  |  | left_only
2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | 2OQ850X | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | [5 empty] | TM 141 / Visual |  |  | left_only
2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | 2OQ850X | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 5.3125 | 5.5625 | TM 141 / Visual |  |  | left_only
2OQ882LI/Lid_12A | 2OQ882LI | Lid_12A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Flip Lid Force | 2OQ882LI/Lid_12A | 1 | both
2OQ882LI/Lid_12A | 2OQ882LI | Lid_12A | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | Flip Lid Force | 1.61 | lb | FDA,2 | 0.75 |  | TM 143 | 2OQ882LI/Lid_12A | 1 | both
2OQ882LI/Lid_12A | 2OQ882LI | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | 2OQ882LI/Lid_12A | 1 | both
2OQ882LI/Lid_12A | 2OQ882LI | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | Bead Diameter | 4.854 | in | FDA,3 | 4.852 | 4.862 | TM 143 / Visual | 2OQ882LI/Lid_12A | 1 | both
2OQ882LI/Lid_12A | 2OQ882LI | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Cap Thickness | Cap Thickness | 0.048 | in | FDA,3 | 0.045 | 0.055 | TM 143 / Visual | 2OQ882LI/Lid_12A | 1 | both
2OQ882LI/Lid_12A | 2OQ882LI | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Dark Green (Done) |  |  |  |  | TM 143 / Visual | 2OQ882LI/Lid_12A | 1 | both
2OQ882LI/Lid_12A | 2OQ882LI | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) |  |  |  |  | TM 143 / Visual | 2OQ882LI/Lid_12A | 1 | both
2OQ882LI/Lid_12A | 2OQ882LI | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Insert Fin Gap (per specification) | Insert Fin Gap | 0.074 | in | FDA,3 | 0.065 | 0.075 | TM 143 / Visual | 2OQ882LI/Lid_12A | 1 | both
2OQ882LI/Lid_12A | 2OQ882LI | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.045 | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual | 2OQ882LI/Lid_12A | 1 | both
2OQ882LI/Lid_12A | 2OQ882LI | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual | 2OQ882LI/Lid_12A | 1 | both
2OQ882LI/Lid_12A | 2OQ882LI | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness (per specification) | Orifice Flap Thickness | 0.021 | in | FDA,3 | 0.018 | 0.028 | TM 143 / Visual | 2OQ882LI/Lid_12A | 1 | both
2OQ882LI/Lid_12A | 2OQ882LI | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.488 | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual | 2OQ882LI/Lid_12A | 1 | both
2OQ882LI/Lid_12A | 2OQ882LI | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Outer Diameter (per specification) | Pop-Up Hole Outer Diameter | 0.168 | in | FDA,3 | 0.16 | 0.17 | TM 143 / Visual | 2OQ882LI/Lid_12A | 1 | both
2OQ882LI/Lid_12A | 2OQ882LI | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 35.675 | g | FDA,1 | 32 | 42 | TM 143 / Visual | 2OQ882LI/Lid_12A | 1 | both
2OU130/Web_NP5115_9.0 | 2OU130 | Web_NP5115_9.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  |  | left_only
2OU130/Web_NP5115_9.0 | 2OU130 | Web_NP5115_9.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Printed/Varnished) | Basis Weight |  | #/Ream | FDA,1 | 55.6 | 68 | TM 101 |  |  | left_only
2OU130/Web_NP5115_9.0 | 2OU130 | Web_NP5115_9.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width |  | in | FDA,4 | 8.9375 | 9.0625 | Measurement |  |  | left_only
2OU130/Web_NP5115_9.0 | 2OU130 | Web_NP5115_9.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@325DegF/40PSI/0.5Sec_Dwell) | Seal Strength (@325°F/40PSI/0.5Sec Dwell) |  | g/in | FDA,0 | 1500 |  | TM 107 |  |  | left_only
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | 2OX465B/Bottle, 26ml | 1 | both
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection | 2OX465B/Bottle, 26ml | 1 | both
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | NY-Analytical | 765 | Canister & Lid Inspection | Bottle Internal Diameter at Flange | Bottle Internal Diameter at Flange | 0.722 | in | FDA,3 | 0.709 | 0.729 | Canister & Lid Inspection | 2OX465B/Bottle, 26ml | 1 | both
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Natural (Done) |  |  |  |  | Canister & Lid Inspection | 2OX465B/Bottle, 26ml | 1 | both
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | NY-Analytical | 765 | Canister & Lid Inspection | Overall Length | Overall Length | 4.855 | in | FDA,3 | 4.825 | 4.875 | Canister & Lid Inspection | 2OX465B/Bottle, 26ml | 1 | both
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | NY-Analytical | 765 | Canister & Lid Inspection | Seal Flange Diameter | Seal Flange Diameter | 0.8735 | in | FDA,3 | 0.864 | 0.874 | Canister & Lid Inspection | 2OX465B/Bottle, 26ml | 1 | both
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 2 | g | FDA,2 |  |  | TM 610 | 2OX465B/Bottle, 26ml | 1 | both
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer |  |  |  |  | TM 610 | 2OX465B/Bottle, 26ml | 1 | both
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610 | 2OX465B/Bottle, 26ml | 1 | both
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 2OX465B/Bottle, 26ml | 1 | both
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | 2OX465B/Bottle, 26ml | 1 | both
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 2OX465B/Bottle, 26ml | 1 | both
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 1 | cfu/device | FDA,1 |  |  | TM 610 | 2OX465B/Bottle, 26ml | 1 | both
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0.0028852 | CFU/g | FDA,1 |  | 15 | TM 610 | 2OX465B/Bottle, 26ml | 1 | both
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 2OX465B/Bottle, 26ml | 1 | both
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Negative Rods (OOS) |  |  |  |  | WI QAS 076 | 2OX465B/Bottle, 26ml | 1 | both
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Mold (OOS) |  |  |  |  | WI QAS 076 | 2OX465B/Bottle, 26ml | 1 | both
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Certificate of Conformance | 2OX510B/Bottle, 5.1ml | 1 | both
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | 2OX510B/Microbiologu - Bioburden_Applicator_NY | 1 | both
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection | 2OX510B/Microbiologu - Bioburden_Applicator_NY | 1 | both
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection | 2OX510B/Bottle, 5.1ml | 1 | both
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 765 | Canister & Lid Inspection | Bottle Internal Diameter at Flange | Bottle Internal Diameter at Flange | 0.4425 | in | FDA,3 | 0.442 | 0.452 | Canister & Lid Inspection | 2OX510B/Microbiologu - Bioburden_Applicator_NY | 1 | both
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | NY-Analytical | 765 | Canister & Lid Inspection | Bottle Internal Diameter at Flange | Bottle Internal Diameter at Flange | 0.443 | in | FDA,3 | 0.442 | 0.452 | Canister & Lid Inspection | 2OX510B/Bottle, 5.1ml | 1 | both
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Natural (Done) |  |  |  |  | Canister & Lid Inspection | 2OX510B/Microbiologu - Bioburden_Applicator_NY | 1 | both
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Natural (Done) |  |  |  |  | Canister & Lid Inspection | 2OX510B/Bottle, 5.1ml | 1 | both
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 765 | Canister & Lid Inspection | Overall Length | Overall Length | 2.8805 | in | FDA,3 | 2.848 | 2.888 | Canister & Lid Inspection | 2OX510B/Microbiologu - Bioburden_Applicator_NY | 1 | both
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | NY-Analytical | 765 | Canister & Lid Inspection | Overall Length | Overall Length | 2.8737 | in | FDA,3 | 2.848 | 2.888 | Canister & Lid Inspection | 2OX510B/Bottle, 5.1ml | 1 | both
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 765 | Canister & Lid Inspection | Seal Flange Diameter | Seal Flange Diameter | 0.55 | in | FDA,3 | 0.544 | 0.554 | Canister & Lid Inspection | 2OX510B/Microbiologu - Bioburden_Applicator_NY | 1 | both
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | NY-Analytical | 765 | Canister & Lid Inspection | Seal Flange Diameter | Seal Flange Diameter | 0.55 | in | FDA,3 | 0.544 | 0.554 | Canister & Lid Inspection | 2OX510B/Bottle, 5.1ml | 1 | both
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight |  | g | FDA,2 |  |  | TM 610 | 2OX510B/Microbiologu - Bioburden_Applicator_NY | 1 | both
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 4.5 | g | FDA,2 |  |  | TM 610 | 2OX510B/Bottle, 5.1ml | 1 | both
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | 2OX510B/Microbiologu - Bioburden_Applicator_NY | 1 | both
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | 2OX510B/Bottle, 5.1ml | 1 | both
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610 | 2OX510B/Microbiologu - Bioburden_Applicator_NY | 1 | both
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610 | 2OX510B/Bottle, 5.1ml | 1 | both
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 2OX510B/Microbiologu - Bioburden_Applicator_NY | 1 | both
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 1.1 | cfu/device | FDA,1 |  |  | TM 610 | 2OX510B/Bottle, 5.1ml | 1 | both
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 100 | TM 610 | 2OX510B/Microbiologu - Bioburden_Applicator_NY | 1 | both
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0.24444 | CFU/g | FDA,1 |  | 100 | TM 610 | 2OX510B/Bottle, 5.1ml | 1 | both
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 2OX510B/Microbiologu - Bioburden_Applicator_NY | 1 | both
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 2OX510B/Bottle, 5.1ml | 1 | both
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 1.1 | cfu/device | FDA,1 |  |  | TM 610 | 2OX510B/Microbiologu - Bioburden_Applicator_NY | 1 | both
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 1 | cfu/device | FDA,1 |  |  | TM 610 | 2OX510B/Bottle, 5.1ml | 1 | both
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0.015027 | CFU/g | FDA,1 |  | 15 | TM 610 | 2OX510B/Microbiologu - Bioburden_Applicator_NY | 1 | both
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0.022222 | CFU/g | FDA,1 |  | 15 | TM 610 | 2OX510B/Bottle, 5.1ml | 1 | both
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 2OX510B/Microbiologu - Bioburden_Applicator_NY | 1 | both
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 1 | cfu | FDA,0 |  |  | TM 610 | 2OX510B/Bottle, 5.1ml | 1 | both
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Positive Rods (Done) |  |  |  |  | WI QAS 076 | 2OX510B/Microbiologu - Bioburden_Applicator_NY | 1 | both
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Positive Rods (Done) |  |  |  |  | WI QAS 076 | 2OX510B/Bottle, 5.1ml | 1 | both
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Mold (OOS) |  |  |  |  | WI QAS 076 | 2OX510B/Microbiologu - Bioburden_Applicator_NY | 1 | both
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Mold (OOS) |  |  |  |  | WI QAS 076 | 2OX510B/Bottle, 5.1ml | 1 | both
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil seal 5.1 ml | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA |  |  | left_only
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil seal 5.1 ml | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight |  | g | FDA,2 |  |  | TM 610 |  |  | left_only
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil seal 5.1 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 100 | TM 610 |  |  | left_only
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil seal 5.1 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) |  | cfu | FDA,0 |  |  | TM 610 |  |  | left_only
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil seal 5.1 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 |  |  | left_only
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil seal 5.1 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 |  |  | left_only
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil seal 5.1 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 15 | TM 610 |  |  | left_only
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil seal 5.1 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  | cfu | FDA,0 |  |  | TM 610 |  |  | left_only
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil seal 5.1 ml | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Negative Rods (OOS) |  |  |  |  | WI QAS 076 |  |  | left_only
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil seal 5.1 ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610 |  |  | left_only
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil seal 5.1 ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 |  |  | left_only
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil seal 5.1 ml | NY-Analytical | 1056 | Functionality | Foil diameter to fit Bottle(Item#2OX510B)opening | Foil diameter to fit Bottle(Item#2OX510B)opening | Pass |  |  |  |  | Visual |  |  | left_only
2PD351/Web_NP5116_4.5 | 2PD351 | Web_NP5116_4.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2PD351/Web_NP5116_4.5 | 5 | both
2PD351/Web_NP5116_4.5 | 2PD351 | Web_NP5116_4.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 4.5 | in | FDA,4 | 4.4375 | 4.5625 | Measurement | 2PD351/Web_NP5116_4.5 | 5 | both
2PD351/Web_NP5116_4.5 | 2PD351 | Web_NP5116_4.5 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Printed/Varnished) | Basis Weight | 59.82 | #/Ream | FDA,1 | 55.8 | 68.2 | TM 101 | 2PD351/Web_NP5116_4.5 | 5 | both
2PD351/Web_NP5116_4.5 | 2PD351 | Web_NP5116_4.5 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/40psi/0.5 sec dewll) | Seal Strength (@300°F/40psi/0.5 sec dewll) | 2163 | g/in | FDA,0 | 1500 |  | TM 107 | 2PD351/Web_NP5116_4.5 | 5 | both
2QB114X/Web_NP3860_5.0 | 2QB114X | Web_NP3860_5.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2QB114X/Web_NP3860_5.0 | 1 | both
2QB114X/Web_NP3860_5.0 | 2QB114X | Web_NP3860_5.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 5 | in | FDA,4 | 4.9375 | 5.0625 | Measurement | 2QB114X/Web_NP3860_5.0 | 1 | both
2QB114X/Web_NP3860_5.0 | 2QB114X | Web_NP3860_5.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 89.29 | #/Ream | FDA,1 | 78.9 | 96.5 | TM 101 | 2QB114X/Web_NP3860_5.0 | 1 | both
2QB114X/Web_NP3860_5.0 | 2QB114X | Web_NP3860_5.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@350DegF/40psi/1.0 sec dwell) | Seal Strength (@350°F/40psi/1.0 sec dwell) | 9928 | g/in | FDA,0 | 7250 |  | TM 107 | 2QB114X/Web_NP3860_5.0 | 1 | both
2QB114X/Web_NP3860_5.0 | 2QB114X | Web_NP3860_5.0 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 5 | mil | FDA,1 | 4.7 | 5.7 | TM 132 | 2QB114X/Web_NP3860_5.0 | 1 | both
2RQ890/Web_NP9508_18.125 | 2RQ890 | Web_NP9508_18.125 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2RQ890/Web_NP9508_18.125 | 2 | both
2RQ890/Web_NP9508_18.125 | 2RQ890 | Web_NP9508_18.125 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 18.125 | in | FDA,4 | 18.0625 | 18.1875 | Measurement | 2RQ890/Web_NP9508_18.125 | 2 | both
2RQ890/Web_NP9508_18.125 | 2RQ890 | Web_NP9508_18.125 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Unprinted) | Basis Weight |  | #/Ream | FDA,1 | 37.8 | 46.2 | TM 101 | 2RQ890/Web_NP9508_18.125 | 2 | both
2RQ890/Web_NP9508_18.125 | 2RQ890 | Web_NP9508_18.125 | NY-Analytical | 762 | Seal Strength | Seal Strength (@300DegF/40psi/0.5 sec dewll) | Seal Strength (@300°F/40psi/0.5 sec dewll) | Sample peeled apart and showed stress whitening when sealed (Done) |  |  |  |  | TM 107 | 2RQ890/Web_NP9508_18.125 | 2 | both
2RQ890/Web_NP9508_18.125 | 2RQ890 | Web_NP9508_18.125 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 2.6 | mil | FDA,1 | 2.3 | 2.8 | TM 132 | 2RQ890/Web_NP9508_18.125 | 2 | both
2RS324/Web_NP3860_12.5 | 2RS324 | Web_NP3860_12.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2RS324/Web_NP3860_12.5 | 1 | both
2RS324/Web_NP3860_12.5 | 2RS324 | Web_NP3860_12.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 12.5 | in | FDA,4 | 12.4375 | 12.5625 | Measurement | 2RS324/Web_NP3860_12.5 | 1 | both
2RS324/Web_NP3860_12.5 | 2RS324 | Web_NP3860_12.5 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 89 | #/Ream | FDA,1 | 78.9 | 96.5 | TM 101 | 2RS324/Web_NP3860_12.5 | 1 | both
2RS324/Web_NP3860_12.5 | 2RS324 | Web_NP3860_12.5 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@350DegF/40psi/1.0 sec dwell) | Seal Strength (@350°F/40psi/1.0 sec dwell) | 9622 | g/in | FDA,0 | 7250 |  | TM 107 | 2RS324/Web_NP3860_12.5 | 1 | both
2RS324/Web_NP3860_12.5 | 2RS324 | Web_NP3860_12.5 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 4.91 | mil | FDA,1 | 4.7 | 5.7 | TM 132 | 2RS324/Web_NP3860_12.5 | 1 | both
2RS407/Web_NP3885_13.0 | 2RS407 | Web_NP3885_13.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2RS407/Web_NP3885_13.0 | 4 | both
2RS407/Web_NP3885_13.0 | 2RS407 | Web_NP3885_13.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 12.9375 | in | FDA,4 | 12.9375 | 13.0625 | Measurement | 2RS407/Web_NP3885_13.0 | 4 | both
2RS407/Web_NP3885_13.0 | 2RS407 | Web_NP3885_13.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 80.69 | #/Ream | FDA,0 | 74 | 90 | TM 101 | 2RS407/Web_NP3885_13.0 | 4 | both
2RS407/Web_NP3885_13.0 | 2RS407 | Web_NP3885_13.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@350DegF/50psi/1.0 sec dwell) | Seal Strength (@350°F/50psi/1.0 sec dwell) | 1941 | g/in | FDA,0 | 700 |  | TM 107 | 2RS407/Web_NP3885_13.0 | 4 | both
2RS419/Web_NP3885_16 | 2RS419 | Web_NP3885_16 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2RS419/Web_NP3885_16 | 2 | both
2RS419/Web_NP3885_16 | 2RS419 | Web_NP3885_16 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 16 | in | FDA,4 | 15.9375 | 16.0625 | Measurement | 2RS419/Web_NP3885_16 | 2 | both
2RS419/Web_NP3885_16 | 2RS419 | Web_NP3885_16 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 81.85 | #/Ream | FDA,0 | 74 | 90 | TM 101 | 2RS419/Web_NP3885_16 | 2 | both
2RS419/Web_NP3885_16 | 2RS419 | Web_NP3885_16 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@350DegF/50psi/1.0 sec dwell) | Seal Strength (@350°F/50psi/1.0 sec dwell) | 2030.5 | g/in | FDA,0 | 700 |  | TM 107 | 2RS419/Web_NP3885_16 | 2 | both
2SS273/Web_NP3860_15.5 | 2SS273 | Web_NP3860_15.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement | 2SS273/Web_NP3860_15.5 | 1 | both
2SS273/Web_NP3860_15.5 | 2SS273 | Web_NP3860_15.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 15.5 | in | FDA,4 | 15.4375 | 15.5625 | Measurement | 2SS273/Web_NP3860_15.5 | 1 | both
2SS273/Web_NP3860_15.5 | 2SS273 | Web_NP3860_15.5 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 87.7 | #/Ream | FDA,1 | 78.9 | 96.5 | TM 101 | 2SS273/Web_NP3860_15.5 | 1 | both
2SS273/Web_NP3860_15.5 | 2SS273 | Web_NP3860_15.5 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@350DegF/40psi/1.0 sec dwell) | Seal Strength (@350°F/40psi/1.0 sec dwell) | 10111.34 | g/in | FDA,0 | 7250 |  | TM 107 | 2SS273/Web_NP3860_15.5 | 1 | both
2SS273/Web_NP3860_15.5 | 2SS273 | Web_NP3860_15.5 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 5.2 | mil | FDA,1 | 4.7 | 5.7 | TM 132 | 2SS273/Web_NP3860_15.5 | 1 | both
2SS428/Web_NP3860_8.0 | 2SS428 | Web_NP3860_8.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Basis Weight | 2SS428/Web_NP3860_8.0 | 2 | both
2SS428/Web_NP3860_8.0 | 2SS428 | Web_NP3860_8.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 87.36 | #/Ream | FDA,1 | 78.9 | 96.5 | TM 101 | 2SS428/Web_NP3860_8.0 | 2 | both
2SS428/Web_NP3860_8.0 | 2SS428 | Web_NP3860_8.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@350DegF/40psi/1.0 sec dwell) | Seal Strength (@350°F/40psi/1.0 sec dwell) | 9981.29 | g/in | FDA,0 | 7250 |  | TM 107 | 2SS428/Web_NP3860_8.0 | 2 | both
2SS428/Web_NP3860_8.0 | 2SS428 | Web_NP3860_8.0 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 5.2 | mil | FDA,1 | 4.7 | 5.7 | TM 132 | 2SS428/Web_NP3860_8.0 | 2 | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3 | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 44 | in | FDA,0 | 44 | 46 | Applicator Measurements | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3 | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 37.5 | in | FDA,1 | 37.3 | 37.7 | Applicator Measurements | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3 | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 55.54 | GSM | FDA,1 | 49.5 | 60.5 | TM 101 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3 | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 1147.63 | g/in | FDA,0 | 692 | 1438 | TM 103 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3 | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 4131.08 | g/in | FDA,0 | 2785 | 5785 | TM 103 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3 | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3 | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 748 | Visual Inspection | Color | Color | White to Off-White (Done) |  |  |  |  | TM 208 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3 | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3 | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 509 | J&J_Pathogens | Biggy or SDA | Biggy or SDA | No Growth_SDA |  |  |  |  | TM 869 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3 | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 509 | J&J_Pathogens | Cetrimide | Cetrimide | No Growth |  |  |  |  | TM 869 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3 | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 524 | J&J Fungal | JJ Fungal_RM | JJ Fungal RM | 0 | CFU/g |  |  | 200 | TM 869 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3 | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 509 | J&J_Pathogens | MacConkey | MacConkey | No Growth |  |  |  |  | TM 869 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3 | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 509 | J&J_Pathogens | MSA or VJ | MSA or VJ | No Growth_MSA |  |  |  |  | TM 869 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3 | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 700 | J&J SDA Controls | Negative Control | Negative Control | No Growth SDA |  |  |  |  | TM 869 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3 | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | NegativeControl | No Growth TSB |  |  |  |  | TM 869 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3 | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 509 | J&J_Pathogens | TSA | TSA | No Growth |  |  |  |  | TM 869 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3 | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ RM | TSA JJ RM | 0 | CFU/g |  |  | 200 | TM 869 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3 | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 517 | J&J_Enrichment | TSB | TSB | Not Turbid |  |  |  |  | TM 869 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3 | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 573 | GramStain_Candida | Gram Stain_J&JCandida | Gram Stain J&JCandida | Non-Subculturable Growth (Done) |  |  |  |  | WI QAS 076 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3 | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 545 | GramStain_J&J_TSA | GramStain_J&J_TSA | GramStain J&J TSA | Gram Positive Rods (Done) |  |  |  |  | WI QAS 076 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3 | both
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 3AQ470-A | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading |  |  |  |  | Color (TM 208) | 3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 1 | both
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 3AQ470-A | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Free of hard wrinkles. Formation uniform (Done) |  |  |  |  | See Analyte for Method # | 3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 1 | both
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 3AQ470-A | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | See Analyte for Method # | 3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 1 | both
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 3AQ470-A | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 49.51 | GSM | FDA,1 | 45 | 55 | See Analyte for Method # | 3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 1 | both
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 3AQ470-A | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (mm) (CofA) | Caliper | 0.59 | mm | FDA,2 | 0.51 | 0.71 | See Analyte for Method # | 3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 1 | both
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 3AQ470-A | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method # | 3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 1 | both
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 3AQ470-A | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | May have slight odor when wet (Done) |  |  |  |  | See Analyte for Method # | 3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 1 | both
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 3AQ470-A | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter | 50 | in | FDA,0 | 45 | 60 | See Analyte for Method # | 3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 1 | both
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 3AQ470-A | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Width (Inches) (Measurement) | Roll Width | 37.75 | in | FDA,2 | 37.5 | 38 | See Analyte for Method # | 3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 1 | both
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 3AQ470-A | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Wet CD (g/in) (TM 103) | Wet CD | 1634 | g/in | FDA,0 | 1100 | 2300 | See Analyte for Method # | 3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 1 | both
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 3AQ470-A | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Wet MD (g/in) (TM 103) | Wet MD | 6147 | g/in | FDA,0 | 3000 | 7000 | See Analyte for Method # | 3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 1 | both
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements |  |  | left_only
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 |  |  | left_only
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208 |  |  | left_only
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208 |  |  | left_only
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight |  | GSM | FDA,1 | 49.5 | 60.5 | TM 101 |  |  | left_only
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper |  | mm | FDA,2 | 0.5 | 0.64 | TM 102 |  |  | left_only
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD |  | g/in | FDA,0 | 1443 | 2463 | TM 103 |  |  | left_only
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD |  | g/in | FDA,0 | 4883 | 7967 | TM 103 |  |  | left_only
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width |  | in | FDA,3 | 7.375 | 7.625 | Applicator Measurements |  |  | left_only
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter |  | in | FDA,0 | 28 | 32 | Applicator Measurements |  |  | left_only
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 |  |  | left_only
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 100 | TM 610 |  |  | left_only
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) |  | cfu | FDA,0 |  |  | TM 610 |  |  | left_only
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 |  |  | left_only
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 15 | TM 610 |  |  | left_only
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  | cfu | FDA,0 |  |  | TM 610 |  |  | left_only
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight |  | g | FDA,2 |  |  | TM 610 |  |  | left_only
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer |  |  |  |  | TM 610 |  |  | left_only
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610 |  |  | left_only
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) |  | cfu | FDA,0 |  |  | TM 610 |  |  | left_only
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight |  | g | FDA,2 |  |  | TM 610 |  |  | left_only
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 |  |  | left_only
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 100 | TM 610 |  |  | left_only
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 |  |  | left_only
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 15 | TM 610 |  |  | left_only
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  | cfu | FDA,0 |  |  | TM 610 |  |  | left_only
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Mold (OOS) |  |  |  |  | WI QAS 076 |  |  | left_only
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | REDUCED TESTING_NY_3BB406 | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight |  | GSM | FDA,1 | 48.9 | 56.3 | See Analyte for Method # |  |  | left_only
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | REDUCED TESTING_NY_3BB406 | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 |  |  | See Analyte for Method # |  |  | left_only
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | REDUCED TESTING_NY_3BB406 | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method # |  |  | left_only
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | REDUCED TESTING_NY_3BB406 | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | May have residual sulfur odor when wet |  |  |  |  | Odor (TM 208) |  |  | left_only
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | REDUCED TESTING_NY_3BB406 | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Meets specifications (Done) |  |  |  |  | See Analyte for Method # |  |  | left_only
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | REDUCED TESTING_NY_3BB406 | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading |  |  |  |  | Color (TM 208) |  |  | left_only
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer |  |  |  |  | TM 610 |  |  | left_only
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610 |  |  | left_only
3BJ141/Applicator_07K | 3BJ141 | Applicator_07K | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements |  |  | left_only
3BJ141/Applicator_07K | 3BJ141 | Applicator_07K | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 |  |  | left_only
3BJ141/Applicator_07K | 3BJ141 | Applicator_07K | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208 |  |  | left_only
3BJ141/Applicator_07K | 3BJ141 | Applicator_07K | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208 |  |  | left_only
3BJ141/Applicator_07K | 3BJ141 | Applicator_07K | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight |  | GSM | FDA,1 | 38.7 | 47.3 | TM 101 |  |  | left_only
3BJ141/Applicator_07K | 3BJ141 | Applicator_07K | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(inch) | Caliper |  | in | FDA,3 | 0.022 | 0.03 | TM 102 |  |  | left_only
3BJ141/Applicator_07K | 3BJ141 | Applicator_07K | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD |  | g/in | FDA,0 | 921 | 1535 | TM 103 |  |  | left_only
3BJ141/Applicator_07K | 3BJ141 | Applicator_07K | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD |  | g/in | FDA,0 | 3623 | 5650 | TM 103 |  |  | left_only
3BJ141/Applicator_07K | 3BJ141 | Applicator_07K | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width |  | in | FDA,2 | 11.5 | 11.75 | Applicator Measurements |  |  | left_only
3DQ595/Applicator_10D | 3DQ595 | Applicator_10D | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements | 3DQ595/Applicator_10D | 1 | both
3DQ595/Applicator_10D | 3DQ595 | Applicator_10D | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 46 | in | FDA,0 | 44 | 47 | Applicator Measurements | 3DQ595/Applicator_10D | 1 | both
3DQ595/Applicator_10D | 3DQ595 | Applicator_10D | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 30 | in | FDA,3 | 29.875 | 30.125 | Applicator Measurements | 3DQ595/Applicator_10D | 1 | both
3DQ595/Applicator_10D | 3DQ595 | Applicator_10D | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 39.18 | GSM | FDA,1 | 38 | 42 | TM 101 | 3DQ595/Applicator_10D | 1 | both
3DQ595/Applicator_10D | 3DQ595 | Applicator_10D | NY-Analytical | 757 | Caliper | Caliper-1_Ply_(mm) | Caliper | 0.36 | mm | FDA,2 | 0.35 | 0.45 | TM 102 | 3DQ595/Applicator_10D | 1 | both
3DQ595/Applicator_10D | 3DQ595 | Applicator_10D | NY-Analytical | 758 | Tensile Strength | Dry_CD_(g/in) | Dry CD | 667.15 | g/in | FDA,0 | 513 | 1130 | TM 103 | 3DQ595/Applicator_10D | 1 | both
3DQ595/Applicator_10D | 3DQ595 | Applicator_10D | NY-Analytical | 758 | Tensile Strength | Dry_MD_(g/in) | Dry CD | 2695.76 | g/in | FDA,0 | 2054 | 4108 | TM 103 | 3DQ595/Applicator_10D | 1 | both
3DQ595/Applicator_10D | 3DQ595 | Applicator_10D | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | 3DQ595/Applicator_10D | 1 | both
3DQ595/Applicator_10D | 3DQ595 | Applicator_10D | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208 | 3DQ595/Applicator_10D | 1 | both
3DQ595/Applicator_10D | 3DQ595 | Applicator_10D | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | 3DQ595/Applicator_10D | 1 | both
3DQ595N/Applicator_10I | 3DQ595N | Applicator_10I | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements | 3DQ595N/Applicator_10I | 1 | both
3DQ595N/Applicator_10I | 3DQ595N | Applicator_10I | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 46 | in | FDA,0 | 44 | 47 | Applicator Measurements | 3DQ595N/Applicator_10I | 1 | both
3DQ595N/Applicator_10I | 3DQ595N | Applicator_10I | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 30 | in | FDA,3 | 29.875 | 30.125 | Applicator Measurements | 3DQ595N/Applicator_10I | 1 | both
3DQ595N/Applicator_10I | 3DQ595N | Applicator_10I | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 39.83 | GSM | FDA,0 | 38 | 42 | TM 101 | 3DQ595N/Applicator_10I | 1 | both
3DQ595N/Applicator_10I | 3DQ595N | Applicator_10I | NY-Analytical | 757 | Caliper | Caliper-1_Ply_(mm) | Caliper | 0.41 | mm | FDA,2 | 0.35 | 0.45 | TM 102 | 3DQ595N/Applicator_10I | 1 | both
3DQ595N/Applicator_10I | 3DQ595N | Applicator_10I | NY-Analytical | 758 | Tensile Strength | Dry_CD_(g/in) | Dry CD | 573.11 | g/in | FDA,0 | 513 | 1130 | TM 103 | 3DQ595N/Applicator_10I | 1 | both
3DQ595N/Applicator_10I | 3DQ595N | Applicator_10I | NY-Analytical | 758 | Tensile Strength | Dry_MD_(g/in) | Dry CD | 3431.99 | g/in | FDA,0 | 2054 | 4108 | TM 103 | 3DQ595N/Applicator_10I | 1 | both
3DQ595N/Applicator_10I | 3DQ595N | Applicator_10I | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | 3DQ595N/Applicator_10I | 1 | both
3DQ595N/Applicator_10I | 3DQ595N | Applicator_10I | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208 | 3DQ595N/Applicator_10I | 1 | both
3DQ595N/Applicator_10I | 3DQ595N | Applicator_10I | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | 3DQ595N/Applicator_10I | 1 | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading |  |  |  |  | Color (TM 208) | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 19 | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading |  |  |  |  | Color (TM 208) | 3OA347/Microbiologu - Bioburden_Applicator_NY | 17 | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | May have residual sulfur odor when wet |  |  |  |  | Odor (TM 208) | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 19 | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | May have residual sulfur odor when wet |  |  |  |  | Odor (TM 208) | 3OA347/Microbiologu - Bioburden_Applicator_NY | 17 | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis | 3OA347/Microbiologu - Bioburden_Applicator_NY | 17 | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator_Inspection | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 19 | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Smooth free of hard wrinkles, formation uniform matches embossing pattern (Done) |  |  |  |  | See Analyte for Method # | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 19 | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Smooth free of hard wrinkles, formation uniform matches embossing pattern (Done) |  |  |  |  | See Analyte for Method # | 3OA347/Microbiologu - Bioburden_Applicator_NY | 17 | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | See Analyte for Method # | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 19 | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | See Analyte for Method # | 3OA347/Microbiologu - Bioburden_Applicator_NY | 17 | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 61.79 | GSM | FDA,1 | 60 | 70 | See Analyte for Method # | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 19 | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 65.66 | GSM | FDA,1 | 60 | 70 | See Analyte for Method # | 3OA347/Microbiologu - Bioburden_Applicator_NY | 17 | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (mm) (TM 102) | Caliper | 0.64 | mm | FDA,2 | 0.54 | 0.66 | See Analyte for Method # | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 19 | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (mm) (TM 102) | Caliper | 0.66 | mm | FDA,2 | 0.54 | 0.66 | See Analyte for Method # | 3OA347/Microbiologu - Bioburden_Applicator_NY | 17 | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Ratio: Wet MD / Wet CD (TM 103) | Ratio: Wet MD / Wet CD | 0 |  | FDA,1 |  | 4.5 | See Analyte for Method # | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 19 | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Ratio: Wet MD / Wet CD (TM 103) | Ratio: Wet MD / Wet CD | 0 |  | FDA,1 |  | 4.5 | See Analyte for Method # | 3OA347/Microbiologu - Bioburden_Applicator_NY | 17 | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter | 45 | in | FDA,0 | 42 | 48 | See Analyte for Method # | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 19 | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter | 43 | in | FDA,0 | 42 | 48 | See Analyte for Method # | 3OA347/Microbiologu - Bioburden_Applicator_NY | 17 | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Width (mm) (Measurement) | Roll Width | 6.062 | in | FDA,2 | 6 | 6.125 | See Analyte for Method # | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 19 | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Roll Width (mm) (Measurement) | Roll Width | 6 | in | FDA,2 | 6 | 6.125 | See Analyte for Method # | 3OA347/Microbiologu - Bioburden_Applicator_NY | 17 | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Wet CD (g/in) (TM 103) | Wet CD | 416 | g/in | FDA,0 | 375 | 575 | See Analyte for Method # | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 19 | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Wet CD (g/in) (TM 103) | Wet CD | 426 | g/in | FDA,0 | 375 | 575 | See Analyte for Method # | 3OA347/Microbiologu - Bioburden_Applicator_NY | 17 | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Wet MD (g/in) (TM 103) | Wet MD | 692 | g/in | FDA,0 | 485 | 915 | See Analyte for Method # | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 19 | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Wet MD (g/in) (TM 103) | Wet MD | 769 | g/in | FDA,0 | 485 | 915 | See Analyte for Method # | 3OA347/Microbiologu - Bioburden_Applicator_NY | 17 | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 67.38 | GSM | FDA,1 | 60 | 70 | TM 101 | 3OA347/Microbiologu - Bioburden_Applicator_NY | 17 | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight |  | GSM | FDA,1 | 60 | 70 | TM 101 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 19 | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper | 0.64 | mm | FDA,2 | 0.54 | 0.66 | TM 102 | 3OA347/Microbiologu - Bioburden_Applicator_NY | 17 | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper | 0.63 | mm | FDA,2 | 0.54 | 0.66 | TM 102 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 19 | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 758 | Tensile Strength | Elongation % Wet CD | Elongation % Wet CD | 43.05 | % | FDA,0 | 15 | 65 | TM 103 | 3OA347/Microbiologu - Bioburden_Applicator_NY | 17 | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 758 | Tensile Strength | Elongation % Wet CD | Elongation % Wet CD | 25.08 | % | FDA,0 | 15 | 65 | TM 103 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 19 | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 758 | Tensile Strength | Elongation % Wet MD | Elongation % Wet MD | 30.3 | % | FDA,0 | 2 | 50 | TM 103 | 3OA347/Microbiologu - Bioburden_Applicator_NY | 17 | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 758 | Tensile Strength | Elongation % Wet MD | Elongation % Wet MD | 17.02 | % | FDA,0 | 2 | 50 | TM 103 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 19 | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 561.7 | g/in | FDA,0 | 375 | 575 | TM 103 | 3OA347/Microbiologu - Bioburden_Applicator_NY | 17 | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 495 | g/in | FDA,0 | 375 | 575 | TM 103 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 19 | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 694.79 | g/in | FDA,0 | 485 | 915 | TM 103 | 3OA347/Microbiologu - Bioburden_Applicator_NY | 17 | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 648 | g/in | FDA,0 | 485 | 915 | TM 103 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 19 | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 19 | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 19 | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 19 | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | 3OA347/Microbiologu - Bioburden_Applicator_NY | 17 | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 2.7 | g | FDA,2 |  |  | TM 610 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 19 | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 19 | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 19 | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 19 | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 19 | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 19 | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 1 | cfu/device | FDA,1 |  |  | TM 610 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 19 | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0.96774 | CFU/g | FDA,1 |  | 15 | TM 610 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 19 | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 19 | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Positive Rods (Done) |  |  |  |  | WI QAS 076 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 19 | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Mold (OOS) |  |  |  |  | WI QAS 076 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 19 | both
3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | 3OA426 | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (mm) (CofA) | Caliper |  | mm | FDA,2 | 0.74 | 0.83 | See Analyte for Method # |  |  | left_only
3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | 3OA426 | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Width (Inches) (Measurement) | Roll Width |  | in | FDA,3 | 7.875 | 8.125 | See Analyte for Method # |  |  | left_only
3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | 3OA426 | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter |  | in | FDA,0 | 43 | 48 | See Analyte for Method # |  |  | left_only
3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | 3OA426 | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight |  | GSM | FDA,0 | 47 | 57 | See Analyte for Method # |  |  | left_only
3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | 3OA426 | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 |  |  | See Analyte for Method # |  |  | left_only
3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | 3OA426 | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Wet CD (g/in) (TM 103) | Wet CD |  | g/in | FDA,0 | 1200 | 2100 | See Analyte for Method # |  |  | left_only
3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | 3OA426 | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Wet MD (g/in) (TM 103) | Wet MD |  | g/in | FDA,0 | 3800 | 6270 | See Analyte for Method # |  |  | left_only
3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | 3OA426 | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method # |  |  | left_only
3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | 3OA426 | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Free of hard wrinkles, formation uniform, large dot embossing pattern. (Done) |  |  |  |  | See Analyte for Method # |  |  | left_only
3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | 3OA426 | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | May have slight odor when wet (Done) |  |  |  |  | See Analyte for Method # |  |  | left_only
3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | 3OA426 | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading |  |  |  |  | Color (TM 208) |  |  | left_only
3OA580/Applicator_38A | 3OA580 | Applicator_38A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements | 3OA580/Applicator_38A | 2 | both
3OA580/Applicator_38A | 3OA580 | Applicator_38A | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 45 | in | FDA,0 | 43 | 47 | Applicator Measurements | 3OA580/Applicator_38A | 2 | both
3OA580/Applicator_38A | 3OA580 | Applicator_38A | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 11.5 | in | FDA,3 | 11.375 | 11.625 | Applicator Measurements | 3OA580/Applicator_38A | 2 | both
3OA580/Applicator_38A | 3OA580 | Applicator_38A | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 44.35 | GSM | FDA,0 | 40 | 50 | TM 101 | 3OA580/Applicator_38A | 2 | both
3OA580/Applicator_38A | 3OA580 | Applicator_38A | NY-Analytical | 757 | Caliper | Caliper_at_Pt_of_Mfr_(micron) | Caliper | 500 | um | FDA,0 | 416 | 534 | TM 102 | 3OA580/Applicator_38A | 2 | both
3OA580/Applicator_38A | 3OA580 | Applicator_38A | NY-Analytical | 758 | Tensile Strength | Dry_CD_CoA_(g/in) | Dry CD (CoA) | 1326 | g/in | FDA,0 | 610 |  | TM 103 | 3OA580/Applicator_38A | 2 | both
3OA580/Applicator_38A | 3OA580 | Applicator_38A | NY-Analytical | 758 | Tensile Strength | Dry_MD_CoA_(g/in) | Dry MD (CoA) | 3086 | g/in | FDA,0 | 1825 |  | TM 103 | 3OA580/Applicator_38A | 2 | both
3OA580/Applicator_38A | 3OA580 | Applicator_38A | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | 3OA580/Applicator_38A | 2 | both
3OA580/Applicator_38A | 3OA580 | Applicator_38A | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208 | 3OA580/Applicator_38A | 2 | both
3OA580/Applicator_38A | 3OA580 | Applicator_38A | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | 3OA580/Applicator_38A | 2 | both
3OA580/Applicator_NY_Skip lot testing | 3OA580 | Applicator_NY_Skip lot testing | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 2.05 | g | FDA,2 |  |  | TM 610 | 3OA580/Applicator_NY_Skip lot testing | 25 | both
3OA580/Applicator_NY_Skip lot testing | 3OA580 | Applicator_NY_Skip lot testing | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer |  |  |  |  | TM 610 | 3OA580/Applicator_NY_Skip lot testing | 25 | both
3OA580/Applicator_NY_Skip lot testing | 3OA580 | Applicator_NY_Skip lot testing | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610 | 3OA580/Applicator_NY_Skip lot testing | 25 | both
3OA580/Applicator_NY_Skip lot testing | 3OA580 | Applicator_NY_Skip lot testing | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3OA580/Applicator_NY_Skip lot testing | 25 | both
3OA580/Applicator_NY_Skip lot testing | 3OA580 | Applicator_NY_Skip lot testing | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | 3OA580/Applicator_NY_Skip lot testing | 25 | both
3OA580/Applicator_NY_Skip lot testing | 3OA580 | Applicator_NY_Skip lot testing | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3OA580/Applicator_NY_Skip lot testing | 25 | both
3OA580/Applicator_NY_Skip lot testing | 3OA580 | Applicator_NY_Skip lot testing | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3OA580/Applicator_NY_Skip lot testing | 25 | both
3OA580/Applicator_NY_Skip lot testing | 3OA580 | Applicator_NY_Skip lot testing | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | 3OA580/Applicator_NY_Skip lot testing | 25 | both
3OA580/Applicator_NY_Skip lot testing | 3OA580 | Applicator_NY_Skip lot testing | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3OA580/Applicator_NY_Skip lot testing | 25 | both
3OA580/Applicator_NY_Skip lot testing | 3OA580 | Applicator_NY_Skip lot testing | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Mold (OOS) |  |  |  |  | WI QAS 076 | 3OA580/Applicator_NY_Skip lot testing | 25 | both
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements |  |  | left_only
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 |  |  | left_only
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208 |  |  | left_only
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208 |  |  | left_only
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight |  | GSM | FDA,0 | 65 | 75 | TM 101 |  |  | left_only
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper |  | mm | FDA,2 | 0.54 | 0.66 | TM 102 |  |  | left_only
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD |  | g/in | FDA,0 | 300 | 500 | TM 103 |  |  | left_only
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD |  | g/in | FDA,0 | 610 | 1040 | TM 103 |  |  | left_only
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width |  | in | FDA,3 | 6 | 6.125 | Applicator Measurements |  |  | left_only
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter |  | in | FDA,0 | 45 | 48 | Applicator Measurements |  |  | left_only
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 |  |  | left_only
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 100 | TM 610 |  |  | left_only
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) |  | cfu | FDA,0 |  |  | TM 610 |  |  | left_only
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 |  |  | left_only
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 15 | TM 610 |  |  | left_only
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  | cfu | FDA,0 |  |  | TM 610 |  |  | left_only
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight |  | g | FDA,2 |  |  | TM 610 |  |  | left_only
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer |  |  |  |  | TM 610 |  |  | left_only
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610 |  |  | left_only
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements |  |  | left_only
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 |  |  | left_only
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208 |  |  | left_only
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208 |  |  | left_only
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight |  | GSM | FDA,0 | 50 | 60 | TM 101 |  |  | left_only
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_CoA_(mm) | Caliper |  | mm | FDA,2 | 0.63 | 0.87 | TM 102 |  |  | left_only
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 758 | Tensile Strength | Dry_CD_(g/in) | Dry CD |  | g/in | FDA,0 | 1036 |  | TM 103 |  |  | left_only
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 758 | Tensile Strength | Dry_MD_(g/in) | Dry CD |  | g/in | FDA,0 | 4404 |  | TM 103 |  |  | left_only
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width |  | in | FDA,3 | 7.9 | 8.025 | Applicator Measurements |  |  | left_only
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 |  |  | left_only
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 100 | TM 610 |  |  | left_only
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) |  | cfu | FDA,0 |  |  | TM 610 |  |  | left_only
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight |  | g | FDA,2 |  |  | TM 610 |  |  | left_only
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 |  |  | left_only
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 15 | TM 610 |  |  | left_only
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  | cfu | FDA,0 |  |  | TM 610 |  |  | left_only
3OA911/Applicator_NY_Skip lot testing | 3OA911 | Applicator_NY_Skip lot testing | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Non-Subculturable Growth (Done) |  |  |  |  | WI QAS 076 |  |  | left_only
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | [5 empty] | TM 610 |  |  | left_only
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | [5 empty] | TM 610 |  |  | left_only
3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | 3OA972 | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements |  |  | left_only
3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | 3OA972 | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 |  |  | left_only
3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | 3OA972 | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208 |  |  | left_only
3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | 3OA972 | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208 |  |  | left_only
3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | 3OA972 | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight |  | GSM | FDA,1 | 35.2 | 44.8 | TM 101 |  |  | left_only
3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | 3OA972 | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 757 | Caliper | Caliper_CoA_(uM) | Caliper |  | micron | FDA,0 | 340 | 540 | TM 102 |  |  | left_only
3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | 3OA972 | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 758 | Tensile Strength | Wet_CD_CoA_(g/in) | Wet CD CoA |  | g/in | FDA,0 | 400 | 1800 | TM 103 |  |  | left_only
3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | 3OA972 | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 758 | Tensile Strength | Wet_MD_CoA_(g/in) | Wet MD CoA |  | g/in | FDA,0 | 1000 | 3400 | TM 103 |  |  | left_only
3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | 3OA972 | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width |  | in | FDA,3 | 8.875 | 9.125 | Applicator Measurements |  |  | left_only
3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | 3OA972 | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter |  | in | FDA,0 | 43 | 47 | Applicator Measurements |  |  | left_only
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | REDUCED TESTING_NY_3OB108 | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading |  |  |  |  | Color (TM 208) | 3OB108/REDUCED TESTING_NY_3OB108 | 1 | both
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | REDUCED TESTING_NY_3OB108 | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Meets specifications (Done) |  |  |  |  | See Analyte for Method # | 3OB108/REDUCED TESTING_NY_3OB108 | 1 | both
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | REDUCED TESTING_NY_3OB108 | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | See Analyte for Method # | 3OB108/REDUCED TESTING_NY_3OB108 | 1 | both
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | REDUCED TESTING_NY_3OB108 | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 54.9 | GSM | FDA,1 | 52.3 | 59.6 | See Analyte for Method # | 3OB108/REDUCED TESTING_NY_3OB108 | 1 | both
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | REDUCED TESTING_NY_3OB108 | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method # | 3OB108/REDUCED TESTING_NY_3OB108 | 1 | both
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | REDUCED TESTING_NY_3OB108 | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | Meets specifications (Done) |  |  |  |  | See Analyte for Method # | 3OB108/REDUCED TESTING_NY_3OB108 | 1 | both
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 1.25 | g | FDA,2 |  |  | TM 610 | 3OB108/REDUCED TESTING_NY_3OB108 | 1 | both
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | 3OB108/REDUCED TESTING_NY_3OB108 | 1 | both
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610 | 3OB108/REDUCED TESTING_NY_3OB108 | 1 | both
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3OB108/REDUCED TESTING_NY_3OB108 | 1 | both
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | 3OB108/REDUCED TESTING_NY_3OB108 | 1 | both
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3OB108/REDUCED TESTING_NY_3OB108 | 1 | both
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3OB108/REDUCED TESTING_NY_3OB108 | 1 | both
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | 3OB108/REDUCED TESTING_NY_3OB108 | 1 | both
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3OB108/REDUCED TESTING_NY_3OB108 | 1 | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading |  |  |  |  | Color (TM 208) | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 1 | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | May have residual sulfur odor when wet |  |  |  |  | Odor (TM 208) | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 1 | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Free of hard wrinkles, and formation uniform, with micro crepe texture (Done) |  |  |  |  | See Analyte for Method # | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 1 | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | See Analyte for Method # | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 1 | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 54.04 | GSM | FDA,0 | 49.5 | 60.5 | See Analyte for Method # | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 1 | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (mm) (TM 102) | Caliper | 0.55 | mm | FDA,2 | 0.52 | 0.74 | See Analyte for Method # | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 1 | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method # | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 1 | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter | 32 | in | FDA,0 | 28 | 36 | See Analyte for Method # | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 1 | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Roll Width (Inches) (Measurement) | Roll Width | 6 | in | FDA,0 | 5.875 | 6.125 | See Analyte for Method # | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 1 | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Wet CD (g/in) (TM 103) | Wet CD | 1004.84 | g/in | FDA,0 | 992 | 2208 | See Analyte for Method # | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 1 | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Wet MD (g/in) (TM 103) | Wet MD | 4435.12 | g/in | FDA,0 | 4392 | 7808 | See Analyte for Method # | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 1 | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 1.5 | g | FDA,2 |  |  | TM 610 | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 1 | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 1 | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610 | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 1 | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 1 | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 1 | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 1 | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 1 | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 1 | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 1 | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading |  |  |  |  | Color (TM 208) | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 1 | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | May have residual sulfur odor when wet |  |  |  |  | Odor (TM 208) | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 1 | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Meets specifications (Done) |  |  |  |  | See Analyte for Method # | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 1 | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 200 |  | FDA,0 |  |  | See Analyte for Method # | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 1 | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (lbs/ream) (TM 101) | Basis Weight | 23.81 | #/Ream | FDA,1 | 21.5 | 25 | See Analyte for Method # | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 1 | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Caliper at point of manufacture (mils) (TM 102) | Caliper | 4.9 | mil | FDA,1 | 4.2 | 5.6 | See Analyte for Method # | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 1 | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method # | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 1 | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter | 32 | in | FDA,0 | 31 | 33 | See Analyte for Method # | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 1 | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Roll Width (Inches) (Measurement) | Roll Width | 7.75 | in | FDA,4 | 7.6875 | 7.8125 | See Analyte for Method # | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 1 | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Wet CD (g/in) (TM 103) | Wet CD | 1261.89 | g/in | FDA,0 | 1000 | 2500 | See Analyte for Method # | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 1 | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Wet MD (g/in) (TM 103) | Wet MD | 2429.95 | g/in | FDA,0 | 1500 | 3850 | See Analyte for Method # | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 1 | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 1.1 | g | FDA,2 |  |  | TM 610 | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 1 | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 1 | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610 | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 1 | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 1 | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 1 | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 1 | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 1 | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 1 | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 1 | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 |  |  | left_only
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | REDUCED TESTING_NY_3OD438 | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading |  |  |  |  | Color (TM 208) | 3OD438/REDUCED TESTING_NY_3OD438 | 1 | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 100 | TM 610 |  |  | left_only
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | REDUCED TESTING_NY_3OD438 | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | May have residual sulfur odor when wet |  |  |  |  | Odor (TM 208) | 3OD438/REDUCED TESTING_NY_3OD438 | 1 | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) |  | cfu | FDA,0 |  |  | TM 610 |  |  | left_only
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | REDUCED TESTING_NY_3OD438 | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Smooth free of hard wrinkles, formation uniform matches embossing pattern (Done) |  |  |  |  | See Analyte for Method # | 3OD438/REDUCED TESTING_NY_3OD438 | 1 | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight |  | g | FDA,2 |  |  | TM 610 |  |  | left_only
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | REDUCED TESTING_NY_3OD438 | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 315 |  | FDA,0 |  |  | See Analyte for Method # | 3OD438/REDUCED TESTING_NY_3OD438 | 1 | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 |  |  | left_only
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | REDUCED TESTING_NY_3OD438 | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 35.09 | GSM | FDA,1 | 31.5 | 38.5 | See Analyte for Method # | 3OD438/REDUCED TESTING_NY_3OD438 | 1 | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 15 | TM 610 |  |  | left_only
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  | cfu | FDA,0 |  |  | TM 610 |  |  | left_only
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | REDUCED TESTING_NY_3OD438 | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method # | 3OD438/REDUCED TESTING_NY_3OD438 | 1 | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (inches) (TM 102) | Caliper |  | in | FDA,4 | 0.006 | 0.0087 | See Analyte for Method # |  |  | left_only
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Roll Width (Inches) (Measurement) | Roll Width |  | in | FDA,2 | 7.75 | 8 | See Analyte for Method # |  |  | left_only
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter |  | in | FDA,0 | 30 | 34 | See Analyte for Method # |  |  | left_only
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight |  | GSM | FDA,1 | 31.5 | 38.5 | See Analyte for Method # |  |  | left_only
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 |  |  | See Analyte for Method # |  |  | left_only
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Wet CD (g/in) (TM 103) | Wet CD |  | g/in | FDA,0 | 250 | 1250 | See Analyte for Method # |  |  | left_only
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 1.1 | g | FDA,2 |  |  | TM 610 | 3OD438/REDUCED TESTING_NY_3OD438 | 1 | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Dry CD (g/in) (TM103) | Dry CD |  | g/in | FDA,0 | 900 | 2200 | See Analyte for Method # |  |  | left_only
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Wet MD (g/in) (TM 103) | Wet MD |  | g/in | FDA,0 | 800 | 2600 | See Analyte for Method # |  |  | left_only
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | 3OD438/REDUCED TESTING_NY_3OD438 | 1 | both
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610 | 3OD438/REDUCED TESTING_NY_3OD438 | 1 | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Dry MD (g/in) (TM 103) | Dry CD |  | g/in | FDA,0 | 2000 | 5000 | See Analyte for Method # |  |  | left_only
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 1 | cfu/device | FDA,1 |  |  | TM 610 | 3OD438/REDUCED TESTING_NY_3OD438 | 1 | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method # |  |  | left_only
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0.18182 | CFU/g | FDA,1 |  | 100 | TM 610 | 3OD438/REDUCED TESTING_NY_3OD438 | 1 | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | May have residual sulfur odor when wet |  |  |  |  | Odor (TM 208) |  |  | left_only
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3OD438/REDUCED TESTING_NY_3OD438 | 1 | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Smooth free of hard wrinkles, formation uniform matches embossing pattern (Done) |  |  |  |  | See Analyte for Method # |  |  | left_only
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3OD438/REDUCED TESTING_NY_3OD438 | 1 | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading |  |  |  |  | Color (TM 208) |  |  | left_only
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | 3OD438/REDUCED TESTING_NY_3OD438 | 1 | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610 |  |  | left_only
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3OD438/REDUCED TESTING_NY_3OD438 | 1 | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 |  |  | left_only
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Positive Rods (Done) |  |  |  |  | WI QAS 076 | 3OD438/REDUCED TESTING_NY_3OD438 | 1 | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading |  |  |  |  | Color (TM 208) | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 1 | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | May have residual sulfur odor when wet |  |  |  |  | Odor (TM 208) | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 1 | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Free of hard wrinkles and formation uniform (Done) |  |  |  |  | See Analyte for Method # | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 1 | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 315 |  | FDA,0 |  |  | See Analyte for Method # | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 1 | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 46.72 | GSM | FDA,1 | 40.5 | 49.5 | See Analyte for Method # | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 1 | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (mm) (TM 102) | Caliper | 0.64 | mm | FDA,2 | 0.55 | 0.67 | See Analyte for Method # | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 1 | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method # | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 1 | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter | 32 | in | FDA,0 | 31 | 33 | See Analyte for Method # | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 1 | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Roll Width (Inches) (Measurement) | Roll Width | 7 | in | FDA,0 | 6.875 | 7.125 | See Analyte for Method # | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 1 | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Wet CD (g/in) (TM 103) | Wet CD | 2598 | g/in | FDA,0 | 1131 | 1593 | See Analyte for Method # | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 1 | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Wet MD (g/in) (TM 103) | Wet MD | 3711 | g/in | FDA,0 | 3906 | 4934 | See Analyte for Method # | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 1 | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 1.75 | g | FDA,2 |  |  | TM 610 | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 1 | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 1 | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610 | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 1 | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 1 | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 1 | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 1 | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 1 | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 1 | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 1 | both
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 |  |  | left_only
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 100 | TM 610 |  |  | left_only
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) |  | cfu | FDA,0 |  |  | TM 610 |  |  | left_only
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight |  | g | FDA,2 |  |  | TM 610 |  |  | left_only
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 |  |  | left_only
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 15 | TM 610 |  |  | left_only
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  | cfu | FDA,0 |  |  | TM 610 |  |  | left_only
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Caliper 1-ply (mm) (TM 102) | Caliper |  | mm | FDA,2 | 0.51 | 0.63 | See Analyte for Method # |  |  | left_only
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Roll Width (Inches) (Measurement) | Roll Width |  | in | FDA,3 | 6.875 | 7.125 | See Analyte for Method # |  |  | left_only
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter |  | in | FDA,0 | 30 | 32 | See Analyte for Method # |  |  | left_only
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight |  | GSM | FDA,1 | 45 | 55 | See Analyte for Method # |  |  | left_only
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 |  |  | See Analyte for Method # |  |  | left_only
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Dry CD (g/in) (TM103) | Dry CD |  | g/in | FDA,0 | 1801 | 2821 | See Analyte for Method # |  |  | left_only
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Dry MD (g/in) (TM 103) | Dry CD |  | g/in | FDA,0 | 6675 | 9447 | See Analyte for Method # |  |  | left_only
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method # |  |  | left_only
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | No objectionable odor detectable wet or dry (Done) |  |  |  |  | See Analyte for Method # |  |  | left_only
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Smooth free of hard wrinkles, formation uniform matches embossing pattern (Done) |  |  |  |  | See Analyte for Method # |  |  | left_only
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading |  |  |  |  | Color (TM 208) |  |  | left_only
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer |  |  |  |  | TM 610 |  |  | left_only
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610 |  |  | left_only
3OH592/Applicator_36A | 3OH592 | Applicator_36A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements |  |  | left_only
3OH592/Applicator_36A | 3OH592 | Applicator_36A | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 |  |  | left_only
3OH592/Applicator_36A | 3OH592 | Applicator_36A | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208 |  |  | left_only
3OH592/Applicator_36A | 3OH592 | Applicator_36A | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208 |  |  | left_only
3OH592/Applicator_36A | 3OH592 | Applicator_36A | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight |  | GSM | FDA,0 | 35 | 45 | TM 101 |  |  | left_only
3OH592/Applicator_36A | 3OH592 | Applicator_36A | NY-Analytical | 757 | Caliper | Caliper_CoA_(uM) | Caliper |  | um | FDA,0 | 340 | 540 | TM 102 |  |  | left_only
3OH592/Applicator_36A | 3OH592 | Applicator_36A | NY-Analytical | 758 | Tensile Strength | Wet_CD_CoA_(g/in) | Wet CD CoA |  | g/in | FDA,0 | 400 |  | TM 103 |  |  | left_only
3OH592/Applicator_36A | 3OH592 | Applicator_36A | NY-Analytical | 758 | Tensile Strength | Wet_MD_CoA_(g/in) | Wet MD CoA |  | g/in | FDA,0 | 1000 |  | TM 103 |  |  | left_only
3OH592/Applicator_36A | 3OH592 | Applicator_36A | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width |  | in | FDA,3 | 7.75 | 8 | Applicator Measurements |  |  | left_only
3OH592/Applicator_36A | 3OH592 | Applicator_36A | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter |  | in | FDA,0 | 30 | 34 | Applicator Measurements |  |  | left_only
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Applicator Measurements | 3OJ213/Applicator_08A | 4 | both
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width |  | in | FDA,3 | 11.875 | 12.125 | Applicator Measurements | 3OJ213/Applicator_08A | 4 | both
3OJ213/Applicator_NY_Skip lot testing | 3OJ213 | Applicator_NY_Skip lot testing | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 12 | in | FDA,3 | 11.875 | 12.125 | Applicator Measurements | 3OJ213/Applicator_NY_Skip lot testing | 6 | both
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 41.76 | GSM | FDA,0 | 36 | 44 | TM 101 | 3OJ213/Applicator_08A | 4 | both
3OJ213/Applicator_NY_Skip lot testing | 3OJ213 | Applicator_NY_Skip lot testing | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 40.9 | GSM | FDA,0 | 36 | 44 | TM 101 | 3OJ213/Applicator_NY_Skip lot testing | 6 | both
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_CoA_(mm) | Caliper | 0.54 | mm | FDA,2 | 0.48 | 0.61 | TM 102 | 3OJ213/Applicator_08A | 4 | both
3OJ213/Applicator_NY_Skip lot testing | 3OJ213 | Applicator_NY_Skip lot testing | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_CoA_(mm) | Caliper | 0.54 | mm | FDA,2 | 0.48 | 0.61 | TM 102 | 3OJ213/Applicator_NY_Skip lot testing | 6 | both
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 1096.49 | g/in | FDA,0 | 1030 | 1546 | TM 103 | 3OJ213/Applicator_08A | 4 | both
3OJ213/Applicator_NY_Skip lot testing | 3OJ213 | Applicator_NY_Skip lot testing | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 865 | g/in | FDA,0 | 1030 | 1546 | TM 103 | 3OJ213/Applicator_NY_Skip lot testing | 6 | both
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 4699.94 | g/in | FDA,0 | 3343 | 4350 | TM 103 | 3OJ213/Applicator_08A | 4 | both
3OJ213/Applicator_NY_Skip lot testing | 3OJ213 | Applicator_NY_Skip lot testing | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 3407 | g/in | FDA,0 | 3343 | 4350 | TM 103 | 3OJ213/Applicator_NY_Skip lot testing | 6 | both
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | 3OJ213/Applicator_08A | 4 | both
3OJ213/Applicator_NY_Skip lot testing | 3OJ213 | Applicator_NY_Skip lot testing | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | 3OJ213/Applicator_NY_Skip lot testing | 6 | both
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208 | 3OJ213/Applicator_08A | 4 | both
3OJ213/Applicator_NY_Skip lot testing | 3OJ213 | Applicator_NY_Skip lot testing | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208 | 3OJ213/Applicator_NY_Skip lot testing | 6 | both
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | 3OJ213/Applicator_08A | 4 | both
3OJ213/Applicator_NY_Skip lot testing | 3OJ213 | Applicator_NY_Skip lot testing | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | 3OJ213/Applicator_NY_Skip lot testing | 6 | both
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight |  | g | FDA,2 |  |  | TM 610 | 3OJ213/Applicator_08A | 4 | both
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | [5 empty] | TM 610 | 3OJ213/Applicator_08A | 4 | both
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | [5 empty] | TM 610 | 3OJ213/Applicator_08A | 4 | both
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 | 3OJ213/Applicator_08A | 4 | both
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 100 | TM 610 | 3OJ213/Applicator_08A | 4 | both
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) |  | cfu | FDA,0 |  |  | TM 610 | 3OJ213/Applicator_08A | 4 | both
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 | 3OJ213/Applicator_08A | 4 | both
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 15 | TM 610 | 3OJ213/Applicator_08A | 4 | both
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  | cfu | FDA,0 |  |  | TM 610 | 3OJ213/Applicator_08A | 4 | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Meets specifications (Done) |  |  |  |  | See Analyte for Method # | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 2 | both
3OM304/Applicator_NY_Skip lot testing | 3OM304 | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Meets specifications (Done) |  |  |  |  | See Analyte for Method # | 3OM304/Applicator_NY_Skip lot testing | 28 | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 315 |  | FDA,0 |  |  | See Analyte for Method # | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 2 | both
3OM304/Applicator_NY_Skip lot testing | 3OM304 | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 315 |  | FDA,0 |  |  | See Analyte for Method # | 3OM304/Applicator_NY_Skip lot testing | 28 | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 47.36 | GSM | FDA,0 | 45 | 55 | See Analyte for Method # | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 2 | both
3OM304/Applicator_NY_Skip lot testing | 3OM304 | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 50.38 | GSM | FDA,0 | 45 | 55 | See Analyte for Method # | 3OM304/Applicator_NY_Skip lot testing | 28 | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Caliper 1-ply (mm) (TM 102) | Caliper | 0.46 | mm | FDA,2 | 0.51 | 0.63 | See Analyte for Method # | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 2 | both
3OM304/Applicator_NY_Skip lot testing | 3OM304 | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Caliper 1-ply (mm) (TM 102) | Caliper | 0.53 | mm | FDA,2 | 0.51 | 0.63 | See Analyte for Method # | 3OM304/Applicator_NY_Skip lot testing | 28 | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method # | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 2 | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | Meets specifications (Done) |  |  |  |  | See Analyte for Method # | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 2 | both
3OM304/Applicator_NY_Skip lot testing | 3OM304 | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | Meets specifications (Done) |  |  |  |  | See Analyte for Method # | 3OM304/Applicator_NY_Skip lot testing | 28 | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Dry CD (g/in) (TM103) | Dry CD | 2142.67 | g/in | FDA,0 | 1801 | 2821 | See Analyte for Method # | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 2 | both
3OM304/Applicator_NY_Skip lot testing | 3OM304 | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Dry CD (g/in) (TM103) | Dry CD | 2395.04 | g/in | FDA,0 | 1801 | 2821 | See Analyte for Method # | 3OM304/Applicator_NY_Skip lot testing | 28 | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Dry MD (g/in) (TM 103) | Dry CD | 7892 | g/in | FDA,0 | 6675 | 9447 | See Analyte for Method # | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 2 | both
3OM304/Applicator_NY_Skip lot testing | 3OM304 | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Dry MD (g/in) (TM 103) | Dry CD | 8931.7 | g/in | FDA,0 | 6675 | 9447 | See Analyte for Method # | 3OM304/Applicator_NY_Skip lot testing | 28 | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | Meets specifications (Done) |  |  |  |  | See Analyte for Method # | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 2 | both
3OM304/Applicator_NY_Skip lot testing | 3OM304 | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | Meets specifications (Done) |  |  |  |  | See Analyte for Method # | 3OM304/Applicator_NY_Skip lot testing | 28 | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (COA) | Roll Diameter (COA) | 43 |  | FDA,0 | 41 | 45 | See Analyte for Method # | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 2 | both
3OM304/Applicator_NY_Skip lot testing | 3OM304 | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (COA) | Roll Diameter (COA) | 43 |  | FDA,0 | 41 | 45 | See Analyte for Method # | 3OM304/Applicator_NY_Skip lot testing | 28 | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Roll width (COA) | Roll width (COA) | 9 | in | FDA,2 | 8.75 | 9.25 | See Analyte for Method # | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 2 | both
3OM304/Applicator_NY_Skip lot testing | 3OM304 | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Roll width (COA) | Roll width (COA) | 9 | in | FDA,2 | 8.75 | 9.25 | See Analyte for Method # | 3OM304/Applicator_NY_Skip lot testing | 28 | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 0.89 | g | FDA,2 |  |  | TM 610 | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 2 | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 2 | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610 | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 2 | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 2 | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 2 | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 2 | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 2 | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 2 | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 2 | both
3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | 3OP065 | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (mm) (CofA) | Caliper |  | mm | FDA,2 | 0.67 | 0.81 | See Analyte for Method # |  |  | left_only
3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | 3OP065 | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Width (Inches) (Measurement) | Roll Width |  | in | FDA,2 | 29.75 | 30.25 | See Analyte for Method # |  |  | left_only
3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | 3OP065 | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter |  | in | FDA,0 | 45 | 47 | See Analyte for Method # |  |  | left_only
3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | 3OP065 | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight |  | GSM | FDA,1 | 54 | 66 | See Analyte for Method # |  |  | left_only
3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | 3OP065 | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 |  |  | See Analyte for Method # |  |  | left_only
3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | 3OP065 | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Wet CD (g/in) (TM 103) | Wet CD |  | g/in | FDA,0 | 1300 | 2000 | See Analyte for Method # |  |  | left_only
3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | 3OP065 | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Wet MD (g/in) (TM 103) | Wet MD |  | g/in | FDA,0 | 4000 | 5500 | See Analyte for Method # |  |  | left_only
3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | 3OP065 | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method # |  |  | left_only
3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | 3OP065 | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Free of hard wrinkles. Formation uniform (Done) |  |  |  |  | See Analyte for Method # |  |  | left_only
3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | 3OP065 | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | Meets specifications (Done) |  |  |  |  | See Analyte for Method # |  |  | left_only
3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | 3OP065 | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading |  |  |  |  | Color (TM 208) |  |  | left_only
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 1 | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter |  | in | FDA,0 | 45 | 47 | Applicator Measurements | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 1 | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width |  | in | FDA,2 | 29.75 | 30.25 | Applicator Measurements | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 1 | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 57.7 | GSM | FDA,1 | 49.5 | 60.5 | TM 101 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 1 | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper | 0.51 | mm | FDA,2 | 0.4 | 0.6 | TM 102 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 1 | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 758 | Tensile Strength | Elongation % Wet CD | Elongation % Wet CD | 103.07 | % | FDA,0 | 80 | 130 | TM 103 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 1 | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 758 | Tensile Strength | Elongation % Wet MD | Elongation % Wet MD | 24.68 | % | FDA,0 | 20 | 35 | TM 103 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 1 | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 893.28 | g/in | FDA,0 | 650 | 1300 | TM 103 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 1 | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 2898.19 | g/in | FDA,0 | 2400 | 3400 | TM 103 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 1 | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | [5 empty] | TM 208 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 1 | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 748 | Visual Inspection | Color | Color | [5 empty] | TM 208 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 1 | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 748 | Visual Inspection | Odor | Odor | [5 empty] | TM 208 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 1 | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 0.84 | g | FDA,2 |  |  | TM 610 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 1 | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer |  |  |  |  | TM 610 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 1 | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 1 | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 1 | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 1 | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 1 | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 1 | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 1 | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 1 | both
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Applicator_NY_Die Cut Pads | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements | 3OP093PD/Applicator_NY_Die Cut Pads | 3 | both
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Applicator_NY_Die Cut Pads | NY-Analytical | 759 | Applicator Measurements | Pad Diameter | Pad Diameter | 2.2188 | in | FDA,2 | 2.19 | 2.31 | Applicator Measurements | 3OP093PD/Applicator_NY_Die Cut Pads | 3 | both
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Applicator_NY_Die Cut Pads | NY-Analytical | 759 | Applicator Measurements | Stack Count | Stack Count | 90 | Units | FDA,0 | 90 | 90 | Applicator Measurements | 3OP093PD/Applicator_NY_Die Cut Pads | 3 | both
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Applicator_NY_Die Cut Pads | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | 3OP093PD/Applicator_NY_Die Cut Pads | 3 | both
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Applicator_NY_Die Cut Pads | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208 | 3OP093PD/Applicator_NY_Die Cut Pads | 3 | both
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Applicator_NY_Die Cut Pads | NY-Analytical | 748 | Visual Inspection | Odor | Odor | No objectionable odor Wet or Dry [Done] |  |  |  |  | TM 208 | 3OP093PD/Applicator_NY_Die Cut Pads | 3 | both
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Applicator_NY_Die Cut Pads | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 2.07 | g | FDA,2 |  |  | TM 610 | 3OP093PD/Applicator_NY_Die Cut Pads | 3 | both
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Applicator_NY_Die Cut Pads | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | 3OP093PD/Applicator_NY_Die Cut Pads | 3 | both
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Applicator_NY_Die Cut Pads | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610 | 3OP093PD/Applicator_NY_Die Cut Pads | 3 | both
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Applicator_NY_Die Cut Pads | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 1 | cfu/device | FDA,1 |  |  | TM 610 | 3OP093PD/Applicator_NY_Die Cut Pads | 3 | both
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Applicator_NY_Die Cut Pads | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0.042735 | CFU/g | FDA,1 |  | 100 | TM 610 | 3OP093PD/Applicator_NY_Die Cut Pads | 3 | both
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Applicator_NY_Die Cut Pads | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3OP093PD/Applicator_NY_Die Cut Pads | 3 | both
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Applicator_NY_Die Cut Pads | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3OP093PD/Applicator_NY_Die Cut Pads | 3 | both
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Applicator_NY_Die Cut Pads | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | 3OP093PD/Applicator_NY_Die Cut Pads | 3 | both
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Applicator_NY_Die Cut Pads | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3OP093PD/Applicator_NY_Die Cut Pads | 3 | both
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Applicator_NY_Die Cut Pads | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Positive Cocci (OOS) |  |  |  |  | WI QAS 076 | 3OP093PD/Applicator_NY_Die Cut Pads | 3 | both
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Applicator_NY_Die Cut Pads | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Non-Subculturable Growth (Done) |  |  |  |  | WI QAS 076 | 3OP093PD/Applicator_NY_Die Cut Pads | 3 | both
3OP113/Applicator_NY_Skip lot testing | 3OP113 | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (mm) (CofA) | Caliper | 0.62 | mm | FDA,2 | 0.5 | 0.62 | See Analyte for Method # | 3OP113/Applicator_NY_Skip lot testing | 3 | both
3OP113/Applicator_NY_Skip lot testing | 3OP113 | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter | 40 | in | FDA,0 | 40 | 46 | See Analyte for Method # | 3OP113/Applicator_NY_Skip lot testing | 3 | both
3OP113/Applicator_NY_Skip lot testing | 3OP113 | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Roll Width (Inches) (Measurement) | Roll Width | 30 | in | FDA,3 | 29.875 | 30.125 | See Analyte for Method # | 3OP113/Applicator_NY_Skip lot testing | 3 | both
3OP113/Applicator_NY_Skip lot testing | 3OP113 | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Wet CD (g/in) (TM 103) | Wet CD | 1469.66 | g/in | FDA,0 | 1142 | 1714 | See Analyte for Method # | 3OP113/Applicator_NY_Skip lot testing | 3 | both
3OP113/Applicator_NY_Skip lot testing | 3OP113 | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Wet MD (g/in) (TM 103) | Wet MD | 4083.47 | g/in | FDA,0 | 3265 | 4897 | See Analyte for Method # | 3OP113/Applicator_NY_Skip lot testing | 3 | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3 | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 40 | in | FDA,0 | 39 | 40 | Applicator Measurements | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3 | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 30 | in | FDA,3 | 29.875 | 30.125 | Applicator Measurements | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3 | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 34.01 | GSM | FDA,1 | 31.5 | 38.5 | TM 101 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3 | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(inch) | Caliper | 0.008 | in | FDA,4 | 0.006 | 0.009 | TM 102 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3 | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_CoA_(mm) | Caliper | 0.0079 | in |  |  |  | TM 102 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3 | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 812.99 | g/in | FDA,0 | 250 | 1250 | TM 103 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3 | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 1821.41 | g/in | FDA,0 | 800 | 2600 | TM 103 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3 | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3 | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 748 | Visual Inspection | Color | Color | White with no off colors or shading [Done] |  |  |  |  | TM 208 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3 | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3 | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 0.97 | g | FDA,2 |  |  | TM 610 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3 | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3 | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3 | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3 | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3 | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3 | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3 | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3 | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3 | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | [5 empty] | WI QAS 076 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3 | both
3OP138/Applicator_36D | 3OP138 | Applicator_36D | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Applicator Measurements | 3OP138/Applicator_36D | 2 | both
3OP138/Applicator_36D | 3OP138 | Applicator_36D | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter |  | in | FDA,0 | 43 | 47 | Applicator Measurements | 3OP138/Applicator_36D | 2 | both
3OP138/Applicator_36D | 3OP138 | Applicator_36D | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 30 | in | FDA,3 | 29.875 | 30.125 | Applicator Measurements | 3OP138/Applicator_36D | 2 | both
3OP138/Applicator_36D | 3OP138 | Applicator_36D | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 40.26 | GSM | FDA,0 | 35 | 45 | TM 101 | 3OP138/Applicator_36D | 2 | both
3OP138/Applicator_36D | 3OP138 | Applicator_36D | NY-Analytical | 757 | Caliper | Caliper_CoA_(uM) | Caliper |  | um | FDA,0 | 340 | 540 | TM 102 | 3OP138/Applicator_36D | 2 | both
3OP138/Applicator_36D | 3OP138 | Applicator_36D | NY-Analytical | 758 | Tensile Strength | Wet_CD_CoA_(g/in) | Wet CD CoA | 1228 | g/in | FDA,0 | 400 |  | TM 103 | 3OP138/Applicator_36D | 2 | both
3OP138/Applicator_36D | 3OP138 | Applicator_36D | NY-Analytical | 758 | Tensile Strength | Wet_MD_CoA_(g/in) | Wet MD CoA | 2545 | g/in | FDA,0 | 1000 |  | TM 103 | 3OP138/Applicator_36D | 2 | both
3OP138/Applicator_36D | 3OP138 | Applicator_36D | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | 3OP138/Applicator_36D | 2 | both
3OP138/Applicator_36D | 3OP138 | Applicator_36D | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208 | 3OP138/Applicator_36D | 2 | both
3OP138/Applicator_36D | 3OP138 | Applicator_36D | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | 3OP138/Applicator_36D | 2 | both
3OP197A/Applicator_NY_Skip lot testing | 3OP197A | Applicator_NY_Skip lot testing | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Bioburden_TMYC | 3OP197A/Applicator_NY_Skip lot testing | 2 | both
3OP197A/Applicator_NY_Skip lot testing | 3OP197A | Applicator_NY_Skip lot testing | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 0.92 | g | FDA,2 |  |  | TM 610 | 3OP197A/Applicator_NY_Skip lot testing | 2 | both
3OP197A/Applicator_NY_Skip lot testing | 3OP197A | Applicator_NY_Skip lot testing | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer |  |  |  |  | TM 610 | 3OP197A/Applicator_NY_Skip lot testing | 2 | both
3OP197A/Applicator_NY_Skip lot testing | 3OP197A | Applicator_NY_Skip lot testing | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610 | 3OP197A/Applicator_NY_Skip lot testing | 2 | both
3OP197A/Applicator_NY_Skip lot testing | 3OP197A | Applicator_NY_Skip lot testing | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3OP197A/Applicator_NY_Skip lot testing | 2 | both
3OP197A/Applicator_NY_Skip lot testing | 3OP197A | Applicator_NY_Skip lot testing | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | 3OP197A/Applicator_NY_Skip lot testing | 2 | both
3OP197A/Applicator_NY_Skip lot testing | 3OP197A | Applicator_NY_Skip lot testing | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3OP197A/Applicator_NY_Skip lot testing | 2 | both
3OP197A/Applicator_NY_Skip lot testing | 3OP197A | Applicator_NY_Skip lot testing | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 1 | cfu/device | FDA,1 |  |  | TM 610 | 3OP197A/Applicator_NY_Skip lot testing | 2 | both
3OP197A/Applicator_NY_Skip lot testing | 3OP197A | Applicator_NY_Skip lot testing | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0.76087 | CFU/g | FDA,1 |  | 15 | TM 610 | 3OP197A/Applicator_NY_Skip lot testing | 2 | both
3OP197A/Applicator_NY_Skip lot testing | 3OP197A | Applicator_NY_Skip lot testing | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 1 | cfu | FDA,0 |  |  | TM 610 | 3OP197A/Applicator_NY_Skip lot testing | 2 | both
3OP197A/Applicator_NY_Skip lot testing | 3OP197A | Applicator_NY_Skip lot testing | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Non-Subculturable Growth (Done) |  |  |  |  | WI QAS 076 | 3OP197A/Applicator_NY_Skip lot testing | 2 | both
3OP197A-SU/Label_NY_NP7200_6.25-4.5 | 3OP197A-SU | Label_NY_NP7200_6.25-4.5 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection | 3OP197A-SU/Label_NY_NP7200_6.25-4.5 | 1 | both
3OP197A-SU/Label_NY_NP7200_6.25-4.5 | 3OP197A-SU | Label_NY_NP7200_6.25-4.5 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | 3OP197A-SU/Label_NY_NP7200_6.25-4.5 | 1 | both
3OP197A-SU/Label_NY_NP7200_6.25-4.5 | 3OP197A-SU | Label_NY_NP7200_6.25-4.5 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | 3OP197A-SU/Label_NY_NP7200_6.25-4.5 | 1 | both
3OP197A-SU/Label_NY_NP7200_6.25-4.5 | 3OP197A-SU | Label_NY_NP7200_6.25-4.5 | NY-Analytical | 770 | Label Inspection | Length | Length | 6.25 | in | FDA,4 | 6.1875 | 6.3125 | TM 143 / Visual | 3OP197A-SU/Label_NY_NP7200_6.25-4.5 | 1 | both
3OP197A-SU/Label_NY_NP7200_6.25-4.5 | 3OP197A-SU | Label_NY_NP7200_6.25-4.5 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | 3OP197A-SU/Label_NY_NP7200_6.25-4.5 | 1 | both
3OP197A-SU/Label_NY_NP7200_6.25-4.5 | 3OP197A-SU | Label_NY_NP7200_6.25-4.5 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | 3OP197A-SU/Label_NY_NP7200_6.25-4.5 | 1 | both
3OP197A-SU/Label_NY_NP7200_6.25-4.5 | 3OP197A-SU | Label_NY_NP7200_6.25-4.5 | NY-Analytical | 770 | Label Inspection | Width | Width | 4.5 | in | FDA,4 | 4.4375 | 4.5625 | TM 143 / Visual | 3OP197A-SU/Label_NY_NP7200_6.25-4.5 | 1 | both
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Basis Weight |  |  | left_only
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 |  |  | left_only
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208 |  |  | left_only
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 748 | Visual Inspection | Color | Color | White with no off colors or shading [Done] |  |  |  |  | TM 208 |  |  | left_only
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight |  | GSM | FDA,1 | 31.5 | 38.5 | TM 101 |  |  | left_only
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper |  | mm | FDA,2 | 0.27 | 0.41 | TM 102 |  |  | left_only
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 758 | Tensile Strength | Wet_CD_CoA_(g/in) | Wet CD CoA |  | g/in | FDA,0 | 300 |  | TM 103 |  |  | left_only
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 758 | Tensile Strength | Wet_MD_CoA_(g/in) | Wet MD CoA |  | g/in | FDA,0 | 2000 | 4400 | TM 103 |  |  | left_only
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width |  | in | FDA,3 | 37.375 | 37.625 | Applicator Measurements |  |  | left_only
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight |  | g | FDA,2 |  |  | TM 610 |  |  | left_only
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) |  | cfu | FDA,0 |  |  | TM 610 |  |  | left_only
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 |  |  | left_only
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 100 | TM 610 |  |  | left_only
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 |  |  | left_only
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 15 | TM 610 |  |  | left_only
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  | cfu | FDA,0 |  |  | TM 610 |  |  | left_only
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Non-Subculturable Growth (Done) |  |  |  |  | WI QAS 076 |  |  | left_only
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610 |  |  | left_only
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 |  |  | left_only
3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | 3OP257 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Basis Weight |  |  | left_only
3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | 3OP257 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 |  |  | left_only
3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | 3OP257 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208 |  |  | left_only
3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | 3OP257 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 748 | Visual Inspection | Odor | Odor | No objectionable odor Wet or Dry [Done] |  |  |  |  | TM 208 |  |  | left_only
3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | 3OP257 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight |  | GSM | FDA,0 | 45 | 55 | TM 101 |  |  | left_only
3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | 3OP257 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper |  | mm | FDA,2 | 0.51 | 0.63 | TM 102 |  |  | left_only
3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | 3OP257 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 758 | Tensile Strength | Dry_CD_(g/in) | Dry CD |  | g/in | FDA,0 | 1801 | 2821 | TM 103 |  |  | left_only
3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | 3OP257 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 758 | Tensile Strength | Dry_MD_(g/in) | Dry CD |  | g/in | FDA,0 | 6675 | 9447 | TM 103 |  |  | left_only
3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | 3OP257 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 759 | Applicator Measurements | Roll width (COA) | Roll width (COA) |  | in | FDA,2 | 37.25 | 37.75 | Applicator Measurements |  |  | left_only
3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | 3OP257 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (COA) | Roll Diameter (COA) |  | in | FDA,0 | 41 | 45 | Applicator Measurements |  |  | left_only
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading |  |  |  |  | Color (TM 208) | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 1 | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 759 | Applicator Measurements | Roll Width (mm) | Roll Width | 953 | mm | FDA,0 | 950 | 956 | Applicator Measurements | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 1 | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Free of hard wrinkles, formation uniform/quilted big dot emboss pattern (Done) |  |  |  |  | See Analyte for Method # | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 1 | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 315 |  | FDA,0 |  |  | See Analyte for Method # | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 1 | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 32.08 | GSM | FDA,1 | 31.5 | 38.5 | See Analyte for Method # | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 1 | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (mm) (TM 102) | Caliper | 0.34 | mm | FDA,2 | 0.26 | 0.4 | See Analyte for Method # | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 1 | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method # | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 1 | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Dry CD (g/in) (TM103) | Dry CD | 1254.21 | g/in | FDA,0 | 1100 | 1700 | See Analyte for Method # | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 1 | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Dry MD (g/in) (TM 103) | Dry CD | 2070.52 | g/in | FDA,0 | 2000 | 3000 | See Analyte for Method # | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 1 | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | May have slight odor when wet (Done) |  |  |  |  | See Analyte for Method # | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 1 | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter | 40 | in | FDA,0 | 40 |  | See Analyte for Method # | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 1 | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Roll Width (Inches) (Measurement) | Roll Width | 37.5 | in | FDA,2 | 37.25 | 37.75 | See Analyte for Method # | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 1 | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 34.44 | GSM | FDA,1 | 31.5 | 38.5 | TM 101 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 1 | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper | 0.29 | mm | FDA,2 | 0.26 | 0.4 | TM 102 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 1 | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 758 | Tensile Strength | Dry_CD_(g/in) | Dry CD | 1143.94 | g/in | FDA,0 | 1100 | 1700 | TM 103 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 1 | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 758 | Tensile Strength | Dry_MD_(g/in) | Dry CD | 2220.05 | g/in | FDA,0 | 2000 | 3000 | TM 103 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 1 | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 1 | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 1 | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 1 | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 0.95 | g | FDA,2 |  |  | TM 610 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 1 | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer |  |  |  |  | TM 610 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 1 | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 1 | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 1 | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 1 | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 1 | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 1 | cfu/device | FDA,1 |  |  | TM 610 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 1 | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0.21053 | CFU/g | FDA,1 |  | 15 | TM 610 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 1 | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 1 | both
3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ222 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Certificate of Analysis | 3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 1 | both
3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis | 3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 1 | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Applicator Measurements | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 19 | both
3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ222 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter |  | in | FDA,0 | 42 | 46 | Applicator Measurements | 3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 1 | both
3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 44 | in | FDA,0 | 42 | 46 | Applicator Measurements | 3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 1 | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter |  | in | FDA,0 | 42 | 46 | Applicator Measurements | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 19 | both
3OQ222/Applicator_NY_Skip lot testing | 3OQ222 | Applicator_NY_Skip lot testing | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter |  | in | FDA,0 | 42 | 46 | Applicator Measurements | 3OQ222/Applicator_NY_Skip lot testing | 46 | both
3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ222 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width |  | in | FDA,3 | 37.375 | 37.625 | Applicator Measurements | 3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 1 | both
3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 36 | in | FDA,3 | 35.875 | 36.125 | Applicator Measurements | 3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 1 | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 30 | in | FDA,3 | 29.875 | 30.125 | Applicator Measurements | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 19 | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 41.12 | GSM | FDA,1 | 35.5 | 44.5 | TM 101 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 19 | both
3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 39.61 | GSM | FDA,1 | 35.5 | 44.5 | TM 101 | 3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 1 | both
3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ222 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight |  | GSM | FDA,1 | 35.5 | 44.5 | TM 101 | 3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 1 | both
3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper | 0.46 | mm | FDA,2 | 0.4 | 0.48 | TM 102 | 3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 1 | both
3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ222 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_CoA_(mm) | Caliper |  | mm | FDA,2 | 0.4 | 0.48 | TM 102 | 3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 1 | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_CoA_(mm) | Caliper | 0.45 | mm | FDA,2 | 0.4 | 0.48 | TM 102 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 19 | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 740.59 | g/in | FDA,0 | 600 | 1200 | TM 103 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 19 | both
3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 521.87 | g/in | FDA,0 | 600 | 1200 | TM 103 | 3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 1 | both
3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ222 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD |  | g/in | FDA,0 | 600 | 1200 | TM 103 | 3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 1 | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD |  | g/in | FDA,0 | 2200 | 3400 | TM 103 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 19 | both
3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 1610.35 | g/in | FDA,0 | 2200 | 3400 | TM 103 | 3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 1 | both
3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ222 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD |  | g/in | FDA,0 | 2200 | 3400 | TM 103 | 3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 1 | both
3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | 3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 1 | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | [5 empty] | TM 208 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 19 | both
3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ222 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | [5 empty] | TM 208 | 3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 1 | both
3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 748 | Visual Inspection | Color | Color | White with no off colors or shading [Done] |  |  |  |  | TM 208 | 3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 1 | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 748 | Visual Inspection | Color | Color | [5 empty] | TM 208 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 19 | both
3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ222 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 748 | Visual Inspection | Color | Color | [5 empty] | TM 208 | 3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 1 | both
3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | 3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 1 | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 748 | Visual Inspection | Odor | Odor | [5 empty] | TM 208 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 19 | both
3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ222 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 748 | Visual Inspection | Odor | Odor | [5 empty] | TM 208 | 3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 1 | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight |  | g | FDA,2 |  |  | TM 610 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 19 | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | [5 empty] | TM 610 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 19 | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | [5 empty] | TM 610 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 19 | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 19 | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 100 | TM 610 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 19 | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) |  | cfu | FDA,0 |  |  | TM 610 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 19 | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 19 | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 15 | TM 610 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 19 | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  | cfu | FDA,0 |  |  | TM 610 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 19 | both
3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ333 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 1.32 | g | FDA,2 |  |  | TM 610 | 3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 31 | both
3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ333 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | 3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 31 | both
3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ333 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610 | 3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 31 | both
3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ333 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 31 | both
3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ333 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | 3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 31 | both
3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ333 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 31 | both
3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ333 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 31 | both
3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ333 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | 3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 31 | both
3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ333 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 31 | both
3OQ444/Applicator_NY_Skip lot testing | 3OQ444 | Applicator_NY_Skip lot testing | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper | 0.41 | mm | FDA,2 | 0.4 | 0.48 | TM 102 | 3OQ444/Applicator_NY_Skip lot testing | 43 | both
3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ444 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight |  | g | FDA,2 |  |  | TM 610 | 3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 19 | both
3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ444 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | [5 empty] | TM 610 | 3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 19 | both
3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ444 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | [5 empty] | TM 610 | 3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 19 | both
3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ444 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 | 3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 19 | both
3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ444 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 100 | TM 610 | 3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 19 | both
3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ444 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) |  | cfu | FDA,0 |  |  | TM 610 | 3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 19 | both
3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ444 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 | 3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 19 | both
3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ444 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 15 | TM 610 | 3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 19 | both
3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ444 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  | cfu | FDA,0 |  |  | TM 610 | 3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 19 | both
3OQ551/Applicator_15E | 3OQ551 | Applicator_15E | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (COA) | Roll Diameter (COA) | 45 | in | FDA,0 | 43 | 47 | Applicator Measurements | 3OQ551/Applicator_15E | 4 | both
3OQ551/Applicator_15E | 3OQ551 | Applicator_15E | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 42 | in | FDA,3 | 41.875 | 42.125 | Applicator Measurements | 3OQ551/Applicator_15E | 4 | both
3OQ551/Applicator_15E | 3OQ551 | Applicator_15E | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Basis Weight | 3OQ551/Applicator_15E | 4 | both
3OQ551/Applicator_15E | 3OQ551 | Applicator_15E | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight |  | GSM | FDA,0 | 35 | 45 | TM 101 | 3OQ551/Applicator_15E | 4 | both
3OQ551/Applicator_15E | 3OQ551 | Applicator_15E | NY-Analytical | 757 | Caliper | Caliper_at_Pt_of_Mfr_(micron) | Caliper | 431.8 | um | FDA,0 | 340 | 540 | TM 102 | 3OQ551/Applicator_15E | 4 | both
3OQ551/Applicator_15E | 3OQ551 | Applicator_15E | NY-Analytical | 758 | Tensile Strength | Wet_CD_CoA_(g/25mm) | Wet CD CoA | 1073 | g/25mm | FDA,0 | 400 |  | TM 103 | 3OQ551/Applicator_15E | 4 | both
3OQ551/Applicator_15E | 3OQ551 | Applicator_15E | NY-Analytical | 758 | Tensile Strength | Wet_MD_CoA_(g/25mm) | Wet MD CoA | 2506 | g/25mm | FDA,0 | 1000 |  | TM 103 | 3OQ551/Applicator_15E | 4 | both
3OQ551/Applicator_15E | 3OQ551 | Applicator_15E | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | 3OQ551/Applicator_15E | 4 | both
3OQ551/Applicator_15E | 3OQ551 | Applicator_15E | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208 | 3OQ551/Applicator_15E | 4 | both
3OQ551/Applicator_15E | 3OQ551 | Applicator_15E | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | 3OQ551/Applicator_15E | 4 | both
3OQ850N/Applicator_01E | 3OQ850N | Applicator_01E | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Applicator Measurements |  |  | left_only
3OQ850N/Applicator_01E | 3OQ850N | Applicator_01E | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | [5 empty] | TM 208 |  |  | left_only
3OQ850N/Applicator_01E | 3OQ850N | Applicator_01E | NY-Analytical | 748 | Visual Inspection | Color | Color | [5 empty] | TM 208 |  |  | left_only
3OQ850N/Applicator_01E | 3OQ850N | Applicator_01E | NY-Analytical | 748 | Visual Inspection | Odor | Odor | [5 empty] | TM 208 |  |  | left_only
3OQ850N/Applicator_01E | 3OQ850N | Applicator_01E | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight |  | GSM | FDA,0 | 38 | 42 | TM 101 |  |  | left_only
3OQ850N/Applicator_01E | 3OQ850N | Applicator_01E | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper |  | mm | FDA,2 | 0.35 | 0.45 | TM 102 |  |  | left_only
3OQ850N/Applicator_01E | 3OQ850N | Applicator_01E | NY-Analytical | 758 | Tensile Strength | Dry_CD_(g/in) | Dry CD |  | g/in | FDA,0 | 513 | 1130 | TM 103 |  |  | left_only
3OQ850N/Applicator_01E | 3OQ850N | Applicator_01E | NY-Analytical | 758 | Tensile Strength | Dry_MD_(g/in) | Dry CD |  | g/in | FDA,0 | 2054 | 4108 | TM 103 |  |  | left_only
3OQ850N/Applicator_01E | 3OQ850N | Applicator_01E | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter |  | in | FDA,0 | 44 | 47 | Applicator Measurements |  |  | left_only
3OQ850N/Applicator_01E | 3OQ850N | Applicator_01E | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width |  | in | FDA,3 | 35.875 | 36.125 | Applicator Measurements |  |  | left_only
3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | 3OQ850N | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (mm) (TM 102) | Caliper |  | mm |  |  |  | See Analyte for Method # |  |  | left_only
3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | 3OQ850N | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Roll Width (Inches) (Measurement) | Roll Width |  | in | FDA,3 | 37.375 | 37.625 | See Analyte for Method # |  |  | left_only
3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | 3OQ850N | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight |  | GSM | FDA,1 | 38 | 42 | See Analyte for Method # |  |  | left_only
3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | 3OQ850N | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter |  | in | FDA,0 | 44 | 47 | See Analyte for Method # |  |  | left_only
3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | 3OQ850N | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 |  |  | See Analyte for Method # |  |  | left_only
3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | 3OQ850N | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Dry CD (g/in) (TM103) | Dry CD |  | g/in | FDA,0 | 513 | 1130 | See Analyte for Method # |  |  | left_only
3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | 3OQ850N | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Dry MD (g/in) (TM 103) | Dry CD |  | g/in | FDA,0 | 2054 | 4108 | See Analyte for Method # |  |  | left_only
3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | 3OQ850N | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method # |  |  | left_only
3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | 3OQ850N | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | No objectionable odor detectable wet or dry (Done) |  |  |  |  | See Analyte for Method # |  |  | left_only
3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | 3OQ850N | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Smooth, free of hard wrinkles and formation uniform (Done) |  |  |  |  | See Analyte for Method # |  |  | left_only
3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | 3OQ850N | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading |  |  |  |  | Color (TM 208) |  |  | left_only
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Swab Measurements | 3OS407/Applicator_25B | 6 | both
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Analytical | 760 | Swab Measurements | Mitt_Length_(Inches) | Mitt Length | 0.99 | in | FDA,3 | 0.97 | 1.03 | TM 109 | 3OS407/Applicator_25B | 6 | both
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Analytical | 760 | Swab Measurements | Mitt_Thickness_(Inches) | Mitt Thickness | 0.314 | in | FDA,3 | 0.265 | 0.385 | TM 109 | 3OS407/Applicator_25B | 6 | both
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Analytical | 760 | Swab Measurements | Mitt_Width_(Inches) | Mitt Width | 0.49 | in | FDA,2 | 0.47 | 0.53 | TM 109 | 3OS407/Applicator_25B | 6 | both
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Analytical | 760 | Swab Measurements | Overall_Swab_Length_(Inches) | Overall Swab Length | 4.77 | in | FDA,3 | 4.69 | 4.81 | TM 109 | 3OS407/Applicator_25B | 6 | both
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Analytical | 760 | Swab Measurements | Scrub_Test | Scrub Test | Meets specifications (Done) |  |  |  |  | TM 109 | 3OS407/Applicator_25B | 6 | both
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Analytical | 760 | Swab Measurements | Tuft_Adhesion_(Lb.) | Tuft Adhesion | 2.59 | lb | FDA,1 | 1.4 |  | TM 109 | 3OS407/Applicator_25B | 6 | both
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Analytical | 760 | Swab Measurements | Use_Test-Solvents | Use Test-Solvents | Meets specifications (Done) |  |  |  |  | TM 109 | 3OS407/Applicator_25B | 6 | both
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 1.2 | g | FDA,2 |  |  | TM 610 | 3OS407/Applicator_25B | 6 | both
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | 3OS407/Applicator_25B | 6 | both
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610 | 3OS407/Applicator_25B | 6 | both
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3OS407/Applicator_25B | 6 | both
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | 3OS407/Applicator_25B | 6 | both
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) |  | cfu | FDA,0 |  |  | TM 610 | 3OS407/Applicator_25B | 6 | both
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3OS407/Applicator_25B | 6 | both
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 15 | TM 610 | 3OS407/Applicator_25B | 6 | both
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  | cfu | FDA,0 |  |  | TM 610 | 3OS407/Applicator_25B | 6 | both
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Positive Rods (Done) |  |  |  |  | WI QAS 076 | 3OS407/Applicator_25B | 6 | both
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Non-Subculturable Growth (Done) |  |  |  |  | WI QAS 076 | 3OS407/Applicator_25B | 6 | both
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Swab Measurements | 3OS419/Applicator_25C | 3 | both
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Analytical | 760 | Swab Measurements | Mitt_Length_(Inches) | Mitt Length | 1.22 | in | FDA,2 | 1.16 | 1.36 | TM 109 | 3OS419/Applicator_25C | 3 | both
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Analytical | 760 | Swab Measurements | Mitt_Thickness_(Inches) | Mitt Thickness | 0.47 | in | FDA,3 | 0.41 | 0.53 | TM 109 | 3OS419/Applicator_25C | 3 | both
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Analytical | 760 | Swab Measurements | Mitt_Width_(Inches) | Mitt Width | 1.19 | in | FDA,2 | 1.1 | 1.3 | TM 109 | 3OS419/Applicator_25C | 3 | both
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Analytical | 760 | Swab Measurements | Overall_Swab_Length_(Inches) | Overall Swab Length | 6.03 | in | FDA,2 | 5.85 | 6.15 | TM 109 | 3OS419/Applicator_25C | 3 | both
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Analytical | 760 | Swab Measurements | Scrub_Test | Scrub Test | Meets specifications (Done) |  |  |  |  | TM 109 | 3OS419/Applicator_25C | 3 | both
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Analytical | 760 | Swab Measurements | Tuft_Adhesion_(Lb.) | Tuft Adhesion | 5.12 | lb | FDA,1 | 1.4 |  | TM 109 | 3OS419/Applicator_25C | 3 | both
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Analytical | 760 | Swab Measurements | Use_Test-Solvents | Use Test-Solvents | Meets specifications (Done) |  |  |  |  | TM 109 | 3OS419/Applicator_25C | 3 | both
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 1.26 | g | FDA,2 |  |  | TM 610 | 3OS419/Applicator_25C | 3 | both
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | [5 empty] | TM 610 | 3OS419/Applicator_25C | 3 | both
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | [5 empty] | TM 610 | 3OS419/Applicator_25C | 3 | both
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 | 3OS419/Applicator_25C | 3 | both
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 100 | TM 610 | 3OS419/Applicator_25C | 3 | both
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) |  | cfu | FDA,0 |  |  | TM 610 | 3OS419/Applicator_25C | 3 | both
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3OS419/Applicator_25C | 3 | both
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | 3OS419/Applicator_25C | 3 | both
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3OS419/Applicator_25C | 3 | both
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Positive Rods (Done) |  |  |  |  | WI QAS 076 | 3OS419/Applicator_25C | 3 | both
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | REDUCED TESTING_NY_3OS428 | NY-Analytical | 760 | Swab Measurements | Tuft_Adhesion_(Lb.) | Tuft Adhesion |  | lb | FDA,1 | 1.4 |  | TM 109 |  |  | left_only
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | REDUCED TESTING_NY_3OS428 | NY-Analytical | 760 | Swab Measurements | AQL Sample Size | AQL Sample Size |  | in |  |  |  | TM 109 |  |  | left_only
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | REDUCED TESTING_NY_3OS428 | NY-Analytical | 760 | Swab Measurements | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | TM 109 |  |  | left_only
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | REDUCED TESTING_NY_3OS428 | NY-Analytical | 760 | Swab Measurements | Scrub_Test | Scrub Test | Pass (Done) |  |  |  |  | TM 109 |  |  | left_only
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 |  |  | left_only
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 100 | TM 610 |  |  | left_only
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) |  | cfu | FDA,0 |  |  | TM 610 |  |  | left_only
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight |  | g | FDA,2 |  |  | TM 610 |  |  | left_only
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 |  |  | left_only
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 15 | TM 610 |  |  | left_only
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  | cfu | FDA,0 |  |  | TM 610 |  |  | left_only
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610 |  |  | left_only
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 |  |  | left_only
3OS484/Applicator_26A | 3OS484 | Applicator_26A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Swab Measurements |  |  | left_only
3OS484/Applicator_26A | 3OS484 | Applicator_26A | NY-Analytical | 760 | Swab Measurements | Tuft_Diameter_(Inches) | Tuft Diameter |  | in | FDA,2 | 0.42 | 0.53 | TM 109 |  |  | left_only
3OS484/Applicator_26A | 3OS484 | Applicator_26A | NY-Analytical | 760 | Swab Measurements | Overall_Swab_Length_(Inches) | Overall Swab Length |  | in | FDA,2 | 7.82 | 7.94 | TM 109 |  |  | left_only
3OS484/Applicator_26A | 3OS484 | Applicator_26A | NY-Analytical | 760 | Swab Measurements | Scrub_Test | Scrub Test | Meets specifications (Done) |  |  |  |  | TM 109 |  |  | left_only
3OS484/Applicator_26A | 3OS484 | Applicator_26A | NY-Analytical | 760 | Swab Measurements | Use_Test-Solvents | Use Test-Solvents | Meets specifications (Done) |  |  |  |  | TM 109 |  |  | left_only
3OS484/Applicator_26A | 3OS484 | Applicator_26A | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) |  | cfu | FDA,0 |  |  | TM 610 |  |  | left_only
3OS484/Applicator_26A | 3OS484 | Applicator_26A | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 100 | TM 610 |  |  | left_only
3OS484/Applicator_26A | 3OS484 | Applicator_26A | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 |  |  | left_only
3OS484/Applicator_26A | 3OS484 | Applicator_26A | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 |  |  | left_only
3OS484/Applicator_26A | 3OS484 | Applicator_26A | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 15 | TM 610 |  |  | left_only
3OS484/Applicator_26A | 3OS484 | Applicator_26A | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  | cfu | FDA,0 |  |  | TM 610 |  |  | left_only
3OS484/Applicator_26A | 3OS484 | Applicator_26A | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight |  | g | FDA,2 |  |  | TM 610 |  |  | left_only
3OS484/Applicator_26A | 3OS484 | Applicator_26A | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Negative Rods (OOS) |  |  |  |  | WI QAS 076 |  |  | left_only
3OS484/Applicator_26A | 3OS484 | Applicator_26A | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610 |  |  | left_only
3OS484/Applicator_26A | 3OS484 | Applicator_26A | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 |  |  | left_only
3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | 3OU252 | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | [5 empty] | See Analyte for Method # |  |  | left_only
3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | 3OU252 | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 |  |  | See Analyte for Method # |  |  | left_only
3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | 3OU252 | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight |  | GSM | FDA,1 | 64.4 | 71.2 | See Analyte for Method # |  |  | left_only
3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | 3OU252 | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (mm) (TM 102) | Caliper |  | mm | FDA,2 | 0.28 | 0.42 | See Analyte for Method # |  |  | left_only
3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | 3OU252 | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | [5 empty] | See Analyte for Method # |  |  | left_only
3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | 3OU252 | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | [5 empty] | See Analyte for Method # |  |  | left_only
3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | 3OU252 | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | [5 empty] | See Analyte for Method # |  |  | left_only
3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | 3OU252 | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (COA) | Roll Diameter (COA) |  |  | FDA,0 | 41 | 45 | See Analyte for Method # |  |  | left_only
3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | 3OU252 | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Roll width (COA) | Roll width (COA) |  | in | FDA,2 | 8.94 | 9.13 | See Analyte for Method # |  |  | left_only
3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | 3OU252 | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Wet CD (g/in) (TM 103) | Wet CD |  | g/in | FDA,0 | 1500 | 2500 | See Analyte for Method # |  |  | left_only
3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | 3OU252 | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Wet MD (g/in) (TM 103) | Wet MD |  | g/in | FDA,0 | 4500 | 6500 | See Analyte for Method # |  |  | left_only
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements |  |  | left_only
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 |  |  | left_only
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208 |  |  | left_only
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208 |  |  | left_only
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight |  | GSM | FDA,1 | 45 | 55 | TM 101 |  |  | left_only
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Analytical | 757 | Caliper | Caliper-1_Ply_(mm) | Caliper |  | mm | FDA,2 | 0.51 | 0.63 | TM 102 |  |  | left_only
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Analytical | 758 | Tensile Strength | Dry_CD_(g/in) | Dry CD |  | g/in | FDA,0 | 1801 | 2821 | TM 103 |  |  | left_only
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Analytical | 758 | Tensile Strength | Dry_MD_(g/in) | Dry CD |  | g/in | FDA,0 | 6675 | 9447 | TM 103 |  |  | left_only
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width |  | in | FDA,2 | 11.25 | 11.75 | Applicator Measurements |  |  | left_only
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter |  | in | FDA,0 | 41 | 45 | Applicator Measurements |  |  | left_only
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 |  |  | left_only
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 100 | TM 610 |  |  | left_only
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) |  | cfu | FDA,0 |  |  | TM 610 |  |  | left_only
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 |  |  | left_only
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 15 | TM 610 |  |  | left_only
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  | cfu | FDA,0 |  |  | TM 610 |  |  | left_only
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight |  | g | FDA,2 |  |  | TM 610 |  |  | left_only
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer |  |  |  |  | TM 610 |  |  | left_only
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610 |  |  | left_only
3OU275/Applicator_36C | 3OU275 | Applicator_36C | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements |  |  | left_only
3OU275/Applicator_36C | 3OU275 | Applicator_36C | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 |  |  | left_only
3OU275/Applicator_36C | 3OU275 | Applicator_36C | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208 |  |  | left_only
3OU275/Applicator_36C | 3OU275 | Applicator_36C | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208 |  |  | left_only
3OU275/Applicator_36C | 3OU275 | Applicator_36C | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight |  | GSM | FDA,0 | 35 | 45 | TM 101 |  |  | left_only
3OU275/Applicator_36C | 3OU275 | Applicator_36C | NY-Analytical | 757 | Caliper | Caliper_CoA_(uM) | Caliper |  | um | FDA,0 | 340 | 540 | TM 102 |  |  | left_only
3OU275/Applicator_36C | 3OU275 | Applicator_36C | NY-Analytical | 758 | Tensile Strength | Wet_CD_CoA_(g/in) | Wet CD CoA |  | g/in | FDA,0 | 400 |  | TM 103 |  |  | left_only
3OU275/Applicator_36C | 3OU275 | Applicator_36C | NY-Analytical | 758 | Tensile Strength | Wet_MD_CoA_(g/in) | Wet MD CoA |  | g/in | FDA,0 | 1000 |  | TM 103 |  |  | left_only
3OU275/Applicator_36C | 3OU275 | Applicator_36C | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width |  | in | FDA,3 | 11.625 | 11.875 | Applicator Measurements |  |  | left_only
3OU275/Applicator_36C | 3OU275 | Applicator_36C | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter |  | in | FDA,0 | 30 | 34 | Applicator Measurements |  |  | left_only
3OU351/Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | 3OU351 | Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA | 3OU351/Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | 3 | both
3OU351/Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | 3OU351 | Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis (ELLI) | Certificate of Analysis | COA Meets Specification (Done) |  |  |  |  | COA | 3OU351/Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | 3 | both
3OU351/Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | 3OU351 | Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 5.968 | in | FDA,3 | 5.875 | 6.125 | TM 143 | 3OU351/Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | 3 | both
3OU351/Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | 3OU351 | Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | NY-Analytical | 397 | Appearance | Appearance/Matches Embossing Pattern | Appearance/Matches Embossing Pattern | Free of Hard wrinkles. Formation uniform (Done) |  |  |  |  | TM 208 | 3OU351/Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | 3 | both
3OU351/Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | 3OU351 | Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | NY-Analytical | 395 | Color | Color | Color | White with no off colors or shading (Done) |  |  |  |  | TM 208 | 3OU351/Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | 3 | both
3OU351/Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | 3OU351 | Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | NY-Analytical | 1038 | KC Part Number | Verify KC Part Number | Verify KC Part Number | SPEC-08468 (Done) | [5 empty] | 3OU351/Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | 3 | both
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Applicatpr_NY_Foam 10.5ml | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Conformance | 3OX105F/Applicatpr_NY_Foam 10.5ml | 1 | both
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Applicatpr_NY_Foam 10.5ml | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Bioburden_TAMC | 3OX105F/Applicatpr_NY_Foam 10.5ml | 1 | both
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 0.91 | g | FDA,2 |  |  | TM 610 | 3OX105F/Applicatpr_NY_Foam 10.5ml | 1 | both
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | 3OX105F/Applicatpr_NY_Foam 10.5ml | 1 | both
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610 | 3OX105F/Applicatpr_NY_Foam 10.5ml | 1 | both
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3OX105F/Applicatpr_NY_Foam 10.5ml | 1 | both
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | 3OX105F/Applicatpr_NY_Foam 10.5ml | 1 | both
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3OX105F/Applicatpr_NY_Foam 10.5ml | 1 | both
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 1 | cfu/device | FDA,1 |  |  | TM 610 | 3OX105F/Applicatpr_NY_Foam 10.5ml | 1 | both
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0.32967 | CFU/g | FDA,1 |  | 15 | TM 610 | 3OX105F/Applicatpr_NY_Foam 10.5ml | 1 | both
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3OX105F/Applicatpr_NY_Foam 10.5ml | 1 | both
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Applicatpr_NY_Foam 10.5ml | NY-Analytical | 397 | Appearance | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | Visual | 3OX105F/Applicatpr_NY_Foam 10.5ml | 1 | both
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Mold (OOS) |  |  |  |  | WI QAS 076 | 3OX105F/Applicatpr_NY_Foam 10.5ml | 1 | both
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA | 3OX465H/Handle, 26 ml | 1 | both
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection | 3OX465H/Handle, 26 ml | 1 | both
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Natural (Done) |  |  |  |  | Canister & Lid Inspection | 3OX465H/Handle, 26 ml | 1 | both
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | NY-Analytical | 765 | Canister & Lid Inspection | Handle Internal Diameter | Handle Internal Diameter | 0.9575 | in | FDA,3 | 0.951 | 0.957 | Canister & Lid Inspection | 3OX465H/Handle, 26 ml | 1 | both
3OX465H/Microbiologu - Bioburden_Applicator_NY | 3OX465H | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 765 | Canister & Lid Inspection | Handle Internal Diameter | Handle Internal Diameter | 0.9545 | in | FDA,3 | 0.951 | 0.957 | Canister & Lid Inspection | 3OX465H/Microbiologu - Bioburden_Applicator_NY | 1 | both
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | NY-Analytical | 765 | Canister & Lid Inspection | Overall Length | Overall Length | 5.23 | in | FDA,3 | 5.2 | 5.25 | Canister & Lid Inspection | 3OX465H/Handle, 26 ml | 1 | both
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 87.86 | g | FDA,2 |  |  | TM 610 | 3OX465H/Handle, 26 ml | 1 | both
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | 3OX465H/Handle, 26 ml | 1 | both
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610 | 3OX465H/Handle, 26 ml | 1 | both
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 3 | cfu/device | FDA,1 |  |  | TM 610 | 3OX465H/Handle, 26 ml | 1 | both
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0.034145 | CFU/g | FDA,1 |  | 100 | TM 610 | 3OX465H/Handle, 26 ml | 1 | both
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3OX465H/Handle, 26 ml | 1 | both
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3OX465H/Handle, 26 ml | 1 | both
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | 3OX465H/Handle, 26 ml | 1 | both
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 13 | cfu | FDA,0 |  |  | TM 610 | 3OX465H/Handle, 26 ml | 1 | both
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Non-Subculturable Growth (Done) |  |  |  |  | WI QAS 076 | 3OX465H/Handle, 26 ml | 1 | both
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Gram Positive Rods (Done) |  |  |  |  | WI QAS 076 | 3OX465H/Handle, 26 ml | 1 | both
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swab_NY_Swabstick, Cotton-Tipped | NY-Analytical | 760 | Swab Measurements | Mitt_Length_(Inches) | Mitt Length |  | in | FDA,3 | 0.593 | 0.781 | Swab Measurements | 3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 1 | both
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swab_NY_Swabstick, Cotton-Tipped | NY-Analytical | 760 | Swab Measurements | Mitt_Width_(Inches) | Mitt Width |  | in | FDA,3 | 0.188 | 0.28 | Swab Measurements | 3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 1 | both
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swab_NY_Swabstick, Cotton-Tipped | NY-Analytical | 760 | Swab Measurements | Overall_Swab_Length_(Inches) | Overall Swab Length |  | in | FDA,3 | 5.875 | 5.999 | Swab Measurements | 3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 1 | both
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swab_NY_Swabstick, Cotton-Tipped | NY-Analytical | 760 | Swab Measurements | Stick Diameter | Stick Diameter |  | in | FDA,3 | 0.093 | 0.107 | Swab Measurements | 3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 1 | both
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swab_NY_Swabstick, Cotton-Tipped | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Certificate of Conformance | 3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 1 | both
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swab_NY_Swabstick, Cotton-Tipped | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | [5 empty] | Bioburden_TAMC | 3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 1 | both
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 0.58 | g | FDA,2 |  |  | TM 610 | 3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 1 | both
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | 3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 1 | both
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610 | 3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 1 | both
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 1 | both
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram) | TAMC (Average per Gram) | 0 | CFU/g |  |  |  | TM 610 | 3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 1 | both
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu |  |  |  | TM 610 | 3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 1 | both
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 1 | both
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | 3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 1 | both
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 1 | both
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA | 3OX510H/Handle, 5.1 ml | 1 | both
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection | 3OX510H/Handle, 5.1 ml | 1 | both
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Natural (Done) |  |  |  |  | Canister & Lid Inspection | 3OX510H/Handle, 5.1 ml | 1 | both
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | NY-Analytical | 765 | Canister & Lid Inspection | Handle Internal Diameter | Handle Internal Diameter | 0.6405 | in | FDA,3 | 0.639 | 0.645 | Canister & Lid Inspection | 3OX510H/Handle, 5.1 ml | 1 | both
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | NY-Analytical | 765 | Canister & Lid Inspection | Overall Length | Overall Length | 3.156 | in | FDA,3 | 3.149 | 3.189 | Canister & Lid Inspection | 3OX510H/Handle, 5.1 ml | 1 | both
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 1.5 | g | FDA,2 |  |  | TM 610 | 3OX510H/Handle, 5.1 ml | 1 | both
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer |  |  |  |  | TM 610 | 3OX510H/Handle, 5.1 ml | 1 | both
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610 | 3OX510H/Handle, 5.1 ml | 1 | both
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3OX510H/Handle, 5.1 ml | 1 | both
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | 3OX510H/Handle, 5.1 ml | 1 | both
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3OX510H/Handle, 5.1 ml | 1 | both
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3OX510H/Handle, 5.1 ml | 1 | both
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | 3OX510H/Handle, 5.1 ml | 1 | both
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3OX510H/Handle, 5.1 ml | 1 | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | 4 | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 39 | in | FDA,0 | 39 | 40 | Applicator Measurements | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | 4 | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 36 | in | FDA,3 | 35.875 | 36.125 | Applicator Measurements | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | 4 | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (inches) (TM 102) | Caliper | 0.0082 | in | FDA,3 | 0.006 | 0.009 | See Analyte for Method # | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | 4 | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 34.44 | GSM | FDA,1 | 31.5 | 38.5 | TM 101 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | 4 | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 842.87 | g/in | FDA,0 | 250 | 1250 | TM 103 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | 4 | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 1755.59 | g/in | FDA,0 | 800 | 2600 | TM 103 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | 4 | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | 4 | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 748 | Visual Inspection | Color | Color | White with no off colors or shading [Done] |  |  |  |  | TM 208 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | 4 | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | 4 | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 1.55 | g | FDA,2 |  |  | TM 610 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | 4 | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | 4 | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | 4 | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 1 | cfu/device | FDA,1 |  |  | TM 610 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | 4 | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0.12903 | CFU/g | FDA,1 |  | 100 | TM 610 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | 4 | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 2 | cfu | FDA,0 |  |  | TM 610 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | 4 | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | 4 | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | 4 | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | 4 | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Negative Rods (OOS) |  |  |  |  | WI QAS 076 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | 4 | both
3XP138/Applicator_36B | 3XP138 | Applicator_36B | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements | 3XP138/Applicator_36B | 1 | both
3XP138/Applicator_36B | 3XP138 | Applicator_36B | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 45 | in | FDA,0 | 43 | 47 | Applicator Measurements | 3XP138/Applicator_36B | 1 | both
3XP138/Applicator_36B | 3XP138 | Applicator_36B | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 37.5 | in | FDA,3 | 37.375 | 37.625 | Applicator Measurements | 3XP138/Applicator_36B | 1 | both
3XP138/Applicator_36B | 3XP138 | Applicator_36B | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 40.04 | GSM | FDA,0 | 35 | 45 | TM 101 | 3XP138/Applicator_36B | 1 | both
3XP138/Applicator_36B | 3XP138 | Applicator_36B | NY-Analytical | 757 | Caliper | Caliper_CoA_(uM) | Caliper | 419.8 | um | FDA,0 | 340 | 540 | TM 102 | 3XP138/Applicator_36B | 1 | both
3XP138/Applicator_36B | 3XP138 | Applicator_36B | NY-Analytical | 758 | Tensile Strength | Wet_CD_CoA_(g/in) | Wet CD CoA | 1069 | g/in | FDA,0 | 400 |  | TM 103 | 3XP138/Applicator_36B | 1 | both
3XP138/Applicator_36B | 3XP138 | Applicator_36B | NY-Analytical | 758 | Tensile Strength | Wet_MD_CoA_(g/in) | Wet MD CoA | 2099 | g/in | FDA,0 | 1000 |  | TM 103 | 3XP138/Applicator_36B | 1 | both
3XP138/Applicator_36B | 3XP138 | Applicator_36B | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | 3XP138/Applicator_36B | 1 | both
3XP138/Applicator_36B | 3XP138 | Applicator_36B | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208 | 3XP138/Applicator_36B | 1 | both
3XP138/Applicator_36B | 3XP138 | Applicator_36B | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | 3XP138/Applicator_36B | 1 | both
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements |  |  | left_only
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 |  |  | left_only
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208 |  |  | left_only
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 748 | Visual Inspection | Color | Color | White with no off colors or shading [Done] |  |  |  |  | TM 208 |  |  | left_only
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight |  | GSM | FDA,1 | 31.5 | 38.5 | TM 101 |  |  | left_only
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_CoA_(mm) | Caliper |  | mm | FDA,2 | 0.27 | 0.41 | TM 102 |  |  | left_only
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD |  | g/in | FDA,0 | 300 |  | TM 103 |  |  | left_only
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD |  | g/in | FDA,0 | 2000 | 4400 | TM 103 |  |  | left_only
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width |  | in | FDA,3 | 29.875 | 30.125 | Applicator Measurements |  |  | left_only
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 100 | TM 610 |  |  | left_only
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 |  |  | left_only
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) |  | cfu | FDA,0 |  |  | TM 610 |  |  | left_only
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 |  |  | left_only
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 15 | TM 610 |  |  | left_only
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  | cfu | FDA,0 |  |  | TM 610 |  |  | left_only
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight |  | g | FDA,2 |  |  | TM 610 |  |  | left_only
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Positive Rods (Done) |  |  |  |  | WI QAS 076 |  |  | left_only
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610 |  |  | left_only
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 |  |  | left_only
3XP540/Applicator_NY_Skip lot testing | 3XP540 | Applicator_NY_Skip lot testing | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (COA) | Roll Diameter (COA) |  | in | FDA,0 | 41 | 45 | Applicator Measurements | 3XP540/Applicator_NY_Skip lot testing | 18 | both
3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 3XP540 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (COA) | Roll Diameter (COA) | 43 | in | FDA,0 | 41 | 45 | Applicator Measurements | 3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 12 | both
3XP540/Applicator_NY_Skip lot testing | 3XP540 | Applicator_NY_Skip lot testing | NY-Analytical | 759 | Applicator Measurements | Roll width (COA) | Roll width (COA) |  | in | FDA,2 | 41.75 | 42.25 | Applicator Measurements | 3XP540/Applicator_NY_Skip lot testing | 18 | both
3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 3XP540 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 759 | Applicator Measurements | Roll width (COA) | Roll width (COA) | 42 | in | FDA,2 | 41.75 | 42.25 | Applicator Measurements | 3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 12 | both
3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 3XP540 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Basis Weight | 3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 12 | both
3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 3XP540 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 52.31 | GSM | FDA,0 | 45 | 55 | TM 101 | 3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 12 | both
3XP540/Applicator_NY_Skip lot testing | 3XP540 | Applicator_NY_Skip lot testing | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 49.551 | GSM | FDA,0 | 45 | 55 | TM 101 | 3XP540/Applicator_NY_Skip lot testing | 18 | both
3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 3XP540 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper | 0.42 | mm | FDA,2 | 0.51 | 0.63 | TM 102 | 3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 12 | both
3XP540/Applicator_NY_Skip lot testing | 3XP540 | Applicator_NY_Skip lot testing | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper | 0.46 | mm | FDA,2 | 0.51 | 0.63 | TM 102 | 3XP540/Applicator_NY_Skip lot testing | 18 | both
3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 3XP540 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 758 | Tensile Strength | Dry_CD_(g/in) | Dry CD | 2372 | g/in | FDA,0 | 1801 | 2821 | TM 103 | 3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 12 | both
3XP540/Applicator_NY_Skip lot testing | 3XP540 | Applicator_NY_Skip lot testing | NY-Analytical | 758 | Tensile Strength | Dry_CD_(g/in) | Dry CD |  | g/in | FDA,0 | 1801 | 2821 | TM 103 | 3XP540/Applicator_NY_Skip lot testing | 18 | both
3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 3XP540 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 758 | Tensile Strength | Dry_MD_(g/in) | Dry CD | 9752 | g/in | FDA,0 | 6675 | 9447 | TM 103 | 3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 12 | both
3XP540/Applicator_NY_Skip lot testing | 3XP540 | Applicator_NY_Skip lot testing | NY-Analytical | 758 | Tensile Strength | Dry_MD_(g/in) | Dry CD |  | g/in | FDA,0 | 6675 | 9447 | TM 103 | 3XP540/Applicator_NY_Skip lot testing | 18 | both
3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 3XP540 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | 3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 12 | both
3XP540/Applicator_NY_Skip lot testing | 3XP540 | Applicator_NY_Skip lot testing | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | 3XP540/Applicator_NY_Skip lot testing | 18 | both
3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 3XP540 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208 | 3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 12 | both
3XP540/Applicator_NY_Skip lot testing | 3XP540 | Applicator_NY_Skip lot testing | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208 | 3XP540/Applicator_NY_Skip lot testing | 18 | both
3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 3XP540 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 748 | Visual Inspection | Odor | Odor | No objectionable odor Wet or Dry [Done] |  |  |  |  | TM 208 | 3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 12 | both
3XP540/Applicator_NY_Skip lot testing | 3XP540 | Applicator_NY_Skip lot testing | NY-Analytical | 748 | Visual Inspection | Odor | Odor | No objectionable odor Wet or Dry [Done] |  |  |  |  | TM 208 | 3XP540/Applicator_NY_Skip lot testing | 18 | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 43 | in | FDA,0 | 41 | 45 | Applicator Measurements | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 1 | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 38.75 | in | FDA,2 | 38.5 | 39 | Applicator Measurements | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 1 | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Basis Weight | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 1 | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 49.94 | GSM | FDA,1 | 45 | 55 | TM 101 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 1 | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_CoA_(mm) | Caliper | 0.56 | mm | FDA,2 | 0.51 | 0.63 | TM 102 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 1 | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 758 | Tensile Strength | Dry_CD_(g/in) | Dry CD | 2635 | g/in | FDA,0 | 1801 | 2821 | TM 103 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 1 | both
3XP667/Applicator_NY_Skip lot testing | 3XP667 | Applicator_NY_Skip lot testing | NY-Analytical | 758 | Tensile Strength | Dry_CD_(g/in) | Dry CD | 1805 | g/in | FDA,0 | 1801 | 2821 | TM 103 | 3XP667/Applicator_NY_Skip lot testing | 19 | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 758 | Tensile Strength | Dry_MD_(g/in) | Dry CD | 8881 | g/in | FDA,0 | 6675 | 9447 | TM 103 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 1 | both
3XP667/Applicator_NY_Skip lot testing | 3XP667 | Applicator_NY_Skip lot testing | NY-Analytical | 758 | Tensile Strength | Dry_MD_(g/in) | Dry CD | 8039 | g/in | FDA,0 | 6675 | 9447 | TM 103 | 3XP667/Applicator_NY_Skip lot testing | 19 | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 1 | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 1 | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 748 | Visual Inspection | Odor | Odor | No objectionable odor Wet or Dry [Done] |  |  |  |  | TM 208 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 1 | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 0.94 | g | FDA,2 |  |  | TM 610 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 1 | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer |  |  |  |  | TM 610 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 1 | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 1 | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 1 | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 1 | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 1 | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 1 | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 1 | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 1 | both
4YBBY12/4YBBY12 | 4YBBY12 | 4YBBY12 | JB-Analytical | 398 | pH | pH | pH |  |  | FDA,1 | 10.8 | 11.3 | TM200 |  |  | left_only
4YBBY12/BL_JB 20X Sodium Benz Base Blend | 4YBBY12 | BL_JB 20X Sodium Benz Base Blend | JB-Analytical | 425 | % Sodium Benzoate | Sodium Benzoate Final | Sodium Benzoate Final |  | % | FDA,2 | 17.86 | 20.68 | TM 395 |  |  | left_only
4YBBY12/BL_JB 20X Sodium Benz Base Blend | 4YBBY12 | BL_JB 20X Sodium Benz Base Blend | JB-Analytical | 425 | % Sodium Benzoate | Sodium Benzoate | Sodium Benzoate |  | % | FDA,2 | 17.86 | 20.68 | TM 395 |  |  | left_only
4YBBY12/4YBBY12 | 4YBBY12 | 4YBBY12 | JB-Analytical | 425 | % Sodium Benzoate | Sodium Benzoate | Sodium Benzoate |  | % | FDA,2 | 8.46 | 10.34 | TM 395 |  |  | left_only
4YBBY12/4YBBY12 | 4YBBY12 | 4YBBY12 | JB-Analytical | 425 | % Sodium Benzoate | Sodium Benzoate Final | Sodium Benzoate Final |  | % | FDA,2 | 8.46 | 10.34 | TM 395 |  |  | left_only
4YBBY12/BL_JB 20X Sodium Benz Base Blend | 4YBBY12 | BL_JB 20X Sodium Benz Base Blend | JB-Analytical | 447 | pH as is | pH | pH |  |  | FDA,2 | 11.2 | 12.6 | TM 200 |  |  | left_only
4YBBY12/BL_JB 20X Sodium Benz Base Blend | 4YBBY12 | BL_JB 20X Sodium Benz Base Blend | JB-Analytical | 748 | Visual Inspection | Color | Color | Colorless to Slight Yellow (Done) |  |  |  |  | TM 208 |  |  | left_only
4YBBY12/BL_JB 20X Sodium Benz Base Blend | 4YBBY12 | BL_JB 20X Sodium Benz Base Blend | JB-Analytical | 748 | Visual Inspection | Contamination | Contamination | None (Done) |  |  |  |  | TM 208 |  |  | left_only
4YBBY12/BL_JB 20X Sodium Benz Base Blend | 4YBBY12 | BL_JB 20X Sodium Benz Base Blend | JB-Analytical | 748 | Visual Inspection | Appearance | Appearance | Slighlty Hazy (Done) |  |  |  |  | TM 208 |  |  | left_only
4YBBY12/BL_JB 20X Sodium Benz Base Blend | 4YBBY12 | BL_JB 20X Sodium Benz Base Blend | JB-Analytical | 748 | Visual Inspection | Odor | Odor | Unscented (Done) |  |  |  |  | TM 208 |  |  | left_only
4YBBY12/4YBBY12 | 4YBBY12 | 4YBBY12 | JB-Analytical | 902 | Bulk Liquid Inspection | Appearance (TM 208) | Appearance | [5 empty] | See Analyte for Method # |  |  | left_only
4YBBY12/4YBBY12 | 4YBBY12 | 4YBBY12 | JB-Analytical | 902 | Bulk Liquid Inspection | Color (TM 208) | Color | [5 empty] | See Analyte for Method # |  |  | left_only
4YBBY12/4YBBY12 | 4YBBY12 | 4YBBY12 | JB-Analytical | 902 | Bulk Liquid Inspection | Contamination (TM 208) | Contamination | [5 empty] | See Analyte for Method # |  |  | left_only
4YBBY12/4YBBY12 | 4YBBY12 | 4YBBY12 | JB-Analytical | 902 | Bulk Liquid Inspection | Odor (TM 208) | Odor | [5 empty] | See Analyte for Method # |  |  | left_only
4YBBY23/4YBBY23 | 4YBBY23 | 4YBBY23 | JB-Analytical | 398 | pH | pH | pH |  |  | FDA,1 | 10.8 | 11.3 | TM 200 |  |  | left_only
4YBBY23/4YBBY23 | 4YBBY23 | 4YBBY23 | JB-Analytical | 425 | % Sodium Benzoate | Sodium Benzoate Final | Sodium Benzoate Final |  | % | FDA,2 | 8.46 | 10.34 | TM 395 |  |  | left_only
4YBBY23/BL_JB 40X Natural Baby W/Vitamin E | 4YBBY23 | BL_JB 40X Natural Baby W/Vitamin E | JB-Analytical | 425 | % Sodium Benzoate | Sodium Benzoate | Sodium Benzoate |  | % | FDA,2 | 17.86 | 20.68 | TM 395 |  |  | left_only
4YBBY23/BL_JB 40X Natural Baby W/Vitamin E | 4YBBY23 | BL_JB 40X Natural Baby W/Vitamin E | JB-Analytical | 425 | % Sodium Benzoate | Sodium Benzoate Final | Sodium Benzoate Final |  | % | FDA,2 | 17.86 | 20.68 | TM 395 |  |  | left_only
4YBBY23/4YBBY23 | 4YBBY23 | 4YBBY23 | JB-Analytical | 425 | % Sodium Benzoate | Sodium Benzoate | Sodium Benzoate |  | % | FDA,2 | 8.46 | 10.34 | TM 395 |  |  | left_only
4YBBY23/BL_JB 40X Natural Baby W/Vitamin E | 4YBBY23 | BL_JB 40X Natural Baby W/Vitamin E | JB-Analytical | 902 | Bulk Liquid Inspection | pH (TM 200) | pH |  |  | FDA,2 | 11.2 | 12.6 | See Analyte for Method # |  |  | left_only
4YBBY23/BL_JB 40X Natural Baby W/Vitamin E | 4YBBY23 | BL_JB 40X Natural Baby W/Vitamin E | JB-Analytical | 902 | Bulk Liquid Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | See Analyte for Method # |  |  | left_only
4YBBY23/BL_JB 40X Natural Baby W/Vitamin E | 4YBBY23 | BL_JB 40X Natural Baby W/Vitamin E | JB-Analytical | 902 | Bulk Liquid Inspection | Appearance (TM 208) | Appearance | Slightly Hazy Liquid (Done) |  |  |  |  | See Analyte for Method # |  |  | left_only
4YBBY23/BL_JB 40X Natural Baby W/Vitamin E | 4YBBY23 | BL_JB 40X Natural Baby W/Vitamin E | JB-Analytical | 902 | Bulk Liquid Inspection | Color (TM 208) | Color | Slightly Yellow Liquid (Done) |  |  |  |  | See Analyte for Method # |  |  | left_only
4YBBY23/BL_JB 40X Natural Baby W/Vitamin E | 4YBBY23 | BL_JB 40X Natural Baby W/Vitamin E | JB-Analytical | 902 | Bulk Liquid Inspection | Odor (TM 208) | Odor | Unscented (Done) |  |  |  |  | See Analyte for Method # |  |  | left_only
4YBBY23/4YBBY23 | 4YBBY23 | 4YBBY23 | JB-Analytical | 902 | Bulk Liquid Inspection | Appearance (TM 208) | Appearance | [5 empty] | See Analyte for Method # |  |  | left_only
4YBBY23/4YBBY23 | 4YBBY23 | 4YBBY23 | JB-Analytical | 902 | Bulk Liquid Inspection | Color (TM 208) | Color | [5 empty] | See Analyte for Method # |  |  | left_only
4YBBY23/4YBBY23 | 4YBBY23 | 4YBBY23 | JB-Analytical | 902 | Bulk Liquid Inspection | Contamination (TM 208) | Contamination | [5 empty] | See Analyte for Method # |  |  | left_only
4YBBY23/4YBBY23 | 4YBBY23 | 4YBBY23 | JB-Analytical | 902 | Bulk Liquid Inspection | Odor (TM 208) | Odor | [5 empty] | See Analyte for Method # |  |  | left_only
5AA347BT/Box_NY_0.020SBS_300_3.375-1.313-2.625 | 5AA347BT | Box_NY_0.020SBS_300_3.375-1.313-2.625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper |  | in | FDA,3 | 0.019 | 0.021 | Box Inspection |  |  | left_only
5AA347BT/Box_NY_0.020SBS_300_3.375-1.313-2.625 | 5AA347BT | Box_NY_0.020SBS_300_3.375-1.313-2.625 | NY-Analytical | 772 | Box Inspection | Width | Width |  | in | FDA,5 | 1.2817 | 1.3442 | Box Inspection |  |  | left_only
5AA347BT/Box_NY_0.020SBS_300_3.375-1.313-2.625 | 5AA347BT | Box_NY_0.020SBS_300_3.375-1.313-2.625 | NY-Analytical | 772 | Box Inspection | Height | Height |  | in | FDA,5 | 2.5938 | 2.6562 | Box Inspection |  |  | left_only
5AA347BT/Box_NY_0.020SBS_300_3.375-1.313-2.625 | 5AA347BT | Box_NY_0.020SBS_300_3.375-1.313-2.625 | NY-Analytical | 772 | Box Inspection | Length | Length |  | in | FDA,5 | 3.3438 | 3.4062 | Box Inspection |  |  | left_only
5AA347BT/Box_NY_0.020SBS_300_3.375-1.313-2.625 | 5AA347BT | Box_NY_0.020SBS_300_3.375-1.313-2.625 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 |  |  | Box Inspection |  |  | left_only
5AA347BT/Box_NY_Skip lot testing | 5AA347BT | Box_NY_Skip lot testing | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Box Inspection | 5AA347BT/Box_NY_Skip lot testing | 1 | both
5AA347BT/Box_NY_0.020SBS_300_3.375-1.313-2.625 | 5AA347BT | Box_NY_0.020SBS_300_3.375-1.313-2.625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection |  |  | left_only
5AA347BT/Box_NY_0.020SBS_300_3.375-1.313-2.625 | 5AA347BT | Box_NY_0.020SBS_300_3.375-1.313-2.625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Open Top/Auto Bottom Tray (Done) |  |  |  |  | Box Inspection |  |  | left_only
5AA347BT/Box_NY_0.020SBS_300_3.375-1.313-2.625 | 5AA347BT | Box_NY_0.020SBS_300_3.375-1.313-2.625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection |  |  | left_only
5AA404S48/Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | 5AA404S48 | Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper |  | in | FDA,3 | 0.017 | 0.019 | Box Inspection |  |  | left_only
5AA404S48/Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | 5AA404S48 | Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | NY-Analytical | 772 | Box Inspection | Width | Width |  | in | FDA,5 | 3.6562 | 3.7188 | Box Inspection |  |  | left_only
5AA404S48/Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | 5AA404S48 | Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | NY-Analytical | 772 | Box Inspection | Height | Height |  | in | FDA,5 | 6.0312 | 6.0938 | Box Inspection |  |  | left_only
5AA404S48/Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | 5AA404S48 | Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | NY-Analytical | 772 | Box Inspection | Length | Length |  | in | FDA,5 | 7.1562 | 7.2188 | Box Inspection |  |  | left_only
5AA404S48/Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | 5AA404S48 | Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection |  |  | left_only
5AA404S48/Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | 5AA404S48 | Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection |  |  | left_only
5AA404S48/Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | 5AA404S48 | Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection |  |  | left_only
5AA404S48/Box_NY_0.018SBS_111_6-5.875-3.625 | 5AA404S48 | Box_NY_0.018SBS_111_6-5.875-3.625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy |  | in |  |  |  | Box Inspection |  |  | left_only
5AA404S48/Box_NY_0.018SBS_111_6-5.875-3.625 | 5AA404S48 | Box_NY_0.018SBS_111_6-5.875-3.625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper |  | in | FDA,3 | 0.017 | 0.019 | Box Inspection |  |  | left_only
5AA404S48/Box_NY_0.018SBS_111_6-5.875-3.625 | 5AA404S48 | Box_NY_0.018SBS_111_6-5.875-3.625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | [5 empty] | Box Inspection |  |  | left_only
5AA404S48/Box_NY_0.018SBS_111_6-5.875-3.625 | 5AA404S48 | Box_NY_0.018SBS_111_6-5.875-3.625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | [5 empty] | Box Inspection |  |  | left_only
5AA404S48/Box_NY_0.018SBS_111_6-5.875-3.625 | 5AA404S48 | Box_NY_0.018SBS_111_6-5.875-3.625 | NY-Analytical | 772 | Box Inspection | Height | Height |  | in | FDA,5 | 3.5938 | 3.6562 | Box Inspection |  |  | left_only
5AA404S48/Box_NY_0.018SBS_111_6-5.875-3.625 | 5AA404S48 | Box_NY_0.018SBS_111_6-5.875-3.625 | NY-Analytical | 772 | Box Inspection | Length | Length |  | in | FDA,5 | 5.9688 | 6.0312 | Box Inspection |  |  | left_only
5AA404S48/Box_NY_0.018SBS_111_6-5.875-3.625 | 5AA404S48 | Box_NY_0.018SBS_111_6-5.875-3.625 | NY-Analytical | 772 | Box Inspection | Width | Width |  | in | FDA,5 | 5.8438 | 5.9062 | Box Inspection |  |  | left_only
5AA617S96/Box_NY_0.018SBS_47B_3.625-3.875-6 | 5AA617S96 | Box_NY_0.018SBS_47B_3.625-3.875-6 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper |  | in | FDA,3 | 0.017 | 0.019 | Box Inspection |  |  | left_only
5AA617S96/Box_NY_0.018SBS_47B_3.625-3.875-6 | 5AA617S96 | Box_NY_0.018SBS_47B_3.625-3.875-6 | NY-Analytical | 772 | Box Inspection | Length | Length |  | in | FDA,5 | 3.5938 | 3.6562 | Box Inspection |  |  | left_only
5AA617S96/Box_NY_0.018SBS_47B_3.625-3.875-6 | 5AA617S96 | Box_NY_0.018SBS_47B_3.625-3.875-6 | NY-Analytical | 772 | Box Inspection | Width | Width |  | in | FDA,5 | 3.8438 | 3.9062 | Box Inspection |  |  | left_only
5AA617S96/Box_NY_0.018SBS_47B_3.625-3.875-6 | 5AA617S96 | Box_NY_0.018SBS_47B_3.625-3.875-6 | NY-Analytical | 772 | Box Inspection | Height | Height |  | in | FDA,5 | 5.9688 | 6.0312 | Box Inspection |  |  | left_only
5AA617S96/Box_NY_0.018SBS_47B_3.625-3.875-6 | 5AA617S96 | Box_NY_0.018SBS_47B_3.625-3.875-6 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection |  |  | left_only
5AA617S96/Box_NY_0.018SBS_47B_3.625-3.875-6 | 5AA617S96 | Box_NY_0.018SBS_47B_3.625-3.875-6 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection |  |  | left_only
5AA617S96/Box_NY_0.018SBS_47B_3.625-3.875-6 | 5AA617S96 | Box_NY_0.018SBS_47B_3.625-3.875-6 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Straight Tuck End - French Style (Done) |  |  |  |  | Box Inspection |  |  | left_only
5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | 5AA632S60 | Box_NY_0.020SBS_101_6-3.625-7 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 3 |  |  |  |  | Box Inspection | 5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | 6 | both
5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | 5AA632S60 | Box_NY_0.020SBS_101_6-3.625-7 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | 6 | both
5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | 5AA632S60 | Box_NY_0.020SBS_101_6-3.625-7 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.021 | in | FDA,3 | 0.019 | 0.021 | Box Inspection | 5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | 6 | both
5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | 5AA632S60 | Box_NY_0.020SBS_101_6-3.625-7 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | 5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | 6 | both
5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | 5AA632S60 | Box_NY_0.020SBS_101_6-3.625-7 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom - French Style (Done) |  |  |  |  | Box Inspection | 5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | 6 | both
5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | 5AA632S60 | Box_NY_0.020SBS_101_6-3.625-7 | NY-Analytical | 772 | Box Inspection | Height | Height | 7 | in | FDA,5 | 6.9688 | 7.0312 | Box Inspection | 5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | 6 | both
5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | 5AA632S60 | Box_NY_0.020SBS_101_6-3.625-7 | NY-Analytical | 772 | Box Inspection | Length | Length | 6 | in | FDA,5 | 5.9688 | 6.0312 | Box Inspection | 5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | 6 | both
5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | 5AA632S60 | Box_NY_0.020SBS_101_6-3.625-7 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.625 | in | FDA,5 | 3.5938 | 3.6562 | Box Inspection | 5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | 6 | both
5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | 5AA911TY | Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom w/perf |  |  |  |  | Box Inspection | 5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | 2 | both
5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | 5AA911TY | Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | Box Inspection | 5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | 2 | both
5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | 5AA911TY | Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | 2 | both
5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | 5AA911TY | Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0195 | in | FDA,3 | 0.019 | 0.021 | Box Inspection | 5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | 2 | both
5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | 5AA911TY | Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | 5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | 2 | both
5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | 5AA911TY | Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | NY-Analytical | 772 | Box Inspection | Height | Height | 7.94 | in | FDA,5 | 7.9062 | 7.9688 | Box Inspection | 5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | 2 | both
5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | 5AA911TY | Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | NY-Analytical | 772 | Box Inspection | Length | Length | 4.51 | in | FDA,5 | 4.4688 | 4.5312 | Box Inspection | 5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | 2 | both
5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | 5AA911TY | Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | NY-Analytical | 772 | Box Inspection | Width | Width | 7.94 | in | FDA,5 | 7.9062 | 7.9688 | Certificate of Conformance | 5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | 2 | both
5AB123CP/Label_NY_Skip lot testing | 5AB123CP | Label_NY_Skip lot testing | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Certificate of Conformance | 5AB123CP/Label_NY_Skip lot testing | 1 | both
5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | 5AB123CP | Label_NY_5AB123CP_1.9375-1.875 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | [5 empty] | Label Inspection | 5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | 1 | both
5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | 5AB123CP | Label_NY_5AB123CP_1.9375-1.875 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | 5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | 1 | both
5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | 5AB123CP | Label_NY_5AB123CP_1.9375-1.875 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | 5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | 1 | both
5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | 5AB123CP | Label_NY_5AB123CP_1.9375-1.875 | NY-Analytical | 770 | Label Inspection | Length | Length | 1.875 | in | FDA,4 | 1.875 | 2 | TM 143 / Visual | 5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | 1 | both
5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | 5AB123CP | Label_NY_5AB123CP_1.9375-1.875 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | 5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | 1 | both
5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | 5AB123CP | Label_NY_5AB123CP_1.9375-1.875 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | 5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | 1 | both
5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | 5AB123CP | Label_NY_5AB123CP_1.9375-1.875 | NY-Analytical | 770 | Label Inspection | Width | Width | 1.9375 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | 5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | 1 | both
5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | 5AB123ST | Box_NY_0.018SBS_107_10.25-2.375-2.75 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | 3 | both
5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | 5AB123ST | Box_NY_0.018SBS_107_10.25-2.375-2.75 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.018 | in | FDA,3 | 0.017 | 0.019 | Box Inspection | 5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | 3 | both
5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | 5AB123ST | Box_NY_0.018SBS_107_10.25-2.375-2.75 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | 5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | 3 | both
5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | 5AB123ST | Box_NY_0.018SBS_107_10.25-2.375-2.75 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | TT/AB w/friction lock (Done) |  |  |  |  | Box Inspection | 5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | 3 | both
5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | 5AB123ST | Box_NY_0.018SBS_107_10.25-2.375-2.75 | NY-Analytical | 772 | Box Inspection | Height | Height | 2.7515 | in | FDA,5 | 2.7188 | 2.7812 | Box Inspection | 5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | 3 | both
5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | 5AB123ST | Box_NY_0.018SBS_107_10.25-2.375-2.75 | NY-Analytical | 772 | Box Inspection | Length | Length | 10.22 | in | FDA,5 | 10.2188 | 10.2812 | Box Inspection | 5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | 3 | both
5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | 5AB123ST | Box_NY_0.018SBS_107_10.25-2.375-2.75 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.351 | in | FDA,5 | 2.3438 | 2.4062 | Box Inspection | 5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | 3 | both
5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | 5AD02424 | Box_NY_0.018SBS_107_4-2.03125-4.5 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top Auto Bottom w/ Tab Lock |  |  |  |  | Box Inspection | 5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | 1 | both
5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | 5AD02424 | Box_NY_0.018SBS_107_4-2.03125-4.5 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | 1 | both
5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | 5AD02424 | Box_NY_0.018SBS_107_4-2.03125-4.5 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0175 | in | FDA,3 | 0.017 | 0.019 | Box Inspection | 5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | 1 | both
5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | 5AD02424 | Box_NY_0.018SBS_107_4-2.03125-4.5 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | 5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | 1 | both
5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | 5AD02424 | Box_NY_0.018SBS_107_4-2.03125-4.5 | NY-Analytical | 772 | Box Inspection | Height | Height | 4.5 | in | FDA,5 | 4.4688 | 4.5312 | Box Inspection | 5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | 1 | both
5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | 5AD02424 | Box_NY_0.018SBS_107_4-2.03125-4.5 | NY-Analytical | 772 | Box Inspection | Length | Length | 4 | in | FDA,5 | 3.9688 | 4.0312 | Box Inspection | 5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | 1 | both
5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | 5AD02424 | Box_NY_0.018SBS_107_4-2.03125-4.5 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.0312 | in | FDA,5 | 2 | 2.0625 | Box Inspection | 5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | 1 | both
5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | 5AH45016 | Box_NY_0.018SBS_101_3.875-2-3.5 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 200 |  |  |  |  | Box Inspection | 5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | 2 | both
5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | 5AH45016 | Box_NY_0.018SBS_101_3.875-2-3.5 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | 2 | both
5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | 5AH45016 | Box_NY_0.018SBS_101_3.875-2-3.5 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.018 | in | FDA,3 | 0.017 | 0.019 | Box Inspection | 5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | 2 | both
5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | 5AH45016 | Box_NY_0.018SBS_101_3.875-2-3.5 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | 5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | 2 | both
5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | 5AH45016 | Box_NY_0.018SBS_101_3.875-2-3.5 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | French Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | 5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | 2 | both
5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | 5AH45016 | Box_NY_0.018SBS_101_3.875-2-3.5 | NY-Analytical | 772 | Box Inspection | Height | Height | 3.5 | in | FDA,5 | 3.4688 | 3.5312 | Box Inspection | 5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | 2 | both
5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | 5AH45016 | Box_NY_0.018SBS_101_3.875-2-3.5 | NY-Analytical | 772 | Box Inspection | Length | Length | 3.875 | in | FDA,5 | 3.8438 | 3.9062 | Box Inspection | 5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | 2 | both
5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | 5AH45016 | Box_NY_0.018SBS_101_3.875-2-3.5 | NY-Analytical | 772 | Box Inspection | Width | Width | 2 | in | FDA,5 | 1.9688 | 2.0312 | Box Inspection | 5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | 2 | both
5AH58195/Box_NY_0.020SBS_127_6.25-2.5625-4.625 | 5AH58195 | Box_NY_0.020SBS_127_6.25-2.5625-4.625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper |  | in | FDA,3 | 0.019 | 0.021 | Box Inspection |  |  | left_only
5AH58195/Box_NY_0.020SBS_127_6.25-2.5625-4.625 | 5AH58195 | Box_NY_0.020SBS_127_6.25-2.5625-4.625 | NY-Analytical | 772 | Box Inspection | Width | Width |  | in | FDA,5 | 2.5312 | 2.5938 | Box Inspection |  |  | left_only
5AH58195/Box_NY_0.020SBS_127_6.25-2.5625-4.625 | 5AH58195 | Box_NY_0.020SBS_127_6.25-2.5625-4.625 | NY-Analytical | 772 | Box Inspection | Height | Height |  | in | FDA,5 | 4.5938 | 4.6562 | Box Inspection |  |  | left_only
5AH58195/Box_NY_0.020SBS_127_6.25-2.5625-4.625 | 5AH58195 | Box_NY_0.020SBS_127_6.25-2.5625-4.625 | NY-Analytical | 772 | Box Inspection | Length | Length |  | in | FDA,5 | 6.2188 | 6.2812 | Box Inspection |  |  | left_only
5AH58195/Box_NY_0.020SBS_127_6.25-2.5625-4.625 | 5AH58195 | Box_NY_0.020SBS_127_6.25-2.5625-4.625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection |  |  | left_only
5AH58195/Box_NY_0.020SBS_127_6.25-2.5625-4.625 | 5AH58195 | Box_NY_0.020SBS_127_6.25-2.5625-4.625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection |  |  | left_only
5AH58195/Box_NY_0.020SBS_127_6.25-2.5625-4.625 | 5AH58195 | Box_NY_0.020SBS_127_6.25-2.5625-4.625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Straight Tuck w/Partition Dispenser Box |  |  |  |  | Box Inspection |  |  | left_only
5AJ130BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | 5AJ130BT | Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 |  |  | Box Inspection |  |  | left_only
5AJ130BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | 5AJ130BT | Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | [5 empty] | Box Inspection |  |  | left_only
5AJ130BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | 5AJ130BT | Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper |  | in | FDA,3 | 0.019 | 0.021 | Box Inspection |  |  | left_only
5AJ130BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | 5AJ130BT | Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | [5 empty] | Box Inspection |  |  | left_only
5AJ130BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | 5AJ130BT | Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | [5 empty] | Box Inspection |  |  | left_only
5AJ130BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | 5AJ130BT | Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | NY-Analytical | 772 | Box Inspection | Height | Height |  | in | FDA,5 | 2.5938 | 2.6562 | Box Inspection |  |  | left_only
5AJ130BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | 5AJ130BT | Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | NY-Analytical | 772 | Box Inspection | Length | Length |  | in | FDA,5 | 3.2812 | 3.3438 | Box Inspection |  |  | left_only
5AJ130BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | 5AJ130BT | Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | NY-Analytical | 772 | Box Inspection | Width | Width |  | in | FDA,5 | 1.4062 | 1.4688 | Box Inspection |  |  | left_only
5AP027TY/Box_57A | 5AP027TY | Box_57A | NY-Analytical | 772 | Box Inspection | Appearance | Appearance | Meets specification (Done) |  |  |  |  | Box Inspection | 5AP027TY/Box_57A | 7 | both
5AP027TY/Box_57A | 5AP027TY | Box_57A | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Box Inspection | 5AP027TY/Box_57A | 7 | both
5AP027TY/Box_57A | 5AP027TY | Box_57A | NY-Analytical | 772 | Box Inspection | Color | Color | Clear (Done) |  |  |  |  | Box Inspection | 5AP027TY/Box_57A | 7 | both
5AP027TY/Box_57A | 5AP027TY | Box_57A | NY-Analytical | 772 | Box Inspection | Functionality | Functionality | Meets specification (Done) |  |  |  |  | Box Inspection | 5AP027TY/Box_57A | 7 | both
5AP027TY/Box_57A | 5AP027TY | Box_57A | NY-Analytical | 772 | Box Inspection | Length | Length | 10.75 | in | FDA,3 | 10.66 | 10.84 | Box Inspection | 5AP027TY/Box_57A | 7 | both
5AP027TY/Box_57A | 5AP027TY | Box_57A | NY-Analytical | 772 | Box Inspection | Odor | Odor | None Detectable (Done) |  |  |  |  | Box Inspection | 5AP027TY/Box_57A | 7 | both
5AP027TY/Box_57A | 5AP027TY | Box_57A | NY-Analytical | 772 | Box Inspection | Width | Width | 5.625 | in | FDA,3 | 5.535 | 5.715 | Box Inspection | 5AP027TY/Box_57A | 7 | both
5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | 5AP07530 | Box_NY_0.020BC_110_6.375-3.4375-4.75 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | Box Inspection | 5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | 1 | both
5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | 5AP07530 | Box_NY_0.020BC_110_6.375-3.4375-4.75 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | 1 | both
5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | 5AP07530 | Box_NY_0.020BC_110_6.375-3.4375-4.75 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.02 | in | FDA,3 | 0.019 | 0.021 | Box Inspection | 5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | 1 | both
5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | 5AP07530 | Box_NY_0.020BC_110_6.375-3.4375-4.75 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | BC (Done) |  |  |  |  | Box Inspection | 5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | 1 | both
5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | 5AP07530 | Box_NY_0.020BC_110_6.375-3.4375-4.75 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | TT/AB w/finger holes (Done) |  |  |  |  | Box Inspection | 5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | 1 | both
5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | 5AP07530 | Box_NY_0.020BC_110_6.375-3.4375-4.75 | NY-Analytical | 772 | Box Inspection | Height | Height | 4.771 | in | FDA,5 | 4.7188 | 4.7812 | Box Inspection | 5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | 1 | both
5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | 5AP07530 | Box_NY_0.020BC_110_6.375-3.4375-4.75 | NY-Analytical | 772 | Box Inspection | Length | Length | 6.395 | in | FDA,5 | 6.3438 | 6.4062 | Box Inspection | 5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | 1 | both
5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | 5AP07530 | Box_NY_0.020BC_110_6.375-3.4375-4.75 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.4225 | in | FDA,5 | 3.4062 | 3.4688 | Box Inspection | 5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | 1 | both
5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | 5AP209SL | Box_NY_0.020BC_220_5.09375-3.6875-4.625 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 30 |  | FDA,0 |  |  | Box Inspection | 5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | 3 | both
5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | 5AP209SL | Box_NY_0.020BC_220_5.09375-3.6875-4.625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | 3 | both
5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | 5AP209SL | Box_NY_0.020BC_220_5.09375-3.6875-4.625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0195 | in | FDA,3 | 0.019 | 0.021 | Box Inspection | 5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | 3 | both
5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | 5AP209SL | Box_NY_0.020BC_220_5.09375-3.6875-4.625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | BC (Done) |  |  |  |  | Box Inspection | 5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | 3 | both
5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | 5AP209SL | Box_NY_0.020BC_220_5.09375-3.6875-4.625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Diecut Glued Sleeve w/ 2-1" fingerholes (Done) |  |  |  |  | Box Inspection | 5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | 3 | both
5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | 5AP209SL | Box_NY_0.020BC_220_5.09375-3.6875-4.625 | NY-Analytical | 772 | Box Inspection | Height | Height | 4.61 | in | FDA,5 | 4.5938 | 4.6562 | Box Inspection | 5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | 3 | both
5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | 5AP209SL | Box_NY_0.020BC_220_5.09375-3.6875-4.625 | NY-Analytical | 772 | Box Inspection | Length | Length | 5.07 | in | FDA,5 | 5.0625 | 5.125 | Box Inspection | 5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | 3 | both
5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | 5AP209SL | Box_NY_0.020BC_220_5.09375-3.6875-4.625 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.7 | in | FDA,5 | 3.6562 | 3.7188 | Box Inspection | 5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | 3 | both
5AP21195/Box_01M | 5AP21195 | Box_01M | NY-Analytical | 772 | Box Inspection | Caliper | Caliper |  | in | FDA,3 | 0.023 | 0.025 | Box Inspection |  |  | left_only
5AP21195/Box_01M | 5AP21195 | Box_01M | NY-Analytical | 772 | Box Inspection | Width | Width |  | in | FDA,5 | 4.4688 | 4.5312 | Box Inspection |  |  | left_only
5AP21195/Box_01M | 5AP21195 | Box_01M | NY-Analytical | 772 | Box Inspection | Length | Length |  | in | FDA,5 | 4.5938 | 4.6562 | Box Inspection |  |  | left_only
5AP21195/Box_01M | 5AP21195 | Box_01M | NY-Analytical | 772 | Box Inspection | Height | Height |  | in | FDA,5 | 5.8438 | 5.9062 | Box Inspection |  |  | left_only
5AP21195/Box_01M | 5AP21195 | Box_01M | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | Bending Chip |  |  |  |  | Box Inspection |  |  | left_only
5AP21195/Box_01M | 5AP21195 | Box_01M | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection |  |  | left_only
5AP21195/Box_01M | 5AP21195 | Box_01M | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection |  |  | left_only
5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | 5AP227TY | Box_NY_0.020SBS_300_4.625-5-3.625 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | Box Inspection | 5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | 2 | both
5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | 5AP227TY | Box_NY_0.020SBS_300_4.625-5-3.625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | 2 | both
5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | 5AP227TY | Box_NY_0.020SBS_300_4.625-5-3.625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.02 | in | FDA,3 | 0.019 | 0.021 | Box Inspection | 5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | 2 | both
5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | 5AP227TY | Box_NY_0.020SBS_300_4.625-5-3.625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | 5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | 2 | both
5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | 5AP227TY | Box_NY_0.020SBS_300_4.625-5-3.625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Open Top/Auto Bottom Tray (Done) |  |  |  |  | Box Inspection | 5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | 2 | both
5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | 5AP227TY | Box_NY_0.020SBS_300_4.625-5-3.625 | NY-Analytical | 772 | Box Inspection | Height | Height | 3.61 | in | FDA,5 | 3.5938 | 3.6562 | Box Inspection | 5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | 2 | both
5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | 5AP227TY | Box_NY_0.020SBS_300_4.625-5-3.625 | NY-Analytical | 772 | Box Inspection | Length | Length | 4.6 | in | FDA,5 | 4.5938 | 4.6562 | Box Inspection | 5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | 2 | both
5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | 5AP227TY | Box_NY_0.020SBS_300_4.625-5-3.625 | NY-Analytical | 772 | Box Inspection | Width | Width | 5.01 | in | FDA,5 | 4.9688 | 5.0312 | Box Inspection | 5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | 2 | both
5AP229TY/Box_NY_0.024SBS_303_8.228-3.661-6.024 | 5AP229TY | Box_NY_0.024SBS_303_8.228-3.661-6.024 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper |  | in | FDA,3 | 0.023 | 0.025 | Box Inspection |  |  | left_only
5AP229TY/Box_NY_0.024SBS_303_8.228-3.661-6.024 | 5AP229TY | Box_NY_0.024SBS_303_8.228-3.661-6.024 | NY-Analytical | 772 | Box Inspection | Width | Width |  | in | FDA,5 | 3.6298 | 3.6923 | Box Inspection |  |  | left_only
5AP229TY/Box_NY_0.024SBS_303_8.228-3.661-6.024 | 5AP229TY | Box_NY_0.024SBS_303_8.228-3.661-6.024 | NY-Analytical | 772 | Box Inspection | Height | Height |  | in | FDA,5 | 5.9928 | 6.0553 | Box Inspection |  |  | left_only
5AP229TY/Box_NY_0.024SBS_303_8.228-3.661-6.024 | 5AP229TY | Box_NY_0.024SBS_303_8.228-3.661-6.024 | NY-Analytical | 772 | Box Inspection | Length | Length |  | in | FDA,5 | 8.1967 | 8.2592 | Box Inspection |  |  | left_only
5AP229TY/Box_NY_0.024SBS_303_8.228-3.661-6.024 | 5AP229TY | Box_NY_0.024SBS_303_8.228-3.661-6.024 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | A/B tray w/tear-off lid (Done) |  |  |  |  | Box Inspection |  |  | left_only
5AP229TY/Box_NY_0.024SBS_303_8.228-3.661-6.024 | 5AP229TY | Box_NY_0.024SBS_303_8.228-3.661-6.024 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection |  |  | left_only
5AP229TY/Box_NY_0.024SBS_303_8.228-3.661-6.024 | 5AP229TY | Box_NY_0.024SBS_303_8.228-3.661-6.024 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection |  |  | left_only
5AP308TY/Box_63A | 5AP308TY | Box_63A | NY-Analytical | 772 | Box Inspection | Appearance | Appearance | Meets specification (Done) |  |  |  |  | Box Inspection | 5AP308TY/Box_63A | 3 | both
5AP308TY/Box_63A | 5AP308TY | Box_63A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Box Inspection | 5AP308TY/Box_63A | 3 | both
5AP308TY/Box_63A | 5AP308TY | Box_63A | NY-Analytical | 772 | Box Inspection | Color | Color | Clear (Done) |  |  |  |  | Box Inspection | 5AP308TY/Box_63A | 3 | both
5AP308TY/Box_63A | 5AP308TY | Box_63A | NY-Analytical | 772 | Box Inspection | Flange - Sides | Flange - Sides | 0.075 | in | FDA,3 | 0.045 | 0.135 | Box Inspection | 5AP308TY/Box_63A | 3 | both
5AP308TY/Box_63A | 5AP308TY | Box_63A | NY-Analytical | 772 | Box Inspection | Flange - Top & Bottom | Flange - Top & Bottom | 0.099 | in | FDA,3 | 0.045 | 0.135 | Box Inspection | 5AP308TY/Box_63A | 3 | both
5AP308TY/Box_63A | 5AP308TY | Box_63A | NY-Analytical | 772 | Box Inspection | Length | Length | 7.2 | in | FDA,3 | 7.156 | 7.22 | Box Inspection | 5AP308TY/Box_63A | 3 | both
5AP308TY/Box_63A | 5AP308TY | Box_63A | NY-Analytical | 772 | Box Inspection | Odor | Odor | None Detectable (Done) |  |  |  |  | Box Inspection | 5AP308TY/Box_63A | 3 | both
5AP308TY/Box_63A | 5AP308TY | Box_63A | NY-Analytical | 772 | Box Inspection | Total Height | Total Height | 2.95 | in | FDA,3 | 2.899 | 2.963 | Box Inspection | 5AP308TY/Box_63A | 3 | both
5AP308TY/Box_63A | 5AP308TY | Box_63A | NY-Analytical | 772 | Box Inspection | Width | Width | 6.1 | in | FDA,3 | 6.093 | 6.157 | Box Inspection | 5AP308TY/Box_63A | 3 | both
5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | 5AU25206 | Box_NY_0.018SBS_100_101.6-44.45-101.6 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 200 |  | FDA,0 |  |  | Box Inspection | 5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | 2 | both
5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | 5AU25206 | Box_NY_0.018SBS_100_101.6-44.45-101.6 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | 2 | both
5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | 5AU25206 | Box_NY_0.018SBS_100_101.6-44.45-101.6 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.018 | in | FDA,3 | 0.017 | 0.019 | Box Inspection | 5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | 2 | both
5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | 5AU25206 | Box_NY_0.018SBS_100_101.6-44.45-101.6 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | 5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | 2 | both
5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | 5AU25206 | Box_NY_0.018SBS_100_101.6-44.45-101.6 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | 5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | 2 | both
5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | 5AU25206 | Box_NY_0.018SBS_100_101.6-44.45-101.6 | NY-Analytical | 772 | Box Inspection | Height | Height | 102 | mm | FDA,1 | 100.8 | 102.4 | Box Inspection | 5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | 2 | both
5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | 5AU25206 | Box_NY_0.018SBS_100_101.6-44.45-101.6 | NY-Analytical | 772 | Box Inspection | Length | Length | 102 | mm | FDA,1 | 100.8 | 102.4 | Box Inspection | 5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | 2 | both
5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | 5AU25206 | Box_NY_0.018SBS_100_101.6-44.45-101.6 | NY-Analytical | 772 | Box Inspection | Width | Width | 45 | mm | FDA,1 | 43.7 | 45.2 | Box Inspection | 5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | 2 | both
5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | 5AU252IN | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Box Inspection |  |  | left_only
5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | 5AU252IN | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper |  | in | FDA,3 | 0.019 | 0.021 | Box Inspection |  |  | left_only
5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | 5AU252IN | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 772 | Box Inspection | Width | Width |  | in | FDA,5 | 3.5938 | 3.6562 | Box Inspection |  |  | left_only
5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | 5AU252IN | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 772 | Box Inspection | Length | Length |  | in | FDA,5 | 4.0625 | 4.125 | Box Inspection |  |  | left_only
5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | 5AU252IN | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 772 | Box Inspection | Height | Height |  | in | FDA,5 | 4.2812 | 4.3438 | Box Inspection |  |  | left_only
5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | 5AU252IN | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 |  |  | Box Inspection |  |  | left_only
5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | 5AU252IN | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Does not meet specifications (OOS) | in |  |  |  | Box Inspection |  |  | left_only
5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | 5AU252IN | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | RT (Done) |  |  |  |  | Box Inspection |  |  | left_only
5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | 5AU252IN | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | TBC (Done) |  |  |  |  | Box Inspection |  |  | left_only
5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | 5AU31450 | Box_NY_0.020SBS_169_6-3.125-5.9375 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | 2 | both
5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | 5AU31450 | Box_NY_0.020SBS_169_6-3.125-5.9375 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.019 | in | FDA,3 | 0.019 | 0.021 | Box Inspection | 5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | 2 | both
5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | 5AU31450 | Box_NY_0.020SBS_169_6-3.125-5.9375 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | 5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | 2 | both
5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | 5AU31450 | Box_NY_0.020SBS_169_6-3.125-5.9375 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Dispenser (Done) |  |  |  |  | Box Inspection | 5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | 2 | both
5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | 5AU31450 | Box_NY_0.020SBS_169_6-3.125-5.9375 | NY-Analytical | 772 | Box Inspection | Height | Height | 5.9375 | in | FDA,5 | 5.9062 | 5.9688 | Box Inspection | 5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | 2 | both
5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | 5AU31450 | Box_NY_0.020SBS_169_6-3.125-5.9375 | NY-Analytical | 772 | Box Inspection | Length | Length | 6 | in | FDA,5 | 5.9688 | 6.0312 | Box Inspection | 5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | 2 | both
5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | 5AU31450 | Box_NY_0.020SBS_169_6-3.125-5.9375 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.125 | in | FDA,5 | 3.0938 | 3.1562 | Box Inspection | 5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | 2 | both
5AU44410/Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | 5AU44410 | Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper |  | in | FDA,3 | 0.015 | 0.017 | Box Inspection |  |  | left_only
5AU44410/Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | 5AU44410 | Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Width | Width |  | in | FDA,5 | 1.8438 | 1.9062 | Box Inspection |  |  | left_only
5AU44410/Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | 5AU44410 | Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Length | Length |  | in | FDA,5 | 2.9062 | 2.9688 | Box Inspection |  |  | left_only
5AU44410/Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | 5AU44410 | Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Height | Height |  | in | FDA,5 | 3.7812 | 3.8438 | Box Inspection |  |  | left_only
5AU44410/Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | 5AU44410 | Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection |  |  | left_only
5AU44410/Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | 5AU44410 | Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection |  |  | left_only
5AU44410/Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | 5AU44410 | Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Straight Tuck End w/1" glued double header (Done) |  |  |  |  | Box Inspection |  |  | left_only
5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | 5AU57314 | Box_NY_0.016CCSBS_100_3-2.5-4 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 200 |  | FDA,0 |  |  | Box Inspection | 5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | 1 | both
5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | 5AU57314 | Box_NY_0.016CCSBS_100_3-2.5-4 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | 1 | both
5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | 5AU57314 | Box_NY_0.016CCSBS_100_3-2.5-4 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0155 | in | FDA,3 | 0.015 | 0.017 | Box Inspection | 5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | 1 | both
5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | 5AU57314 | Box_NY_0.016CCSBS_100_3-2.5-4 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | Clay Coated SBS (Done) |  |  |  |  | Box Inspection | 5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | 1 | both
5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | 5AU57314 | Box_NY_0.016CCSBS_100_3-2.5-4 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | 5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | 1 | both
5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | 5AU57314 | Box_NY_0.016CCSBS_100_3-2.5-4 | NY-Analytical | 772 | Box Inspection | Height | Height | 4 | in | FDA,5 | 3.9688 | 4.0312 | Box Inspection | 5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | 1 | both
5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | 5AU57314 | Box_NY_0.016CCSBS_100_3-2.5-4 | NY-Analytical | 772 | Box Inspection | Length | Length | 3 | in | FDA,5 | 2.9688 | 3.0312 | Box Inspection | 5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | 1 | both
5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | 5AU57314 | Box_NY_0.016CCSBS_100_3-2.5-4 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.5 | in | FDA,5 | 2.4688 | 2.5312 | Box Inspection | 5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | 1 | both
5AU66522/Box_NY_0.016BMPB_122_3.625-3-3.8125 | 5AU66522 | Box_NY_0.016BMPB_122_3.625-3-3.8125 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper |  | in | FDA,3 | 0.015 | 0.017 | Box Inspection |  |  | left_only
5AU66522/Box_NY_0.016BMPB_122_3.625-3-3.8125 | 5AU66522 | Box_NY_0.016BMPB_122_3.625-3-3.8125 | NY-Analytical | 772 | Box Inspection | Width | Width |  | in | FDA,5 | 2.9688 | 3.0312 | Box Inspection |  |  | left_only
5AU66522/Box_NY_0.016BMPB_122_3.625-3-3.8125 | 5AU66522 | Box_NY_0.016BMPB_122_3.625-3-3.8125 | NY-Analytical | 772 | Box Inspection | Length | Length |  | in | FDA,5 | 3.5938 | 3.6562 | Box Inspection |  |  | left_only
5AU66522/Box_NY_0.016BMPB_122_3.625-3-3.8125 | 5AU66522 | Box_NY_0.016BMPB_122_3.625-3-3.8125 | NY-Analytical | 772 | Box Inspection | Height | Height |  | in | FDA,5 | 3.7812 | 3.8438 | Box Inspection |  |  | left_only
5AU66522/Box_NY_0.016BMPB_122_3.625-3-3.8125 | 5AU66522 | Box_NY_0.016BMPB_122_3.625-3-3.8125 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | Bright Metalized Poly Board (Done) |  |  |  |  | Box Inspection |  |  | left_only
5AU66522/Box_NY_0.016BMPB_122_3.625-3-3.8125 | 5AU66522 | Box_NY_0.016BMPB_122_3.625-3-3.8125 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection |  |  | left_only
5AU66522/Box_NY_0.016BMPB_122_3.625-3-3.8125 | 5AU66522 | Box_NY_0.016BMPB_122_3.625-3-3.8125 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Straight Tuck End (Done) |  |  |  |  | Box Inspection |  |  | left_only
5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | 5BU19101 | Box_NY_0.020SBS_129_5.75-3-3.875 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | 1 | both
5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | 5BU19101 | Box_NY_0.020SBS_129_5.75-3-3.875 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0207 | in | FDA,3 | 0.019 | 0.021 | Box Inspection | 5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | 1 | both
5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | 5BU19101 | Box_NY_0.020SBS_129_5.75-3-3.875 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | 5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | 1 | both
5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | 5BU19101 | Box_NY_0.020SBS_129_5.75-3-3.875 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Straight Tuck End French Style Dipsenser w/False Bottom |  |  |  |  | Box Inspection | 5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | 1 | both
5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | 5BU19101 | Box_NY_0.020SBS_129_5.75-3-3.875 | NY-Analytical | 772 | Box Inspection | Height | Height | 3.875 | in | FDA,5 | 3.8438 | 3.9062 | Box Inspection | 5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | 1 | both
5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | 5BU19101 | Box_NY_0.020SBS_129_5.75-3-3.875 | NY-Analytical | 772 | Box Inspection | Length | Length | 5.75 | in | FDA,5 | 5.7188 | 5.7812 | Box Inspection | 5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | 1 | both
5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | 5BU19101 | Box_NY_0.020SBS_129_5.75-3-3.875 | NY-Analytical | 772 | Box Inspection | Width | Width | 3 | in | FDA,5 | 2.9688 | 3.0312 | Box Inspection | 5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | 1 | both
5BU26595/Box_NY_0.024SBS_102_4.390625-4.390625-12 | 5BU26595 | Box_NY_0.024SBS_102_4.390625-4.390625-12 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper |  | in | FDA,3 | 0.023 | 0.025 | Box Inspection |  |  | left_only
5BU26595/Box_NY_0.024SBS_102_4.390625-4.390625-12 | 5BU26595 | Box_NY_0.024SBS_102_4.390625-4.390625-12 | NY-Analytical | 772 | Box Inspection | Width | Width |  | in | FDA,5 | 4.3594 | 4.4219 | Box Inspection |  |  | left_only
5BU26595/Box_NY_0.024SBS_102_4.390625-4.390625-12 | 5BU26595 | Box_NY_0.024SBS_102_4.390625-4.390625-12 | NY-Analytical | 772 | Box Inspection | Length | Length |  | in | FDA,5 | 4.3594 | 4.4219 | Box Inspection |  |  | left_only
5BU26595/Box_NY_0.024SBS_102_4.390625-4.390625-12 | 5BU26595 | Box_NY_0.024SBS_102_4.390625-4.390625-12 | NY-Analytical | 772 | Box Inspection | Height | Height |  | in | FDA,5 | 11.9688 | 12.0312 | Box Inspection |  |  | left_only
5BU26595/Box_NY_0.024SBS_102_4.390625-4.390625-12 | 5BU26595 | Box_NY_0.024SBS_102_4.390625-4.390625-12 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection |  |  | left_only
5BU26595/Box_NY_0.024SBS_102_4.390625-4.390625-12 | 5BU26595 | Box_NY_0.024SBS_102_4.390625-4.390625-12 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection |  |  | left_only
5BU26595/Box_NY_0.024SBS_102_4.390625-4.390625-12 | 5BU26595 | Box_NY_0.024SBS_102_4.390625-4.390625-12 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom Dispenser Box |  |  |  |  | Box Inspection |  |  | left_only
5BU27500/Box_NY_0.024SBS_102_4.3906-4.3906-12 | 5BU27500 | Box_NY_0.024SBS_102_4.3906-4.3906-12 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper |  | in | FDA,3 | 0.023 | 0.025 | Box Inspection |  |  | left_only
5BU27500/Box_NY_0.024SBS_102_4.3906-4.3906-12 | 5BU27500 | Box_NY_0.024SBS_102_4.3906-4.3906-12 | NY-Analytical | 772 | Box Inspection | Length | Length |  | in | FDA,5 | 4.3594 | 4.4219 | Box Inspection |  |  | left_only
5BU27500/Box_NY_0.024SBS_102_4.3906-4.3906-12 | 5BU27500 | Box_NY_0.024SBS_102_4.3906-4.3906-12 | NY-Analytical | 772 | Box Inspection | Width | Width |  | in | FDA,5 | 4.3594 | 4.4219 | Box Inspection |  |  | left_only
5BU27500/Box_NY_0.024SBS_102_4.3906-4.3906-12 | 5BU27500 | Box_NY_0.024SBS_102_4.3906-4.3906-12 | NY-Analytical | 772 | Box Inspection | Height | Height |  | in | FDA,5 | 11.9688 | 12.0312 | Box Inspection |  |  | left_only
5BU27500/Box_NY_0.024SBS_102_4.3906-4.3906-12 | 5BU27500 | Box_NY_0.024SBS_102_4.3906-4.3906-12 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection |  |  | left_only
5BU27500/Box_NY_0.024SBS_102_4.3906-4.3906-12 | 5BU27500 | Box_NY_0.024SBS_102_4.3906-4.3906-12 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection |  |  | left_only
5BU27500/Box_NY_0.024SBS_102_4.3906-4.3906-12 | 5BU27500 | Box_NY_0.024SBS_102_4.3906-4.3906-12 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom Dispenser Box |  |  |  |  | Box Inspection |  |  | left_only
5BU37230/Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | 5BU37230 | Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper |  | in | FDA,3 | 0.017 | 0.019 | Box Inspection |  |  | left_only
5BU37230/Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | 5BU37230 | Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | NY-Analytical | 772 | Box Inspection | Height | Height |  | in | FDA,5 | 2.9062 | 2.9688 | Box Inspection |  |  | left_only
5BU37230/Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | 5BU37230 | Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | NY-Analytical | 772 | Box Inspection | Width | Width |  | in | FDA,5 | 3.3438 | 3.4062 | Box Inspection |  |  | left_only
5BU37230/Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | 5BU37230 | Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | NY-Analytical | 772 | Box Inspection | Length | Length |  | in | FDA,5 | 3.7812 | 3.8438 | Box Inspection |  |  | left_only
5BU37230/Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | 5BU37230 | Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | [5 empty] | Box Inspection |  |  | left_only
5BU37230/Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | 5BU37230 | Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection |  |  | left_only
5BU37230/Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | 5BU37230 | Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection |  |  | left_only
5BU37230/Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | 5BU37230 | Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom - French Style (Done) |  |  |  |  | Box Inspection |  |  | left_only
5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | 5BU55510 | Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | 2 | both
5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | 5BU55510 | Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0155 | in | FDA,3 | 0.015 | 0.017 | Box Inspection | 5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | 2 | both
5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | 5BU55510 | Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | Bright Metalized Poly Board (Done) |  |  |  |  | Box Inspection | 5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | 2 | both
5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | 5BU55510 | Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Straight Tuck End w/glued header (Done) |  |  |  |  | Box Inspection | 5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | 2 | both
5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | 5BU55510 | Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Height | Height | 3.8125 | in | FDA,5 | 3.7812 | 3.8438 | Box Inspection | 5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | 2 | both
5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | 5BU55510 | Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Length | Length | 2.9375 | in | FDA,5 | 2.9062 | 2.9688 | Box Inspection | 5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | 2 | both
5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | 5BU55510 | Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Width | Width | 1.875 | in | FDA,5 | 1.8438 | 1.9062 | Box Inspection | 5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | 2 | both
5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | 5BU71230 | Box_NY_0.018SBS_122_4.6875-3-3.8125 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | 2 | both
5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | 5BU71230 | Box_NY_0.018SBS_122_4.6875-3-3.8125 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.018 | in | FDA,3 | 0.017 | 0.019 | Box Inspection | 5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | 2 | both
5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | 5BU71230 | Box_NY_0.018SBS_122_4.6875-3-3.8125 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | 5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | 2 | both
5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | 5BU71230 | Box_NY_0.018SBS_122_4.6875-3-3.8125 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Straight Tuck End (Done) |  |  |  |  | Box Inspection | 5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | 2 | both
5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | 5BU71230 | Box_NY_0.018SBS_122_4.6875-3-3.8125 | NY-Analytical | 772 | Box Inspection | Height | Height | 3.82 | in | FDA,5 | 3.7812 | 3.8438 | Box Inspection | 5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | 2 | both
5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | 5BU71230 | Box_NY_0.018SBS_122_4.6875-3-3.8125 | NY-Analytical | 772 | Box Inspection | Length | Length | 4.7 | in | FDA,5 | 4.6562 | 4.7188 | Box Inspection | 5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | 2 | both
5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | 5BU71230 | Box_NY_0.018SBS_122_4.6875-3-3.8125 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.01 | in | FDA,5 | 2.9688 | 3.0312 | Box Inspection | 5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | 2 | both
5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | 5CA284TR | Box_NY_0.018SBS_111_6-5.3125-3.6875 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | 1 | both
5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | 5CA284TR | Box_NY_0.018SBS_111_6-5.3125-3.6875 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0175 | in | FDA,3 | 0.017 | 0.019 | Box Inspection | 5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | 1 | both
5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | 5CA284TR | Box_NY_0.018SBS_111_6-5.3125-3.6875 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | 5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | 1 | both
5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | 5CA284TR | Box_NY_0.018SBS_111_6-5.3125-3.6875 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom French Style w/Perfs (Done) |  |  |  |  | Box Inspection | 5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | 1 | both
5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | 5CA284TR | Box_NY_0.018SBS_111_6-5.3125-3.6875 | NY-Analytical | 772 | Box Inspection | Height | Height | 3.6875 | in | FDA,5 | 3.6562 | 3.7188 | Box Inspection | 5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | 1 | both
5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | 5CA284TR | Box_NY_0.018SBS_111_6-5.3125-3.6875 | NY-Analytical | 772 | Box Inspection | Length | Length | 6 | in | FDA,5 | 5.9688 | 6.0312 | Box Inspection | 5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | 1 | both
5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | 5CA284TR | Box_NY_0.018SBS_111_6-5.3125-3.6875 | NY-Analytical | 772 | Box Inspection | Width | Width | 5.3125 | in | FDA,5 | 5.2812 | 5.3438 | Box Inspection | 5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | 1 | both
5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 5CB19600 | Box_NY_0.018SBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | Box Inspection | 5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 5 | both
5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 5CB19600 | Box_NY_0.018SBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 5 | both
5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 5CB19600 | Box_NY_0.018SBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.018 | in | FDA,3 | 0.017 | 0.019 | Box Inspection | 5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 5 | both
5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 5CB19600 | Box_NY_0.018SBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | 5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 5 | both
5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 5CB19600 | Box_NY_0.018SBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | 5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 5 | both
5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 5CB19600 | Box_NY_0.018SBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Height | Height | 2.0938 | in | FDA,5 | 2.0625 | 2.125 | Box Inspection | 5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 5 | both
5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 5CB19600 | Box_NY_0.018SBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Length | Length | 5.125 | in | FDA,5 | 5.0938 | 5.1562 | Box Inspection | 5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 5 | both
5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 5CB19600 | Box_NY_0.018SBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.25 | in | FDA,5 | 2.2188 | 2.2812 | Box Inspection | 5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 5 | both
5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | 5CB51200 | Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | UCSBS |  |  |  |  | Box Inspection | 5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | 2 | both
5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | 5CB51200 | Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | 2 | both
5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | 5CB51200 | Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0155 | in | FDA,3 | 0.015 | 0.017 | Box Inspection | 5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | 2 | both
5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | 5CB51200 | Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | 5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | 2 | both
5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | 5CB51200 | Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | NY-Analytical | 772 | Box Inspection | Height | Height | 2.375 | in | FDA,5 | 2.3438 | 2.4062 | Box Inspection | 5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | 2 | both
5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | 5CB51200 | Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | NY-Analytical | 772 | Box Inspection | Length | Length | 6.125 | in | FDA,5 | 6.0938 | 6.1562 | Box Inspection | 5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | 2 | both
5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | 5CB51200 | Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.75 | in | FDA,5 | 2.7188 | 2.7812 | Box Inspection | 5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | 2 | both
5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | 5CC80100 | Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | UCSBS |  |  |  |  | Box Inspection | 5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | 3 | both
5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | 5CC80100 | Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | Box Inspection | 5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | 3 | both
5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | 5CC80100 | Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | 3 | both
5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | 5CC80100 | Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.015 | in | FDA,3 | 0.015 | 0.017 | Box Inspection | 5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | 3 | both
5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | 5CC80100 | Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | 5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | 3 | both
5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | 5CC80100 | Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Height | Height | 3.07 | in | FDA,5 | 3.0312 | 3.0938 | Box Inspection | 5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | 3 | both
5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | 5CC80100 | Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Length | Length | 5.07 | in | FDA,5 | 5.0312 | 5.0938 | Box Inspection | 5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | 3 | both
5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | 5CC80100 | Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Width | Width | 1.73 | in | FDA,5 | 1.7188 | 1.7812 | Box Inspection | 5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | 3 | both
5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 5CD35185 | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | UCSBS |  |  |  |  | Box Inspection | 5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 4 | both
5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 5CD35185 | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | Box Inspection | 5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 4 | both
5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 5CD35185 | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 4 | both
5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 5CD35185 | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.016 | in | FDA,3 | 0.015 | 0.017 | Box Inspection | 5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 4 | both
5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 5CD35185 | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | 5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 4 | both
5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 5CD35185 | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 772 | Box Inspection | Height | Height | 6.078 | in | FDA,5 | 6.0312 | 6.0938 | Box Inspection | 5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 4 | both
5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 5CD35185 | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 772 | Box Inspection | Length | Length | 6.5 | in | FDA,5 | 6.4688 | 6.5312 | Box Inspection | 5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 4 | both
5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 5CD35185 | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.312 | in | FDA,5 | 2.2812 | 2.3438 | Box Inspection | 5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 4 | both
5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 5CS42050 | Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | UCSBS |  |  |  |  | Box Inspection | 5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 1 | both
5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 5CS42050 | Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | Box Inspection | 5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 1 | both
5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 5CS42050 | Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 1 | both
5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 5CS42050 | Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.016 | in | FDA,3 | 0.015 | 0.017 | Box Inspection | 5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 1 | both
5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 5CS42050 | Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | 5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 1 | both
5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 5CS42050 | Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | NY-Analytical | 772 | Box Inspection | Height | Height | 5.81 | in | FDA,5 | 5.7812 | 5.8438 | Box Inspection | 5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 1 | both
5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 5CS42050 | Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | NY-Analytical | 772 | Box Inspection | Length | Length | 6.51 | in | FDA,5 | 6.4688 | 6.5312 | Box Inspection | 5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 1 | both
5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 5CS42050 | Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.32 | in | FDA,5 | 3.2812 | 3.3438 | Box Inspection | 5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 1 | both
5DA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | 5DA941TY | Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper |  | in | FDA,3 | 0.023 | 0.025 | Box Inspection |  |  | left_only
5DA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | 5DA941TY | Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | NY-Analytical | 772 | Box Inspection | Front Height | Front Height |  | in | FDA,5 | 3.4688 | 3.5312 | Box Inspection |  |  | left_only
5DA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | 5DA941TY | Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | NY-Analytical | 772 | Box Inspection | Length | Length |  | in | FDA,5 | 5.2188 | 5.2812 | Box Inspection |  |  | left_only
5DA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | 5DA941TY | Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | NY-Analytical | 772 | Box Inspection | Width | Width |  | in | FDA,5 | 6.2188 | 6.2812 | Box Inspection |  |  | left_only
5DA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | 5DA941TY | Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | NY-Analytical | 772 | Box Inspection | Back Height | Back Height |  | in | FDA,5 | 8.5938 | 8.6562 | Box Inspection |  |  | left_only
5DA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | 5DA941TY | Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection |  |  | left_only
5DA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | 5DA941TY | Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom French Style w/Perfs (Done) |  |  |  |  | Box Inspection |  |  | left_only
5DA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | 5DA941TY | Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | WCK (Done) |  |  |  |  | Box Inspection |  |  | left_only
5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | 5DH12824 | Box_NY_0.018SBS_100_4-2.125-6.125 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 30 |  | FDA,0 |  |  | Box Inspection | 5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | 2 | both
5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | 5DH12824 | Box_NY_0.018SBS_100_4-2.125-6.125 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | 2 | both
5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | 5DH12824 | Box_NY_0.018SBS_100_4-2.125-6.125 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.017 | in | FDA,3 | 0.017 | 0.019 | Box Inspection | 5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | 2 | both
5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | 5DH12824 | Box_NY_0.018SBS_100_4-2.125-6.125 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | 5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | 2 | both
5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | 5DH12824 | Box_NY_0.018SBS_100_4-2.125-6.125 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | 5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | 2 | both
5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | 5DH12824 | Box_NY_0.018SBS_100_4-2.125-6.125 | NY-Analytical | 772 | Box Inspection | Height | Height | 6.125 | in | FDA,5 | 6.0938 | 6.1562 | Box Inspection | 5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | 2 | both
5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | 5DH12824 | Box_NY_0.018SBS_100_4-2.125-6.125 | NY-Analytical | 772 | Box Inspection | Length | Length | 4 | in | FDA,5 | 3.9688 | 4.0312 | Box Inspection | 5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | 2 | both
5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | 5DH12824 | Box_NY_0.018SBS_100_4-2.125-6.125 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.125 | in | FDA,5 | 2.0938 | 2.1562 | Box Inspection | 5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | 2 | both
5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | 5EH21130 | Box_NY_0.016SBS_100_4.75-2.375-3.25 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | Box Inspection | 5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | 1 | both
5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | 5EH21130 | Box_NY_0.016SBS_100_4.75-2.375-3.25 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | 1 | both
5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | 5EH21130 | Box_NY_0.016SBS_100_4.75-2.375-3.25 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.015 | in | FDA,3 | 0.015 | 0.017 | Box Inspection | 5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | 1 | both
5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | 5EH21130 | Box_NY_0.016SBS_100_4.75-2.375-3.25 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | 5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | 1 | both
5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | 5EH21130 | Box_NY_0.016SBS_100_4.75-2.375-3.25 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | 5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | 1 | both
5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | 5EH21130 | Box_NY_0.016SBS_100_4.75-2.375-3.25 | NY-Analytical | 772 | Box Inspection | Height | Height | 3.26 | in | FDA,5 | 3.2188 | 3.2812 | Box Inspection | 5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | 1 | both
5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | 5EH21130 | Box_NY_0.016SBS_100_4.75-2.375-3.25 | NY-Analytical | 772 | Box Inspection | Length | Length | 4.75 | in | FDA,5 | 4.7188 | 4.7812 | Box Inspection | 5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | 1 | both
5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | 5EH21130 | Box_NY_0.016SBS_100_4.75-2.375-3.25 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.36 | in | FDA,5 | 2.3438 | 2.4062 | Certificate of Analysis | 5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | 1 | both
5FD41900/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 5FD41900 | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Box Inspection | 5FD41900/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 1 | both
5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | 5FH07050 | Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | 1 | both
5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | 5FH07050 | Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.02 | in | FDA,3 | 0.019 | 0.021 | Box Inspection | 5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | 1 | both
5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | 5FH07050 | Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | 5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | 1 | both
5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | 5FH07050 | Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Straight Tuck End w/Glued Internal Dividers&15/16" Dummy Platform Bottom (Done) |  |  |  |  | Box Inspection | 5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | 1 | both
5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | 5FH07050 | Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | NY-Analytical | 772 | Box Inspection | Height | Height | 6 | in | FDA,5 | 5.9844 | 6.0469 | Box Inspection | 5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | 1 | both
5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | 5FH07050 | Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | NY-Analytical | 772 | Box Inspection | Length | Length | 6.14 | in | FDA,5 | 6.0938 | 6.1562 | Box Inspection | 5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | 1 | both
5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | 5FH07050 | Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.44 | in | FDA,5 | 2.4062 | 2.4688 | Box Inspection | 5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | 1 | both
5GB10300/Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | 5GB10300 | Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper |  | in | FDA,3 | 0.015 | 0.017 | Box Inspection |  |  | left_only
5GB10300/Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | 5GB10300 | Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | NY-Analytical | 772 | Box Inspection | Height | Height |  | in | FDA,5 | 2.3438 | 2.4062 | Box Inspection |  |  | left_only
5GB10300/Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | 5GB10300 | Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | NY-Analytical | 772 | Box Inspection | Width | Width |  | in | FDA,5 | 2.7188 | 2.7812 | Box Inspection |  |  | left_only
5GB10300/Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | 5GB10300 | Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | NY-Analytical | 772 | Box Inspection | Length | Length |  | in | FDA,5 | 6.0938 | 6.1562 | Box Inspection |  |  | left_only
5GB10300/Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | 5GB10300 | Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection |  |  | left_only
5GB10300/Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | 5GB10300 | Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection |  |  | left_only
5GB10300/Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | 5GB10300 | Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | UCSBS |  |  |  |  | Box Inspection |  |  | left_only
5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 5HB59800 | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | Box Inspection | 5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 1 | both
5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 5HB59800 | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 1 | both
5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 5HB59800 | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | Back Flip | Back Flip | 5.158 | in | FDA,5 | 4.6875 | 5.3125 | Box Inspection | 5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 1 | both
5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 5HB59800 | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.001 | in | FDA,4 | 0.0008 | 0.0028 | Box Inspection | 5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 1 | both
5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 5HB59800 | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | LDPE (Done) |  |  |  |  | Box Inspection | 5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 1 | both
5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 5HB59800 | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Poly Bag w/fold over top on 2” wicketts (Done) |  |  |  |  | Box Inspection | 5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 1 | both
5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 5HB59800 | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | Gussett | Gussett | 3.8 | in | FDA,5 | 3.6875 | 4.3125 | Box Inspection | 5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 1 | both
5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 5HB59800 | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | Height | Height | 5.76 | in | FDA,5 | 5.1875 | 5.8125 | Box Inspection | 5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 1 | both
5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 5HB59800 | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | Width | Width | 11.812 | in | FDA,5 | 11.6875 | 12.3125 | Box Inspection | 5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 1 | both
5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | 5HC12400 | Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | UCSBS |  |  |  |  | Box Inspection | 5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | 1 | both
5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | 5HC12400 | Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 20 |  | FDA,0 |  |  | Box Inspection | 5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | 1 | both
5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | 5HC12400 | Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | 1 | both
5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | 5HC12400 | Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0169 | in | FDA,3 | 0.015 | 0.017 | Box Inspection | 5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | 1 | both
5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | 5HC12400 | Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | 5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | 1 | both
5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | 5HC12400 | Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | NY-Analytical | 772 | Box Inspection | Height | Height | 3.0625 | in | FDA,5 | 3.0312 | 3.0938 | Box Inspection | 5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | 1 | both
5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | 5HC12400 | Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | NY-Analytical | 772 | Box Inspection | Length | Length | 5.0625 | in | FDA,5 | 5.0312 | 5.0938 | Box Inspection | 5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | 1 | both
5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | 5HC12400 | Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | NY-Analytical | 772 | Box Inspection | Width | Width | 4.25 | in | FDA,5 | 4.2188 | 4.2812 | Box Inspection | 5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | 1 | both
5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | 5HC22324 | Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | 5 | both
5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | 5HC22324 | Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0198 | in | FDA,3 | 0.019 | 0.021 | Box Inspection | 5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | 5 | both
5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | 5HC22324 | Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | WCCN (Done) |  |  |  |  | Box Inspection | 5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | 5 | both
5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | 5HC22324 | Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | 5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | 5 | both
5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | 5HC22324 | Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Height | Height | 3.0625 | in | FDA,5 | 3.0312 | 3.0938 | Box Inspection | 5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | 5 | both
5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | 5HC22324 | Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Length | Length | 5.0625 | in | FDA,5 | 5.0312 | 5.0938 | Box Inspection | 5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | 5 | both
5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | 5HC22324 | Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Width | Width | 1.75 | in | FDA,5 | 1.7188 | 1.7812 | Box Inspection | 5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | 5 | both
5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | 5JB16400 | Box_NY_0.016UCSBS_100_4-3.125-2.375 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | UCSBS |  |  |  |  | Box Inspection | 5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | 4 | both
5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | 5JB16400 | Box_NY_0.016UCSBS_100_4-3.125-2.375 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | 4 | both
5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | 5JB16400 | Box_NY_0.016UCSBS_100_4-3.125-2.375 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.015 | in | FDA,3 | 0.015 | 0.017 | Box Inspection | 5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | 4 | both
5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | 5JB16400 | Box_NY_0.016UCSBS_100_4-3.125-2.375 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | 5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | 4 | both
5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | 5JB16400 | Box_NY_0.016UCSBS_100_4-3.125-2.375 | NY-Analytical | 772 | Box Inspection | Height | Height | 2.375 | in | FDA,5 | 2.3438 | 2.4062 | Box Inspection | 5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | 4 | both
5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | 5JB16400 | Box_NY_0.016UCSBS_100_4-3.125-2.375 | NY-Analytical | 772 | Box Inspection | Length | Length | 4 | in | FDA,5 | 3.9688 | 4.0312 | Box Inspection | 5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | 4 | both
5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | 5JB16400 | Box_NY_0.016UCSBS_100_4-3.125-2.375 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.125 | in | FDA,5 | 3.0938 | 3.1562 | Box Inspection | 5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | 4 | both
5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | 5LB60302 | Box_NY_0.016UCSBS_100_6.125-4-2.375 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | UCSBS |  |  |  |  | Box Inspection | 5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | 7 | both
5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | 5LB60302 | Box_NY_0.016UCSBS_100_6.125-4-2.375 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | Box Inspection | 5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | 7 | both
5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | 5LB60302 | Box_NY_0.016UCSBS_100_6.125-4-2.375 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | 7 | both
5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | 5LB60302 | Box_NY_0.016UCSBS_100_6.125-4-2.375 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.016 | in | FDA,3 | 0.015 | 0.017 | Box Inspection | 5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | 7 | both
5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | 5LB60302 | Box_NY_0.016UCSBS_100_6.125-4-2.375 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | 5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | 7 | both
5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | 5LB60302 | Box_NY_0.016UCSBS_100_6.125-4-2.375 | NY-Analytical | 772 | Box Inspection | Height | Height | 2.375 | in | FDA,5 | 2.3438 | 2.4062 | Box Inspection | 5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | 7 | both
5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | 5LB60302 | Box_NY_0.016UCSBS_100_6.125-4-2.375 | NY-Analytical | 772 | Box Inspection | Length | Length | 6.0938 | in | FDA,5 | 6.0938 | 6.1562 | Box Inspection | 5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | 7 | both
5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | 5LB60302 | Box_NY_0.016UCSBS_100_6.125-4-2.375 | NY-Analytical | 772 | Box Inspection | Width | Width | 4 | in | FDA,5 | 3.9688 | 4.0312 | Box Inspection | 5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | 7 | both
5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | 5MS48050 | Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | UCSBS |  |  |  |  | Box Inspection | 5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | 1 | both
5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | 5MS48050 | Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | Box Inspection | 5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | 1 | both
5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | 5MS48050 | Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | 1 | both
5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | 5MS48050 | Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0155 | in | FDA,3 | 0.015 | 0.017 | Box Inspection | 5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | 1 | both
5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | 5MS48050 | Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | 5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | 1 | both
5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | 5MS48050 | Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | NY-Analytical | 772 | Box Inspection | Height | Height | 5.8125 | in | FDA,5 | 5.7812 | 5.8438 | Box Inspection | 5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | 1 | both
5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | 5MS48050 | Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | NY-Analytical | 772 | Box Inspection | Length | Length | 6.5 | in | FDA,5 | 6.4688 | 6.5312 | Box Inspection | 5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | 1 | both
5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | 5MS48050 | Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.3125 | in | FDA,5 | 3.2812 | 3.3438 | Box Inspection | 5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | 1 | both
5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | 5NB40600 | Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | UCSBS |  |  |  |  | Box Inspection | 5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | 3 | both
5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | 5NB40600 | Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | Box Inspection | 5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | 3 | both
5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | 5NB40600 | Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | 3 | both
5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | 5NB40600 | Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0155 | in | FDA,3 | 0.015 | 0.017 | Box Inspection | 5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | 3 | both
5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | 5NB40600 | Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | 5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | 3 | both
5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | 5NB40600 | Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Height | Height | 2.07 | in | FDA,5 | 2.0625 | 2.125 | Box Inspection | 5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | 3 | both
5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | 5NB40600 | Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Length | Length | 5.13 | in | FDA,5 | 5.0938 | 5.1562 | Box Inspection | 5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | 3 | both
5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | 5NB40600 | Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.24 | in | FDA,5 | 2.2188 | 2.2812 | Box Inspection | 5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | 3 | both
5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | 5NU87250 | Box_NY_0.024SBS_127_4.375-4.375-12 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom Dispenser Box |  |  |  |  | Box Inspection | 5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | 1 | both
5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | 5NU87250 | Box_NY_0.024SBS_127_4.375-4.375-12 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | 1 | both
5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | 5NU87250 | Box_NY_0.024SBS_127_4.375-4.375-12 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.024 | in | FDA,3 | 0.023 | 0.025 | Box Inspection | 5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | 1 | both
5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | 5NU87250 | Box_NY_0.024SBS_127_4.375-4.375-12 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | 5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | 1 | both
5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | 5NU87250 | Box_NY_0.024SBS_127_4.375-4.375-12 | NY-Analytical | 772 | Box Inspection | Height | Height | 12 | in | FDA,5 | 11.9688 | 12.0312 | Box Inspection | 5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | 1 | both
5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | 5NU87250 | Box_NY_0.024SBS_127_4.375-4.375-12 | NY-Analytical | 772 | Box Inspection | Length | Length | 4.39 | in | FDA,5 | 4.3438 | 4.4062 | Box Inspection | 5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | 1 | both
5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | 5NU87250 | Box_NY_0.024SBS_127_4.375-4.375-12 | NY-Analytical | 772 | Box Inspection | Width | Width | 4.39 | in | FDA,5 | 4.3438 | 4.4062 | Box Inspection | 5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | 1 | both
5OD125PB/Box_80A | 5OD125PB | Box_80A | NY-Analytical | 772 | Box Inspection | Caliper | Caliper |  | in | FDA,4 | 0.0005 | 0.0025 | Box Inspection |  |  | left_only
5OD125PB/Box_80A | 5OD125PB | Box_80A | NY-Analytical | 772 | Box Inspection | Outside Flap | Outside Flap |  | in | FDA,5 | 2.2188 | 2.2812 | Box Inspection |  |  | left_only
5OD125PB/Box_80A | 5OD125PB | Box_80A | NY-Analytical | 772 | Box Inspection | Length | Length |  | in | FDA,5 | 6.0625 | 6.5 | Box Inspection |  |  | left_only
5OD125PB/Box_80A | 5OD125PB | Box_80A | NY-Analytical | 772 | Box Inspection | Width | Width |  | in | FDA,5 | 8.25 | 8.5 | Box Inspection |  |  | left_only
5OD125PB/Box_80A | 5OD125PB | Box_80A | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | LDPE (Done) |  |  |  |  | Box Inspection |  |  | left_only
5OD125PB/Box_80A | 5OD125PB | Box_80A | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection |  |  | left_only
5OD125PB/Box_80A | 5OD125PB | Box_80A | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Poly Bag w/fold over top on 4” wicketts (Done) |  |  |  |  | Box Inspection |  |  | left_only
5OD239PD/Box_51A | 5OD239PD | Box_51A | NY-Analytical | 772 | Box Inspection | Caliper | Caliper |  | in | FDA,3 |  |  | Box Inspection |  |  | left_only
5OD239PD/Box_51A | 5OD239PD | Box_51A | NY-Analytical | 772 | Box Inspection | Width | Width |  | in | FDA,5 | 5.4688 | 5.5312 | Box Inspection |  |  | left_only
5OD239PD/Box_51A | 5OD239PD | Box_51A | NY-Analytical | 772 | Box Inspection | Length | Length |  | in | FDA,5 | 15.3438 | 15.4062 | Box Inspection |  |  | left_only
5OD239PD/Box_51A | 5OD239PD | Box_51A | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | Chip Board (Done) |  |  |  |  | Box Inspection |  |  | left_only
5OD239PD/Box_51A | 5OD239PD | Box_51A | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Layer Pad (Done) |  |  |  |  | Box Inspection |  |  | left_only
5OD239PD/Box_51A | 5OD239PD | Box_51A | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection |  |  | left_only
5OH382PD/Box_50D | 5OH382PD | Box_50D | NY-Analytical | 772 | Box Inspection | Caliper | Caliper |  | in | FDA,3 |  |  | Box Inspection |  |  | left_only
5OH382PD/Box_50D | 5OH382PD | Box_50D | NY-Analytical | 772 | Box Inspection | Width | Width |  | in | FDA,5 | 9.4688 | 9.5312 | Box Inspection |  |  | left_only
5OH382PD/Box_50D | 5OH382PD | Box_50D | NY-Analytical | 772 | Box Inspection | Length | Length |  | in | FDA,5 | 11.5938 | 11.6562 | Box Inspection |  |  | left_only
5OH382PD/Box_50D | 5OH382PD | Box_50D | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | Any Board (Done) |  |  |  |  | Box Inspection |  |  | left_only
5OH382PD/Box_50D | 5OH382PD | Box_50D | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Cut Pad (Done) |  |  |  |  | Box Inspection |  |  | left_only
5OH382PD/Box_50D | 5OH382PD | Box_50D | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection |  |  | left_only
5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | 5OP067BT | Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 50 |  | FDA,0 |  |  | Box Inspection | 5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | 1 | both
5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | 5OP067BT | Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | 1 | both
5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | 5OP067BT | Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.016 | in | FDA,3 | 0.015 | 0.017 | Box Inspection | 5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | 1 | both
5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | 5OP067BT | Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS C1S (Done) |  |  |  |  | Box Inspection | 5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | 1 | both
5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | 5OP067BT | Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Open Top/Tuck Bottom (Done) |  |  |  |  | Box Inspection | 5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | 1 | both
5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | 5OP067BT | Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Height | Height | 2.44 | in | FDA,5 | 2.4062 | 2.4688 | Box Inspection | 5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | 1 | both
5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | 5OP067BT | Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Length | Length | 3.2 | in | FDA,5 | 3.1488 | 3.2113 | Box Inspection | 5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | 1 | both
5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | 5OP067BT | Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.19 | in | FDA,5 | 2.1562 | 2.2188 | Box Inspection | 5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | 1 | both
5OP073BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | 5OP073BT | Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper |  | in | FDA,3 | 0.015 | 0.017 | Box Inspection |  |  | left_only
5OP073BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | 5OP073BT | Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Width | Width |  | in | FDA,5 | 2.1562 | 2.2188 | Box Inspection |  |  | left_only
5OP073BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | 5OP073BT | Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Height | Height |  | in | FDA,5 | 2.4062 | 2.4688 | Box Inspection |  |  | left_only
5OP073BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | 5OP073BT | Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Length | Length |  | in | FDA,5 | 3.1488 | 3.2113 | Box Inspection |  |  | left_only
5OP073BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | 5OP073BT | Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 |  |  | Box Inspection |  |  | left_only
5OP073BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | 5OP073BT | Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection |  |  | left_only
5OP073BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | 5OP073BT | Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Open Top/Tuck Bottom (Done) |  |  |  |  | Box Inspection |  |  | left_only
5OP073BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | 5OP073BT | Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS C1S (Done) |  |  |  |  | Box Inspection |  |  | left_only
5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | 5OP092BT | Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 315 |  | FDA,0 |  |  | Box Inspection | 5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | 1 | both
5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | 5OP092BT | Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | 1 | both
5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | 5OP092BT | Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.016 | in | FDA,3 | 0.015 | 0.017 | Box Inspection | 5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | 1 | both
5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | 5OP092BT | Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS C1S (Done) |  |  |  |  | Box Inspection | 5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | 1 | both
5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | 5OP092BT | Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Open Top/Tuck Bottom (Done) |  |  |  |  | Box Inspection | 5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | 1 | both
5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | 5OP092BT | Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | NY-Analytical | 772 | Box Inspection | Height | Height | 2.6875 | in | FDA,5 | 2.6562 | 2.7188 | Box Inspection | 5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | 1 | both
5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | 5OP092BT | Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | NY-Analytical | 772 | Box Inspection | Length | Length | 3.1875 | in | FDA,5 | 3.1488 | 3.2113 | Box Inspection | 5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | 1 | both
5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | 5OP092BT | Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | NY-Analytical | 772 | Box Inspection | Width | Width | 1.4375 | in | FDA,5 | 1.4062 | 1.4688 | Box Inspection | 5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | 1 | both
5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | 5OP092PB | Box_NY_LDPE_407_13.5-6.5-6-4 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | Box Inspection | 5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | 1 | both
5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | 5OP092PB | Box_NY_LDPE_407_13.5-6.5-6-4 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | 1 | both
5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | 5OP092PB | Box_NY_LDPE_407_13.5-6.5-6-4 | NY-Analytical | 772 | Box Inspection | Back Flip | Back Flip | 5.75 | in | FDA,5 | 5.6875 | 6.3125 | Box Inspection | 5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | 1 | both
5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | 5OP092PB | Box_NY_LDPE_407_13.5-6.5-6-4 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | LDPE (Done) |  |  |  |  | Box Inspection | 5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | 1 | both
5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | 5OP092PB | Box_NY_LDPE_407_13.5-6.5-6-4 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Poly Bag w/fold over top on 2” wicketts (Done) |  |  |  |  | Box Inspection | 5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | 1 | both
5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | 5OP092PB | Box_NY_LDPE_407_13.5-6.5-6-4 | NY-Analytical | 772 | Box Inspection | Gussett | Gussett | 4 | in | FDA,5 | 3.6875 | 4.3125 | Box Inspection | 5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | 1 | both
5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | 5OP092PB | Box_NY_LDPE_407_13.5-6.5-6-4 | NY-Analytical | 772 | Box Inspection | Height | Height | 6.5 | in | FDA,5 | 6.1875 | 6.8125 | Box Inspection | 5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | 1 | both
5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | 5OP092PB | Box_NY_LDPE_407_13.5-6.5-6-4 | NY-Analytical | 772 | Box Inspection | Width | Width | 13.1875 | in | FDA,5 | 13.1875 | 13.8125 | Box Inspection | 5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | 1 | both
5OP308PD/Box_50C | 5OP308PD | Box_50C | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5OP308PD/Box_50C | 1 | both
5OP308PD/Box_50C | 5OP308PD | Box_50C | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.02 | in | FDA,3 | 0.018 | 0.02 | Box Inspection | 5OP308PD/Box_50C | 1 | both
5OP308PD/Box_50C | 5OP308PD | Box_50C | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | Chip Board (Done) |  |  |  |  | Box Inspection | 5OP308PD/Box_50C | 1 | both
5OP308PD/Box_50C | 5OP308PD | Box_50C | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Cut Pad (Done) |  |  |  |  | Box Inspection | 5OP308PD/Box_50C | 1 | both
5OP308PD/Box_50C | 5OP308PD | Box_50C | NY-Analytical | 772 | Box Inspection | Length | Length | 7 | in | FDA,5 | 6.9688 | 7.0312 | Box Inspection | 5OP308PD/Box_50C | 1 | both
5OP308PD/Box_50C | 5OP308PD | Box_50C | NY-Analytical | 772 | Box Inspection | Width | Width | 12.28 | in | FDA,5 | 12.2188 | 12.2812 | Box Inspection | 5OP308PD/Box_50C | 1 | both
5PB114IS/Box_NY_NA50#_402_2.5-3.25 | 5PB114IS | Box_NY_NA50#_402_2.5-3.25 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 500 |  | FDA,0 |  |  | Box Inspection | 5PB114IS/Box_NY_NA50#_402_2.5-3.25 | 1 | both
5PB114IS/Box_NY_NA50#_402_2.5-3.25 | 5PB114IS | Box_NY_NA50#_402_2.5-3.25 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5PB114IS/Box_NY_NA50#_402_2.5-3.25 | 1 | both
5PB114IS/Box_NY_NA50#_402_2.5-3.25 | 5PB114IS | Box_NY_NA50#_402_2.5-3.25 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | 50# White Paper Stock (Done) |  |  |  |  | Box Inspection | 5PB114IS/Box_NY_NA50#_402_2.5-3.25 | 1 | both
5PB114IS/Box_NY_NA50#_402_2.5-3.25 | 5PB114IS | Box_NY_NA50#_402_2.5-3.25 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Insert (Done) |  |  |  |  | Box Inspection | 5PB114IS/Box_NY_NA50#_402_2.5-3.25 | 1 | both
5PB114IS/Box_NY_NA50#_402_2.5-3.25 | 5PB114IS | Box_NY_NA50#_402_2.5-3.25 | NY-Analytical | 772 | Box Inspection | Length | Length | 2.5 | in | FDA,5 | 2.4688 | 2.5312 | Box Inspection | 5PB114IS/Box_NY_NA50#_402_2.5-3.25 | 1 | both
5PB114IS/Box_NY_NA50#_402_2.5-3.25 | 5PB114IS | Box_NY_NA50#_402_2.5-3.25 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.25 | in | FDA,5 | 3.2188 | 3.2812 | Box Inspection | 5PB114IS/Box_NY_NA50#_402_2.5-3.25 | 1 | both
5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | 5PS40750 | Box_NY_0.018SBS_100_7.125-3-6.625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | 2 | both
5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | 5PS40750 | Box_NY_0.018SBS_100_7.125-3-6.625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0175 | in | FDA,3 | 0.017 | 0.019 | Box Inspection | 5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | 2 | both
5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | 5PS40750 | Box_NY_0.018SBS_100_7.125-3-6.625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | 5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | 2 | both
5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | 5PS40750 | Box_NY_0.018SBS_100_7.125-3-6.625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | 5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | 2 | both
5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | 5PS40750 | Box_NY_0.018SBS_100_7.125-3-6.625 | NY-Analytical | 772 | Box Inspection | Height | Height | 6.625 | in | FDA,5 | 6.5938 | 6.6562 | Box Inspection | 5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | 2 | both
5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | 5PS40750 | Box_NY_0.018SBS_100_7.125-3-6.625 | NY-Analytical | 772 | Box Inspection | Length | Length | 7.125 | in | FDA,5 | 7.0938 | 7.1562 | Box Inspection | 5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | 2 | both
5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | 5PS40750 | Box_NY_0.018SBS_100_7.125-3-6.625 | NY-Analytical | 772 | Box Inspection | Width | Width | 3 | in | FDA,5 | 2.9688 | 3.0312 | Box Inspection | 5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | 2 | both
5QS273IS/Box_NY_NA50#_402_5-3.25 | 5QS273IS | Box_NY_NA50#_402_5-3.25 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 500 |  | FDA,0 |  |  | Box Inspection | 5QS273IS/Box_NY_NA50#_402_5-3.25 | 1 | both
5QS273IS/Box_NY_NA50#_402_5-3.25 | 5QS273IS | Box_NY_NA50#_402_5-3.25 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5QS273IS/Box_NY_NA50#_402_5-3.25 | 1 | both
5QS273IS/Box_NY_NA50#_402_5-3.25 | 5QS273IS | Box_NY_NA50#_402_5-3.25 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | 50# White Paper Stock (Done) |  |  |  |  | Box Inspection | 5QS273IS/Box_NY_NA50#_402_5-3.25 | 1 | both
5QS273IS/Box_NY_NA50#_402_5-3.25 | 5QS273IS | Box_NY_NA50#_402_5-3.25 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Insert (Done) |  |  |  |  | Box Inspection | 5QS273IS/Box_NY_NA50#_402_5-3.25 | 1 | both
5QS273IS/Box_NY_NA50#_402_5-3.25 | 5QS273IS | Box_NY_NA50#_402_5-3.25 | NY-Analytical | 772 | Box Inspection | Length | Length | 5 | in | FDA,5 | 4.9688 | 5.0312 | Box Inspection | 5QS273IS/Box_NY_NA50#_402_5-3.25 | 1 | both
5QS273IS/Box_NY_NA50#_402_5-3.25 | 5QS273IS | Box_NY_NA50#_402_5-3.25 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.2188 | in | FDA,5 | 3.2188 | 3.2812 | Box Inspection | 5QS273IS/Box_NY_NA50#_402_5-3.25 | 1 | both
5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | 5QS41950 | Box_NY_0.018SBS_100_6.625-4.125-8.125 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | 5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | 2 | both
5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | 5QS41950 | Box_NY_0.018SBS_100_6.625-4.125-8.125 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.017 | in | FDA,3 | 0.017 | 0.019 | Box Inspection | 5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | 2 | both
5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | 5QS41950 | Box_NY_0.018SBS_100_6.625-4.125-8.125 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | 5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | 2 | both
5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | 5QS41950 | Box_NY_0.018SBS_100_6.625-4.125-8.125 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | 5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | 2 | both
5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | 5QS41950 | Box_NY_0.018SBS_100_6.625-4.125-8.125 | NY-Analytical | 772 | Box Inspection | Height | Height | 8.125 | in | FDA,5 | 8.0938 | 8.1562 | Box Inspection | 5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | 2 | both
5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | 5QS41950 | Box_NY_0.018SBS_100_6.625-4.125-8.125 | NY-Analytical | 772 | Box Inspection | Length | Length | 6.625 | in | FDA,5 | 6.5938 | 6.6562 | Box Inspection | 5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | 2 | both
5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | 5QS41950 | Box_NY_0.018SBS_100_6.625-4.125-8.125 | NY-Analytical | 772 | Box Inspection | Width | Width | 4.125 | in | FDA,5 | 4.0938 | 4.1562 | Box Inspection | 5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | 2 | both
6AA34724/Shipper_NY_Skip lot testing | 6AA34724 | Shipper_NY_Skip lot testing | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Shipper Inspection | 6AA34724/Shipper_NY_Skip lot testing | 2 | both
6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | 6AA34724 | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height |  | in | FDA,4 | 6.0625 | 6.4375 | TM 141 / Visual |  |  | left_only
6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | 6AA34724 | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 6.8125 | 7.1875 | TM 141 / Visual |  |  | left_only
6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | 6AA34724 | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual |  |  | left_only
6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | 6AA34724 | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 |  |  | TM 141 / Visual |  |  | left_only
6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | 6AA34724 | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual |  |  | left_only
6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | 6AA34724 | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | 6AA34724 | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | 6AA34724 | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | 6AA34724 | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | 6AA617S96 | SHIP_NY_44K_100_12.75-11.625-7.875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | 2 | both
6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | 6AA617S96 | SHIP_NY_44K_100_12.75-11.625-7.875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | 6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | 2 | both
6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | 6AA617S96 | SHIP_NY_44K_100_12.75-11.625-7.875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | 2 | both
6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | 6AA617S96 | SHIP_NY_44K_100_12.75-11.625-7.875 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7.75 | in | FDA,4 | 7.6875 | 8.0625 | TM 141 / Visual | 6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | 2 | both
6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | 6AA617S96 | SHIP_NY_44K_100_12.75-11.625-7.875 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 12.75 | in | FDA,4 | 12.5625 | 12.9375 | TM 141 / Visual | 6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | 2 | both
6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | 6AA617S96 | SHIP_NY_44K_100_12.75-11.625-7.875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | 2 | both
6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | 6AA617S96 | SHIP_NY_44K_100_12.75-11.625-7.875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | 2 | both
6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | 6AA617S96 | SHIP_NY_44K_100_12.75-11.625-7.875 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 11.5 | in | FDA,4 | 11.4375 | 11.8125 | TM 141 / Visual | 6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | 2 | both
6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | 6AA632TY | SHIP_NY_44K_231_18.5-7.375-14 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | 7 | both
6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | 6AA632TY | SHIP_NY_44K_231_18.5-7.375-14 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | 6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | 7 | both
6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | 6AA632TY | SHIP_NY_44K_231_18.5-7.375-14 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | 7 | both
6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | 6AA632TY | SHIP_NY_44K_231_18.5-7.375-14 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 14 | in | FDA,4 | 13.875 | 14.125 | TM 141 / Visual | 6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | 7 | both
6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | 6AA632TY | SHIP_NY_44K_231_18.5-7.375-14 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 18.5 | in | FDA,4 | 18.375 | 18.625 | TM 141 / Visual | 6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | 7 | both
6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | 6AA632TY | SHIP_NY_44K_231_18.5-7.375-14 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | 7 | both
6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | 6AA632TY | SHIP_NY_44K_231_18.5-7.375-14 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | D/C (Done) |  |  |  |  | TM 141 / Visual | 6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | 7 | both
6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | 6AA632TY | SHIP_NY_44K_231_18.5-7.375-14 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.5 | in | FDA,4 | 7.25 | 7.5 | TM 141 / Visual | 6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | 7 | both
6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 6AA709S48 | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 200 |  | FDA,0 |  |  | TM 141 / Visual | 6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 5 | both
6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 6AA709S48 | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 5 | both
6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 6AA709S48 | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 40 | Lbs./In. | FDA,0 | 40 |  | TM 141 / Visual | 6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 5 | both
6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 6AA709S48 | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 5 | both
6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 6AA709S48 | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6.125 | in | FDA,4 | 5.9375 | 6.3125 | TM 141 / Visual | 6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 5 | both
6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 6AA709S48 | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15 | in | FDA,4 | 14.8125 | 15.1875 | TM 141 / Visual | 6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 5 | both
6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 6AA709S48 | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 5 | both
6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 6AA709S48 | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 5 | both
6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 6AA709S48 | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 11 | in | FDA,4 | 10.8125 | 11.1875 | TM 141 / Visual | 6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 5 | both
6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 6AA73530 | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual | 6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 1 | both
6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 6AA73530 | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 1 | both
6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 6AA73530 | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 1 | both
6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 6AA73530 | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 1 | both
6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 6AA73530 | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 8.25 | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual | 6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 1 | both
6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 6AA73530 | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 10 | in | FDA,4 | 9.8125 | 10.1875 | TM 141 / Visual | 6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 1 | both
6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 6AA73530 | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 1 | both
6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 6AA73530 | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 1 | both
6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 6AA73530 | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 8.375 | in | FDA,4 | 8.1875 | 8.5625 | TM 141 / Visual | 6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 1 | both
6AA76930/SHIP_NY_44K_100_9-9-4.75 | 6AA76930 | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual | 6AA76930/SHIP_NY_44K_100_9-9-4.75 | 1 | both
6AA76930/SHIP_NY_44K_100_9-9-4.75 | 6AA76930 | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6AA76930/SHIP_NY_44K_100_9-9-4.75 | 1 | both
6AA76930/SHIP_NY_44K_100_9-9-4.75 | 6AA76930 | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | 6AA76930/SHIP_NY_44K_100_9-9-4.75 | 1 | both
6AA76930/SHIP_NY_44K_100_9-9-4.75 | 6AA76930 | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6AA76930/SHIP_NY_44K_100_9-9-4.75 | 1 | both
6AA76930/SHIP_NY_44K_100_9-9-4.75 | 6AA76930 | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 4.75 | in | FDA,4 | 4.625 | 4.875 | TM 141 / Visual | 6AA76930/SHIP_NY_44K_100_9-9-4.75 | 1 | both
6AA76930/SHIP_NY_44K_100_9-9-4.75 | 6AA76930 | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 9 | in | FDA,4 | 8.875 | 9.125 | TM 141 / Visual | 6AA76930/SHIP_NY_44K_100_9-9-4.75 | 1 | both
6AA76930/SHIP_NY_44K_100_9-9-4.75 | 6AA76930 | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6AA76930/SHIP_NY_44K_100_9-9-4.75 | 1 | both
6AA76930/SHIP_NY_44K_100_9-9-4.75 | 6AA76930 | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6AA76930/SHIP_NY_44K_100_9-9-4.75 | 1 | both
6AA76930/SHIP_NY_44K_100_9-9-4.75 | 6AA76930 | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 9 | in | FDA,4 | 8.875 | 9.125 | TM 141 / Visual | 6AA76930/SHIP_NY_44K_100_9-9-4.75 | 1 | both
6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 6AB16400 | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual | 6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 4 | both
6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 6AB16400 | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 4 | both
6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 6AB16400 | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 4 | both
6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 6AB16400 | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 4 | both
6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 6AB16400 | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 5.25 | in | FDA,4 | 4.9375 | 5.3125 | TM 141 / Visual | 6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 4 | both
6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 6AB16400 | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 16.125 | in | FDA,4 | 15.8125 | 16.1875 | TM 141 / Visual | 6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 4 | both
6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 6AB16400 | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 4 | both
6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 6AB16400 | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 4 | both
6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 6AB16400 | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 4.25 | in | FDA,4 | 4.0625 | 4.4375 | TM 141 / Visual | 6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 4 | both
6AD11055/SHIP_NY_32K_100_10.625-6.375-11.25 | 6AD11055 | SHIP_NY_32K_100_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 6.1875 | 6.5625 | TM 141 / Visual |  |  | left_only
6AD11055/SHIP_NY_32K_100_10.625-6.375-11.25 | 6AD11055 | SHIP_NY_32K_100_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 10.4375 | 10.8125 | TM 141 / Visual |  |  | left_only
6AD11055/SHIP_NY_32K_100_10.625-6.375-11.25 | 6AD11055 | SHIP_NY_32K_100_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height |  | in | FDA,4 | 11.0625 | 11.4375 | TM 141 / Visual |  |  | left_only
6AD11055/SHIP_NY_32K_100_10.625-6.375-11.25 | 6AD11055 | SHIP_NY_32K_100_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual |  |  | left_only
6AD11055/SHIP_NY_32K_100_10.625-6.375-11.25 | 6AD11055 | SHIP_NY_32K_100_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AD11055/SHIP_NY_32K_100_10.625-6.375-11.25 | 6AD11055 | SHIP_NY_32K_100_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AD11055/SHIP_NY_32K_100_10.625-6.375-11.25 | 6AD11055 | SHIP_NY_32K_100_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AD11055/SHIP_NY_32K_100_10.625-6.375-11.25 | 6AD11055 | SHIP_NY_32K_100_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AD41900/SHIP_NY_32K_100_13.25-12-6.125 | 6AD41900 | SHIP_NY_32K_100_13.25-12-6.125 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height |  | in | FDA,4 | 5.9375 | 6.3125 | TM 141 / Visual |  |  | left_only
6AD41900/SHIP_NY_32K_100_13.25-12-6.125 | 6AD41900 | SHIP_NY_32K_100_13.25-12-6.125 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 11.8125 | 12.1875 | TM 141 / Visual |  |  | left_only
6AD41900/SHIP_NY_32K_100_13.25-12-6.125 | 6AD41900 | SHIP_NY_32K_100_13.25-12-6.125 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 13.0625 | 13.4375 | TM 141 / Visual |  |  | left_only
6AD41900/SHIP_NY_32K_100_13.25-12-6.125 | 6AD41900 | SHIP_NY_32K_100_13.25-12-6.125 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual |  |  | left_only
6AD41900/SHIP_NY_32K_100_13.25-12-6.125 | 6AD41900 | SHIP_NY_32K_100_13.25-12-6.125 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AD41900/SHIP_NY_32K_100_13.25-12-6.125 | 6AD41900 | SHIP_NY_32K_100_13.25-12-6.125 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AD41900/SHIP_NY_32K_100_13.25-12-6.125 | 6AD41900 | SHIP_NY_32K_100_13.25-12-6.125 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AD41900/SHIP_NY_32K_100_13.25-12-6.125 | 6AD41900 | SHIP_NY_32K_100_13.25-12-6.125 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AH58195/SHIP_NY_32K_100_13.4375-6.25-9.6875 | 6AH58195 | SHIP_NY_32K_100_13.4375-6.25-9.6875 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 6.0625 | 6.4375 | TM 141 / Visual |  |  | left_only
6AH58195/SHIP_NY_32K_100_13.4375-6.25-9.6875 | 6AH58195 | SHIP_NY_32K_100_13.4375-6.25-9.6875 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height |  | in | FDA,4 | 9.5 | 9.875 | TM 141 / Visual |  |  | left_only
6AH58195/SHIP_NY_32K_100_13.4375-6.25-9.6875 | 6AH58195 | SHIP_NY_32K_100_13.4375-6.25-9.6875 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 13.25 | 13.625 | TM 141 / Visual |  |  | left_only
6AH58195/SHIP_NY_32K_100_13.4375-6.25-9.6875 | 6AH58195 | SHIP_NY_32K_100_13.4375-6.25-9.6875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual |  |  | left_only
6AH58195/SHIP_NY_32K_100_13.4375-6.25-9.6875 | 6AH58195 | SHIP_NY_32K_100_13.4375-6.25-9.6875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AH58195/SHIP_NY_32K_100_13.4375-6.25-9.6875 | 6AH58195 | SHIP_NY_32K_100_13.4375-6.25-9.6875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AH58195/SHIP_NY_32K_100_13.4375-6.25-9.6875 | 6AH58195 | SHIP_NY_32K_100_13.4375-6.25-9.6875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AH58195/SHIP_NY_32K_100_13.4375-6.25-9.6875 | 6AH58195 | SHIP_NY_32K_100_13.4375-6.25-9.6875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 6AJ19025 | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual | 6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 1 | both
6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 6AJ19025 | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 1 | both
6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 6AJ19025 | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 1 | both
6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 6AJ19025 | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 1 | both
6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 6AJ19025 | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6.75 | in | FDA,4 | 6.5625 | 6.9375 | TM 141 / Visual | 6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 1 | both
6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 6AJ19025 | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 12 | in | FDA,4 | 11.8125 | 12.1875 | TM 141 / Visual | 6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 1 | both
6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 6AJ19025 | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 1 | both
6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 6AJ19025 | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 1 | both
6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 6AJ19025 | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.25 | in | FDA,4 | 7.0625 | 7.4375 | TM 141 / Visual | 6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 1 | both
6AM30472/SHIP_NY_44K_100_15.5-9-12 | 6AM30472 | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual | 6AM30472/SHIP_NY_44K_100_15.5-9-12 | 3 | both
6AM30472/SHIP_NY_44K_100_15.5-9-12 | 6AM30472 | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6AM30472/SHIP_NY_44K_100_15.5-9-12 | 3 | both
6AM30472/SHIP_NY_44K_100_15.5-9-12 | 6AM30472 | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | 6AM30472/SHIP_NY_44K_100_15.5-9-12 | 3 | both
6AM30472/SHIP_NY_44K_100_15.5-9-12 | 6AM30472 | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6AM30472/SHIP_NY_44K_100_15.5-9-12 | 3 | both
6AM30472/SHIP_NY_44K_100_15.5-9-12 | 6AM30472 | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 12.125 | in | FDA,4 | 11.8125 | 12.1875 | TM 141 / Visual | 6AM30472/SHIP_NY_44K_100_15.5-9-12 | 3 | both
6AM30472/SHIP_NY_44K_100_15.5-9-12 | 6AM30472 | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.375 | in | FDA,4 | 15.3125 | 15.6875 | TM 141 / Visual | 6AM30472/SHIP_NY_44K_100_15.5-9-12 | 3 | both
6AM30472/SHIP_NY_44K_100_15.5-9-12 | 6AM30472 | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6AM30472/SHIP_NY_44K_100_15.5-9-12 | 3 | both
6AM30472/SHIP_NY_44K_100_15.5-9-12 | 6AM30472 | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6AM30472/SHIP_NY_44K_100_15.5-9-12 | 3 | both
6AM30472/SHIP_NY_44K_100_15.5-9-12 | 6AM30472 | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 8.875 | in | FDA,4 | 8.8125 | 9.1875 | TM 141 / Visual | 6AM30472/SHIP_NY_44K_100_15.5-9-12 | 3 | both
6AP03707/SHIP_NY_32K_100_13-7.3125-4 | 6AP03707 | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height |  | in | FDA,4 | 3.8125 | 4.1875 | TM 141 / Visual |  |  | left_only
6AP03707/SHIP_NY_32K_100_13-7.3125-4 | 6AP03707 | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 7.125 | 7.5 | TM 141 / Visual |  |  | left_only
6AP03707/SHIP_NY_32K_100_13-7.3125-4 | 6AP03707 | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 12.8125 | 13.1875 | TM 141 / Visual |  |  | left_only
6AP03707/SHIP_NY_32K_100_13-7.3125-4 | 6AP03707 | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual |  |  | left_only
6AP03707/SHIP_NY_32K_100_13-7.3125-4 | 6AP03707 | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 |  |  | TM 141 / Visual |  |  | left_only
6AP03707/SHIP_NY_32K_100_13-7.3125-4 | 6AP03707 | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP03707/SHIP_NY_32K_100_13-7.3125-4 | 6AP03707 | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP03707/SHIP_NY_32K_100_13-7.3125-4 | 6AP03707 | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP03707/SHIP_NY_32K_100_13-7.3125-4 | 6AP03707 | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | 6AP06528 | SHIP_NY_32K_100_13.25-7.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | 2 | both
6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | 6AP06528 | SHIP_NY_32K_100_13.25-7.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | 2 | both
6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | 6AP06528 | SHIP_NY_32K_100_13.25-7.75-5.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | 2 | both
6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | 6AP06528 | SHIP_NY_32K_100_13.25-7.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 5.9375 | in | FDA,4 | 5.5625 | 5.9375 | TM 141 / Visual | 6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | 2 | both
6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | 6AP06528 | SHIP_NY_32K_100_13.25-7.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 13.25 | in | FDA,4 | 13.0625 | 13.4375 | TM 141 / Visual | 6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | 2 | both
6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | 6AP06528 | SHIP_NY_32K_100_13.25-7.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | 2 | both
6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | 6AP06528 | SHIP_NY_32K_100_13.25-7.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | 2 | both
6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | 6AP06528 | SHIP_NY_32K_100_13.25-7.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.75 | in | FDA,4 | 7.5625 | 7.9375 | TM 141 / Visual | 6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | 2 | both
6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | 6AP06550 | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height |  | in | FDA,4 | 10.375 | 10.625 | TM 141 / Visual |  |  | left_only
6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | 6AP06550 | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 11.9375 | 12.1875 | TM 141 / Visual |  |  | left_only
6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | 6AP06550 | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 17.25 | 17.5 | TM 141 / Visual |  |  | left_only
6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | 6AP06550 | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 |  |  | TM 141 / Visual |  |  | left_only
6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | 6AP06550 | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual |  |  | left_only
6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | 6AP06550 | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | 6AP06550 | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | 6AP06550 | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | 6AP06550 | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 6AP06728 | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual | 6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 1 | both
6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 6AP06728 | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 1 | both
6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 6AP06728 | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 1 | both
6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 6AP06728 | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 1 | both
6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 6AP06728 | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6 | in | FDA,4 | 5.875 | 6.25 | TM 141 / Visual | 6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 1 | both
6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 6AP06728 | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 7 | in | FDA,4 | 6.9375 | 7.3125 | TM 141 / Visual | 6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 1 | both
6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 6AP06728 | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 1 | both
6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 6AP06728 | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 1 | both
6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 6AP06728 | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.125 | in | FDA,4 | 6.8125 | 7.1875 | TM 141 / Visual | 6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 1 | both
6AP07207/Simple Kind To Skin facial 7ct Shipper Canada | 6AP07207 | Simple Kind To Skin facial 7ct Shipper Canada | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6AP07207/Simple Kind To Skin facial 7ct Shipper Canada | 1 | both
6AP07207/Simple Kind To Skin facial 7ct Shipper Canada | 6AP07207 | Simple Kind To Skin facial 7ct Shipper Canada | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6AP07207/Simple Kind To Skin facial 7ct Shipper Canada | 1 | both
6AP07207/Simple Kind To Skin facial 7ct Shipper Canada | 6AP07207 | Simple Kind To Skin facial 7ct Shipper Canada | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6AP07207/Simple Kind To Skin facial 7ct Shipper Canada | 1 | both
6AP07207/Simple Kind To Skin facial 7ct Shipper Canada | 6AP07207 | Simple Kind To Skin facial 7ct Shipper Canada | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 4 | in | FDA,4 | 3.9375 | 4.0625 | TM 141 / Visual | 6AP07207/Simple Kind To Skin facial 7ct Shipper Canada | 1 | both
6AP07207/Simple Kind To Skin facial 7ct Shipper Canada | 6AP07207 | Simple Kind To Skin facial 7ct Shipper Canada | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 13 | in | FDA,4 | 12.9375 | 13.0625 | TM 141 / Visual | 6AP07207/Simple Kind To Skin facial 7ct Shipper Canada | 1 | both
6AP07207/Simple Kind To Skin facial 7ct Shipper Canada | 6AP07207 | Simple Kind To Skin facial 7ct Shipper Canada | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6AP07207/Simple Kind To Skin facial 7ct Shipper Canada | 1 | both
6AP07207/Simple Kind To Skin facial 7ct Shipper Canada | 6AP07207 | Simple Kind To Skin facial 7ct Shipper Canada | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6AP07207/Simple Kind To Skin facial 7ct Shipper Canada | 1 | both
6AP07207/Simple Kind To Skin facial 7ct Shipper Canada | 6AP07207 | Simple Kind To Skin facial 7ct Shipper Canada | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.375 | in | FDA,4 | 7.25 | 7.375 | TM 141 / Visual | 6AP07207/Simple Kind To Skin facial 7ct Shipper Canada | 1 | both
6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 6AP09215 | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual | 6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 1 | both
6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 6AP09215 | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 1 | both
6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 6AP09215 | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 1 | both
6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 6AP09215 | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 1 | both
6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 6AP09215 | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 5.75 | in | FDA,4 | 5.5625 | 5.9375 | TM 141 / Visual | 6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 1 | both
6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 6AP09215 | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 13.5 | in | FDA,4 | 13.25 | 13.625 | TM 141 / Visual | 6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 1 | both
6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 6AP09215 | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 1 | both
6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 6AP09215 | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 1 | both
6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 6AP09215 | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 9.25 | in | FDA,4 | 9.0625 | 9.4375 | TM 141 / Visual | 6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 1 | both
6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 6AP09384 | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual | 6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 9 | both
6AP09384/Shipper_NY_Skip lot testing | 6AP09384 | Shipper_NY_Skip lot testing | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual | 6AP09384/Shipper_NY_Skip lot testing | 1 | both
6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 6AP09384 | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 9 | both
6AP09384/Shipper_NY_Skip lot testing | 6AP09384 | Shipper_NY_Skip lot testing | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6AP09384/Shipper_NY_Skip lot testing | 1 | both
6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 6AP09384 | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 9 | both
6AP09384/Shipper_NY_Skip lot testing | 6AP09384 | Shipper_NY_Skip lot testing | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6AP09384/Shipper_NY_Skip lot testing | 1 | both
6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 6AP09384 | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 9 | both
6AP09384/Shipper_NY_Skip lot testing | 6AP09384 | Shipper_NY_Skip lot testing | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6AP09384/Shipper_NY_Skip lot testing | 1 | both
6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 6AP09384 | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 3.8125 | in | FDA,4 | 3.5625 | 3.8125 | TM 141 / Visual | 6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 9 | both
6AP09384/Shipper_NY_Skip lot testing | 6AP09384 | Shipper_NY_Skip lot testing | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 3.75 | in | FDA,4 | 3.5625 | 3.8125 | TM 141 / Visual | 6AP09384/Shipper_NY_Skip lot testing | 1 | both
6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 6AP09384 | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 8.5 | in | FDA,4 | 8.375 | 8.625 | TM 141 / Visual | 6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 9 | both
6AP09384/Shipper_NY_Skip lot testing | 6AP09384 | Shipper_NY_Skip lot testing | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 8.5 | in | FDA,4 | 8.375 | 8.625 | TM 141 / Visual | 6AP09384/Shipper_NY_Skip lot testing | 1 | both
6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 6AP09384 | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 9 | both
6AP09384/Shipper_NY_Skip lot testing | 6AP09384 | Shipper_NY_Skip lot testing | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6AP09384/Shipper_NY_Skip lot testing | 1 | both
6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 6AP09384 | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 9 | both
6AP09384/Shipper_NY_Skip lot testing | 6AP09384 | Shipper_NY_Skip lot testing | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6AP09384/Shipper_NY_Skip lot testing | 1 | both
6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 6AP09384 | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 5.75 | in | FDA,4 | 5.625 | 5.875 | TM 141 / Visual | 6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 9 | both
6AP09384/Shipper_NY_Skip lot testing | 6AP09384 | Shipper_NY_Skip lot testing | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 5.75 | in | FDA,4 | 5.625 | 5.875 | TM 141 / Visual | 6AP09384/Shipper_NY_Skip lot testing | 1 | both
6AP11325/SHIP_NY_44K_100_10.625-9.75-7.75 | 6AP11325 | SHIP_NY_44K_100_10.625-9.75-7.75 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height |  | in | FDA,4 | 7.5625 | 7.9375 | TM 141 / Visual |  |  | left_only
6AP11325/SHIP_NY_44K_100_10.625-9.75-7.75 | 6AP11325 | SHIP_NY_44K_100_10.625-9.75-7.75 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 9.5625 | 9.9375 | TM 141 / Visual |  |  | left_only
6AP11325/SHIP_NY_44K_100_10.625-9.75-7.75 | 6AP11325 | SHIP_NY_44K_100_10.625-9.75-7.75 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 10.4375 | 10.8125 | TM 141 / Visual |  |  | left_only
6AP11325/SHIP_NY_44K_100_10.625-9.75-7.75 | 6AP11325 | SHIP_NY_44K_100_10.625-9.75-7.75 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual |  |  | left_only
6AP11325/SHIP_NY_44K_100_10.625-9.75-7.75 | 6AP11325 | SHIP_NY_44K_100_10.625-9.75-7.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP11325/SHIP_NY_44K_100_10.625-9.75-7.75 | 6AP11325 | SHIP_NY_44K_100_10.625-9.75-7.75 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP11325/SHIP_NY_44K_100_10.625-9.75-7.75 | 6AP11325 | SHIP_NY_44K_100_10.625-9.75-7.75 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP11325/SHIP_NY_44K_100_10.625-9.75-7.75 | 6AP11325 | SHIP_NY_44K_100_10.625-9.75-7.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | 6AP13472 | SHIP_NY_32K_100_16.3125-12.25-7.5625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | 5 | both
6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | 6AP13472 | SHIP_NY_32K_100_16.3125-12.25-7.5625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | 5 | both
6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | 6AP13472 | SHIP_NY_32K_100_16.3125-12.25-7.5625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | 5 | both
6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | 6AP13472 | SHIP_NY_32K_100_16.3125-12.25-7.5625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7.75 | in | FDA,4 | 7.375 | 7.75 | TM 141 / Visual | 6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | 5 | both
6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | 6AP13472 | SHIP_NY_32K_100_16.3125-12.25-7.5625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 16.25 | in | FDA,4 | 16.125 | 16.5 | TM 141 / Visual | 6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | 5 | both
6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | 6AP13472 | SHIP_NY_32K_100_16.3125-12.25-7.5625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | 5 | both
6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | 6AP13472 | SHIP_NY_32K_100_16.3125-12.25-7.5625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | 5 | both
6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | 6AP13472 | SHIP_NY_32K_100_16.3125-12.25-7.5625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 12.25 | in | FDA,4 | 12.0625 | 12.4375 | TM 141 / Visual | 6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | 5 | both
6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | 6AP15984 | SHIP_NY_44K_100_15.25-10.375-7.6875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | 2 | both
6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | 6AP15984 | SHIP_NY_44K_100_15.25-10.375-7.6875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | 6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | 2 | both
6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | 6AP15984 | SHIP_NY_44K_100_15.25-10.375-7.6875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | 2 | both
6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | 6AP15984 | SHIP_NY_44K_100_15.25-10.375-7.6875 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7.75 | in | FDA,4 | 7.5 | 7.875 | TM 141 / Visual | 6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | 2 | both
6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | 6AP15984 | SHIP_NY_44K_100_15.25-10.375-7.6875 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.125 | in | FDA,4 | 15.0625 | 15.4375 | TM 141 / Visual | 6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | 2 | both
6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | 6AP15984 | SHIP_NY_44K_100_15.25-10.375-7.6875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | 2 | both
6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | 6AP15984 | SHIP_NY_44K_100_15.25-10.375-7.6875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | 2 | both
6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | 6AP15984 | SHIP_NY_44K_100_15.25-10.375-7.6875 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 10.5 | in | FDA,4 | 10.1875 | 10.5625 | TM 141 / Visual | 6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | 2 | both
6AP22325/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | 6AP22325 | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Shipper Inspection | 6AP22325/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | 1 | both
6AP22907/SHIP_NY_32K_100_8.5625-7.8125-6.4375 | 6AP22907 | SHIP_NY_32K_100_8.5625-7.8125-6.4375 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height |  | in | FDA,4 | 6.25 | 6.625 | TM 141 / Visual |  |  | left_only
6AP22907/SHIP_NY_32K_100_8.5625-7.8125-6.4375 | 6AP22907 | SHIP_NY_32K_100_8.5625-7.8125-6.4375 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 7.625 | 8 | TM 141 / Visual |  |  | left_only
6AP22907/SHIP_NY_32K_100_8.5625-7.8125-6.4375 | 6AP22907 | SHIP_NY_32K_100_8.5625-7.8125-6.4375 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 8.375 | 8.75 | TM 141 / Visual |  |  | left_only
6AP22907/SHIP_NY_32K_100_8.5625-7.8125-6.4375 | 6AP22907 | SHIP_NY_32K_100_8.5625-7.8125-6.4375 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual |  |  | left_only
6AP22907/SHIP_NY_32K_100_8.5625-7.8125-6.4375 | 6AP22907 | SHIP_NY_32K_100_8.5625-7.8125-6.4375 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 |  |  | TM 141 / Visual |  |  | left_only
6AP22907/SHIP_NY_32K_100_8.5625-7.8125-6.4375 | 6AP22907 | SHIP_NY_32K_100_8.5625-7.8125-6.4375 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP22907/SHIP_NY_32K_100_8.5625-7.8125-6.4375 | 6AP22907 | SHIP_NY_32K_100_8.5625-7.8125-6.4375 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP22907/SHIP_NY_32K_100_8.5625-7.8125-6.4375 | 6AP22907 | SHIP_NY_32K_100_8.5625-7.8125-6.4375 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP22907/SHIP_NY_32K_100_8.5625-7.8125-6.4375 | 6AP22907 | SHIP_NY_32K_100_8.5625-7.8125-6.4375 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP23025/SHIP_NY_32K_100_11.25-5.8125-4.3125 | 6AP23025 | SHIP_NY_32K_100_11.25-5.8125-4.3125 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height |  | in | FDA,4 | 4.125 | 4.5 | TM 141 / Visual |  |  | left_only
6AP23025/SHIP_NY_32K_100_11.25-5.8125-4.3125 | 6AP23025 | SHIP_NY_32K_100_11.25-5.8125-4.3125 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 5.625 | 6 | TM 141 / Visual |  |  | left_only
6AP23025/SHIP_NY_32K_100_11.25-5.8125-4.3125 | 6AP23025 | SHIP_NY_32K_100_11.25-5.8125-4.3125 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 11.0625 | 11.4375 | TM 141 / Visual |  |  | left_only
6AP23025/SHIP_NY_32K_100_11.25-5.8125-4.3125 | 6AP23025 | SHIP_NY_32K_100_11.25-5.8125-4.3125 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual |  |  | left_only
6AP23025/SHIP_NY_32K_100_11.25-5.8125-4.3125 | 6AP23025 | SHIP_NY_32K_100_11.25-5.8125-4.3125 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 |  |  | TM 141 / Visual |  |  | left_only
6AP23025/SHIP_NY_32K_100_11.25-5.8125-4.3125 | 6AP23025 | SHIP_NY_32K_100_11.25-5.8125-4.3125 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP23025/SHIP_NY_32K_100_11.25-5.8125-4.3125 | 6AP23025 | SHIP_NY_32K_100_11.25-5.8125-4.3125 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP23025/SHIP_NY_32K_100_11.25-5.8125-4.3125 | 6AP23025 | SHIP_NY_32K_100_11.25-5.8125-4.3125 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP23025/SHIP_NY_32K_100_11.25-5.8125-4.3125 | 6AP23025 | SHIP_NY_32K_100_11.25-5.8125-4.3125 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | 6AP24284 | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | 1 | both
6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | 6AP24284 | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 40 | Lbs./In. | FDA,0 | 40 |  | TM 141 / Visual | 6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | 1 | both
6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | 6AP24284 | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | 1 | both
6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | 6AP24284 | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 9.625 | in | FDA,4 | 9.4375 | 9.8125 | TM 141 / Visual | 6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | 1 | both
6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | 6AP24284 | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.375 | in | FDA,4 | 15.1875 | 15.5625 | TM 141 / Visual | 6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | 1 | both
6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | 6AP24284 | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | 1 | both
6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | 6AP24284 | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | 1 | both
6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | 6AP24284 | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 10.5625 | in | FDA,4 | 10.375 | 10.75 | TM 141 / Visual | 6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | 1 | both
6AP24372/SHIP_NY_44K_100_15.25-10.375-15.4375 | 6AP24372 | SHIP_NY_44K_100_15.25-10.375-15.4375 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 10.1875 | 10.5625 | TM 141 / Visual |  |  | left_only
6AP24372/SHIP_NY_44K_100_15.25-10.375-15.4375 | 6AP24372 | SHIP_NY_44K_100_15.25-10.375-15.4375 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 15.0625 | 15.4375 | TM 141 / Visual |  |  | left_only
6AP24372/SHIP_NY_44K_100_15.25-10.375-15.4375 | 6AP24372 | SHIP_NY_44K_100_15.25-10.375-15.4375 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height |  | in | FDA,4 | 15.25 | 15.625 | TM 141 / Visual |  |  | left_only
6AP24372/SHIP_NY_44K_100_15.25-10.375-15.4375 | 6AP24372 | SHIP_NY_44K_100_15.25-10.375-15.4375 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual |  |  | left_only
6AP24372/SHIP_NY_44K_100_15.25-10.375-15.4375 | 6AP24372 | SHIP_NY_44K_100_15.25-10.375-15.4375 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP24372/SHIP_NY_44K_100_15.25-10.375-15.4375 | 6AP24372 | SHIP_NY_44K_100_15.25-10.375-15.4375 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP24372/SHIP_NY_44K_100_15.25-10.375-15.4375 | 6AP24372 | SHIP_NY_44K_100_15.25-10.375-15.4375 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP24372/SHIP_NY_44K_100_15.25-10.375-15.4375 | 6AP24372 | SHIP_NY_44K_100_15.25-10.375-15.4375 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP24684/SHIP_NY_40K_100_15.375-10.5625-9.625 | 6AP24684 | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual | 6AP24684/SHIP_NY_40K_100_15.375-10.5625-9.625 | 1 | both
6AP27284/SHIP_NY_44K_100_15.375-10.5625-9.625 | 6AP27284 | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height |  | in | FDA,4 | 9.4375 | 9.8125 | TM 141 / Visual |  |  | left_only
6AP27284/SHIP_NY_44K_100_15.375-10.5625-9.625 | 6AP27284 | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 10.375 | 10.75 | TM 141 / Visual |  |  | left_only
6AP27284/SHIP_NY_44K_100_15.375-10.5625-9.625 | 6AP27284 | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 15.1875 | 15.5625 | TM 141 / Visual |  |  | left_only
6AP27284/SHIP_NY_44K_100_15.375-10.5625-9.625 | 6AP27284 | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual |  |  | left_only
6AP27284/SHIP_NY_44K_100_15.375-10.5625-9.625 | 6AP27284 | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP27284/SHIP_NY_44K_100_15.375-10.5625-9.625 | 6AP27284 | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP27284/SHIP_NY_44K_100_15.375-10.5625-9.625 | 6AP27284 | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP27284/SHIP_NY_44K_100_15.375-10.5625-9.625 | 6AP27284 | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP30895/SHIP_32_1BI | 6AP30895 | SHIP_32_1BI | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 7.0625 | 7.4375 | TM 141 / Visual |  |  | left_only
6AP30895/SHIP_32_1BI | 6AP30895 | SHIP_32_1BI | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height |  | in | FDA,4 | 7.6875 | 8.0625 | TM 141 / Visual |  |  | left_only
6AP30895/SHIP_32_1BI | 6AP30895 | SHIP_32_1BI | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 12.3125 | 12.6875 | TM 141 / Visual |  |  | left_only
6AP30895/SHIP_32_1BI | 6AP30895 | SHIP_32_1BI | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual |  |  | left_only
6AP30895/SHIP_32_1BI | 6AP30895 | SHIP_32_1BI | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP30895/SHIP_32_1BI | 6AP30895 | SHIP_32_1BI | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP30895/SHIP_32_1BI | 6AP30895 | SHIP_32_1BI | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP30895/SHIP_32_1BI | 6AP30895 | SHIP_32_1BI | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP32825/SHIP_32_1P | 6AP32825 | SHIP_32_1P | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6AP32825/SHIP_32_1P | 1 | both
6AP32825/SHIP_32_1P | 6AP32825 | SHIP_32_1P | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6AP32825/SHIP_32_1P | 1 | both
6AP32825/SHIP_32_1P | 6AP32825 | SHIP_32_1P | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6AP32825/SHIP_32_1P | 1 | both
6AP32825/SHIP_32_1P | 6AP32825 | SHIP_32_1P | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 4.25 | in | FDA,4 | 4.125 | 4.5 | TM 141 / Visual | 6AP32825/SHIP_32_1P | 1 | both
6AP32825/SHIP_32_1P | 6AP32825 | SHIP_32_1P | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.1875 | in | FDA,4 | 11.0625 | 11.4375 | TM 141 / Visual | 6AP32825/SHIP_32_1P | 1 | both
6AP32825/SHIP_32_1P | 6AP32825 | SHIP_32_1P | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6AP32825/SHIP_32_1P | 1 | both
6AP32825/SHIP_32_1P | 6AP32825 | SHIP_32_1P | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6AP32825/SHIP_32_1P | 1 | both
6AP32825/SHIP_32_1P | 6AP32825 | SHIP_32_1P | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6 | in | FDA,4 | 5.625 | 6 | TM 141 / Visual | 6AP32825/SHIP_32_1P | 1 | both
6AP71310/SHIP_NY_44K_100_18.875-11.625-9 | 6AP71310 | SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height |  | in | FDA,4 | 8.8125 | 9.1875 | TM 141 / Visual |  |  | left_only
6AP71310/SHIP_NY_44K_100_18.875-11.625-9 | 6AP71310 | SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 11.4375 | 11.8125 | TM 141 / Visual |  |  | left_only
6AP71310/SHIP_NY_44K_100_18.875-11.625-9 | 6AP71310 | SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 18.6875 | 19.0625 | TM 141 / Visual |  |  | left_only
6AP71310/SHIP_NY_44K_100_18.875-11.625-9 | 6AP71310 | SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual |  |  | left_only
6AP71310/SHIP_NY_44K_100_18.875-11.625-9 | 6AP71310 | SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 |  |  | TM 141 / Visual |  |  | left_only
6AP71310/SHIP_NY_44K_100_18.875-11.625-9 | 6AP71310 | SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP71310/SHIP_NY_44K_100_18.875-11.625-9 | 6AP71310 | SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP71310/SHIP_NY_44K_100_18.875-11.625-9 | 6AP71310 | SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP71310/SHIP_NY_44K_100_18.875-11.625-9 | 6AP71310 | SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP80625/SHIP_NY_32K_100_11.25-5.8125-8.625 | 6AP80625 | SHIP_NY_32K_100_11.25-5.8125-8.625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 5.625 | 6 | TM 141 / Visual |  |  | left_only
6AP80625/SHIP_NY_32K_100_11.25-5.8125-8.625 | 6AP80625 | SHIP_NY_32K_100_11.25-5.8125-8.625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height |  | in | FDA,4 | 8.4375 | 8.8125 | TM 141 / Visual |  |  | left_only
6AP80625/SHIP_NY_32K_100_11.25-5.8125-8.625 | 6AP80625 | SHIP_NY_32K_100_11.25-5.8125-8.625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 11.0625 | 11.4375 | TM 141 / Visual |  |  | left_only
6AP80625/SHIP_NY_32K_100_11.25-5.8125-8.625 | 6AP80625 | SHIP_NY_32K_100_11.25-5.8125-8.625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual |  |  | left_only
6AP80625/SHIP_NY_32K_100_11.25-5.8125-8.625 | 6AP80625 | SHIP_NY_32K_100_11.25-5.8125-8.625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 |  |  | TM 141 / Visual |  |  | left_only
6AP80625/SHIP_NY_32K_100_11.25-5.8125-8.625 | 6AP80625 | SHIP_NY_32K_100_11.25-5.8125-8.625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP80625/SHIP_NY_32K_100_11.25-5.8125-8.625 | 6AP80625 | SHIP_NY_32K_100_11.25-5.8125-8.625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP80625/SHIP_NY_32K_100_11.25-5.8125-8.625 | 6AP80625 | SHIP_NY_32K_100_11.25-5.8125-8.625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AP80625/SHIP_NY_32K_100_11.25-5.8125-8.625 | 6AP80625 | SHIP_NY_32K_100_11.25-5.8125-8.625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AS27350/SHIP_44_2O | 6AS27350 | SHIP_44_2O | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6AS27350/SHIP_44_2O | 8 | both
6AS27350/SHIP_44_2O | 6AS27350 | SHIP_44_2O | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | 6AS27350/SHIP_44_2O | 8 | both
6AS27350/SHIP_44_2O | 6AS27350 | SHIP_44_2O | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6AS27350/SHIP_44_2O | 8 | both
6AS27350/SHIP_44_2O | 6AS27350 | SHIP_44_2O | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 9 | in | FDA,4 | 8.8125 | 9.1875 | TM 141 / Visual | 6AS27350/SHIP_44_2O | 8 | both
6AS27350/SHIP_44_2O | 6AS27350 | SHIP_44_2O | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.375 | in | FDA,4 | 15.1875 | 15.5625 | TM 141 / Visual | 6AS27350/SHIP_44_2O | 8 | both
6AS27350/SHIP_44_2O | 6AS27350 | SHIP_44_2O | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6AS27350/SHIP_44_2O | 8 | both
6AS27350/SHIP_44_2O | 6AS27350 | SHIP_44_2O | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6AS27350/SHIP_44_2O | 8 | both
6AS27350/SHIP_44_2O | 6AS27350 | SHIP_44_2O | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 11.375 | in | FDA,4 | 11.1875 | 11.5625 | TM 141 / Visual | 6AS27350/SHIP_44_2O | 8 | both
6AS40750/SHIP_32_1AP | 6AS40750 | SHIP_32_1AP | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6AS40750/SHIP_32_1AP | 3 | both
6AS40750/SHIP_32_1AP | 6AS40750 | SHIP_32_1AP | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6AS40750/SHIP_32_1AP | 3 | both
6AS40750/SHIP_32_1AP | 6AS40750 | SHIP_32_1AP | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6AS40750/SHIP_32_1AP | 3 | both
6AS40750/SHIP_32_1AP | 6AS40750 | SHIP_32_1AP | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 13.4375 | in | FDA,4 | 13.25 | 13.625 | TM 141 / Visual | 6AS40750/SHIP_32_1AP | 3 | both
6AS40750/SHIP_32_1AP | 6AS40750 | SHIP_32_1AP | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.625 | in | FDA,4 | 15.4375 | 15.8125 | TM 141 / Visual | 6AS40750/SHIP_32_1AP | 3 | both
6AS40750/SHIP_32_1AP | 6AS40750 | SHIP_32_1AP | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6AS40750/SHIP_32_1AP | 3 | both
6AS40750/SHIP_32_1AP | 6AS40750 | SHIP_32_1AP | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6AS40750/SHIP_32_1AP | 3 | both
6AS40750/SHIP_32_1AP | 6AS40750 | SHIP_32_1AP | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.25 | in | FDA,4 | 7.0625 | 7.4375 | TM 141 / Visual | 6AS40750/SHIP_32_1AP | 3 | both
6AS41950/SHIP_40_1O | 6AS41950 | SHIP_40_1O | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6AS41950/SHIP_40_1O | 2 | both
6AS41950/SHIP_40_1O | 6AS41950 | SHIP_40_1O | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 40 | Lbs./In. | FDA,0 | 40 |  | TM 141 / Visual | 6AS41950/SHIP_40_1O | 2 | both
6AS41950/SHIP_40_1O | 6AS41950 | SHIP_40_1O | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6AS41950/SHIP_40_1O | 2 | both
6AS41950/SHIP_40_1O | 6AS41950 | SHIP_40_1O | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 8.3125 | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual | 6AS41950/SHIP_40_1O | 2 | both
6AS41950/SHIP_40_1O | 6AS41950 | SHIP_40_1O | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 21.125 | in | FDA,4 | 20.9375 | 21.3125 | TM 141 / Visual | 6AS41950/SHIP_40_1O | 2 | both
6AS41950/SHIP_40_1O | 6AS41950 | SHIP_40_1O | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6AS41950/SHIP_40_1O | 2 | both
6AS41950/SHIP_40_1O | 6AS41950 | SHIP_40_1O | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6AS41950/SHIP_40_1O | 2 | both
6AS41950/SHIP_40_1O | 6AS41950 | SHIP_40_1O | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 13.5625 | in | FDA,4 | 13.375 | 13.75 | TM 141 / Visual | 6AS41950/SHIP_40_1O | 2 | both
6AU12310/SHIP_32_1U | 6AU12310 | SHIP_32_1U | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6AU12310/SHIP_32_1U | 6 | both
6AU12310/SHIP_32_1U | 6AU12310 | SHIP_32_1U | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6AU12310/SHIP_32_1U | 6 | both
6AU12310/SHIP_32_1U | 6AU12310 | SHIP_32_1U | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6AU12310/SHIP_32_1U | 6 | both
6AU12310/SHIP_32_1U | 6AU12310 | SHIP_32_1U | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 8 | in | FDA,4 | 7.8125 | 8.1875 | TM 141 / Visual | 6AU12310/SHIP_32_1U | 6 | both
6AU12310/SHIP_32_1U | 6AU12310 | SHIP_32_1U | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 12.125 | in | FDA,4 | 11.8125 | 12.1875 | TM 141 / Visual | 6AU12310/SHIP_32_1U | 6 | both
6AU12310/SHIP_32_1U | 6AU12310 | SHIP_32_1U | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6AU12310/SHIP_32_1U | 6 | both
6AU12310/SHIP_32_1U | 6AU12310 | SHIP_32_1U | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6AU12310/SHIP_32_1U | 6 | both
6AU12310/SHIP_32_1U | 6AU12310 | SHIP_32_1U | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6 | in | FDA,4 | 5.9375 | 6.3125 | TM 141 / Visual | 6AU12310/SHIP_32_1U | 6 | both
6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | 6AU130TY | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height |  | in | FDA,4 | 4.375 | 4.625 | TM 141 / Visual |  |  | left_only
6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | 6AU130TY | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 6.375 | 6.625 | TM 141 / Visual |  |  | left_only
6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | 6AU130TY | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 9.5 | 9.75 | TM 141 / Visual |  |  | left_only
6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | 6AU130TY | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual |  |  | left_only
6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | 6AU130TY | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 |  |  | TM 141 / Visual |  |  | left_only
6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | 6AU130TY | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | 6AU130TY | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | Fold Over Tray w/header (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | 6AU130TY | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | 6AU130TY | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | White (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | 6AU25206 | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height |  | in | FDA,4 | 4.375 | 4.75 | TM 141 / Visual |  |  | left_only
6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | 6AU25206 | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 8.25 | 8.625 | TM 141 / Visual |  |  | left_only
6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | 6AU25206 | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 11.0625 | 11.4375 | TM 141 / Visual |  |  | left_only
6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | 6AU25206 | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual |  |  | left_only
6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | 6AU25206 | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 |  |  | TM 141 / Visual |  |  | left_only
6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | 6AU25206 | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | 6AU25206 | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | 6AU25206 | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | No Go - Does not meet Specifications (OOS) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | 6AU25206 | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 6AU53601 | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 50 |  | FDA,0 |  |  | TM 141 / Visual | 6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 1 | both
6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 6AU53601 | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 1 | both
6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 6AU53601 | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 1 | both
6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 6AU53601 | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 1 | both
6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 6AU53601 | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 8.5 | in | FDA,4 | 8.3125 | 8.6875 | TM 141 / Visual | 6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 1 | both
6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 6AU53601 | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.875 | in | FDA,4 | 11.625 | 12 | TM 141 / Visual | 6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 1 | both
6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 6AU53601 | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 1 | both
6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 6AU53601 | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 1 | both
6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 6AU53601 | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 9.5 | in | FDA,4 | 9.1875 | 9.5625 | TM 141 / Visual | 6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 1 | both
6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 6AU56400 | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 50 |  | FDA,0 |  |  | TM 141 / Visual | 6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 2 | both
6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 6AU56400 | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 2 | both
6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 6AU56400 | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 2 | both
6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 6AU56400 | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 2 | both
6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 6AU56400 | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7.8125 | in | FDA,4 | 7.5005 | 7.8755 | TM 141 / Visual | 6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 2 | both
6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 6AU56400 | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.75 | in | FDA,4 | 11.5625 | 11.9375 | TM 141 / Visual | 6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 2 | both
6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 6AU56400 | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 2 | both
6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 6AU56400 | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 2 | both
6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 6AU56400 | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 9.625 | in | FDA,4 | 9.4375 | 9.8125 | TM 141 / Visual | 6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 2 | both
6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | 6AU57314 | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 6.0625 | 6.4375 | TM 141 / Visual |  |  | left_only
6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | 6AU57314 | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 7.5625 | 7.9375 | TM 141 / Visual |  |  | left_only
6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | 6AU57314 | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height |  | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual |  |  | left_only
6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | 6AU57314 | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual |  |  | left_only
6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | 6AU57314 | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 |  |  | TM 141 / Visual |  |  | left_only
6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | 6AU57314 | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | 6AU57314 | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | 6AU57314 | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | 6AU57314 | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 6AU71230 | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 50 |  | FDA,0 |  |  | TM 141 / Visual | 6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 1 | both
6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 6AU71230 | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 1 | both
6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 6AU71230 | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 1 | both
6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 6AU71230 | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 1 | both
6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 6AU71230 | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7.75 | in | FDA,4 | 7.5625 | 7.9375 | TM 141 / Visual | 6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 1 | both
6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 6AU71230 | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 9.5625 | in | FDA,4 | 9.3755 | 9.7505 | TM 141 / Visual | 6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 1 | both
6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 6AU71230 | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 1 | both
6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 6AU71230 | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 1 | both
6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 6AU71230 | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 9.188 | in | FDA,4 | 9.0005 | 9.3755 | TM 141 / Visual | 6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 1 | both
6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 6BA709SWN | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual | 6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 3 | both
6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 6BA709SWN | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 3 | both
6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 6BA709SWN | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 40 | Lbs./In. | FDA,0 | 40 |  | TM 141 / Visual | 6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 3 | both
6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 6BA709SWN | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 3 | both
6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 6BA709SWN | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6.125 | in | FDA,4 | 5.9375 | 6.3125 | TM 141 / Visual | 6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 3 | both
6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 6BA709SWN | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 14.75 | in | FDA,4 | 14.5625 | 14.9375 | TM 141 / Visual | 6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 3 | both
6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 6BA709SWN | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 3 | both
6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 6BA709SWN | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 3 | both
6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 6BA709SWN | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 9.375 | in | FDA,4 | 9.1875 | 9.5625 | TM 141 / Visual | 6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 3 | both
6BB10800/SHIP_32_1S | 6BB10800 | SHIP_32_1S | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6BB10800/SHIP_32_1S | 6 | both
6BB10800/SHIP_32_1S | 6BB10800 | SHIP_32_1S | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6BB10800/SHIP_32_1S | 6 | both
6BB10800/SHIP_32_1S | 6BB10800 | SHIP_32_1S | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6BB10800/SHIP_32_1S | 6 | both
6BB10800/SHIP_32_1S | 6BB10800 | SHIP_32_1S | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 5.3125 | in | FDA,4 | 5.125 | 5.5 | TM 141 / Visual | 6BB10800/SHIP_32_1S | 6 | both
6BB10800/SHIP_32_1S | 6BB10800 | SHIP_32_1S | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.875 | in | FDA,4 | 11.6875 | 12.0625 | TM 141 / Visual | 6BB10800/SHIP_32_1S | 6 | both
6BB10800/SHIP_32_1S | 6BB10800 | SHIP_32_1S | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6BB10800/SHIP_32_1S | 6 | both
6BB10800/SHIP_32_1S | 6BB10800 | SHIP_32_1S | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6BB10800/SHIP_32_1S | 6 | both
6BB10800/SHIP_32_1S | 6BB10800 | SHIP_32_1S | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.25 | in | FDA,4 | 6.0625 | 6.4375 | TM 141 / Visual | 6BB10800/SHIP_32_1S | 6 | both
6BD17385/SHIP_44_2AB | 6BD17385 | SHIP_44_2AB | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6BD17385/SHIP_44_2AB | 3 | both
6BD17385/SHIP_44_2AB | 6BD17385 | SHIP_44_2AB | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | 6BD17385/SHIP_44_2AB | 3 | both
6BD17385/SHIP_44_2AB | 6BD17385 | SHIP_44_2AB | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6BD17385/SHIP_44_2AB | 3 | both
6BD17385/SHIP_44_2AB | 6BD17385 | SHIP_44_2AB | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6.5 | in | FDA,4 | 6.3125 | 6.6875 | TM 141 / Visual | 6BD17385/SHIP_44_2AB | 3 | both
6BD17385/SHIP_44_2AB | 6BD17385 | SHIP_44_2AB | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 20 | in | FDA,4 | 19.8125 | 20.1875 | TM 141 / Visual | 6BD17385/SHIP_44_2AB | 3 | both
6BD17385/SHIP_44_2AB | 6BD17385 | SHIP_44_2AB | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6BD17385/SHIP_44_2AB | 3 | both
6BD17385/SHIP_44_2AB | 6BD17385 | SHIP_44_2AB | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6BD17385/SHIP_44_2AB | 3 | both
6BD17385/SHIP_44_2AB | 6BD17385 | SHIP_44_2AB | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 12 | in | FDA,4 | 11.8125 | 12.1875 | TM 141 / Visual | 6BD17385/SHIP_44_2AB | 3 | both
6BP21195/SHIP_NY_44K_100_14.375-12.1875-9.25 | 6BP21195 | SHIP_NY_44K_100_14.375-12.1875-9.25 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height |  | in | FDA,4 | 9.0625 | 9.4375 | TM 141 / Visual |  |  | left_only
6BP21195/SHIP_NY_44K_100_14.375-12.1875-9.25 | 6BP21195 | SHIP_NY_44K_100_14.375-12.1875-9.25 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 12 | 12.375 | TM 141 / Visual |  |  | left_only
6BP21195/SHIP_NY_44K_100_14.375-12.1875-9.25 | 6BP21195 | SHIP_NY_44K_100_14.375-12.1875-9.25 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 14.1875 | 14.5625 | TM 141 / Visual |  |  | left_only
6BP21195/SHIP_NY_44K_100_14.375-12.1875-9.25 | 6BP21195 | SHIP_NY_44K_100_14.375-12.1875-9.25 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual |  |  | left_only
6BP21195/SHIP_NY_44K_100_14.375-12.1875-9.25 | 6BP21195 | SHIP_NY_44K_100_14.375-12.1875-9.25 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6BP21195/SHIP_NY_44K_100_14.375-12.1875-9.25 | 6BP21195 | SHIP_NY_44K_100_14.375-12.1875-9.25 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6BP21195/SHIP_NY_44K_100_14.375-12.1875-9.25 | 6BP21195 | SHIP_NY_44K_100_14.375-12.1875-9.25 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6BP21195/SHIP_NY_44K_100_14.375-12.1875-9.25 | 6BP21195 | SHIP_NY_44K_100_14.375-12.1875-9.25 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 6BP25372 | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 3 | both
6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 6BP25372 | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | TM 141 / Visual | 6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 3 | both
6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 6BP25372 | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | 6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 3 | both
6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 6BP25372 | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 3 | both
6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 6BP25372 | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 15.25 | in | FDA,4 | 15 | 15.375 | TM 141 / Visual | 6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 3 | both
6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 6BP25372 | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.125 | in | FDA,4 | 15 | 15.375 | TM 141 / Visual | 6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 3 | both
6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 6BP25372 | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 3 | both
6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 6BP25372 | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 3 | both
6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 6BP25372 | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 10.25 | in | FDA,4 | 10 | 10.375 | TM 141 / Visual | 6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 3 | both
6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | 6BP30895 | SHIP_NY_32K_100_12.5-7.25-7.875 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual | 6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | 2 | both
6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | 6BP30895 | SHIP_NY_32K_100_12.5-7.25-7.875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | 2 | both
6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | 6BP30895 | SHIP_NY_32K_100_12.5-7.25-7.875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | 2 | both
6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | 6BP30895 | SHIP_NY_32K_100_12.5-7.25-7.875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | 2 | both
6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | 6BP30895 | SHIP_NY_32K_100_12.5-7.25-7.875 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7.6875 | in | FDA,4 | 7.6875 | 8.0625 | TM 141 / Visual | 6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | 2 | both
6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | 6BP30895 | SHIP_NY_32K_100_12.5-7.25-7.875 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 12.375 | in | FDA,4 | 12.3125 | 12.6875 | TM 141 / Visual | 6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | 2 | both
6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | 6BP30895 | SHIP_NY_32K_100_12.5-7.25-7.875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | 2 | both
6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | 6BP30895 | SHIP_NY_32K_100_12.5-7.25-7.875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | 2 | both
6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | 6BP30895 | SHIP_NY_32K_100_12.5-7.25-7.875 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.125 | in | FDA,4 | 7.0625 | 7.4375 | TM 141 / Visual | 6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | 2 | both
6BP54072/SHIP_44_2L | 6BP54072 | SHIP_44_2L | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6BP54072/SHIP_44_2L | 1 | both
6BP54072/SHIP_44_2L | 6BP54072 | SHIP_44_2L | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | 6BP54072/SHIP_44_2L | 1 | both
6BP54072/SHIP_44_2L | 6BP54072 | SHIP_44_2L | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6BP54072/SHIP_44_2L | 1 | both
6BP54072/SHIP_44_2L | 6BP54072 | SHIP_44_2L | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 15.4375 | in | FDA,4 | 15.25 | 15.625 | TM 141 / Visual | 6BP54072/SHIP_44_2L | 1 | both
6BP54072/SHIP_44_2L | 6BP54072 | SHIP_44_2L | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.25 | in | FDA,4 | 15.0625 | 15.4375 | TM 141 / Visual | 6BP54072/SHIP_44_2L | 1 | both
6BP54072/SHIP_44_2L | 6BP54072 | SHIP_44_2L | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6BP54072/SHIP_44_2L | 1 | both
6BP54072/SHIP_44_2L | 6BP54072 | SHIP_44_2L | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6BP54072/SHIP_44_2L | 1 | both
6BP54072/SHIP_44_2L | 6BP54072 | SHIP_44_2L | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 10.375 | in | FDA,4 | 10.1875 | 10.5625 | TM 141 / Visual | 6BP54072/SHIP_44_2L | 1 | both
6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | 6BU13195 | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 4.3125 | 4.6875 | TM 141 / Visual |  |  | left_only
6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | 6BU13195 | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height |  | in | FDA,4 | 12.0625 | 12.4375 | TM 141 / Visual |  |  | left_only
6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | 6BU13195 | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 13.4375 | 13.8125 | TM 141 / Visual |  |  | left_only
6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | 6BU13195 | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual |  |  | left_only
6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | 6BU13195 | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 |  |  | TM 141 / Visual |  |  | left_only
6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | 6BU13195 | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | 6BU13195 | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | 6BU13195 | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | 6BU13195 | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | 6BU25206 | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height |  | in | FDA,4 | 4.4375 | 4.8125 | TM 141 / Visual |  |  | left_only
6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | 6BU25206 | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 8.25 | 8.625 | TM 141 / Visual |  |  | left_only
6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | 6BU25206 | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 |  |  | TM 141 / Visual |  |  | left_only
6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | 6BU25206 | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 11.125 | 11.5 | TM 141 / Visual |  |  | left_only
6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | 6BU25206 | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual |  |  | left_only
6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | 6BU25206 | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | 6BU25206 | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | 6BU25206 | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | 6BU25206 | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | 6CA972CK | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual |  |  | left_only
6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | 6CA972CK | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height |  | in | FDA,4 | 10.0625 | 10.4375 | TM 141 / Visual |  |  | left_only
6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | 6CA972CK | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 17.8125 | 18.1875 | TM 141 / Visual |  |  | left_only
6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | 6CA972CK | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 |  |  | TM 141 / Visual |  |  | left_only
6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | 6CA972CK | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual |  |  | left_only
6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | 6CA972CK | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | 6CA972CK | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | 6CA972CK | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | 6CA972CK | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | 6CJ21396 | SHIP_NY_44K_100_16.625-9.375-7 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | 1 | both
6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | 6CJ21396 | SHIP_NY_44K_100_16.625-9.375-7 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | 6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | 1 | both
6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | 6CJ21396 | SHIP_NY_44K_100_16.625-9.375-7 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | 1 | both
6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | 6CJ21396 | SHIP_NY_44K_100_16.625-9.375-7 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7.1875 | in | FDA,4 | 6.8125 | 7.1875 | TM 141 / Visual | 6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | 1 | both
6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | 6CJ21396 | SHIP_NY_44K_100_16.625-9.375-7 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 16.75 | in | FDA,4 | 16.4375 | 16.8125 | TM 141 / Visual | 6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | 1 | both
6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | 6CJ21396 | SHIP_NY_44K_100_16.625-9.375-7 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | 1 | both
6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | 6CJ21396 | SHIP_NY_44K_100_16.625-9.375-7 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | 1 | both
6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | 6CJ21396 | SHIP_NY_44K_100_16.625-9.375-7 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 9.5 | in | FDA,4 | 9.1875 | 9.5625 | TM 141 / Visual | 6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | 1 | both
6CP02724/SHIP_NY_32K_100_11-5.8125-4.0625 | 6CP02724 | SHIP_NY_32K_100_11-5.8125-4.0625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6CP02724/SHIP_NY_32K_100_11-5.8125-4.0625 | 2 | both
6CP02724/SHIP_NY_32K_100_11-5.8125-4.0625 | 6CP02724 | SHIP_NY_32K_100_11-5.8125-4.0625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6CP02724/SHIP_NY_32K_100_11-5.8125-4.0625 | 2 | both
6CP02724/SHIP_NY_32K_100_11-5.8125-4.0625 | 6CP02724 | SHIP_NY_32K_100_11-5.8125-4.0625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6CP02724/SHIP_NY_32K_100_11-5.8125-4.0625 | 2 | both
6CP02724/SHIP_NY_32K_100_11-5.8125-4.0625 | 6CP02724 | SHIP_NY_32K_100_11-5.8125-4.0625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 4.25 | in | FDA,4 | 3.875 | 4.25 | TM 141 / Visual | 6CP02724/SHIP_NY_32K_100_11-5.8125-4.0625 | 2 | both
6CP02724/SHIP_NY_32K_100_11-5.8125-4.0625 | 6CP02724 | SHIP_NY_32K_100_11-5.8125-4.0625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11 | in | FDA,4 | 10.8125 | 11.1875 | TM 141 / Visual | 6CP02724/SHIP_NY_32K_100_11-5.8125-4.0625 | 2 | both
6CP02724/SHIP_NY_32K_100_11-5.8125-4.0625 | 6CP02724 | SHIP_NY_32K_100_11-5.8125-4.0625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6CP02724/SHIP_NY_32K_100_11-5.8125-4.0625 | 2 | both
6CP02724/SHIP_NY_32K_100_11-5.8125-4.0625 | 6CP02724 | SHIP_NY_32K_100_11-5.8125-4.0625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6CP02724/SHIP_NY_32K_100_11-5.8125-4.0625 | 2 | both
6CP02724/SHIP_NY_32K_100_11-5.8125-4.0625 | 6CP02724 | SHIP_NY_32K_100_11-5.8125-4.0625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6 | in | FDA,4 | 5.625 | 6 | TM 141 / Visual | 6CP02724/SHIP_NY_32K_100_11-5.8125-4.0625 | 2 | both
6CP22884/SHIP_NY_44K_100_15.375-10.5625-9.625 | 6CP22884 | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Shipper Inspection | 6CP22884/SHIP_NY_44K_100_15.375-10.5625-9.625 | 2 | both
6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 6DC22370 | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual | 6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 3 | both
6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 6DC22370 | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 3 | both
6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 6DC22370 | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 3 | both
6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 6DC22370 | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 3 | both
6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 6DC22370 | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 10.375 | in | FDA,4 | 10.1875 | 10.5625 | TM 141 / Visual | 6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 3 | both
6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 6DC22370 | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11 | in | FDA,4 | 10.9375 | 11.3125 | TM 141 / Visual | 6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 3 | both
6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 6DC22370 | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 3 | both
6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 6DC22370 | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 3 | both
6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 6DC22370 | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.25 | in | FDA,4 | 6.1875 | 6.5625 | TM 141 / Visual | 6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 3 | both
6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | 6DD33333 | SHIP_NY_44K_100_21-12.5-10.5 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | 1 | both
6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | 6DD33333 | SHIP_NY_44K_100_21-12.5-10.5 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | 6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | 1 | both
6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | 6DD33333 | SHIP_NY_44K_100_21-12.5-10.5 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | 1 | both
6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | 6DD33333 | SHIP_NY_44K_100_21-12.5-10.5 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 10.5 | in | FDA,4 | 10.3125 | 10.6875 | TM 141 / Visual | 6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | 1 | both
6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | 6DD33333 | SHIP_NY_44K_100_21-12.5-10.5 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 21 | in | FDA,4 | 20.8125 | 21.1875 | TM 141 / Visual | 6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | 1 | both
6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | 6DD33333 | SHIP_NY_44K_100_21-12.5-10.5 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | 1 | both
6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | 6DD33333 | SHIP_NY_44K_100_21-12.5-10.5 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | 1 | both
6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | 6DD33333 | SHIP_NY_44K_100_21-12.5-10.5 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 12.625 | in | FDA,4 | 12.3125 | 12.6875 | TM 141 / Visual | 6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | 1 | both
6DP63884/SHIP_NY_44K_100_15.375-10.5625-9.625 | 6DP63884 | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual | 6DP63884/SHIP_NY_44K_100_15.375-10.5625-9.625 | 1 | both
6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | 6EP44584 | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height |  | in | FDA,4 | 9.4375 | 9.8125 | TM 141 / Visual |  |  | left_only
6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | 6EP44584 | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 10 | 10.375 | TM 141 / Visual |  |  | left_only
6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | 6EP44584 | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 15 | 15.375 | TM 141 / Visual |  |  | left_only
6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | 6EP44584 | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual |  |  | left_only
6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | 6EP44584 | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 |  |  | TM 141 / Visual |  |  | left_only
6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | 6EP44584 | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | 6EP44584 | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | 6EP44584 | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | 6EP44584 | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 6FB71200 | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 50 |  | FDA,0 |  |  | TM 141 / Visual | 6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 1 | both
6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 6FB71200 | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 1 | both
6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 6FB71200 | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 1 | both
6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 6FB71200 | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 1 | both
6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 6FB71200 | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 5.125 | in | FDA,4 | 4.8125 | 5.1875 | TM 141 / Visual | 6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 1 | both
6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 6FB71200 | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 14.75 | in | FDA,4 | 14.4375 | 14.8125 | TM 141 / Visual | 6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 1 | both
6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 6FB71200 | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 1 | both
6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 6FB71200 | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC w/ shyed flaps (Done) |  |  |  |  | TM 141 / Visual | 6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 1 | both
6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 6FB71200 | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.344 | in | FDA,4 | 6.1875 | 6.5625 | TM 141 / Visual | 6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 1 | both
6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | 6FJ14108 | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height |  | in | FDA,4 | 6.8125 | 7.1875 | TM 141 / Visual |  |  | left_only
6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | 6FJ14108 | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 8.4375 | 8.8125 | TM 141 / Visual |  |  | left_only
6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | 6FJ14108 | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 16.4375 | 16.8125 | TM 141 / Visual |  |  | left_only
6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | 6FJ14108 | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual |  |  | left_only
6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | 6FJ14108 | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 |  |  | TM 141 / Visual |  |  | left_only
6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | 6FJ14108 | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | 6FJ14108 | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | 6FJ14108 | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | 6FJ14108 | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6FJ14143/SHIP_NY_44K_100_16.5-10.875-5.5 | 6FJ14143 | SHIP_NY_44K_100_16.5-10.875-5.5 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height |  | in | FDA,4 | 5.3125 | 5.6875 | TM 141 / Visual |  |  | left_only
6FJ14143/SHIP_NY_44K_100_16.5-10.875-5.5 | 6FJ14143 | SHIP_NY_44K_100_16.5-10.875-5.5 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 10.6875 | 11.0625 | TM 141 / Visual |  |  | left_only
6FJ14143/SHIP_NY_44K_100_16.5-10.875-5.5 | 6FJ14143 | SHIP_NY_44K_100_16.5-10.875-5.5 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 16.3125 | 16.6875 | TM 141 / Visual |  |  | left_only
6FJ14143/SHIP_NY_44K_100_16.5-10.875-5.5 | 6FJ14143 | SHIP_NY_44K_100_16.5-10.875-5.5 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual |  |  | left_only
6FJ14143/SHIP_NY_44K_100_16.5-10.875-5.5 | 6FJ14143 | SHIP_NY_44K_100_16.5-10.875-5.5 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6FJ14143/SHIP_NY_44K_100_16.5-10.875-5.5 | 6FJ14143 | SHIP_NY_44K_100_16.5-10.875-5.5 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6FJ14143/SHIP_NY_44K_100_16.5-10.875-5.5 | 6FJ14143 | SHIP_NY_44K_100_16.5-10.875-5.5 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6FJ14143/SHIP_NY_44K_100_16.5-10.875-5.5 | 6FJ14143 | SHIP_NY_44K_100_16.5-10.875-5.5 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | 6GB10300 | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height |  | in | FDA,4 | 4.8125 | 5.1875 | TM 141 / Visual |  |  | left_only
6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | 6GB10300 | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 6.1875 | 6.5625 | TM 141 / Visual |  |  | left_only
6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | 6GB10300 | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 14.4375 | 14.8125 | TM 141 / Visual |  |  | left_only
6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | 6GB10300 | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual |  |  | left_only
6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | 6GB10300 | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 |  |  | TM 141 / Visual |  |  | left_only
6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | 6GB10300 | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | 6GB10300 | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | 6GB10300 | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | 6GB10300 | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC w/ shyed flaps (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 6GD35185 | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 50 |  | FDA,0 |  |  | TM 141 / Visual | 6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 3 | both
6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 6GD35185 | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 3 | both
6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 6GD35185 | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 3 | both
6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 6GD35185 | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 3 | both
6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 6GD35185 | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 12.4375 | in | FDA,4 | 12.25 | 12.625 | TM 141 / Visual | 6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 3 | both
6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 6GD35185 | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 13.25 | in | FDA,4 | 13.0625 | 13.4375 | TM 141 / Visual | 6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 3 | both
6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 6GD35185 | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 3 | both
6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 6GD35185 | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 3 | both
6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 6GD35185 | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 12 | in | FDA,4 | 11.8125 | 12.1875 | TM 141 / Visual | 6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 3 | both
6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | 6GS48450 | SHIP_NY_29K_100_17.125-6.6875-11.75 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | 1 | both
6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | 6GS48450 | SHIP_NY_29K_100_17.125-6.6875-11.75 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 29 | Lbs./In. | FDA,0 | 29 |  | TM 141 / Visual | 6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | 1 | both
6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | 6GS48450 | SHIP_NY_29K_100_17.125-6.6875-11.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | 1 | both
6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | 6GS48450 | SHIP_NY_29K_100_17.125-6.6875-11.75 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 11.875 | in | FDA,4 | 11.5625 | 11.9375 | TM 141 / Visual | 6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | 1 | both
6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | 6GS48450 | SHIP_NY_29K_100_17.125-6.6875-11.75 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 17.125 | in | FDA,4 | 16.9375 | 17.3125 | TM 141 / Visual | 6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | 1 | both
6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | 6GS48450 | SHIP_NY_29K_100_17.125-6.6875-11.75 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | 1 | both
6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | 6GS48450 | SHIP_NY_29K_100_17.125-6.6875-11.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | 1 | both
6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | 6GS48450 | SHIP_NY_29K_100_17.125-6.6875-11.75 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.75 | in | FDA,4 | 6.5 | 6.875 | TM 141 / Visual | 6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | 1 | both
6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | 6HB60307 | SHIP_NY_40K_100_16.5-12.3125-6.1875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | 4 | both
6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | 6HB60307 | SHIP_NY_40K_100_16.5-12.3125-6.1875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 40 | Lbs./In. | FDA,0 | 40 |  | TM 141 / Visual | 6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | 4 | both
6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | 6HB60307 | SHIP_NY_40K_100_16.5-12.3125-6.1875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | 4 | both
6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | 6HB60307 | SHIP_NY_40K_100_16.5-12.3125-6.1875 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6.312 | in | FDA,4 | 6 | 6.375 | TM 141 / Visual | 6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | 4 | both
6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | 6HB60307 | SHIP_NY_40K_100_16.5-12.3125-6.1875 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 16.5 | in | FDA,4 | 16.3125 | 16.6875 | TM 141 / Visual | 6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | 4 | both
6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | 6HB60307 | SHIP_NY_40K_100_16.5-12.3125-6.1875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | 4 | both
6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | 6HB60307 | SHIP_NY_40K_100_16.5-12.3125-6.1875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | 4 | both
6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | 6HB60307 | SHIP_NY_40K_100_16.5-12.3125-6.1875 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 12.25 | in | FDA,4 | 12.125 | 12.5 | TM 141 / Visual | 6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | 4 | both
6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 6HC69900 | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 50 |  | FDA,0 |  |  | TM 141 / Visual | 6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 5 | both
6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 6HC69900 | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 5 | both
6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 6HC69900 | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 5 | both
6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 6HC69900 | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 5 | both
6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 6HC69900 | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 5.25 | in | FDA,4 | 5 | 5.375 | TM 141 / Visual | 6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 5 | both
6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 6HC69900 | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.75 | in | FDA,4 | 15.5625 | 15.9375 | TM 141 / Visual | 6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 5 | both
6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 6HC69900 | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 5 | both
6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 6HC69900 | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 5 | both
6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 6HC69900 | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 8.625 | in | FDA,4 | 8.4375 | 8.8125 | TM 141 / Visual | 6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 5 | both
6JA580FW/Shipper_NY_Skip lot testing | 6JA580FW | Shipper_NY_Skip lot testing | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Shipper Inspection | 6JA580FW/Shipper_NY_Skip lot testing | 9 | both
6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | 6JA580FW | SHIP_NY_44K_100_16.750-8.125-10.750 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | 1 | both
6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | 6JA580FW | SHIP_NY_44K_100_16.750-8.125-10.750 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | 6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | 1 | both
6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | 6JA580FW | SHIP_NY_44K_100_16.750-8.125-10.750 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | 1 | both
6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | 6JA580FW | SHIP_NY_44K_100_16.750-8.125-10.750 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 10.875 | in | FDA,4 | 10.5625 | 10.9375 | TM 141 / Visual | 6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | 1 | both
6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | 6JA580FW | SHIP_NY_44K_100_16.750-8.125-10.750 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 16.9375 | in | FDA,4 | 16.5625 | 16.9375 | TM 141 / Visual | 6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | 1 | both
6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | 6JA580FW | SHIP_NY_44K_100_16.750-8.125-10.750 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | 1 | both
6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | 6JA580FW | SHIP_NY_44K_100_16.750-8.125-10.750 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | 1 | both
6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | 6JA580FW | SHIP_NY_44K_100_16.750-8.125-10.750 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 8.1875 | in | FDA,4 | 7.9375 | 8.3125 | TM 141 / Visual | 6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | 1 | both
6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | 6LB40600 | SHIP_NY_29K_100_12.375-5.5-4.625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | 2 | both
6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | 6LB40600 | SHIP_NY_29K_100_12.375-5.5-4.625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 29 | Lbs./In. | FDA,0 | 29 |  | TM 141 / Visual | 6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | 2 | both
6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | 6LB40600 | SHIP_NY_29K_100_12.375-5.5-4.625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | 2 | both
6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | 6LB40600 | SHIP_NY_29K_100_12.375-5.5-4.625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 4.75 | in | FDA,4 | 4.4375 | 4.8125 | TM 141 / Visual | 6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | 2 | both
6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | 6LB40600 | SHIP_NY_29K_100_12.375-5.5-4.625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 12.437 | in | FDA,4 | 12.1875 | 12.5625 | TM 141 / Visual | 6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | 2 | both
6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | 6LB40600 | SHIP_NY_29K_100_12.375-5.5-4.625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | 2 | both
6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | 6LB40600 | SHIP_NY_29K_100_12.375-5.5-4.625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | 2 | both
6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | 6LB40600 | SHIP_NY_29K_100_12.375-5.5-4.625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 5.5 | in | FDA,4 | 5.3125 | 5.6875 | TM 141 / Visual | 6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | 2 | both
6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | 6OA404S48 | SHIP_NY_44K_100_15.625-12.5-6.1875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | 2 | both
6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | 6OA404S48 | SHIP_NY_44K_100_15.625-12.5-6.1875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | 6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | 2 | both
6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | 6OA404S48 | SHIP_NY_44K_100_15.625-12.5-6.1875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | 2 | both
6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | 6OA404S48 | SHIP_NY_44K_100_15.625-12.5-6.1875 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6 | in | FDA,4 | 6 | 6.375 | TM 141 / Visual | 6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | 2 | both
6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | 6OA404S48 | SHIP_NY_44K_100_15.625-12.5-6.1875 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.625 | in | FDA,4 | 15.4375 | 15.8125 | TM 141 / Visual | 6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | 2 | both
6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | 6OA404S48 | SHIP_NY_44K_100_15.625-12.5-6.1875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | 2 | both
6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | 6OA404S48 | SHIP_NY_44K_100_15.625-12.5-6.1875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | 2 | both
6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | 6OA404S48 | SHIP_NY_44K_100_15.625-12.5-6.1875 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 12.5625 | in | FDA,4 | 12.3125 | 12.6875 | TM 141 / Visual | 6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | 2 | both
6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 6OA540S48 | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual | 6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 1 | both
6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 6OA540S48 | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 1 | both
6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 6OA540S48 | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 1 | both
6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 6OA540S48 | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 1 | both
6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 6OA540S48 | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 12.625 | in | FDA,4 | 12.4375 | 12.8125 | TM 141 / Visual | 6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 1 | both
6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 6OA540S48 | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.875 | in | FDA,4 | 11.6875 | 12.0625 | TM 141 / Visual | 6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 1 | both
6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 6OA540S48 | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 1 | both
6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 6OA540S48 | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 1 | both
6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 6OA540S48 | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.75 | in | FDA,4 | 7.5625 | 7.9375 | TM 141 / Visual | 6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 1 | both
6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 6OA632S60 | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual | 6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 4 | both
6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 6OA632S60 | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 4 | both
6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 6OA632S60 | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 4 | both
6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 6OA632S60 | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 4 | both
6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 6OA632S60 | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 14.25 | in | FDA,4 | 13.9375 | 14.3125 | TM 141 / Visual | 6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 4 | both
6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 6OA632S60 | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 19.5 | in | FDA,4 | 19.1875 | 19.5625 | TM 141 / Visual | 6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 4 | both
6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 6OA632S60 | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 4 | both
6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 6OA632S60 | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | HSC (Done) |  |  |  |  | TM 141 / Visual | 6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 4 | both
6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 6OA632S60 | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 8.25 | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual | 6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 4 | both
6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 6OB123ST | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual | 6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 3 | both
6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 6OB123ST | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 3 | both
6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 6OB123ST | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 3 | both
6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 6OB123ST | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 3 | both
6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 6OB123ST | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 5.875 | in | FDA,4 | 5.5625 | 5.9375 | TM 141 / Visual | 6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 3 | both
6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 6OB123ST | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 12.5 | in | FDA,4 | 12.1875 | 12.5625 | TM 141 / Visual | 6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 3 | both
6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 6OB123ST | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 3 | both
6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 6OB123ST | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 3 | both
6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 6OB123ST | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 10.75 | in | FDA,4 | 10.5625 | 10.9375 | TM 141 / Visual | 6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 3 | both
6OB31400/SHIP_32_1AL | 6OB31400 | SHIP_32_1AL | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6OB31400/SHIP_32_1AL | 2 | both
6OB31400/SHIP_32_1AL | 6OB31400 | SHIP_32_1AL | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6OB31400/SHIP_32_1AL | 2 | both
6OB31400/SHIP_32_1AL | 6OB31400 | SHIP_32_1AL | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6OB31400/SHIP_32_1AL | 2 | both
6OB31400/SHIP_32_1AL | 6OB31400 | SHIP_32_1AL | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 5 | in | FDA,4 | 4.8125 | 5.1875 | TM 141 / Visual | 6OB31400/SHIP_32_1AL | 2 | both
6OB31400/SHIP_32_1AL | 6OB31400 | SHIP_32_1AL | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 14.625 | in | FDA,4 | 14.4375 | 14.8125 | TM 141 / Visual | 6OB31400/SHIP_32_1AL | 2 | both
6OB31400/SHIP_32_1AL | 6OB31400 | SHIP_32_1AL | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6OB31400/SHIP_32_1AL | 2 | both
6OB31400/SHIP_32_1AL | 6OB31400 | SHIP_32_1AL | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6OB31400/SHIP_32_1AL | 2 | both
6OB31400/SHIP_32_1AL | 6OB31400 | SHIP_32_1AL | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.375 | in | FDA,4 | 6.1875 | 6.5625 | TM 141 / Visual | 6OB31400/SHIP_32_1AL | 2 | both
6OB389PD/Ship_NA_3A | 6OB389PD | Ship_NA_3A | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 6.0625 | 6.4375 | TM 141 / Visual |  |  | left_only
6OB389PD/Ship_NA_3A | 6OB389PD | Ship_NA_3A | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 14.3125 | 14.6875 | TM 141 / Visual |  |  | left_only
6OB389PD/Ship_NA_3A | 6OB389PD | Ship_NA_3A | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6OB389PD/Ship_NA_3A | 6OB389PD | Ship_NA_3A | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | Layer Pad (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6OC22324/SHIP_32_1N | 6OC22324 | SHIP_32_1N | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 6.1875 | 6.5625 | TM 141 / Visual |  |  | left_only
6OC22324/SHIP_32_1N | 6OC22324 | SHIP_32_1N | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height |  | in | FDA,4 | 10.1875 | 10.5625 | TM 141 / Visual |  |  | left_only
6OC22324/SHIP_32_1N | 6OC22324 | SHIP_32_1N | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 10.9375 | 11.3125 | TM 141 / Visual |  |  | left_only
6OC22324/SHIP_32_1N | 6OC22324 | SHIP_32_1N | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual |  |  | left_only
6OC22324/SHIP_32_1N | 6OC22324 | SHIP_32_1N | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6OC22324/SHIP_32_1N | 6OC22324 | SHIP_32_1N | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6OC22324/SHIP_32_1N | 6OC22324 | SHIP_32_1N | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6OC22324/SHIP_32_1N | 6OC22324 | SHIP_32_1N | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | 6OD01524 | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height |  | in | FDA,4 | 4.5 | 4.875 | TM 141 / Visual |  |  | left_only
6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | 6OD01524 | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 6.0625 | 6.4375 | TM 141 / Visual |  |  | left_only
6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | 6OD01524 | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual |  |  | left_only
6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | 6OD01524 | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual |  |  | left_only
6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | 6OD01524 | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 |  |  | TM 141 / Visual |  |  | left_only
6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | 6OD01524 | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | 6OD01524 | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | 6OD01524 | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | 6OD01524 | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6OD13500A/SHIP_32_1AD | 6OD13500A | SHIP_32_1AD | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6OD13500A/SHIP_32_1AD | 1 | both
6OD13500A/SHIP_32_1AD | 6OD13500A | SHIP_32_1AD | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6OD13500A/SHIP_32_1AD | 1 | both
6OD13500A/SHIP_32_1AD | 6OD13500A | SHIP_32_1AD | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6OD13500A/SHIP_32_1AD | 1 | both
6OD13500A/SHIP_32_1AD | 6OD13500A | SHIP_32_1AD | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6.25 | in | FDA,4 | 5.9375 | 6.3125 | TM 141 / Visual | 6OD13500A/SHIP_32_1AD | 1 | both
6OD13500A/SHIP_32_1AD | 6OD13500A | SHIP_32_1AD | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 13.25 | in | FDA,4 | 13.0625 | 13.4375 | TM 141 / Visual | 6OD13500A/SHIP_32_1AD | 1 | both
6OD13500A/SHIP_32_1AD | 6OD13500A | SHIP_32_1AD | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6OD13500A/SHIP_32_1AD | 1 | both
6OD13500A/SHIP_32_1AD | 6OD13500A | SHIP_32_1AD | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6OD13500A/SHIP_32_1AD | 1 | both
6OD13500A/SHIP_32_1AD | 6OD13500A | SHIP_32_1AD | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 12 | in | FDA,4 | 11.8125 | 12.1875 | TM 141 / Visual | 6OD13500A/SHIP_32_1AD | 1 | both
6OD239DR/SHIP_44_1A | 6OD239DR | SHIP_44_1A | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 7.4375 | 7.8125 | TM 141 / Visual |  |  | left_only
6OD239DR/SHIP_44_1A | 6OD239DR | SHIP_44_1A | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 15.1875 | 15.5625 | TM 141 / Visual |  |  | left_only
6OD239DR/SHIP_44_1A | 6OD239DR | SHIP_44_1A | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual |  |  | left_only
6OD239DR/SHIP_44_1A | 6OD239DR | SHIP_44_1A | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6OD239DR/SHIP_44_1A | 6OD239DR | SHIP_44_1A | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | Divider |  |  |  |  | Shipper Style |  |  | left_only
6OD239DR/SHIP_44_1A | 6OD239DR | SHIP_44_1A | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6OD239DR/SHIP_44_1A | 6OD239DR | SHIP_44_1A | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | 6OD36055 | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 6.1875 | 6.5625 | TM 141 / Visual |  |  | left_only
6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | 6OD36055 | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 10.4375 | 10.8125 | TM 141 / Visual |  |  | left_only
6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | 6OD36055 | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height |  | in | FDA,4 | 11.0625 | 11.4375 | TM 141 / Visual |  |  | left_only
6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | 6OD36055 | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual |  |  | left_only
6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | 6OD36055 | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 |  |  | TM 141 / Visual |  |  | left_only
6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | 6OD36055 | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | 6OD36055 | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Does not meet specifications (OOS) |  |  |  |  | TM 141 / Visual |  |  | left_only
6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | 6OD36055 | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | 6OD36055 | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC w/ shyed flaps (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6OH14324/SHIP_32_1Y | 6OH14324 | SHIP_32_1Y | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6OH14324/SHIP_32_1Y | 2 | both
6OH14324/SHIP_32_1Y | 6OH14324 | SHIP_32_1Y | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6OH14324/SHIP_32_1Y | 2 | both
6OH14324/SHIP_32_1Y | 6OH14324 | SHIP_32_1Y | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6OH14324/SHIP_32_1Y | 2 | both
6OH14324/SHIP_32_1Y | 6OH14324 | SHIP_32_1Y | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6.25 | in | FDA,4 | 6.0625 | 6.4375 | TM 141 / Visual | 6OH14324/SHIP_32_1Y | 2 | both
6OH14324/SHIP_32_1Y | 6OH14324 | SHIP_32_1Y | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 12.4375 | in | FDA,4 | 12.25 | 12.625 | TM 141 / Visual | 6OH14324/SHIP_32_1Y | 2 | both
6OH14324/SHIP_32_1Y | 6OH14324 | SHIP_32_1Y | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6OH14324/SHIP_32_1Y | 2 | both
6OH14324/SHIP_32_1Y | 6OH14324 | SHIP_32_1Y | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6OH14324/SHIP_32_1Y | 2 | both
6OH14324/SHIP_32_1Y | 6OH14324 | SHIP_32_1Y | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 9 | in | FDA,4 | 8.8125 | 9.1875 | TM 141 / Visual | 6OH14324/SHIP_32_1Y | 2 | both
6OH21112/SHIP_32_1F | 6OH21112 | SHIP_32_1F | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6OH21112/SHIP_32_1F | 1 | both
6OH21112/SHIP_32_1F | 6OH21112 | SHIP_32_1F | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6OH21112/SHIP_32_1F | 1 | both
6OH21112/SHIP_32_1F | 6OH21112 | SHIP_32_1F | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6OH21112/SHIP_32_1F | 1 | both
6OH21112/SHIP_32_1F | 6OH21112 | SHIP_32_1F | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6.75 | in | FDA,4 | 6.5625 | 6.9375 | TM 141 / Visual | 6OH21112/SHIP_32_1F | 1 | both
6OH21112/SHIP_32_1F | 6OH21112 | SHIP_32_1F | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 9.75 | in | FDA,4 | 9.5625 | 9.9375 | TM 141 / Visual | 6OH21112/SHIP_32_1F | 1 | both
6OH21112/SHIP_32_1F | 6OH21112 | SHIP_32_1F | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6OH21112/SHIP_32_1F | 1 | both
6OH21112/SHIP_32_1F | 6OH21112 | SHIP_32_1F | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6OH21112/SHIP_32_1F | 1 | both
6OH21112/SHIP_32_1F | 6OH21112 | SHIP_32_1F | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.5 | in | FDA,4 | 7.4375 | 7.8125 | TM 141 / Visual | 6OH21112/SHIP_32_1F | 1 | both
6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | 6OH45016 | SHIP_NY_32K_100_8.25-6.875-7.375 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | 1 | both
6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | 6OH45016 | SHIP_NY_32K_100_8.25-6.875-7.375 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | 1 | both
6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | 6OH45016 | SHIP_NY_32K_100_8.25-6.875-7.375 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | 1 | both
6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | 6OH45016 | SHIP_NY_32K_100_8.25-6.875-7.375 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7.25 | in | FDA,4 | 7.1875 | 7.5625 | TM 141 / Visual | 6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | 1 | both
6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | 6OH45016 | SHIP_NY_32K_100_8.25-6.875-7.375 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 8.1562 | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual | 6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | 1 | both
6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | 6OH45016 | SHIP_NY_32K_100_8.25-6.875-7.375 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | 1 | both
6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | 6OH45016 | SHIP_NY_32K_100_8.25-6.875-7.375 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | 1 | both
6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | 6OH45016 | SHIP_NY_32K_100_8.25-6.875-7.375 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7 | in | FDA,4 | 6.6875 | 7.0625 | TM 141 / Visual | 6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | 1 | both
6OK75595/SHIP_32_1R | 6OK75595 | SHIP_32_1R | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6OK75595/SHIP_32_1R | 1 | both
6OK75595/SHIP_32_1R | 6OK75595 | SHIP_32_1R | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6OK75595/SHIP_32_1R | 1 | both
6OK75595/SHIP_32_1R | 6OK75595 | SHIP_32_1R | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6OK75595/SHIP_32_1R | 1 | both
6OK75595/SHIP_32_1R | 6OK75595 | SHIP_32_1R | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7.75 | in | FDA,4 | 7.5 | 7.875 | TM 141 / Visual | 6OK75595/SHIP_32_1R | 1 | both
6OK75595/SHIP_32_1R | 6OK75595 | SHIP_32_1R | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.75 | in | FDA,4 | 11.5625 | 11.9375 | TM 141 / Visual | 6OK75595/SHIP_32_1R | 1 | both
6OK75595/SHIP_32_1R | 6OK75595 | SHIP_32_1R | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6OK75595/SHIP_32_1R | 1 | both
6OK75595/SHIP_32_1R | 6OK75595 | SHIP_32_1R | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6OK75595/SHIP_32_1R | 1 | both
6OK75595/SHIP_32_1R | 6OK75595 | SHIP_32_1R | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 9.5 | in | FDA,4 | 9.4375 | 9.8125 | TM 141 / Visual | 6OK75595/SHIP_32_1R | 1 | both
6OM960CB/SHIP_NY_NA_200_2.5-2.5-33.75 | 6OM960CB | SHIP_NY_NA_200_2.5-2.5-33.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6OM960CB/SHIP_NY_NA_200_2.5-2.5-33.75 | 1 | both
6OM960CB/SHIP_NY_NA_200_2.5-2.5-33.75 | 6OM960CB | SHIP_NY_NA_200_2.5-2.5-33.75 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 33.625 | in | FDA,4 | 33.5625 | 33.9375 | TM 141 / Visual | 6OM960CB/SHIP_NY_NA_200_2.5-2.5-33.75 | 1 | both
6OM960CB/SHIP_NY_NA_200_2.5-2.5-33.75 | 6OM960CB | SHIP_NY_NA_200_2.5-2.5-33.75 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 2.5 | in | FDA,4 | 2.3125 | 2.6875 | TM 141 / Visual | 6OM960CB/SHIP_NY_NA_200_2.5-2.5-33.75 | 1 | both
6OM960CB/SHIP_NY_NA_200_2.5-2.5-33.75 | 6OM960CB | SHIP_NY_NA_200_2.5-2.5-33.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | ITW STACK MASTER 2600 (Done) |  |  |  |  | TM 141 / Visual | 6OM960CB/SHIP_NY_NA_200_2.5-2.5-33.75 | 1 | both
6OM960CB/SHIP_NY_NA_200_2.5-2.5-33.75 | 6OM960CB | SHIP_NY_NA_200_2.5-2.5-33.75 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 2.5 | in | FDA,4 | 2.3125 | 2.6875 | TM 141 / Visual | 6OM960CB/SHIP_NY_NA_200_2.5-2.5-33.75 | 1 | both
6OP027LP/SHIP_32_13A | 6OP027LP | SHIP_32_13A | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | Layer Pad |  |  |  |  | Shipper Style | 6OP027LP/SHIP_32_13A | 9 | both
6OP027LP/SHIP_32_13A | 6OP027LP | SHIP_32_13A | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6OP027LP/SHIP_32_13A | 9 | both
6OP027LP/SHIP_32_13A | 6OP027LP | SHIP_32_13A | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6OP027LP/SHIP_32_13A | 9 | both
6OP027LP/SHIP_32_13A | 6OP027LP | SHIP_32_13A | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6OP027LP/SHIP_32_13A | 9 | both
6OP027LP/SHIP_32_13A | 6OP027LP | SHIP_32_13A | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 10.75 | in | FDA,4 | 10.5625 | 10.9375 | TM 141 / Visual | 6OP027LP/SHIP_32_13A | 9 | both
6OP027LP/SHIP_32_13A | 6OP027LP | SHIP_32_13A | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6OP027LP/SHIP_32_13A | 9 | both
6OP027LP/SHIP_32_13A | 6OP027LP | SHIP_32_13A | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 5.5625 | in | FDA,4 | 5.375 | 5.75 | TM 141 / Visual | 6OP027LP/SHIP_32_13A | 9 | both
6OP081LP/SHIP_NY_32NA_250_17.250-11.938 | 6OP081LP | SHIP_NY_32NA_250_17.250-11.938 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,3 | 11.813 | 12.063 | TM 141 / Visual |  |  | left_only
6OP081LP/SHIP_NY_32NA_250_17.250-11.938 | 6OP081LP | SHIP_NY_32NA_250_17.250-11.938 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,3 | 17.125 | 17.375 | TM 141 / Visual |  |  | left_only
6OP081LP/SHIP_NY_32NA_250_17.250-11.938 | 6OP081LP | SHIP_NY_32NA_250_17.250-11.938 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual |  |  | left_only
6OP081LP/SHIP_NY_32NA_250_17.250-11.938 | 6OP081LP | SHIP_NY_32NA_250_17.250-11.938 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 |  |  | TM 141 / Visual |  |  | left_only
6OP081LP/SHIP_NY_32NA_250_17.250-11.938 | 6OP081LP | SHIP_NY_32NA_250_17.250-11.938 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6OP081LP/SHIP_NY_32NA_250_17.250-11.938 | 6OP081LP | SHIP_NY_32NA_250_17.250-11.938 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6OP081LP/SHIP_NY_32NA_250_17.250-11.938 | 6OP081LP | SHIP_NY_32NA_250_17.250-11.938 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | Inserts (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6OP34524/SHIP_44_2G | 6OP34524 | SHIP_44_2G | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6OP34524/SHIP_44_2G | 1 | both
6OP34524/SHIP_44_2G | 6OP34524 | SHIP_44_2G | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | 6OP34524/SHIP_44_2G | 1 | both
6OP34524/SHIP_44_2G | 6OP34524 | SHIP_44_2G | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6OP34524/SHIP_44_2G | 1 | both
6OP34524/SHIP_44_2G | 6OP34524 | SHIP_44_2G | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 8.25 | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual | 6OP34524/SHIP_44_2G | 1 | both
6OP34524/SHIP_44_2G | 6OP34524 | SHIP_44_2G | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 13.5 | in | FDA,4 | 13.3125 | 13.6875 | TM 141 / Visual | 6OP34524/SHIP_44_2G | 1 | both
6OP34524/SHIP_44_2G | 6OP34524 | SHIP_44_2G | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6OP34524/SHIP_44_2G | 1 | both
6OP34524/SHIP_44_2G | 6OP34524 | SHIP_44_2G | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6OP34524/SHIP_44_2G | 1 | both
6OP34524/SHIP_44_2G | 6OP34524 | SHIP_44_2G | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 11.75 | in | FDA,4 | 11.5625 | 11.9375 | TM 141 / Visual | 6OP34524/SHIP_44_2G | 1 | both
6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 6OP75510 | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 32 |  | FDA,0 |  |  | TM 141 / Visual | 6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 1 | both
6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 6OP75510 | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 1 | both
6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 6OP75510 | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | 6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 1 | both
6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 6OP75510 | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 1 | both
6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 6OP75510 | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 10.5625 | in | FDA,4 | 10.375 | 10.75 | TM 141 / Visual | 6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 1 | both
6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 6OP75510 | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.375 | in | FDA,4 | 15.1875 | 15.5625 | TM 141 / Visual | 6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 1 | both
6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 6OP75510 | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 1 | both
6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 6OP75510 | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 1 | both
6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 6OP75510 | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 11.5625 | in | FDA,4 | 11.375 | 11.75 | TM 141 / Visual | 6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 1 | both
6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | 6OQ22272 | SHIP_NY_44K_100_15.188-10.188-15.188 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | 3 | both
6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | 6OQ22272 | SHIP_NY_44K_100_15.188-10.188-15.188 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | 6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | 3 | both
6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | 6OQ22272 | SHIP_NY_44K_100_15.188-10.188-15.188 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | 3 | both
6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | 6OQ22272 | SHIP_NY_44K_100_15.188-10.188-15.188 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 15.188 | in | FDA,4 | 15.0005 | 15.3755 | TM 141 / Visual | 6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | 3 | both
6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | 6OQ22272 | SHIP_NY_44K_100_15.188-10.188-15.188 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.188 | in | FDA,4 | 15.0005 | 15.3755 | TM 141 / Visual | 6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | 3 | both
6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | 6OQ22272 | SHIP_NY_44K_100_15.188-10.188-15.188 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | 3 | both
6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | 6OQ22272 | SHIP_NY_44K_100_15.188-10.188-15.188 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | 3 | both
6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | 6OQ22272 | SHIP_NY_44K_100_15.188-10.188-15.188 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 10.188 | in | FDA,4 | 10.0005 | 10.3755 | TM 141 / Visual | 6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | 3 | both
6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 6OS12350 | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual | 6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 1 | both
6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 6OS12350 | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 1 | both
6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 6OS12350 | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | TM 141 / Visual | 6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 1 | both
6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 6OS12350 | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 1 | both
6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 6OS12350 | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 1 | both
6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 6OS12350 | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 5.938 | in | FDA,4 | 5.7505 | 6.1255 | TM 141 / Visual | 6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 1 | both
6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 6OS12350 | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.75 | in | FDA,4 | 11.4375 | 11.8125 | TM 141 / Visual | 6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 1 | both
6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 6OS12350 | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 1 | both
6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 6OS12350 | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 1 | both
6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 6OS12350 | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.375 | in | FDA,4 | 6.3125 | 6.6875 | TM 141 / Visual | 6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 1 | both
6OS42850/SHIP_44_2N | 6OS42850 | SHIP_44_2N | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6OS42850/SHIP_44_2N | 2 | both
6OS42850/SHIP_44_2N | 6OS42850 | SHIP_44_2N | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | 6OS42850/SHIP_44_2N | 2 | both
6OS42850/SHIP_44_2N | 6OS42850 | SHIP_44_2N | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6OS42850/SHIP_44_2N | 2 | both
6OS42850/SHIP_44_2N | 6OS42850 | SHIP_44_2N | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 5.625 | in | FDA,4 | 5.3125 | 5.6875 | TM 141 / Visual | 6OS42850/SHIP_44_2N | 2 | both
6OS42850/SHIP_44_2N | 6OS42850 | SHIP_44_2N | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.375 | in | FDA,4 | 15.1875 | 15.5625 | TM 141 / Visual | 6OS42850/SHIP_44_2N | 2 | both
6OS42850/SHIP_44_2N | 6OS42850 | SHIP_44_2N | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6OS42850/SHIP_44_2N | 2 | both
6OS42850/SHIP_44_2N | 6OS42850 | SHIP_44_2N | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6OS42850/SHIP_44_2N | 2 | both
6OS42850/SHIP_44_2N | 6OS42850 | SHIP_44_2N | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 11.375 | in | FDA,4 | 11.1875 | 11.5625 | TM 141 / Visual | 6OS42850/SHIP_44_2N | 2 | both
6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 6OS45625 | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual | 6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 1 | both
6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 6OS45625 | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 1 | both
6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 6OS45625 | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | TM 141 / Visual | 6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 1 | both
6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 6OS45625 | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 1 | both
6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 6OS45625 | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 1 | both
6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 6OS45625 | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6 | in | FDA,4 | 5.7505 | 6.1255 | TM 141 / Visual | 6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 1 | both
6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 6OS45625 | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 6.75 | in | FDA,4 | 6.6875 | 7.0625 | TM 141 / Visual | 6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 1 | both
6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 6OS45625 | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 1 | both
6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 6OS45625 | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 1 | both
6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 6OS45625 | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.5 | in | FDA,4 | 6.3125 | 6.6875 | TM 141 / Visual | 6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 1 | both
6OS46250/SHIP_29_1J | 6OS46250 | SHIP_29_1J | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6OS46250/SHIP_29_1J | 1 | both
6OS46250/SHIP_29_1J | 6OS46250 | SHIP_29_1J | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 29 | Lbs./In. | FDA,0 | 29 |  | TM 141 / Visual | 6OS46250/SHIP_29_1J | 1 | both
6OS46250/SHIP_29_1J | 6OS46250 | SHIP_29_1J | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6OS46250/SHIP_29_1J | 1 | both
6OS46250/SHIP_29_1J | 6OS46250 | SHIP_29_1J | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 11.75 | in | FDA,4 | 11.5625 | 11.9375 | TM 141 / Visual | 6OS46250/SHIP_29_1J | 1 | both
6OS46250/SHIP_29_1J | 6OS46250 | SHIP_29_1J | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 17.125 | in | FDA,4 | 16.9375 | 17.3125 | TM 141 / Visual | 6OS46250/SHIP_29_1J | 1 | both
6OS46250/SHIP_29_1J | 6OS46250 | SHIP_29_1J | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6OS46250/SHIP_29_1J | 1 | both
6OS46250/SHIP_29_1J | 6OS46250 | SHIP_29_1J | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6OS46250/SHIP_29_1J | 1 | both
6OS46250/SHIP_29_1J | 6OS46250 | SHIP_29_1J | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.6875 | in | FDA,4 | 6.5 | 6.875 | TM 141 / Visual | 6OS46250/SHIP_29_1J | 1 | both
6OU12066/SHIP_40_1M | 6OU12066 | SHIP_40_1M | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height |  | in | FDA,4 | 8.8125 | 9.1875 | TM 141 / Visual |  |  | left_only
6OU12066/SHIP_40_1M | 6OU12066 | SHIP_40_1M | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 14.9375 | 15.3125 | TM 141 / Visual |  |  | left_only
6OU12066/SHIP_40_1M | 6OU12066 | SHIP_40_1M | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 17.625 | 18 | TM 141 / Visual |  |  | left_only
6OU12066/SHIP_40_1M | 6OU12066 | SHIP_40_1M | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 40 |  | TM 141 / Visual |  |  | left_only
6OU12066/SHIP_40_1M | 6OU12066 | SHIP_40_1M | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6OU12066/SHIP_40_1M | 6OU12066 | SHIP_40_1M | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6OU12066/SHIP_40_1M | 6OU12066 | SHIP_40_1M | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6OU12066/SHIP_40_1M | 6OU12066 | SHIP_40_1M | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6OU21470/SHIP_40_1J | 6OU21470 | SHIP_40_1J | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height |  | in | FDA,4 | 9.0625 | 9.4375 | TM 141 / Visual |  |  | left_only
6OU21470/SHIP_40_1J | 6OU21470 | SHIP_40_1J | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 11.1875 | 11.5625 | TM 141 / Visual |  |  | left_only
6OU21470/SHIP_40_1J | 6OU21470 | SHIP_40_1J | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 15.1875 | 15.5625 | TM 141 / Visual |  |  | left_only
6OU21470/SHIP_40_1J | 6OU21470 | SHIP_40_1J | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 40 |  | TM 141 / Visual |  |  | left_only
6OU21470/SHIP_40_1J | 6OU21470 | SHIP_40_1J | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6OU21470/SHIP_40_1J | 6OU21470 | SHIP_40_1J | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6OU21470/SHIP_40_1J | 6OU21470 | SHIP_40_1J | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6OU21470/SHIP_40_1J | 6OU21470 | SHIP_40_1J | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 6OU31450 | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual | 6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 1 | both
6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 6OU31450 | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 1 | both
6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 6OU31450 | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | 6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 1 | both
6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 6OU31450 | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 1 | both
6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 6OU31450 | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6.125 | in | FDA,4 | 5.9375 | 6.3125 | TM 141 / Visual | 6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 1 | both
6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 6OU31450 | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 18.5 | in | FDA,4 | 18.3125 | 18.6875 | TM 141 / Visual | 6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 1 | both
6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 6OU31450 | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 1 | both
6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 6OU31450 | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 1 | both
6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 6OU31450 | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 13.5 | in | FDA,4 | 13.3125 | 13.6875 | TM 141 / Visual | 6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 1 | both
6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 6OU35199 | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual | 6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 1 | both
6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 6OU35199 | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 1 | both
6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 6OU35199 | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 1 | both
6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 6OU35199 | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 1 | both
6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 6OU35199 | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 11.5 | in | FDA,4 | 11.25 | 11.5 | TM 141 / Visual | 6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 1 | both
6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 6OU35199 | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 19.75 | in | FDA,4 | 19.625 | 19.875 | TM 141 / Visual | 6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 1 | both
6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 6OU35199 | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 1 | both
6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 6OU35199 | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 1 | both
6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 6OU35199 | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 12 | in | FDA,4 | 11.875 | 12.125 | TM 141 / Visual | 6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 1 | both
6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | 6OU351LP | SHIP_NY_32K_NA_19.500X11.750 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | Pad |  |  |  |  | Shipper Style | 6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | 2 | both
6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | 6OU351LP | SHIP_NY_32K_NA_19.500X11.750 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual | 6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | 2 | both
6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | 6OU351LP | SHIP_NY_32K_NA_19.500X11.750 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | 2 | both
6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | 6OU351LP | SHIP_NY_32K_NA_19.500X11.750 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | 2 | both
6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | 6OU351LP | SHIP_NY_32K_NA_19.500X11.750 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | 2 | both
6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | 6OU351LP | SHIP_NY_32K_NA_19.500X11.750 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 19.5 | in | FDA,4 | 19.375 | 19.625 | TM 141 / Visual | 6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | 2 | both
6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | 6OU351LP | SHIP_NY_32K_NA_19.500X11.750 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | 2 | both
6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | 6OU351LP | SHIP_NY_32K_NA_19.500X11.750 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 11.75 | in | FDA,4 | 11.625 | 11.875 | TM 141 / Visual | 6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | 2 | both
6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 6OU351TY | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | D/C Tray |  |  |  |  | Shipper Style | 6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 2 | both
6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 6OU351TY | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual | 6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 2 | both
6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 6OU351TY | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 2 | both
6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 6OU351TY | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 2 | both
6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 6OU351TY | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 2 | both
6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 6OU351TY | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 3.5 | in | FDA,4 | 3.4375 | 3.5625 | TM 141 / Visual | 6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 2 | both
6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 6OU351TY | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 19.062 | in | FDA,4 | 19 | 19.125 | TM 141 / Visual | 6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 2 | both
6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 6OU351TY | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 2 | both
6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 6OU351TY | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 2.812 | in | FDA,4 | 2.75 | 2.875 | TM 141 / Visual | 6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 2 | both
6OU38510/SHIP_32_1BG | 6OU38510 | SHIP_32_1BG | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6OU38510/SHIP_32_1BG | 1 | both
6OU38510/SHIP_32_1BG | 6OU38510 | SHIP_32_1BG | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6OU38510/SHIP_32_1BG | 1 | both
6OU38510/SHIP_32_1BG | 6OU38510 | SHIP_32_1BG | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6OU38510/SHIP_32_1BG | 1 | both
6OU38510/SHIP_32_1BG | 6OU38510 | SHIP_32_1BG | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6 | in | FDA,4 | 5.8125 | 6.1875 | TM 141 / Visual | 6OU38510/SHIP_32_1BG | 1 | both
6OU38510/SHIP_32_1BG | 6OU38510 | SHIP_32_1BG | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 8 | in | FDA,4 | 7.8125 | 8.1875 | TM 141 / Visual | 6OU38510/SHIP_32_1BG | 1 | both
6OU38510/SHIP_32_1BG | 6OU38510 | SHIP_32_1BG | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6OU38510/SHIP_32_1BG | 1 | both
6OU38510/SHIP_32_1BG | 6OU38510 | SHIP_32_1BG | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6OU38510/SHIP_32_1BG | 1 | both
6OU38510/SHIP_32_1BG | 6OU38510 | SHIP_32_1BG | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.125 | in | FDA,4 | 5.9375 | 6.3125 | TM 141 / Visual | 6OU38510/SHIP_32_1BG | 1 | both
6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 6OU55510 | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual | 6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 3 | both
6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 6OU55510 | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 3 | both
6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 6OU55510 | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 3 | both
6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 6OU55510 | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 3 | both
6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 6OU55510 | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 8 | in | FDA,4 | 7.8125 | 8.1875 | TM 141 / Visual | 6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 3 | both
6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 6OU55510 | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.75 | in | FDA,4 | 11.5625 | 11.9375 | TM 141 / Visual | 6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 3 | both
6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 6OU55510 | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 3 | both
6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 6OU55510 | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 3 | both
6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 6OU55510 | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.375 | in | FDA,4 | 6.0625 | 6.4375 | TM 141 / Visual | 6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 3 | both
6OU62620/SHIP_NY_32K_100_15-9.5-8 | 6OU62620 | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual | 6OU62620/SHIP_NY_32K_100_15-9.5-8 | 3 | both
6OU62620/SHIP_NY_32K_100_15-9.5-8 | 6OU62620 | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6OU62620/SHIP_NY_32K_100_15-9.5-8 | 3 | both
6OU62620/SHIP_NY_32K_100_15-9.5-8 | 6OU62620 | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6OU62620/SHIP_NY_32K_100_15-9.5-8 | 3 | both
6OU62620/SHIP_NY_32K_100_15-9.5-8 | 6OU62620 | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6OU62620/SHIP_NY_32K_100_15-9.5-8 | 3 | both
6OU62620/SHIP_NY_32K_100_15-9.5-8 | 6OU62620 | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 8 | in | FDA,4 | 7.8125 | 8.1875 | TM 141 / Visual | 6OU62620/SHIP_NY_32K_100_15-9.5-8 | 3 | both
6OU62620/SHIP_NY_32K_100_15-9.5-8 | 6OU62620 | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.125 | in | FDA,4 | 14.8125 | 15.1875 | TM 141 / Visual | 6OU62620/SHIP_NY_32K_100_15-9.5-8 | 3 | both
6OU62620/SHIP_NY_32K_100_15-9.5-8 | 6OU62620 | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6OU62620/SHIP_NY_32K_100_15-9.5-8 | 3 | both
6OU62620/SHIP_NY_32K_100_15-9.5-8 | 6OU62620 | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | 6OU62620/SHIP_NY_32K_100_15-9.5-8 | 3 | both
6OU62620/SHIP_NY_32K_100_15-9.5-8 | 6OU62620 | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 9.375 | in | FDA,4 | 9.3125 | 9.6875 | TM 141 / Visual | 6OU62620/SHIP_NY_32K_100_15-9.5-8 | 3 | both
6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 6OX46500 | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual | 6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 1 | both
6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 6OX46500 | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | 6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 1 | both
6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 6OX46500 | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | 6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 1 | both
6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 6OX46500 | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | 6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 1 | both
6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 6OX46500 | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7.5 | in | FDA,4 | 7.3125 | 7.6875 | TM 141 / Visual | 6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 1 | both
6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 6OX46500 | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 23.75 | in | FDA,4 | 23.5005 | 23.8755 | TM 141 / Visual | 6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 1 | both
6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 6OX46500 | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | 6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 1 | both
6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 6OX46500 | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | Shipper (Done) |  |  |  |  | TM 141 / Visual | 6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 1 | both
6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 6OX46500 | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 15.6 | in | FDA,4 | 15.5005 | 15.8755 | TM 141 / Visual | 6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 1 | both
6RQ55172/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 6RQ55172 | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Shipper Inspection | 6RQ55172/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 6 | both
6ZS505PD/SHIP_32_14A | 6ZS505PD | SHIP_32_14A | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 5.5625 | 5.9375 | TM 141 / Visual |  |  | left_only
6ZS505PD/SHIP_32_14A | 6ZS505PD | SHIP_32_14A | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 11.1875 | 11.5625 | TM 141 / Visual |  |  | left_only
6ZS505PD/SHIP_32_14A | 6ZS505PD | SHIP_32_14A | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual |  |  | left_only
6ZS505PD/SHIP_32_14A | 6ZS505PD | SHIP_32_14A | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6ZS505PD/SHIP_32_14A | 6ZS505PD | SHIP_32_14A | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | Cut Pad - Sheet (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
6ZS505PD/SHIP_32_14A | 6ZS505PD | SHIP_32_14A | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual |  |  | left_only
7A4AA580/Chem_NY_Disodium Cocoamphodipropinate | 7A4AA580 | Chem_NY_Disodium Cocoamphodipropinate | IN-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 |  |  | left_only
7A4AA580/Chem_NY_Disodium Cocoamphodipropinate | 7A4AA580 | Chem_NY_Disodium Cocoamphodipropinate | IN-Analytical | 447 | pH as is | pH | pH |  |  | FDA,1 | 9 | 10 | TM 200 |  |  | left_only
7A4AA580/Chem_NY_Disodium Cocoamphodipropinate | 7A4AA580 | Chem_NY_Disodium Cocoamphodipropinate | IN-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA |  |  | left_only
7A4AA580/Chem_NY_Disodium Cocoamphodipropinate | 7A4AA580 | Chem_NY_Disodium Cocoamphodipropinate | IN-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear to Hazy liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7A4AA580/Chem_NY_Disodium Cocoamphodipropinate | 7A4AA580 | Chem_NY_Disodium Cocoamphodipropinate | IN-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to Yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7A4AA580/Chem_NY_Disodium Cocoamphodipropinate | 7A4AA580 | Chem_NY_Disodium Cocoamphodipropinate | IN-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7A4AA580/Chem_NY_Disodium Cocoamphodipropinate | 7A4AA580 | Chem_NY_Disodium Cocoamphodipropinate | IN-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7A4AL930/Chem_08A | 7A4AL930 | Chem_08A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection | 7A4AL930/Chem_08A | 6 | both
7A4AL930/Chem_08A | 7A4AL930 | Chem_08A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Pellets (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4AL930/Chem_08A | 6 | both
7A4AL930/Chem_08A | 7A4AL930 | Chem_08A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White/Grey (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4AL930/Chem_08A | 6 | both
7A4AL930/Chem_08A | 7A4AL930 | Chem_08A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4AL930/Chem_08A | 6 | both
7A4AL930/Chem_08A | 7A4AL930 | Chem_08A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4AL930/Chem_08A | 6 | both
7A4AL930/Chem_08A | 7A4AL930 | Chem_08A | NY-Analytical | 390 | Identification | Identification for Potassium (USP <191>) | Identification for Potassium | Pass (Done) |  |  |  |  | TM400/408b/421/422/USP | 7A4AL930/Chem_08A | 6 | both
7A4AS162/Chem_09A | 7A4AS162 | Chem_09A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C | 7A4AS162/Chem_09A | 4 | both
7A4AS162/Chem_09A | 7A4AS162 | Chem_09A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.2846 |  | FDA,3 | 1.249 |  | TM 202 | 7A4AS162/Chem_09A | 4 | both
7A4AS162/Chem_09A | 7A4AS162 | Chem_09A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7A4AS162/Chem_09A | 4 | both
7A4AS162/Chem_09A | 7A4AS162 | Chem_09A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, syrupy liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4AS162/Chem_09A | 4 | both
7A4AS162/Chem_09A | 7A4AS162 | Chem_09A | NY-Analytical | 792 | Chemical Inspection | Color (USP) | Color | Meet USP requirements (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4AS162/Chem_09A | 4 | both
7A4AS162/Chem_09A | 7A4AS162 | Chem_09A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4AS162/Chem_09A | 4 | both
7A4AS162/Chem_09A | 7A4AS162 | Chem_09A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4AS162/Chem_09A | 4 | both
7A4AS162/Chem_09A | 7A4AS162 | Chem_09A | NY-Analytical | 816 | Related Compounds | Individual Impurity | Individual Impurity | 0.01 | % | FDA,3 |  | 0.1 | USP <621> | 7A4AS162/Chem_09A | 4 | both
7A4AS162/Chem_09A | 7A4AS162 | Chem_09A | NY-Analytical | 816 | Related Compounds | Total Impurities | Total Impurities | 0.02 | % | FDA,3 |  | 1 | USP <621> | 7A4AS162/Chem_09A | 4 | both
7A4BL625/Chem_11A | 7A4BL625 | Chem_11A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification | 7A4BL625/Chem_11A | 1 | both
7A4BL625/Chem_11A | 7A4BL625 | Chem_11A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7A4BL625/Chem_11A | 1 | both
7A4BL625/Chem_11A | 7A4BL625 | Chem_11A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4BL625/Chem_11A | 1 | both
7A4BL625/Chem_11A | 7A4BL625 | Chem_11A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Light Yellow to Yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4BL625/Chem_11A | 1 | both
7A4BL625/Chem_11A | 7A4BL625 | Chem_11A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4BL625/Chem_11A | 1 | both
7A4BL625/Chem_11A | 7A4BL625 | Chem_11A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Lard-Like (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4BL625/Chem_11A | 1 | both
7A4CD419/Chem_13A | 7A4CD419 | Chem_13A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification | 7A4CD419/Chem_13A | 2 | both
7A4CD419/Chem_13A | 7A4CD419 | Chem_13A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7A4CD419/Chem_13A | 2 | both
7A4CD419/Chem_13A | 7A4CD419 | Chem_13A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Crystalline Powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4CD419/Chem_13A | 2 | both
7A4CD419/Chem_13A | 7A4CD419 | Chem_13A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White to off White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4CD419/Chem_13A | 2 | both
7A4CD419/Chem_13A | 7A4CD419 | Chem_13A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4CD419/Chem_13A | 2 | both
7A4CD419/Chem_13A | 7A4CD419 | Chem_13A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Faint phenolic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4CD419/Chem_13A | 2 | both
7A4FD110/Chem_Sodium Chloride USP/NF | 7A4FD110 | Chem_Sodium Chloride USP/NF | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection | 7A4FD110/Chem_Sodium Chloride USP/NF | 4 | both
7A4FD110/Chem_Sodium Chloride USP/NF | 7A4FD110 | Chem_Sodium Chloride USP/NF | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Crystals/crystalline powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4FD110/Chem_Sodium Chloride USP/NF | 4 | both
7A4FD110/Chem_Sodium Chloride USP/NF | 7A4FD110 | Chem_Sodium Chloride USP/NF | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to white (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4FD110/Chem_Sodium Chloride USP/NF | 4 | both
7A4FD110/Chem_Sodium Chloride USP/NF | 7A4FD110 | Chem_Sodium Chloride USP/NF | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4FD110/Chem_Sodium Chloride USP/NF | 4 | both
7A4FD110/Chem_Sodium Chloride USP/NF | 7A4FD110 | Chem_Sodium Chloride USP/NF | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4FD110/Chem_Sodium Chloride USP/NF | 4 | both
7A4FD110/Chem_Sodium Chloride USP/NF | 7A4FD110 | Chem_Sodium Chloride USP/NF | NY-Analytical | 390 | Identification | Identification for Chloride (USP <191>) | Identification for Chloride | Meets USP (Done) |  |  |  |  | TM400/408b/421/422/USP | 7A4FD110/Chem_Sodium Chloride USP/NF | 4 | both
7A4FD110/Chem_Sodium Chloride USP/NF | 7A4FD110 | Chem_Sodium Chloride USP/NF | NY-Analytical | 390 | Identification | Identification for Sodium (USP <191>) | Identification for Sodium | Meets USP (Done) |  |  |  |  | TM400/408b/421/422/USP | 7A4FD110/Chem_Sodium Chloride USP/NF | 4 | both
7A4FD110/Chem_Sodium Chloride USP/NF | 7A4FD110 | Chem_Sodium Chloride USP/NF | NY-Analytical | 802 | Loss on Drying | Loss on Drying | Loss on Drying |  | % | FDA,1 |  | 0.5 | USP | 7A4FD110/Chem_Sodium Chloride USP/NF | 4 | both
7A4FD110/Chem_Sodium Chloride USP/NF | 7A4FD110 | Chem_Sodium Chloride USP/NF | NY-Analytical | 802 | Loss on Drying | Loss on Drying Final | Loss on Drying Final |  | % | FDA,1 |  | 0.5 | USP | 7A4FD110/Chem_Sodium Chloride USP/NF | 4 | both
7A4FD110/Chem_Sodium Chloride USP/NF | 7A4FD110 | Chem_Sodium Chloride USP/NF | NY-Analytical | 796 | Assay | Assay on Dried Basis | Assay on Dried Basis | 99.7 | % | FDA,1 | 99 | 100.5 | USP <541> | 7A4FD110/Chem_Sodium Chloride USP/NF | 4 | both
7A4FD110/Chem_Sodium Chloride USP/NF | 7A4FD110 | Chem_Sodium Chloride USP/NF | NY-Analytical | 796 | Assay | Assay on Dried Basis Final | Assay on Dried Basis Final | 99.7 | % | FDA,1 | 99 | 100.5 | USP <541> | 7A4FD110/Chem_Sodium Chloride USP/NF | 4 | both
7A4OA347/Chem_NY_Sodium-L(+) Lactate 60% | 7A4OA347 | Chem_NY_Sodium-L(+) Lactate 60% | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Comparison of scans of sample vs standard (Done) |  |  |  |  | TM 411 |  |  | left_only
7A4OA347/Chem_NY_Sodium-L(+) Lactate 60% | 7A4OA347 | Chem_NY_Sodium-L(+) Lactate 60% | NY-Analytical | 398 | pH | pH (1 in 20 solution) | pH (1 in 20 solution) |  |  | FDA,1 | 6.5 | 7.5 | TM 200 |  |  | left_only
7A4OA347/Chem_NY_Sodium-L(+) Lactate 60% | 7A4OA347 | Chem_NY_Sodium-L(+) Lactate 60% | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH |  |  | left_only
7A4OA347/Chem_NY_Sodium-L(+) Lactate 60% | 7A4OA347 | Chem_NY_Sodium-L(+) Lactate 60% | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7A4OA347/Chem_NY_Sodium-L(+) Lactate 60% | 7A4OA347 | Chem_NY_Sodium-L(+) Lactate 60% | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7A4OA347/Chem_NY_Sodium-L(+) Lactate 60% | 7A4OA347 | Chem_NY_Sodium-L(+) Lactate 60% | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless or with a slight Characteristic Odor (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7A4OA347/Chem_NY_Sodium-L(+) Lactate 60% | 7A4OA347 | Chem_NY_Sodium-L(+) Lactate 60% | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Transparent liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7A4OA709/Chem_NY_Chamomile Extract # 11-19-2 | 7A4OA709 | Chem_NY_Chamomile Extract # 11-19-2 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH | 7A4OA709/Chem_NY_Chamomile Extract # 11-19-2 | 2 | both
7A4OA709/Chem_NY_Chamomile Extract # 11-19-2 | 7A4OA709 | Chem_NY_Chamomile Extract # 11-19-2 | NY-Analytical | 398 | pH | pH | pH | 4.92 |  | FDA,1 | 4 | 6.5 | TM 200 | 7A4OA709/Chem_NY_Chamomile Extract # 11-19-2 | 2 | both
7A4OA709/Chem_NY_Chamomile Extract # 11-19-2 | 7A4OA709 | Chem_NY_Chamomile Extract # 11-19-2 | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.0482 |  | FDA,3 | 1.043 | 1.053 | TM 202 | 7A4OA709/Chem_NY_Chamomile Extract # 11-19-2 | 2 | both
7A4OA709/Chem_NY_Chamomile Extract # 11-19-2 | 7A4OA709 | Chem_NY_Chamomile Extract # 11-19-2 | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.412 |  | FDA,3 | 1.409 | 1.419 | TM 204 | 7A4OA709/Chem_NY_Chamomile Extract # 11-19-2 | 2 | both
7A4OA709/Chem_NY_Chamomile Extract # 11-19-2 | 7A4OA709 | Chem_NY_Chamomile Extract # 11-19-2 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None- free of foreign particles (Done) |  |  |  |  | TM 208 | 7A4OA709/Chem_NY_Chamomile Extract # 11-19-2 | 2 | both
7A4OA709/Chem_NY_Chamomile Extract # 11-19-2 | 7A4OA709 | Chem_NY_Chamomile Extract # 11-19-2 | NY-Analytical | 792 | Chemical Inspection | Description (TM 208) | Description | Clear Amber to brown liquid having a characteristic odor (Done) |  |  |  |  | TM 208 | 7A4OA709/Chem_NY_Chamomile Extract # 11-19-2 | 2 | both
7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | 7A4OA911 | Chem_NY_Fragrance Hugs & Kisses Sheer1 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Refractive Index at 25°C | 7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | 1 | both
7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | 7A4OA911 | Chem_NY_Fragrance Hugs & Kisses Sheer1 | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.4537 |  | FDA,4 | 1.4536 | 1.4636 | TM 204 | 7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | 1 | both
7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | 7A4OA911 | Chem_NY_Fragrance Hugs & Kisses Sheer1 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Liquid (Done) |  |  |  |  | TM 208 | 7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | 1 | both
7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | 7A4OA911 | Chem_NY_Fragrance Hugs & Kisses Sheer1 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to Slightly Yellow (Done) |  |  |  |  | TM 208 | 7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | 1 | both
7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | 7A4OA911 | Chem_NY_Fragrance Hugs & Kisses Sheer1 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 | 7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | 1 | both
7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | 7A4OA911 | Chem_NY_Fragrance Hugs & Kisses Sheer1 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208 | 7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | 1 | both
7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | 7A4OA911 | Chem_NY_Fragrance Hugs & Kisses Sheer1 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411 | 7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | 1 | both
7A4OB223/Chem_189A | 7A4OB223 | Chem_189A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification | 7A4OB223/Chem_189A | 1 | both
7A4OB223/Chem_189A | 7A4OB223 | Chem_189A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7A4OB223/Chem_189A | 1 | both
7A4OB223/Chem_189A | 7A4OB223 | Chem_189A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Translucent crystals or crystalline powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OB223/Chem_189A | 1 | both
7A4OB223/Chem_189A | 7A4OB223 | Chem_189A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to white (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OB223/Chem_189A | 1 | both
7A4OB223/Chem_189A | 7A4OB223 | Chem_189A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OB223/Chem_189A | 1 | both
7A4OB223/Chem_189A | 7A4OB223 | Chem_189A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OB223/Chem_189A | 1 | both
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 125 | both
7A4OB314/Chem_125A | 7A4OB314 | Chem_125A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | [5 empty] | TM 411 |  |  | left_only
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | 1 | both
7A4OB314/Chem_125A | 7A4OB314 | Chem_125A | NY-Analytical | 398 | pH | pH (10% Solution) | pH (10% Solution) |  |  | FDA,1 | 6 | 8 | TM 200 |  |  | left_only
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.7837 |  | FDA,3 | 0.783 | 0.787 | TM 202 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 125 | both
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.7839 |  | FDA,3 | 0.783 | 0.787 | TM 202 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | 1 | both
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C |  |  | FDA,3 | 1.376 | 1.378 | TM 204 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 125 | both
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, mobile liquid (Done) |  |  |  |  | TM 208 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | 1 | both
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM 208 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | 1 | both
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | 1 | both
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | IPA - Characteristic (Done) |  |  |  |  | TM 208 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | 1 | both
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 797 | Nonvolatile Residue | End Weight (g) | End Weight |  | g | FDA,4 |  |  | TM 226 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 125 | both
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 797 | Nonvolatile Residue | Nonvolatile Residue (calc) | Nonvolatile Residue |  | % | FDA,3 |  | 0.005 | TM 226 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 125 | both
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 797 | Nonvolatile Residue | Nonvolatile Residue Final (calc) | Nonvolatile Residue Final |  | % | FDA,3 |  | 0.005 | TM 226 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 125 | both
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 797 | Nonvolatile Residue | Start Weight (g) | Start Weight |  | g | FDA,4 |  |  | TM 226 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 125 | both
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 797 | Nonvolatile Residue | Weight of Residue (g) | Weight of Residue |  | g | FDA,4 |  |  | TM 226 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 125 | both
7A4OB314/Chem_125A | 7A4OB314 | Chem_125A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Certificate of Analysis |  |  | left_only
7A4OB314/Chem_125A | 7A4OB314 | Chem_125A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear viscous liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 795 | Acidity | Acidity (calc) | Acidity |  | mL | FDA,2 |  | 0.7 | TM 375 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 125 | both
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 795 | Acidity | Acidity Final (Calc) | Acidity Final (Calc) |  | mL | FDA,2 |  | 0.7 | TM 375 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 125 | both
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 795 | Acidity | Titrant End Volume (mL) | Titrant End Volume |  | mL | FDA,2 |  |  | TM 375 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 125 | both
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 795 | Acidity | Titrant Start Volume (mL) | Titrant Start Volume |  | mL | FDA,2 |  |  | TM 375 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 125 | both
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 796 | Assay | Assay | Assay |  | % | FDA,1 | 99 |  | TM 376 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 125 | both
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 796 | Assay | Assay | Assay | 99.96 | % | FDA,1 | 99 |  | TM 376 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | 1 | both
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 796 | Assay | Assay Final | Assay Final |  | % | FDA,1 | 99 |  | TM 376 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 125 | both
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 796 | Assay | Assay Final | Assay Final |  | % | FDA,1 | 99 |  | TM 376 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | 1 | both
7A4OB314/Chem_125A | 7A4OB314 | Chem_125A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Match standard, per USP (Done) |  |  |  |  | TM 411 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | 1 | both
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 390 | Identification | IR Identification | IR Identification | [5 empty] | TM 411 | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 125 | both
7A4OB314/Chem_125A | 7A4OB314 | Chem_125A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | [5 empty] | TM208/220/230/USP/Kao LM |  |  | left_only
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | [5 empty] | TM208/220/230/USP/Kao LM | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 125 | both
7A4OB314/Chem_125A | 7A4OB314 | Chem_125A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | [5 empty] | TM208/220/230/USP/Kao LM |  |  | left_only
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | [5 empty] | TM208/220/230/USP/Kao LM | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 125 | both
7A4OB314/Chem_125A | 7A4OB314 | Chem_125A | NY-Analytical | 796 | Assay | Assay | Assay |  | % | FDA,1 | 76 | 84 | TM 301/434 |  |  | left_only
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | [5 empty] | TM208/220/230/USP/Kao LM | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 125 | both
7A4OB314/Chem_125A | 7A4OB314 | Chem_125A | NY-Analytical | 796 | Assay | Assay Final | Assay Final |  | % | FDA,1 | 76 | 84 | TM 301/434 |  |  | left_only
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | [5 empty] | TM208/220/230/USP/Kao LM | 7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 125 | both
7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB339 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.376 |  | FDA,3 | 1.376 | 1.378 | TM 204 | 7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 138 | both
7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB339 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 797 | Nonvolatile Residue | End Weight (g) | End Weight | 92.0885 | g | FDA,4 |  |  | TM 226 | 7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 138 | both
7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB339 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 797 | Nonvolatile Residue | Nonvolatile Residue (calc) | Nonvolatile Residue | 0 | % | FDA,3 |  | 0.005 | TM 226 | 7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 138 | both
7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB339 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 797 | Nonvolatile Residue | Nonvolatile Residue Final (calc) | Nonvolatile Residue Final | 0 | % | FDA,3 |  | 0.005 | TM 226 | 7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 138 | both
7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB339 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 797 | Nonvolatile Residue | Start Weight (g) | Start Weight | 92.0885 | g | FDA,4 |  |  | TM 226 | 7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 138 | both
7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB339 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 797 | Nonvolatile Residue | Weight of Residue (g) | Weight of Residue | 0 | g | FDA,4 |  |  | TM 226 | 7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 138 | both
7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB339 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 795 | Acidity | Acidity (calc) | Acidity | 0.4 | mL | FDA,2 |  | 0.7 | TM 375 | 7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 138 | both
7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB339 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 795 | Acidity | Acidity Final (Calc) | Acidity Final (Calc) | 0.425 | mL | FDA,2 |  | 0.7 | TM 375 | 7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 138 | both
7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB339 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 795 | Acidity | Titrant End Volume (mL) | Titrant End Volume | 0.5 | mL | FDA,2 |  |  | TM 375 | 7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 138 | both
7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB339 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 795 | Acidity | Titrant Start Volume (mL) | Titrant Start Volume | 0 | mL | FDA,2 |  |  | TM 375 | 7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 138 | both
7A4OB397/Chem_04A | 7A4OB397 | Chem_04A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Assay | 7A4OB397/Chem_04A | 2 | both
7A4OB397/Chem_04A | 7A4OB397 | Chem_04A | NY-Analytical | 796 | Assay | Assay | Assay | 55.85 | % | FDA,1 | 47.5 | 52.5 | TM 301/434 | 7A4OB397/Chem_04A | 2 | both
7A4OB397/Chem_04A | 7A4OB397 | Chem_04A | NY-Analytical | 796 | Assay | Assay Final | Assay Final | 55.77 | % | FDA,1 | 47.5 | 52.5 | TM 301/434 | 7A4OB397/Chem_04A | 2 | both
7A4OB397/Chem_04A | 7A4OB397 | Chem_04A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear viscous liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OB397/Chem_04A | 2 | both
7A4OB397/Chem_04A | 7A4OB397 | Chem_04A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to light yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OB397/Chem_04A | 2 | both
7A4OB397/Chem_04A | 7A4OB397 | Chem_04A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OB397/Chem_04A | 2 | both
7A4OB397/Chem_04A | 7A4OB397 | Chem_04A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Mild almond (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OB397/Chem_04A | 2 | both
7A4OB397/Chem_04A | 7A4OB397 | Chem_04A | NY-Analytical | 390 | Identification | Identification (TM 400) | Identification | Formation of precipitate per test (Done) |  |  |  |  | TM400/408b/421/422/USP | 7A4OB397/Chem_04A | 2 | both
7A4OB690/Chem_01A | 7A4OB690 | Chem_01A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification | 7A4OB690/Chem_01A | 5 | both
7A4OB690/Chem_01A | 7A4OB690 | Chem_01A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7A4OB690/Chem_01A | 5 | both
7A4OB690/Chem_01A | 7A4OB690 | Chem_01A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, mobile liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OB690/Chem_01A | 5 | both
7A4OB690/Chem_01A | 7A4OB690 | Chem_01A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OB690/Chem_01A | 5 | both
7A4OB690/Chem_01A | 7A4OB690 | Chem_01A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OB690/Chem_01A | 5 | both
7A4OB690/Chem_01A | 7A4OB690 | Chem_01A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Solvent (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OB690/Chem_01A | 5 | both
7A4OF602/Chem_19A | 7A4OF602 | Chem_19A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.968 |  | FDA,3 | 0.964 | 0.968 | TM 202 | 7A4OF602/Chem_19A | 1 | both
7A4OF602/Chem_19A | 7A4OF602 | Chem_19A | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.458 |  | FDA,3 | 1.455 | 1.461 | TM 204 | 7A4OF602/Chem_19A | 1 | both
7A4OF602/Chem_19A | 7A4OF602 | Chem_19A | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411 | 7A4OF602/Chem_19A | 1 | both
7A4OF602/Chem_19A | 7A4OF602 | Chem_19A | NY-Analytical | 792 | Chemical Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OF602/Chem_19A | 1 | both
7A4OF602/Chem_19A | 7A4OF602 | Chem_19A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OF602/Chem_19A | 1 | both
7A4OG169/Chem_20A | 7A4OG169 | Chem_20A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification | 7A4OG169/Chem_20A | 1 | both
7A4OG169/Chem_20A | 7A4OG169 | Chem_20A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7A4OG169/Chem_20A | 1 | both
7A4OG169/Chem_20A | 7A4OG169 | Chem_20A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OG169/Chem_20A | 1 | both
7A4OG169/Chem_20A | 7A4OG169 | Chem_20A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Amber (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OG169/Chem_20A | 1 | both
7A4OG169/Chem_20A | 7A4OG169 | Chem_20A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OG169/Chem_20A | 1 | both
7A4OH247/Chem_21A | 7A4OH247 | Chem_21A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection | 7A4OH247/Chem_21A | 5 | both
7A4OH247/Chem_21A | 7A4OH247 | Chem_21A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | White granular material (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OH247/Chem_21A | 5 | both
7A4OH247/Chem_21A | 7A4OH247 | Chem_21A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Match reference (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OH247/Chem_21A | 5 | both
7A4OH247/Chem_21A | 7A4OH247 | Chem_21A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OH247/Chem_21A | 5 | both
7A4OH247/Chem_21A | 7A4OH247 | Chem_21A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match reference (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OH247/Chem_21A | 5 | both
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | Chem_NY_Salicylic Acid USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection | 7A4OK459/Chem_NY_Salicylic Acid USP | 3 | both
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | Chem_NY_Salicylic Acid USP | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | White crystalline powder (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | 7A4OK459/Chem_NY_Salicylic Acid USP | 3 | both
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | Chem_NY_Salicylic Acid USP | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | 7A4OK459/Chem_NY_Salicylic Acid USP | 3 | both
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | Chem_NY_Salicylic Acid USP | NY-Analytical | 832 | Chloride | Chloride | Chloride | <0.014 | % | FDA,3 |  | 0.014 | USP | 7A4OK459/Chem_NY_Salicylic Acid USP | 3 | both
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | Chem_NY_Salicylic Acid USP | NY-Analytical | 833 | Heavy Metals | Heavy Metals | Heavy Metals | <20 | ug/g | FDA,0 |  | <20 | USP | 7A4OK459/Chem_NY_Salicylic Acid USP | 3 | both
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | Chem_NY_Salicylic Acid USP | NY-Analytical | 816 | Related Compounds | Impurity | Impurity | 0 | % | FDA,4 |  | 0.05 | USP | 7A4OK459/Chem_NY_Salicylic Acid USP | 3 | both
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | Chem_NY_Salicylic Acid USP | NY-Analytical | 802 | Loss on Drying | Loss on Drying | Loss on Drying | 0.03 | % | FDA,3 |  | 0.5 | USP | 7A4OK459/Chem_NY_Salicylic Acid USP | 3 | both
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | Chem_NY_Salicylic Acid USP | NY-Analytical | 816 | Related Compounds | Phenol | Phenol | 0 | % | FDA,4 |  | 0.01 | USP | 7A4OK459/Chem_NY_Salicylic Acid USP | 3 | both
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | Chem_NY_Salicylic Acid USP | NY-Analytical | 816 | Related Compounds | Related Compound A | Related Compound A | 0.043 | % | FDA,4 |  | 0.1 | USP | 7A4OK459/Chem_NY_Salicylic Acid USP | 3 | both
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | Chem_NY_Salicylic Acid USP | NY-Analytical | 816 | Related Compounds | Related Compound B | Related Compound B | 0.037 | % | FDA,4 |  | 0.05 | USP | 7A4OK459/Chem_NY_Salicylic Acid USP | 3 | both
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | Chem_NY_Salicylic Acid USP | NY-Analytical | 834 | Residue on Ignition | Residue on Ignition | Residue on Ignition | 0.02 | % | FDA,2 |  | 0.05 | USP | 7A4OK459/Chem_NY_Salicylic Acid USP | 3 | both
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | Chem_NY_Salicylic Acid USP | NY-Analytical | 835 | Sulfate Content | Sulfate Content | Sulfate Content | <0.020 | % | FDA,3 |  | 0.02 | USP | 7A4OK459/Chem_NY_Salicylic Acid USP | 3 | both
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | Chem_NY_Salicylic Acid USP | NY-Analytical | 816 | Related Compounds | Total Impurities | Total Impurities | 0.081 | % | FDA,4 |  | 0.2 | USP | 7A4OK459/Chem_NY_Salicylic Acid USP | 3 | both
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | Chem_NY_Salicylic Acid USP | NY-Analytical | 796 | Assay | Assay | Assay | 99.7 | % | FDA,1 | 98 | 102 | USP/NF | 7A4OK459/Chem_NY_Salicylic Acid USP | 3 | both
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | Chem_NY_Salicylic Acid USP | NY-Analytical | 455 | Identification A | Identification A | Identification A | Passes Test (Done) |  |  |  |  | USP/NF | 7A4OK459/Chem_NY_Salicylic Acid USP | 3 | both
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | Chem_NY_Salicylic Acid USP | NY-Analytical | 456 | Identification B | Identification B | Identification B | Passes Test (Done) |  |  |  |  | USP/NF | 7A4OK459/Chem_NY_Salicylic Acid USP | 3 | both
7A4OM002/Chem_NY_ Dissolvine GL-47-S | 7A4OM002 | Chem_NY_ Dissolvine GL-47-S | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411 |  |  | left_only
7A4OM002/Chem_NY_ Dissolvine GL-47-S | 7A4OM002 | Chem_NY_ Dissolvine GL-47-S | NY-Analytical | 398 | pH | pH (1% w/v aqueous solution) | pH (1% w/v aqueous solution) |  |  | FDA,1 | 11 | 11.8 | TM 200 |  |  | left_only
7A4OM002/Chem_NY_ Dissolvine GL-47-S | 7A4OM002 | Chem_NY_ Dissolvine GL-47-S | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH |  |  | left_only
7A4OM002/Chem_NY_ Dissolvine GL-47-S | 7A4OM002 | Chem_NY_ Dissolvine GL-47-S | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Characteristic (Done) |  |  |  |  | TM 208 |  |  | left_only
7A4OM002/Chem_NY_ Dissolvine GL-47-S | 7A4OM002 | Chem_NY_ Dissolvine GL-47-S | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) |  |  |  |  | TM 208 |  |  | left_only
7A4OM002/Chem_NY_ Dissolvine GL-47-S | 7A4OM002 | Chem_NY_ Dissolvine GL-47-S | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM 208 |  |  | left_only
7A4OM002/Chem_NY_ Dissolvine GL-47-S | 7A4OM002 | Chem_NY_ Dissolvine GL-47-S | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 |  |  | left_only
7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP027 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection | 7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 1 | both
7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP027 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | A clear colorless to pale straw colored liquid. (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | 7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 1 | both
7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP027 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 792 | Chemical Inspection | Odor (J&J TM-003188) | Odor | Passes Test (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | 7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 1 | both
7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP027 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | 7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 1 | both
7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP027 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 | 7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 1 | both
7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP027 | MASTER - Microbiology - J&J Chemical Testing) | NY-Microbiology | 524 | J&J Fungal | Fungal | Fungal | 0 | CFU/g | FDA,0 | 0 | 200 | TM 869 | 7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 1 | both
7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP027 | MASTER - Microbiology - J&J Chemical Testing) | NY-Microbiology | 700 | J&J SDA Controls | Negative Control | Negative Control | [5 empty] | TM 869 | 7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 1 | both
7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP027 | MASTER - Microbiology - J&J Chemical Testing) | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | NegativeControl | No Growth TSB |  |  |  |  | TM 869 | 7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 1 | both
7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP027 | MASTER - Microbiology - J&J Chemical Testing) | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | TSA JJ | 0 | cfu/mL | FDA,0 | 0 | 500 | TM 869 | 7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 1 | both
7A4OP065/Chem_NY_ Fragrance Plaisir Gourmand 23 | 7A4OP065 | Chem_NY_ Fragrance Plaisir Gourmand 23 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411 |  |  | left_only
7A4OP065/Chem_NY_ Fragrance Plaisir Gourmand 23 | 7A4OP065 | Chem_NY_ Fragrance Plaisir Gourmand 23 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  |  | left_only
7A4OP065/Chem_NY_ Fragrance Plaisir Gourmand 23 | 7A4OP065 | Chem_NY_ Fragrance Plaisir Gourmand 23 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM 208 |  |  | left_only
7A4OP065/Chem_NY_ Fragrance Plaisir Gourmand 23 | 7A4OP065 | Chem_NY_ Fragrance Plaisir Gourmand 23 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208 |  |  | left_only
7A4OP065/Chem_NY_ Fragrance Plaisir Gourmand 23 | 7A4OP065 | Chem_NY_ Fragrance Plaisir Gourmand 23 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 |  |  | left_only
7A4OP065/Chem_NY_ Fragrance Plaisir Gourmand 23 | 7A4OP065 | Chem_NY_ Fragrance Plaisir Gourmand 23 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Pale Yellow to Yellow (Done) |  |  |  |  | TM 208 |  |  | left_only
7A4OP073/Chem_NY_Fragrance Bali Leia 5 | 7A4OP073 | Chem_NY_Fragrance Bali Leia 5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection | 7A4OP073/Chem_NY_Fragrance Bali Leia 5 | 1 | both
7A4OP073/Chem_NY_Fragrance Bali Leia 5 | 7A4OP073 | Chem_NY_Fragrance Bali Leia 5 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Liquid (Done) |  |  |  |  | TM 208 | 7A4OP073/Chem_NY_Fragrance Bali Leia 5 | 1 | both
7A4OP073/Chem_NY_Fragrance Bali Leia 5 | 7A4OP073 | Chem_NY_Fragrance Bali Leia 5 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Dark yellow to lemon yellow (Done) |  |  |  |  | TM 208 | 7A4OP073/Chem_NY_Fragrance Bali Leia 5 | 1 | both
7A4OP073/Chem_NY_Fragrance Bali Leia 5 | 7A4OP073 | Chem_NY_Fragrance Bali Leia 5 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 | 7A4OP073/Chem_NY_Fragrance Bali Leia 5 | 1 | both
7A4OP073/Chem_NY_Fragrance Bali Leia 5 | 7A4OP073 | Chem_NY_Fragrance Bali Leia 5 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208 | 7A4OP073/Chem_NY_Fragrance Bali Leia 5 | 1 | both
7A4OP073/Chem_NY_Fragrance Bali Leia 5 | 7A4OP073 | Chem_NY_Fragrance Bali Leia 5 | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411 | 7A4OP073/Chem_NY_Fragrance Bali Leia 5 | 1 | both
7A4OP081/Chem_NY_Fragrance SC 2705 LF | 7A4OP081 | Chem_NY_Fragrance SC 2705 LF | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection | 7A4OP081/Chem_NY_Fragrance SC 2705 LF | 2 | both
7A4OP081/Chem_NY_Fragrance SC 2705 LF | 7A4OP081 | Chem_NY_Fragrance SC 2705 LF | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Liquid (Done) |  |  |  |  | TM 208 | 7A4OP081/Chem_NY_Fragrance SC 2705 LF | 2 | both
7A4OP081/Chem_NY_Fragrance SC 2705 LF | 7A4OP081 | Chem_NY_Fragrance SC 2705 LF | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to very light yellow (Done) |  |  |  |  | TM 208 | 7A4OP081/Chem_NY_Fragrance SC 2705 LF | 2 | both
7A4OP081/Chem_NY_Fragrance SC 2705 LF | 7A4OP081 | Chem_NY_Fragrance SC 2705 LF | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 | 7A4OP081/Chem_NY_Fragrance SC 2705 LF | 2 | both
7A4OP081/Chem_NY_Fragrance SC 2705 LF | 7A4OP081 | Chem_NY_Fragrance SC 2705 LF | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208 | 7A4OP081/Chem_NY_Fragrance SC 2705 LF | 2 | both
7A4OP081/Chem_NY_Fragrance SC 2705 LF | 7A4OP081 | Chem_NY_Fragrance SC 2705 LF | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411 | 7A4OP081/Chem_NY_Fragrance SC 2705 LF | 2 | both
7A4OP093/Chem_NY_Essential Oil 27A Fragrance | 7A4OP093 | Chem_NY_Essential Oil 27A Fragrance | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection | 7A4OP093/Chem_NY_Essential Oil 27A Fragrance | 1 | both
7A4OP093/Chem_NY_Essential Oil 27A Fragrance | 7A4OP093 | Chem_NY_Essential Oil 27A Fragrance | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Liquid (Done) |  |  |  |  | TM 208 | 7A4OP093/Chem_NY_Essential Oil 27A Fragrance | 1 | both
7A4OP093/Chem_NY_Essential Oil 27A Fragrance | 7A4OP093 | Chem_NY_Essential Oil 27A Fragrance | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Amber (Done) |  |  |  |  | TM 208 | 7A4OP093/Chem_NY_Essential Oil 27A Fragrance | 1 | both
7A4OP093/Chem_NY_Essential Oil 27A Fragrance | 7A4OP093 | Chem_NY_Essential Oil 27A Fragrance | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 | 7A4OP093/Chem_NY_Essential Oil 27A Fragrance | 1 | both
7A4OP093/Chem_NY_Essential Oil 27A Fragrance | 7A4OP093 | Chem_NY_Essential Oil 27A Fragrance | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208 | 7A4OP093/Chem_NY_Essential Oil 27A Fragrance | 1 | both
7A4OP093/Chem_NY_Essential Oil 27A Fragrance | 7A4OP093 | Chem_NY_Essential Oil 27A Fragrance | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411 | 7A4OP093/Chem_NY_Essential Oil 27A Fragrance | 1 | both
7A4OP113/Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | 7A4OP113 | Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | NY-Analytical | 398 | pH | pH | pH |  |  | FDA,1 | 2.5 | 6 | TM 200 |  |  | left_only
7A4OP113/Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | 7A4OP113 | Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH |  |  | left_only
7A4OP113/Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | 7A4OP113 | Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) |  |  |  |  | TM 208 |  |  | left_only
7A4OP113/Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | 7A4OP113 | Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear to slightly cloudy liquid (Done) |  |  |  |  | TM 208 |  |  | left_only
7A4OP113/Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | 7A4OP113 | Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM 208 |  |  | left_only
7A4OP113/Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | 7A4OP113 | Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 |  |  | left_only
7A4OP197/Chem_NY_D-Panthenol 75W_01 | 7A4OP197 | Chem_NY_D-Panthenol 75W_01 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH | 7A4OP197/Chem_NY_D-Panthenol 75W_01 | 12 | both
7A4OP197/Chem_NY_D-Panthenol 75W_01 | 7A4OP197 | Chem_NY_D-Panthenol 75W_01 | NY-Analytical | 398 | pH | pH | pH | 6.96 |  | FDA,1 | 6 | 7.5 | TM 200 | 7A4OP197/Chem_NY_D-Panthenol 75W_01 | 12 | both
7A4OP197/Chem_NY_D-Panthenol 75W_01 | 7A4OP197 | Chem_NY_D-Panthenol 75W_01 | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.466 |  | FDA,3 | 1.466 | 1.475 | TM 204 | 7A4OP197/Chem_NY_D-Panthenol 75W_01 | 12 | both
7A4OP197/Chem_NY_D-Panthenol 75W_01 | 7A4OP197 | Chem_NY_D-Panthenol 75W_01 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM 208 | 7A4OP197/Chem_NY_D-Panthenol 75W_01 | 12 | both
7A4OP197/Chem_NY_D-Panthenol 75W_01 | 7A4OP197 | Chem_NY_D-Panthenol 75W_01 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Slightly Yellowish or Colorless (Done) |  |  |  |  | TM 208 | 7A4OP197/Chem_NY_D-Panthenol 75W_01 | 12 | both
7A4OP197/Chem_NY_D-Panthenol 75W_01 | 7A4OP197 | Chem_NY_D-Panthenol 75W_01 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 | 7A4OP197/Chem_NY_D-Panthenol 75W_01 | 12 | both
7A4OP197/Chem_NY_D-Panthenol 75W_01 | 7A4OP197 | Chem_NY_D-Panthenol 75W_01 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208 | 7A4OP197/Chem_NY_D-Panthenol 75W_01 | 12 | both
7A4OP197/Chem_NY_D-Panthenol 75W_01 | 7A4OP197 | Chem_NY_D-Panthenol 75W_01 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411 | 7A4OP197/Chem_NY_D-Panthenol 75W_01 | 12 | both
7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | 7A4OP205 | Chem_NY_Cucumis Sativus Fruit Extract | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411 |  |  | left_only
7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | 7A4OP205 | Chem_NY_Cucumis Sativus Fruit Extract | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  |  | left_only
7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | 7A4OP205 | Chem_NY_Cucumis Sativus Fruit Extract | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram) | TAMC (Average per Gram) |  | CFU/g | FDA,1 |  | 100 | TM 610 |  |  | left_only
7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | 7A4OP205 | Chem_NY_Cucumis Sativus Fruit Extract | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) |  | cfu | FDA,0 |  |  | TM 610 |  |  | left_only
7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | 7A4OP205 | Chem_NY_Cucumis Sativus Fruit Extract | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram) | TMYC (Average per Gram) |  | CFU/g | FDA,1 |  | 100 | TM 610 |  |  | left_only
7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | 7A4OP205 | Chem_NY_Cucumis Sativus Fruit Extract | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  | cfu | FDA,0 |  |  | TM 610 |  |  | left_only
7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | 7A4OP205 | Chem_NY_Cucumis Sativus Fruit Extract | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to pale yellow (Done) |  |  |  |  | TM 208 |  |  | left_only
7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | 7A4OP205 | Chem_NY_Cucumis Sativus Fruit Extract | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208 |  |  | left_only
7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | 7A4OP205 | Chem_NY_Cucumis Sativus Fruit Extract | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 |  |  | left_only
7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | 7A4OP205 | Chem_NY_Cucumis Sativus Fruit Extract | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Transparent Liquid to slightly Turbid (Done) |  |  |  |  | TM 208 |  |  | left_only
7A4OP345/Chem_195A | 7A4OP345 | Chem_195A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 |  |  | left_only
7A4OP345/Chem_195A | 7A4OP345 | Chem_195A | NY-Analytical | 447 | pH as is | pH | pH |  |  | FDA,1 | 3.5 | 4.5 | TM 200 |  |  | left_only
7A4OP345/Chem_195A | 7A4OP345 | Chem_195A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH as is |  |  | left_only
7A4OP345/Chem_195A | 7A4OP345 | Chem_195A | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram) | TAMC (Average per Gram) |  | CFU/g | FDA,1 |  |  | TM 610 |  |  | left_only
7A4OP345/Chem_195A | 7A4OP345 | Chem_195A | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) |  | cfu | FDA,0 |  |  | TM 610 |  |  | left_only
7A4OP345/Chem_195A | 7A4OP345 | Chem_195A | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram) | TMYC (Average per Gram) |  | CFU/g | FDA,1 |  |  | TM 610 |  |  | left_only
7A4OP345/Chem_195A | 7A4OP345 | Chem_195A | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  | cfu | FDA,0 |  |  | TM 610 |  |  | left_only
7A4OP345/Chem_195A | 7A4OP345 | Chem_195A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7A4OP345/Chem_195A | 7A4OP345 | Chem_195A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Hazy liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7A4OP345/Chem_195A | 7A4OP345 | Chem_195A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Light green/amber (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7A4OP345/Chem_195A | 7A4OP345 | Chem_195A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7A4OP448/Chem_208A | 7A4OP448 | Chem_208A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis | 7A4OP448/Chem_208A | 2 | both
7A4OP448/Chem_197A | 7A4OP448 | Chem_197A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH | 7A4OP448/Chem_197A | 1 | both
7A4OP448/Chem_197A | 7A4OP448 | Chem_197A | NY-Analytical | 398 | pH | pH (10% in water) | pH (10% in water) | 5.9 |  | FDA,1 | 9 | 11 | TM 200 | 7A4OP448/Chem_197A | 1 | both
7A4OP448/Chem_208A | 7A4OP448 | Chem_208A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.9801 |  | FDA,4 | 0.9694 | 0.9894 | TM 202 | 7A4OP448/Chem_208A | 2 | both
7A4OP448/Chem_208A | 7A4OP448 | Chem_208A | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.469 |  | FDA,4 | 1.4655 | 1.4855 | TM 204 | 7A4OP448/Chem_208A | 2 | both
7A4OP448/Chem_197A | 7A4OP448 | Chem_197A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7A4OP448/Chem_197A | 1 | both
7A4OP448/Chem_208A | 7A4OP448 | Chem_208A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7A4OP448/Chem_208A | 2 | both
7A4OP448/Chem_197A | 7A4OP448 | Chem_197A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OP448/Chem_197A | 1 | both
7A4OP448/Chem_208A | 7A4OP448 | Chem_208A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear to pale yellow liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OP448/Chem_208A | 2 | both
7A4OP448/Chem_197A | 7A4OP448 | Chem_197A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OP448/Chem_197A | 1 | both
7A4OP448/Chem_197A | 7A4OP448 | Chem_197A | NY-Analytical | 792 | Chemical Inspection | Color (TM 230) | Color |  |  | FDA,1 | 0 | 3 | TM208/220/230/USP/Kao LM | 7A4OP448/Chem_197A | 1 | both
7A4OP448/Chem_197A | 7A4OP448 | Chem_197A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OP448/Chem_197A | 1 | both
7A4OP448/Chem_208A | 7A4OP448 | Chem_208A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OP448/Chem_208A | 2 | both
7A4OP448/Chem_197A | 7A4OP448 | Chem_197A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Fatty odor (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OP448/Chem_197A | 1 | both
7A4OP448/Chem_208A | 7A4OP448 | Chem_208A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Fresh Bouquet (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OP448/Chem_208A | 2 | both
7A4OP567/Chem_NY_Surfonic L12-6_Surfactant | 7A4OP567 | Chem_NY_Surfonic L12-6_Surfactant | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA | 7A4OP567/Chem_NY_Surfonic L12-6_Surfactant | 5 | both
7A4OP567/Chem_NY_Surfonic L12-6_Surfactant | 7A4OP567 | Chem_NY_Surfonic L12-6_Surfactant | NY-Analytical | 398 | pH | pH (1%) in 10:6 IPA: H2O | pH (1%) in 10:6 IPA: H2O | 6.9 |  | FDA,1 | 6.5 | 8 | TM 200 | 7A4OP567/Chem_NY_Surfonic L12-6_Surfactant | 5 | both
7A4OP567/Chem_NY_Surfonic L12-6_Surfactant | 7A4OP567 | Chem_NY_Surfonic L12-6_Surfactant | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear to slightly turbid, free of suspended matter (Done) |  |  |  |  | TM 208 | 7A4OP567/Chem_NY_Surfonic L12-6_Surfactant | 5 | both
7A4OP567/Chem_NY_Surfonic L12-6_Surfactant | 7A4OP567 | Chem_NY_Surfonic L12-6_Surfactant | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to Slightly Yellow (Done) |  |  |  |  | TM 208 | 7A4OP567/Chem_NY_Surfonic L12-6_Surfactant | 5 | both
7A4OP567/Chem_NY_Surfonic L12-6_Surfactant | 7A4OP567 | Chem_NY_Surfonic L12-6_Surfactant | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 | 7A4OP567/Chem_NY_Surfonic L12-6_Surfactant | 5 | both
7A4OP567/Chem_NY_Surfonic L12-6_Surfactant | 7A4OP567 | Chem_NY_Surfonic L12-6_Surfactant | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | To match standard (Done) |  |  |  |  | TM 208 | 7A4OP567/Chem_NY_Surfonic L12-6_Surfactant | 5 | both
7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | 7A4OP601 | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 |  |  | left_only
7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | 7A4OP601 | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 398 | pH | pH (5% Solution in DI Water:Methanol 1:1) | pH (5% Solution in DI Water:Methanol 1:1) |  |  | FDA,1 | 6.5 | 7.5 | J&J TMWW78014 |  |  | left_only
7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | 7A4OP601 | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Iodine Value |  |  | left_only
7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | 7A4OP601 | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Light yellow pasty fluid (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | 7A4OP601 | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | 7A4OP601 | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 801 | % Water | Water Content | Water Content |  | % | FDA,2 | 33 | 37 | J&J TMWW78017 |  |  | left_only
7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | 7A4OP601 | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 810 | Saponification Value | Saponification Value Final | Saponification Value Final |  | KOH/g | FDA,1 | 47 | 57 | J&J TMWW78016 |  |  | left_only
7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | 7A4OP601 | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 810 | Saponification Value | Saponification Value | Saponification Value |  | KOH/g | FDA,1 | 47 | 57 | J&J TMWW78016 |  |  | left_only
7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | 7A4OP601 | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 828 | Iodine Value | Iodine Value | Iodine Value |  | g I/100g | FDA,2 |  | 5 | J&J TM-003029 |  |  | left_only
7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | 7A4OP601 | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 828 | Iodine Value | Iodine Value Final | Iodine Value Final |  | g I/100g | FDA,2 |  | 5 | J&J TM-003029 |  |  | left_only
7A4OP806/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP806 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C |  |  | FDA,4 | 1.11 | 1.12 | J&J TMWW78012 |  |  | left_only
7A4OP806/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP806 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7A4OP806/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP806 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C |  |  | FDA,4 | 1.386 | 1.396 | J&J TMWW78015 |  |  | left_only
7A4OP806/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP806 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 825 | Limit of Ethylene Glycol and Diethylene Glycol | Diethylene Glycol | Diethylene Glycol |  | % | FDA,2 |  | 0.1 | USP |  |  | left_only
7A4OP806/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP806 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 825 | Limit of Ethylene Glycol and Diethylene Glycol | Ethylene Glycol | Ethylene Glycol |  | % | FDA,2 |  | 0.1 | USP |  |  | left_only
7A4OQ028/Chem_NY_Emulgin VL-75 | 7A4OQ028 | Chem_NY_Emulgin VL-75 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH as is | 7A4OQ028/Chem_NY_Emulgin VL-75 | 3 | both
7A4OQ028/Chem_NY_Emulgin VL-75 | 7A4OQ028 | Chem_NY_Emulgin VL-75 | NY-Analytical | 447 | pH as is | pH | pH | 5.76 |  | FDA,1 | 5 | 6 | TM 200 | 7A4OQ028/Chem_NY_Emulgin VL-75 | 3 | both
7A4OQ028/Chem_NY_Emulgin VL-75 | 7A4OQ028 | Chem_NY_Emulgin VL-75 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7A4OQ028/Chem_NY_Emulgin VL-75 | 3 | both
7A4OQ028/Chem_NY_Emulgin VL-75 | 7A4OQ028 | Chem_NY_Emulgin VL-75 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear to slightly Turbid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OQ028/Chem_NY_Emulgin VL-75 | 3 | both
7A4OQ028/Chem_NY_Emulgin VL-75 | 7A4OQ028 | Chem_NY_Emulgin VL-75 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Yellowish Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OQ028/Chem_NY_Emulgin VL-75 | 3 | both
7A4OQ028/Chem_NY_Emulgin VL-75 | 7A4OQ028 | Chem_NY_Emulgin VL-75 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OQ028/Chem_NY_Emulgin VL-75 | 3 | both
7A4OQ028/Chem_NY_Emulgin VL-75 | 7A4OQ028 | Chem_NY_Emulgin VL-75 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OQ028/Chem_NY_Emulgin VL-75 | 3 | both
7A4OQ084/Chem_23A | 7A4OQ084 | Chem_23A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification | 7A4OQ084/Chem_23A | 14 | both
7A4OQ084/Chem_23A | 7A4OQ084 | Chem_23A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7A4OQ084/Chem_23A | 14 | both
7A4OQ084/Chem_23A | 7A4OQ084 | Chem_23A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OQ084/Chem_23A | 14 | both
7A4OQ084/Chem_23A | 7A4OQ084 | Chem_23A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | None or Pale Straw (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OQ084/Chem_23A | 14 | both
7A4OQ084/Chem_23A | 7A4OQ084 | Chem_23A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OQ084/Chem_23A | 14 | both
7A4OQ084/Chem_23A | 7A4OQ084 | Chem_23A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Little or no distinguisable odor (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OQ084/Chem_23A | 14 | both
7A4OQ132/Chem_25A | 7A4OQ132 | Chem_25A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification | 7A4OQ132/Chem_25A | 5 | both
7A4OQ132/Chem_25A | 7A4OQ132 | Chem_25A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7A4OQ132/Chem_25A | 5 | both
7A4OQ132/Chem_25A | 7A4OQ132 | Chem_25A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Transparent liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OQ132/Chem_25A | 5 | both
7A4OQ132/Chem_25A | 7A4OQ132 | Chem_25A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to pale yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OQ132/Chem_25A | 5 | both
7A4OQ132/Chem_25A | 7A4OQ132 | Chem_25A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OQ132/Chem_25A | 5 | both
7A4OQ132/Chem_25A | 7A4OQ132 | Chem_25A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Mild (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OQ132/Chem_25A | 5 | both
7A4OQ203/MASTER - Microbiology - J&J Chemical Testing) | 7A4OQ203 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 747 | Viscosity | Viscosity (Spindle 7@20RPM@30 sec 0.5% solution) | Viscosity (Spindle 7@20RPM@30 sec 0.5% solution) |  | cps | FDA,0 | 45000 | 65000 | J&J TM7423 |  |  | left_only
7A4OQ203/MASTER - Microbiology - J&J Chemical Testing) | 7A4OQ203 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 802 | Loss on Drying | Loss on Drying | Loss on Drying |  | % | FDA,2 |  | 2 | J&J TM7221 |  |  | left_only
7A4OQ203/MASTER - Microbiology - J&J Chemical Testing) | 7A4OQ203 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 802 | Loss on Drying | Loss on Drying Final | Loss on Drying Final |  | % | FDA,2 |  | 2 | J&J TM7221 |  |  | left_only
7A4OQ229/Chem_Trisodium Citrate Dihydrate USP | 7A4OQ229 | Chem_Trisodium Citrate Dihydrate USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA | 7A4OQ229/Chem_Trisodium Citrate Dihydrate USP | 16 | both
7A4OQ229/Chem_NY_Witch Hazel | 7A4OQ229 | Chem_NY_Witch Hazel | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C | 7A4OQ229/Chem_NY_Witch Hazel | 1 | both
7A4OQ229/Chem_NY_Witch Hazel | 7A4OQ229 | Chem_NY_Witch Hazel | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.981 |  | FDA,3 | 0.979 | 0.983 | TM 202 | 7A4OQ229/Chem_NY_Witch Hazel | 1 | both
7A4OQ229/Chem_Trisodium Citrate Dihydrate USP | 7A4OQ229 | Chem_Trisodium Citrate Dihydrate USP | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Fine granular (Done) |  |  |  |  | TM 208 | 7A4OQ229/Chem_Trisodium Citrate Dihydrate USP | 16 | both
7A4OQ229/Chem_Trisodium Citrate Dihydrate USP | 7A4OQ229 | Chem_Trisodium Citrate Dihydrate USP | NY-Analytical | 792 | Chemical Inspection | Color | Color | White (Done) |  |  |  |  | TM 208 | 7A4OQ229/Chem_Trisodium Citrate Dihydrate USP | 16 | both
7A4OQ229/Chem_Trisodium Citrate Dihydrate USP | 7A4OQ229 | Chem_Trisodium Citrate Dihydrate USP | NY-Analytical | 792 | Chemical Inspection | Contamination | Contamination | None (Done) |  |  |  |  | TM 208 | 7A4OQ229/Chem_Trisodium Citrate Dihydrate USP | 16 | both
7A4OQ229/Chem_Trisodium Citrate Dihydrate USP | 7A4OQ229 | Chem_Trisodium Citrate Dihydrate USP | NY-Analytical | 792 | Chemical Inspection | Odor | Odor | Odorless (Done) |  |  |  |  | TM 208 | 7A4OQ229/Chem_Trisodium Citrate Dihydrate USP | 16 | both
7A4OQ229/Chem_Trisodium Citrate Dihydrate USP | 7A4OQ229 | Chem_Trisodium Citrate Dihydrate USP | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7A4OQ229/Chem_Trisodium Citrate Dihydrate USP | 16 | both
7A4OQ229/Chem_NY_Witch Hazel | 7A4OQ229 | Chem_NY_Witch Hazel | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OQ229/Chem_NY_Witch Hazel | 1 | both
7A4OQ229/Chem_NY_Witch Hazel | 7A4OQ229 | Chem_NY_Witch Hazel | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OQ229/Chem_NY_Witch Hazel | 1 | both
7A4OQ229/Chem_NY_Witch Hazel | 7A4OQ229 | Chem_NY_Witch Hazel | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OQ229/Chem_NY_Witch Hazel | 1 | both
7A4OQ229/Chem_NY_Witch Hazel | 7A4OQ229 | Chem_NY_Witch Hazel | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic witch hazel, free from acetous odor (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OQ229/Chem_NY_Witch Hazel | 1 | both
7A4OQ229/Chem_NY_Witch Hazel | 7A4OQ229 | Chem_NY_Witch Hazel | NY-Analytical | 797 | Nonvolatile Residue | End Weight (g) | End Weight | 240.3127 | g | FDA,0 |  |  | USP | 7A4OQ229/Chem_NY_Witch Hazel | 1 | both
7A4OQ229/Chem_NY_Witch Hazel | 7A4OQ229 | Chem_NY_Witch Hazel | NY-Analytical | 797 | Nonvolatile Residue | Start Weight (g) | Start Weight | 240.3107 | g | FDA,0 |  |  | USP | 7A4OQ229/Chem_NY_Witch Hazel | 1 | both
7A4OQ229/Chem_NY_Witch Hazel | 7A4OQ229 | Chem_NY_Witch Hazel | NY-Analytical | 797 | Nonvolatile Residue | Total non-volatile reside (calc) | Total non-volatile reside | 2 | mg/100 mL | FDA,0 |  | 25 | USP | 7A4OQ229/Chem_NY_Witch Hazel | 1 | both
7A4OQ229/Chem_NY_Witch Hazel | 7A4OQ229 | Chem_NY_Witch Hazel | NY-Analytical | 447 | pH as is | pH | pH | 3.64 |  | FDA,1 | 3 | 5 | USP <791> | 7A4OQ229/Chem_NY_Witch Hazel | 1 | both
7A4OQ260/Chem_29A | 7A4OQ260 | Chem_29A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 |  |  | left_only
7A4OQ260/Chem_29A | 7A4OQ260 | Chem_29A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C |  |  | FDA,3 | 0.954 | 0.974 | TM 202 |  |  | left_only
7A4OQ260/Chem_29A | 7A4OQ260 | Chem_29A | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C |  |  | FDA,3 | 1.464 | 1.484 | TM 204 |  |  | left_only
7A4OQ260/Chem_29A | 7A4OQ260 | Chem_29A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C |  |  | left_only
7A4OQ260/Chem_29A | 7A4OQ260 | Chem_29A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7A4OQ260/Chem_29A | 7A4OQ260 | Chem_29A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to pale yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7A4OQ260/Chem_29A | 7A4OQ260 | Chem_29A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7A4OQ260/Chem_29A | 7A4OQ260 | Chem_29A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Powder Rose (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7A4OQ518/Chem_31A | 7A4OQ518 | Chem_31A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH | 7A4OQ518/Chem_31A | 24 | both
7A4OQ518/Chem_31A | 7A4OQ518 | Chem_31A | NY-Analytical | 398 | pH | pH (10% Solution) | pH (10% Solution) | 7.57 |  | FDA,1 | 6 | 9 | TM 200 | 7A4OQ518/Chem_31A | 24 | both
7A4OQ518/Chem_31A | 7A4OQ518 | Chem_31A | NY-Analytical | 796 | Assay | Assay | Assay | 51.97 | % | FDA,1 | 50 | 54 | TM 301/434 | 7A4OQ518/Chem_31A | 24 | both
7A4OQ518/Chem_31A | 7A4OQ518 | Chem_31A | NY-Analytical | 796 | Assay | Assay Final | Assay Final | 51.98 | % | FDA,1 | 50 | 54 | TM 301/434 | 7A4OQ518/Chem_31A | 24 | both
7A4OQ518/Chem_31A | 7A4OQ518 | Chem_31A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear viscous liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OQ518/Chem_31A | 24 | both
7A4OQ518/Chem_31A | 7A4OQ518 | Chem_31A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to light straw (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OQ518/Chem_31A | 24 | both
7A4OQ518/Chem_31A | 7A4OQ518 | Chem_31A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OQ518/Chem_31A | 24 | both
7A4OQ518/Chem_31A | 7A4OQ518 | Chem_31A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Pleasant almond (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OQ518/Chem_31A | 24 | both
7A4OQ890/Chem_34A | 7A4OQ890 | Chem_34A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection | 7A4OQ890/Chem_34A | 6 | both
7A4OQ890/Chem_34A | 7A4OQ890 | Chem_34A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Granular Powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OQ890/Chem_34A | 6 | both
7A4OQ890/Chem_34A | 7A4OQ890 | Chem_34A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OQ890/Chem_34A | 6 | both
7A4OQ890/Chem_34A | 7A4OQ890 | Chem_34A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OQ890/Chem_34A | 6 | both
7A4OQ890/Chem_34A | 7A4OQ890 | Chem_34A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OQ890/Chem_34A | 6 | both
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection | 7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 14 | both
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 792 | Chemical Inspection | Description | Description | Transparent, colorless, mobile, volatile liquid (Done) |  |  |  |  | ORGANOLEPTIC | 7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 14 | both
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 795 | Acidity | Acidity (calc) | Acidity | 0.4 | mL | FDA,2 |  | 0.7 | USP | 7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 14 | both
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 795 | Acidity | Acidity Final (Calc) | Acidity Final (Calc) | 0.45 | mL | FDA,2 |  | 0.7 | USP | 7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 14 | both
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 796 | Assay | Assay | Assay | 100 | % | FDA,1 | 99 |  | USP | 7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 14 | both
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 796 | Assay | Assay Final | Assay Final | 99.975 | % | FDA,1 | 99 |  | USP | 7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 14 | both
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 797 | Nonvolatile Residue | End Weight (g) | End Weight | 90.0835 | g | FDA,4 |  |  | USP | 7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 14 | both
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 797 | Nonvolatile Residue | Nonvolatile Residue (calc) | Nonvolatile Residue | 0 | % | FDA,3 |  | 0.005 | USP | 7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 14 | both
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 797 | Nonvolatile Residue | Nonvolatile Residue Final (calc) | Nonvolatile Residue Final | 0.0008 | % | FDA,3 |  | 0.005 | USP | 7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 14 | both
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.376 |  | FDA,3 | 1.376 | 1.378 | USP | 7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 14 | both
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 797 | Nonvolatile Residue | Start Weight (g) | Start Weight | 90.7647 | g | FDA,4 |  |  | USP | 7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 14 | both
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 795 | Acidity | Titrant End Volume (mL) | Titrant End Volume | 0.9 | mL | FDA,2 |  |  | USP | 7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 14 | both
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 795 | Acidity | Titrant Start Volume (mL) | Titrant Start Volume | 0.5 | mL | FDA,2 |  |  | USP | 7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 14 | both
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 797 | Nonvolatile Residue | Weight of Residue (g) | Weight of Residue | 0.0004 | g | FDA,4 |  |  | USP | 7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 14 | both
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR conform to standard (Done) |  |  |  |  | USP <197F> | 7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 14 | both
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.7832 |  | FDA,3 | 0.783 | 0.787 | USP/NF <841> | 7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 14 | both
7A4OS411/Chem_36A | 7A4OS411 | Chem_36A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection | 7A4OS411/Chem_36A | 4 | both
7A4OS411/Chem_36A | 7A4OS411 | Chem_36A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Free Flowing Powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OS411/Chem_36A | 4 | both
7A4OS411/Chem_36A | 7A4OS411 | Chem_36A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Brownish-Red (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OS411/Chem_36A | 4 | both
7A4OS411/Chem_36A | 7A4OS411 | Chem_36A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OS411/Chem_36A | 4 | both
7A4OS411/Chem_36A | 7A4OS411 | Chem_36A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7A4OS411/Chem_36A | 4 | both
7A4OS411/Chem_36A | 7A4OS411 | Chem_36A | NY-Analytical | 796 | Assay | Assay as is (Without Drying) | Assay as is (Without Drying) | 10.17 | % | FDA,1 |  |  | USP | 7A4OS411/Chem_36A | 4 | both
7A4OS411/Chem_36A | 7A4OS411 | Chem_36A | NY-Analytical | 796 | Assay | Assay as is Final (Without Drying) | Assay as is Final (Without Drying) | 10.15 | % | FDA,1 |  |  | USP | 7A4OS411/Chem_36A | 4 | both
7A4OS411/Chem_36A | 7A4OS411 | Chem_36A | NY-Analytical | 796 | Assay | Assay on Dried Basis | Assay on Dried Basis | 10.59 | % | FDA,1 | 9 | 12 | USP | 7A4OS411/Chem_36A | 4 | both
7A4OS411/Chem_36A | 7A4OS411 | Chem_36A | NY-Analytical | 796 | Assay | Assay on Dried Basis Final | Assay on Dried Basis Final | 10.54 | % | FDA,1 | 9 | 12 | USP | 7A4OS411/Chem_36A | 4 | both
7A4OS411/Chem_36A | 7A4OS411 | Chem_36A | NY-Analytical | 802 | Loss on Drying | Loss on Drying | Loss on Drying | 5.55 | % | FDA,1 |  | 8 | USP | 7A4OS411/Chem_36A | 4 | both
7A4OS411/Chem_36A | 7A4OS411 | Chem_36A | NY-Analytical | 802 | Loss on Drying | Loss on Drying Final | Loss on Drying Final | 5.55 | % | FDA,1 |  | 8 | USP | 7A4OS411/Chem_36A | 4 | both
7A4OU351/Chem_NY_Dehydrated Alcohol USP | 7A4OU351 | Chem_NY_Dehydrated Alcohol USP | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Compare to reference Standard (Done) |  |  |  |  | TM 411 |  |  | left_only
7A4OU351/Chem_NY_Dehydrated Alcohol USP | 7A4OU351 | Chem_NY_Dehydrated Alcohol USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis (ELLI) | Certificate of Analysis | COA Meets Specification (Done) |  |  |  |  | COA |  |  | left_only
7A4OU351/Chem_NY_Dehydrated Alcohol USP | 7A4OU351 | Chem_NY_Dehydrated Alcohol USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA |  |  | left_only
7A4OU351/Chem_NY_Dehydrated Alcohol USP | 7A4OU351 | Chem_NY_Dehydrated Alcohol USP | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Alcohol (Done) |  |  |  |  | TM 208 |  |  | left_only
7A4OU351/Chem_NY_Dehydrated Alcohol USP | 7A4OU351 | Chem_NY_Dehydrated Alcohol USP | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM 208 |  |  | left_only
7A4OU351/Chem_NY_Dehydrated Alcohol USP | 7A4OU351 | Chem_NY_Dehydrated Alcohol USP | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Clear to Colorless and Match standard (Done) |  |  |  |  | TM 208 |  |  | left_only
7A4OU351/Chem_NY_Dehydrated Alcohol USP | 7A4OU351 | Chem_NY_Dehydrated Alcohol USP | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 |  |  | left_only
7A4OU351/Chem_NY_Dehydrated Alcohol USP | 7A4OU351 | Chem_NY_Dehydrated Alcohol USP | NY-Analytical | 822 | Specific Gravity @ 15.56°C/ 60° F | Specific Gravity @ 15.56°C | Specific Gravity at 15.56°C |  |  | FDA,4 |  | 0.7962 | TM 202 |  |  | left_only
7A4OU555/Chem_NY_Phenylephrine Reduced | 7A4OU555 | Chem_NY_Phenylephrine Reduced | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis | 7A4OU555/Chem_NY_Phenylephrine Reduced | 2 | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection | 7A4OU555/Chem_NY_Phenylephrine HCL | 2 | both
7A4OU555/Chem_NY_Phenylephrine Reduced | 7A4OU555 | Chem_NY_Phenylephrine Reduced | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Crystalline Powder or Crystals (Done) |  |  |  |  | TM 208 | 7A4OU555/Chem_NY_Phenylephrine Reduced | 2 | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Crystalline Powder or Crystals (Done) |  |  |  |  | TM 208 | 7A4OU555/Chem_NY_Phenylephrine HCL | 2 | both
7A4OU555/Chem_NY_Phenylephrine Reduced | 7A4OU555 | Chem_NY_Phenylephrine Reduced | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White to almost white (Done) |  |  |  |  | TM 208 | 7A4OU555/Chem_NY_Phenylephrine Reduced | 2 | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White to almost white (Done) |  |  |  |  | TM 208 | 7A4OU555/Chem_NY_Phenylephrine HCL | 2 | both
7A4OU555/Chem_NY_Phenylephrine Reduced | 7A4OU555 | Chem_NY_Phenylephrine Reduced | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 | 7A4OU555/Chem_NY_Phenylephrine Reduced | 2 | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 | 7A4OU555/Chem_NY_Phenylephrine HCL | 2 | both
7A4OU555/Chem_NY_Phenylephrine Reduced | 7A4OU555 | Chem_NY_Phenylephrine Reduced | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208 | 7A4OU555/Chem_NY_Phenylephrine Reduced | 2 | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208 | 7A4OU555/Chem_NY_Phenylephrine HCL | 2 | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7A4OU555/Chem_NY_Phenylephrine HCL | 2 | both
7A4OU555/Chem_NY_Phenylephrine Reduced | 7A4OU555 | Chem_NY_Phenylephrine Reduced | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7A4OU555/Chem_NY_Phenylephrine Reduced | 2 | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 812 | Organic Impurites | Any individual unspecified impuritv | Any individual unspecified impuritv |  | % | FDA,2 |  | 0.1 | USP | 7A4OU555/Chem_NY_Phenylephrine HCL | 2 | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 796 | Assay | Assay | Assay | 99.4 | % | FDA,1 | 98 | 102 | USP | 7A4OU555/Chem_NY_Phenylephrine HCL | 2 | both
7A4OU555/Chem_NY_Phenylephrine Reduced | 7A4OU555 | Chem_NY_Phenylephrine Reduced | NY-Analytical | 796 | Assay | Assay | Assay | 99.9 | % | FDA,1 | 98 | 102 | USP | 7A4OU555/Chem_NY_Phenylephrine Reduced | 2 | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 802 | Loss on Drying | Loss on Drying | Loss on Drying |  | % | FDA,1 |  | 1 | USP | 7A4OU555/Chem_NY_Phenylephrine HCL | 2 | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 802 | Loss on Drying | Loss on Drying Final | Loss on Drying Final |  | % | FDA,1 |  | 1 | USP | 7A4OU555/Chem_NY_Phenylephrine HCL | 2 | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 812 | Organic Impurites | Norphenylephrine | Norphenylephrine |  | % | FDA,2 |  | 0.1 | USP | 7A4OU555/Chem_NY_Phenylephrine HCL | 2 | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 1043 | Optical Rotation | Optical Rotation | Optical Rotation |  | ° | FDA,0 | -47 | -43 | USP | 7A4OU555/Chem_NY_Phenylephrine HCL | 2 | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 812 | Organic Impurites | Phenylephrine Related Compound C | Phenylephrine Related Compound C |  | % | FDA,2 |  | 0.1 | USP | 7A4OU555/Chem_NY_Phenylephrine HCL | 2 | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 812 | Organic Impurites | Phenylephrine Related Compound D | Phenylephrine Related Compound D |  | % | FDA,2 |  | 0.1 | USP | 7A4OU555/Chem_NY_Phenylephrine HCL | 2 | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 812 | Organic Impurites | Phenylephrine Related Compound E | Phenylephrine Related Compound E |  | % | FDA,2 |  | 0.1 | USP | 7A4OU555/Chem_NY_Phenylephrine HCL | 2 | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 834 | Residue on Ignition | Residue on Ignition | Residue on Ignition |  | % | FDA,1 |  | 0.2 | USP | 7A4OU555/Chem_NY_Phenylephrine HCL | 2 | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 835 | Sulfate Content | Sulfate | Sulfate |  | % | FDA,2 |  | 0.2 | USP | 7A4OU555/Chem_NY_Phenylephrine HCL | 2 | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 812 | Organic Impurites | Total Impurities | Total Impurities |  | % | FDA,2 |  | 0.2 | USP | 7A4OU555/Chem_NY_Phenylephrine HCL | 2 | both
7A4OX110/Chem_39A | 7A4OX110 | Chem_39A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification | 7A4OX110/Chem_39A | 3 | both
7A4OX110/Chem_39A | 7A4OX110 | Chem_39A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | [5 empty] | TM 411 | 7A4OX110/Chem_39A | 3 | both
7A4OX110/Chem_39A | 7A4OX110 | Chem_39A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | [5 empty] | TM208/220/230/USP/Kao LM | 7A4OX110/Chem_39A | 3 | both
7A4OX110/Chem_39A | 7A4OX110 | Chem_39A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | [5 empty] | TM208/220/230/USP/Kao LM | 7A4OX110/Chem_39A | 3 | both
7A4OX110/Chem_39A | 7A4OX110 | Chem_39A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | [5 empty] | TM208/220/230/USP/Kao LM | 7A4OX110/Chem_39A | 3 | both
7A4OX110/Chem_39A | 7A4OX110 | Chem_39A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | [5 empty] | TM208/220/230/USP/Kao LM | 7A4OX110/Chem_39A | 3 | both
7A4OX231/Chem_NY_D&C Yellow 10 | 7A4OX231 | Chem_NY_D&C Yellow 10 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Appearance | 7A4OX231/Chem_NY_D&C Yellow 10 | 2 | both
7A4OX231/Chem_NY_D&C Yellow 10 | 7A4OX231 | Chem_NY_D&C Yellow 10 | NY-Analytical | 397 | Appearance | Physical Properties | Physical Properties | Yellow Powder (Done) |  |  |  |  | TM 208 | 7A4OX231/Chem_NY_D&C Yellow 10 | 2 | both
7A4OX231/Chem_NY_D&C Yellow 10 | 7A4OX231 | Chem_NY_D&C Yellow 10 | NY-Analytical | 390 | Identification | Spectrophotometric | Spectrophotometric | 410.88 | λ | FDA,0 | 408 | 416 | TM 3014 | 7A4OX231/Chem_NY_D&C Yellow 10 | 2 | both
7A4OX231/Chem_NY_D&C Yellow 10 | 7A4OX231 | Chem_NY_D&C Yellow 10 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram) | TAMC (Average per Gram) | 0 | CFU/g | FDA,1 |  |  | TM 610 | 7A4OX231/Chem_NY_D&C Yellow 10 | 2 | both
7A4OX231/Chem_NY_D&C Yellow 10 | 7A4OX231 | Chem_NY_D&C Yellow 10 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 7A4OX231/Chem_NY_D&C Yellow 10 | 2 | both
7A4OX231/Chem_NY_D&C Yellow 10 | 7A4OX231 | Chem_NY_D&C Yellow 10 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram) | TMYC (Average per Gram) | 0 | CFU/g | FDA,1 |  |  | TM 610 | 7A4OX231/Chem_NY_D&C Yellow 10 | 2 | both
7A4OX231/Chem_NY_D&C Yellow 10 | 7A4OX231 | Chem_NY_D&C Yellow 10 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 7A4OX231/Chem_NY_D&C Yellow 10 | 2 | both
7ACHE61P/Chem_NY_Sodium Benzoate | 7ACHE61P | Chem_NY_Sodium Benzoate | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection | 7ACHE61P/Chem_NY_Sodium Benzoate | 3 | both
7ACHE61P/Chem_NY_Sodium Benzoate | 7ACHE61P | Chem_NY_Sodium Benzoate | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Extruded Pellets (Done) |  |  |  |  | TM 208 | 7ACHE61P/Chem_NY_Sodium Benzoate | 3 | both
7ACHE61P/Chem_NY_Sodium Benzoate | 7ACHE61P | Chem_NY_Sodium Benzoate | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) |  |  |  |  | TM 208 | 7ACHE61P/Chem_NY_Sodium Benzoate | 3 | both
7ACHE61P/Chem_NY_Sodium Benzoate | 7ACHE61P | Chem_NY_Sodium Benzoate | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 | 7ACHE61P/Chem_NY_Sodium Benzoate | 3 | both
7ACHE61P/Chem_NY_Sodium Benzoate | 7ACHE61P | Chem_NY_Sodium Benzoate | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless (Done) |  |  |  |  | TM 208 | 7ACHE61P/Chem_NY_Sodium Benzoate | 3 | both
7ACHE61P/Chem_NY_Sodium Benzoate | 7ACHE61P | Chem_NY_Sodium Benzoate | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7ACHE61P/Chem_NY_Sodium Benzoate | 3 | both
7B4AA580/Chem_NY_Acetic Acid - USP | 7B4AA580 | Chem_NY_Acetic Acid - USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection | 7B4AA580/Chem_NY_Acetic Acid - USP | 1 | both
7B4AA580/Chem_NY_Acetic Acid - USP | 7B4AA580 | Chem_NY_Acetic Acid - USP | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4AA580/Chem_NY_Acetic Acid - USP | 1 | both
7B4AA580/Chem_NY_Acetic Acid - USP | 7B4AA580 | Chem_NY_Acetic Acid - USP | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4AA580/Chem_NY_Acetic Acid - USP | 1 | both
7B4AA580/Chem_NY_Acetic Acid - USP | 7B4AA580 | Chem_NY_Acetic Acid - USP | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4AA580/Chem_NY_Acetic Acid - USP | 1 | both
7B4AA580/Chem_NY_Acetic Acid - USP | 7B4AA580 | Chem_NY_Acetic Acid - USP | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Pungent (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4AA580/Chem_NY_Acetic Acid - USP | 1 | both
7B4AL930/Chem_42A | 7B4AL930 | Chem_42A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH as is | 7B4AL930/Chem_42A | 2 | both
7B4AL930/Chem_42A | 7B4AL930 | Chem_42A | NY-Analytical | 447 | pH as is | pH | pH |  |  | FDA,1 | 1.1 | 1.5 | TM 200 | 7B4AL930/Chem_42A | 2 | both
7B4AL930/Chem_42A | 7B4AL930 | Chem_42A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.2444 |  | FDA,2 | 1.21 | 1.27 | TM 202 | 7B4AL930/Chem_42A | 2 | both
7B4AL930/Chem_42A | 7B4AL930 | Chem_42A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7B4AL930/Chem_42A | 2 | both
7B4AL930/Chem_42A | 7B4AL930 | Chem_42A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4AL930/Chem_42A | 2 | both
7B4AL930/Chem_42A | 7B4AL930 | Chem_42A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Light Yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4AL930/Chem_42A | 2 | both
7B4AL930/Chem_42A | 7B4AL930 | Chem_42A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4AL930/Chem_42A | 2 | both
7B4AL930/Chem_42A | 7B4AL930 | Chem_42A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4AL930/Chem_42A | 2 | both
7B4FD110/Chem_44A | 7B4FD110 | Chem_44A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 |  |  | left_only
7B4FD110/Chem_44A | 7B4FD110 | Chem_44A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  |  | left_only
7B4FD110/Chem_44A | 7B4FD110 | Chem_44A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7B4FD110/Chem_44A | 7B4FD110 | Chem_44A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7B4FD110/Chem_44A | 7B4FD110 | Chem_44A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | White granular powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7B4GB598/Chem_45A | 7B4GB598 | Chem_45A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  |  | left_only
7B4GB598/Chem_45A | 7B4GB598 | Chem_45A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7B4GB598/Chem_45A | 7B4GB598 | Chem_45A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Grain Powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7B4GB598/Chem_45A | 7B4GB598 | Chem_45A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7B4GB598/Chem_45A | 7B4GB598 | Chem_45A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Off-White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7B4OA242/Chem_48A | 7B4OA242 | Chem_48A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification | 7B4OA242/Chem_48A | 9 | both
7B4OA242/Chem_48A | 7B4OA242 | Chem_48A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7B4OA242/Chem_48A | 9 | both
7B4OA242/Chem_48A | 7B4OA242 | Chem_48A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Crystals/crystalline powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4OA242/Chem_48A | 9 | both
7B4OA242/Chem_48A | 7B4OA242 | Chem_48A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to white (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4OA242/Chem_48A | 9 | both
7B4OA242/Chem_48A | 7B4OA242 | Chem_48A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4OA242/Chem_48A | 9 | both
7B4OA242/Chem_48A | 7B4OA242 | Chem_48A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4OA242/Chem_48A | 9 | both
7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | 7B4OA347 | Chem_NY_Calming Waters INCI w/symfresh | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C | 7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | 2 | both
7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | 7B4OA347 | Chem_NY_Calming Waters INCI w/symfresh | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.964 |  | FDA,3 | 0.959 | 0.969 | TM 202 | 7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | 2 | both
7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | 7B4OA347 | Chem_NY_Calming Waters INCI w/symfresh | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.472 |  | FDA,3 | 1.465 | 1.475 | TM 204 | 7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | 2 | both
7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | 7B4OA347 | Chem_NY_Calming Waters INCI w/symfresh | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM 208 | 7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | 2 | both
7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | 7B4OA347 | Chem_NY_Calming Waters INCI w/symfresh | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to pale yellow (Done) |  |  |  |  | TM 208 | 7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | 2 | both
7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | 7B4OA347 | Chem_NY_Calming Waters INCI w/symfresh | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 | 7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | 2 | both
7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | 7B4OA347 | Chem_NY_Calming Waters INCI w/symfresh | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208 | 7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | 2 | both
7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | 7B4OA347 | Chem_NY_Calming Waters INCI w/symfresh | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411 | 7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | 2 | both
7B4OA709/Chem_NY_Bell fragrance | 7B4OA709 | Chem_NY_Bell fragrance | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C | 7B4OA709/Chem_NY_Bell fragrance | 2 | both
7B4OA709/Chem_NY_Bell fragrance | 7B4OA709 | Chem_NY_Bell fragrance | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.0639 |  | FDA,3 | 1.06 | 1.07 | TM 202 | 7B4OA709/Chem_NY_Bell fragrance | 2 | both
7B4OA709/Chem_NY_Bell fragrance | 7B4OA709 | Chem_NY_Bell fragrance | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.468 |  | FDA,3 | 1.467 | 1.477 | TM 204 | 7B4OA709/Chem_NY_Bell fragrance | 2 | both
7B4OA709/Chem_NY_Bell fragrance | 7B4OA709 | Chem_NY_Bell fragrance | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM 208 | 7B4OA709/Chem_NY_Bell fragrance | 2 | both
7B4OA709/Chem_NY_Bell fragrance | 7B4OA709 | Chem_NY_Bell fragrance | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Pale to light yellow (Done) |  |  |  |  | TM 208 | 7B4OA709/Chem_NY_Bell fragrance | 2 | both
7B4OA709/Chem_NY_Bell fragrance | 7B4OA709 | Chem_NY_Bell fragrance | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 | 7B4OA709/Chem_NY_Bell fragrance | 2 | both
7B4OA709/Chem_NY_Bell fragrance | 7B4OA709 | Chem_NY_Bell fragrance | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Conforms to Standard (Done) |  |  |  |  | TM 208 | 7B4OA709/Chem_NY_Bell fragrance | 2 | both
7B4OB690/Chem_50A | 7B4OB690 | Chem_50A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  |  | left_only
7B4OB690/Chem_50A | 7B4OB690 | Chem_50A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, mobile liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7B4OB690/Chem_50A | 7B4OB690 | Chem_50A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7B4OB690/Chem_50A | 7B4OB690 | Chem_50A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7B4OB690/Chem_50A | 7B4OB690 | Chem_50A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Solvent (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | CHE ANTIFOAM AF EMULSION | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 |  |  | left_only
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | CHE ANTIFOAM AF EMULSION | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  |  | left_only
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | CHE ANTIFOAM AF EMULSION | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per 10mL) | TAMC (Average per 10mL) |  | cfu/10mL | FDA,1 |  | 0 | TM 610 |  |  | left_only
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | CHE ANTIFOAM AF EMULSION | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per mL - Undiluted) | TAMC (Average per mL - Undiluted) |  | cfu/mL |  |  | 0 | TM 610 |  |  | left_only
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | CHE ANTIFOAM AF EMULSION | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) |  | cfu | FDA,0 |  | 0 | TM 610 |  |  | left_only
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | CHE ANTIFOAM AF EMULSION | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per 10mL) | TMYC (Average per 10mL) |  | cfu/10mL | FDA,1 |  | 0 | TM 610 |  |  | left_only
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | CHE ANTIFOAM AF EMULSION | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per mL - Undiluted) | TMYC (Average per mL - Undiluted) |  | cfu/mL |  |  | 0 | TM 610 |  |  | left_only
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | CHE ANTIFOAM AF EMULSION | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  | cfu | FDA,0 |  | 0 | TM 610 |  |  | left_only
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | CHE ANTIFOAM AF EMULSION | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | [5 empty] | WI QAS 076 |  |  | left_only
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | CHE ANTIFOAM AF EMULSION | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | CHE ANTIFOAM AF EMULSION | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Homogeneous Liquid, Free of Dark Particles (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | CHE ANTIFOAM AF EMULSION | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | CHE ANTIFOAM AF EMULSION | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White to off White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7B4OD232/Chem_52A | 7B4OD232 | Chem_52A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C | 7B4OD232/Chem_52A | 11 | both
7B4OD232/Chem_52A | 7B4OD232 | Chem_52A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.0361 |  | FDA,3 | 1.035 | 1.037 | TM 202 | 7B4OD232/Chem_52A | 11 | both
7B4OD232/Chem_52A | 7B4OD232 | Chem_52A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7B4OD232/Chem_52A | 11 | both
7B4OD232/Chem_52A | 7B4OD232 | Chem_52A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear viscous liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4OD232/Chem_52A | 11 | both
7B4OD232/Chem_52A | 7B4OD232 | Chem_52A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4OD232/Chem_52A | 11 | both
7B4OD232/Chem_52A | 7B4OD232 | Chem_52A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4OD232/Chem_52A | 11 | both
7B4OD232/Chem_52A | 7B4OD232 | Chem_52A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4OD232/Chem_52A | 11 | both
7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | 7B4OM799 | Chem_NY_Lipotec_Cucumber Extract in Glycerin | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C | 7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | 1 | both
7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | 7B4OM799 | Chem_NY_Lipotec_Cucumber Extract in Glycerin | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.2627 |  | FDA,3 | 1.25 | 1.29 | TM 202 | 7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | 1 | both
7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | 7B4OM799 | Chem_NY_Lipotec_Cucumber Extract in Glycerin | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.473 |  | FDA,3 | 1.45 | 1.49 | TM 204 | 7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | 1 | both
7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | 7B4OM799 | Chem_NY_Lipotec_Cucumber Extract in Glycerin | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear viscous liquid (Done) |  |  |  |  | TM 208 | 7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | 1 | both
7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | 7B4OM799 | Chem_NY_Lipotec_Cucumber Extract in Glycerin | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Light to Medium Yellow (Done) |  |  |  |  | TM 208 | 7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | 1 | both
7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | 7B4OM799 | Chem_NY_Lipotec_Cucumber Extract in Glycerin | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 | 7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | 1 | both
7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | 7B4OM799 | Chem_NY_Lipotec_Cucumber Extract in Glycerin | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) |  |  |  |  | TM 208 | 7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | 1 | both
7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | 7B4OM799 | Chem_NY_Lipotec_Cucumber Extract in Glycerin | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | 1 | both
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Chem_NY_Abil B88183 PH | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 |  |  | left_only
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Chem_NY_Abil B88183 PH | NY-Microbiology | 509 | J&J_Pathogens | Cetrimide | Cetrimide | No Growth |  |  |  |  | TM 869 |  |  | left_only
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Chem_NY_Abil B88183 PH | NY-Microbiology | 509 | J&J_Pathogens | MacConkey | MacConkey | No Growth |  |  |  |  | TM 869 |  |  | left_only
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Chem_NY_Abil B88183 PH | NY-Microbiology | 509 | J&J_Pathogens | TSA | TSA | No Growth |  |  |  |  | TM 869 |  |  | left_only
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Chem_NY_Abil B88183 PH | NY-Microbiology | 509 | J&J_Pathogens | MSA or VJ | MSA or VJ | No Growth_MSA |  |  |  |  | TM 869 |  |  | left_only
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Chem_NY_Abil B88183 PH | NY-Microbiology | 509 | J&J_Pathogens | Biggy or SDA | Biggy or SDA | No Growth_SDA |  |  |  |  | TM 869 |  |  | left_only
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Chem_NY_Abil B88183 PH | NY-Microbiology | 517 | J&J_Enrichment | TSB | TSB | Not Turbid |  |  |  |  | TM 869 |  |  | left_only
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Chem_NY_Abil B88183 PH | NY-Microbiology | 524 | J&J Fungal | Fungal | Fungal |  | CFU/g | FDA,0 | 0 | 200 | TM 869 |  |  | left_only
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Chem_NY_Abil B88183 PH | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | TSA JJ |  | cfu/mL | FDA,0 | 0 | 200 | TM 869 |  |  | left_only
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Chem_NY_Abil B88183 PH | NY-Microbiology | 700 | J&J SDA Controls | Negative Control | Negative Control | No Growth SDA |  |  |  |  | TM 869 |  |  | left_only
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Chem_NY_Abil B88183 PH | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | NegativeControl | No Growth TSB |  |  |  |  | TM 869 |  |  | left_only
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Chem_NY_Abil B88183 PH | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  |  | left_only
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Chem_NY_Abil B88183 PH | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear to slightly turbid.Colorless to pale yellow,slightly viscous liquid (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Chem_NY_Abil B88183 PH | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7B4OP065/Chem_NY_Chamomile Flower Extract | 7B4OP065 | Chem_NY_Chamomile Flower Extract | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection | 7B4OP065/Chem_NY_Chamomile Flower Extract | 2 | both
7B4OP065/Chem_NY_Chamomile Flower Extract | 7B4OP065 | Chem_NY_Chamomile Flower Extract | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Liquid (Done) |  |  |  |  | TM 208 | 7B4OP065/Chem_NY_Chamomile Flower Extract | 2 | both
7B4OP065/Chem_NY_Chamomile Flower Extract | 7B4OP065 | Chem_NY_Chamomile Flower Extract | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Greenish Yellow to Yellow Orange (Done) |  |  |  |  | TM 208 | 7B4OP065/Chem_NY_Chamomile Flower Extract | 2 | both
7B4OP065/Chem_NY_Chamomile Flower Extract | 7B4OP065 | Chem_NY_Chamomile Flower Extract | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 | 7B4OP065/Chem_NY_Chamomile Flower Extract | 2 | both
7B4OP065/Chem_NY_Chamomile Flower Extract | 7B4OP065 | Chem_NY_Chamomile Flower Extract | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208 | 7B4OP065/Chem_NY_Chamomile Flower Extract | 2 | both
7B4OP065/Chem_NY_Chamomile Flower Extract | 7B4OP065 | Chem_NY_Chamomile Flower Extract | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411 | 7B4OP065/Chem_NY_Chamomile Flower Extract | 2 | both
7B4OP065/Chem_NY_Chamomile Flower Extract | 7B4OP065 | Chem_NY_Chamomile Flower Extract | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram) | TAMC (Average per Gram) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | 7B4OP065/Chem_NY_Chamomile Flower Extract | 2 | both
7B4OP065/Chem_NY_Chamomile Flower Extract | 7B4OP065 | Chem_NY_Chamomile Flower Extract | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 7B4OP065/Chem_NY_Chamomile Flower Extract | 2 | both
7B4OP065/Chem_NY_Chamomile Flower Extract | 7B4OP065 | Chem_NY_Chamomile Flower Extract | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram) | TMYC (Average per Gram) |  | CFU/g | FDA,1 |  | 100 | TM 610 | 7B4OP065/Chem_NY_Chamomile Flower Extract | 2 | both
7B4OP065/Chem_NY_Chamomile Flower Extract | 7B4OP065 | Chem_NY_Chamomile Flower Extract | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  | cfu | FDA,0 |  |  | TM 610 | 7B4OP065/Chem_NY_Chamomile Flower Extract | 2 | both
7B4OP065/Chem_NY_Chamomile Flower Extract | 7B4OP065 | Chem_NY_Chamomile Flower Extract | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Non-Subculturable Growth (Done) |  |  |  |  | WI QAS 076 | 7B4OP065/Chem_NY_Chamomile Flower Extract | 2 | both
7B4OP093/Chem_NY_PEG-4 | 7B4OP093 | Chem_NY_PEG-4 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH | 7B4OP093/Chem_NY_PEG-4 | 10 | both
7B4OP093/Chem_NY_PEG-4 | 7B4OP093 | Chem_NY_PEG-4 | NY-Analytical | 398 | pH | pH (5% aqueous @25 degree C) | pH (5% aqueous @25 degree C) | 6.01 |  | FDA,1 | 4.5 | 7.5 | TM 200 | 7B4OP093/Chem_NY_PEG-4 | 10 | both
7B4OP093/Chem_NY_PEG-4 | 7B4OP093 | Chem_NY_PEG-4 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Not More Than Slightly Hazy (Done) |  |  |  |  | TM 208 | 7B4OP093/Chem_NY_PEG-4 | 10 | both
7B4OP093/Chem_NY_PEG-4 | 7B4OP093 | Chem_NY_PEG-4 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM 208 | 7B4OP093/Chem_NY_PEG-4 | 10 | both
7B4OP093/Chem_NY_PEG-4 | 7B4OP093 | Chem_NY_PEG-4 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 | 7B4OP093/Chem_NY_PEG-4 | 10 | both
7B4OP093/Chem_NY_PEG-4 | 7B4OP093 | Chem_NY_PEG-4 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208 | 7B4OP093/Chem_NY_PEG-4 | 10 | both
7B4OP093/Chem_NY_PEG-4 | 7B4OP093 | Chem_NY_PEG-4 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7B4OP093/Chem_NY_PEG-4 | 10 | both
7B4OP113/Chem_NY_Mentha Piperita Extract_01 | 7B4OP113 | Chem_NY_Mentha Piperita Extract_01 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH | 7B4OP113/Chem_NY_Mentha Piperita Extract_01 | 1 | both
7B4OP113/Chem_NY_Mentha Piperita Extract_01 | 7B4OP113 | Chem_NY_Mentha Piperita Extract_01 | NY-Analytical | 398 | pH | pH | pH | 4.24 |  | FDA,1 | 4 | 5 | TM 200 | 7B4OP113/Chem_NY_Mentha Piperita Extract_01 | 1 | both
7B4OP113/Chem_NY_Mentha Piperita Extract_01 | 7B4OP113 | Chem_NY_Mentha Piperita Extract_01 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear aqueous solution (Done) |  |  |  |  | TM 208 | 7B4OP113/Chem_NY_Mentha Piperita Extract_01 | 1 | both
7B4OP113/Chem_NY_Mentha Piperita Extract_01 | 7B4OP113 | Chem_NY_Mentha Piperita Extract_01 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Amber (Done) |  |  |  |  | TM 208 | 7B4OP113/Chem_NY_Mentha Piperita Extract_01 | 1 | both
7B4OP113/Chem_NY_Mentha Piperita Extract_01 | 7B4OP113 | Chem_NY_Mentha Piperita Extract_01 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 | 7B4OP113/Chem_NY_Mentha Piperita Extract_01 | 1 | both
7B4OP113/Chem_NY_Mentha Piperita Extract_01 | 7B4OP113 | Chem_NY_Mentha Piperita Extract_01 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) |  |  |  |  | TM 208 | 7B4OP113/Chem_NY_Mentha Piperita Extract_01 | 1 | both
7B4OP203/Chem_NY_Diazolidinyl Urea – RI-1157 | 7B4OP203 | Chem_NY_Diazolidinyl Urea – RI-1157 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 |  |  | left_only
7B4OP203/Chem_NY_Diazolidinyl Urea – RI-1157 | 7B4OP203 | Chem_NY_Diazolidinyl Urea – RI-1157 | NY-Analytical | 398 | pH | pH (1% w/v aqueous solution) | pH (1% w/v aqueous solution) |  |  | FDA,1 | 6.5 | 8 | TM 200 |  |  | left_only
7B4OP203/Chem_NY_Diazolidinyl Urea – RI-1157 | 7B4OP203 | Chem_NY_Diazolidinyl Urea – RI-1157 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH |  |  | left_only
7B4OP203/Chem_NY_Diazolidinyl Urea – RI-1157 | 7B4OP203 | Chem_NY_Diazolidinyl Urea – RI-1157 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Free flowing powder, free of foreign material (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7B4OP203/Chem_NY_Diazolidinyl Urea – RI-1157 | 7B4OP203 | Chem_NY_Diazolidinyl Urea – RI-1157 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Similar in character and intensity to standard, slight pungent odor (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7B4OP203/Chem_NY_Diazolidinyl Urea – RI-1157 | 7B4OP203 | Chem_NY_Diazolidinyl Urea – RI-1157 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7B4OP203/Chem_NY_Diazolidinyl Urea – RI-1157 | 7B4OP203 | Chem_NY_Diazolidinyl Urea – RI-1157 | NY-Analytical | 891 | Solutions and Solubility | Solubility (1% w/v aqueous solution) | Solubility (1% w/v aqueous solution) | Clear and Colorless (Done) |  |  |  |  | TM 222 |  |  | left_only
7B4OP209/MASTER - Microbiology - J&J Chemical Testing) | 7B4OP209 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 398 | pH | pH (1% Solution) | pH (1% Solution) |  |  | FDA,1 | 4.5 | 7 | J&J TMWW78014 |  |  | left_only
7B4OP209/MASTER - Microbiology - J&J Chemical Testing) | 7B4OP209 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 808 | Absorption | Absorption (1%/1cm) | Absorption (1%/1cm) |  |  | FDA,1 | 401.8 |  | J&J TM-002747 |  |  | left_only
7B4OP230/Chem_NY_Cosvat (Elestab CPN Ultra Pure) | 7B4OP230 | Chem_NY_Cosvat (Elestab CPN Ultra Pure) | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 |  |  | left_only
7B4OP230/Chem_NY_Cosvat (Elestab CPN Ultra Pure) | 7B4OP230 | Chem_NY_Cosvat (Elestab CPN Ultra Pure) | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  |  | left_only
7B4OP230/Chem_NY_Cosvat (Elestab CPN Ultra Pure) | 7B4OP230 | Chem_NY_Cosvat (Elestab CPN Ultra Pure) | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | White to off - white powder or crystalline aggregates. (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7B4OP230/Chem_NY_Cosvat (Elestab CPN Ultra Pure) | 7B4OP230 | Chem_NY_Cosvat (Elestab CPN Ultra Pure) | NY-Analytical | 796 | Assay | Assay | Assay |  | % | FDA,1 | 99 |  | TM-002291 |  |  | left_only
7B4OP230/Chem_NY_Cosvat (Elestab CPN Ultra Pure) | 7B4OP230 | Chem_NY_Cosvat (Elestab CPN Ultra Pure) | NY-Analytical | 796 | Assay | Assay Final | Assay Final |  | % | FDA,1 | 99 |  | TM-002291 |  |  | left_only
7B4OP230/Chem_NY_Cosvat (Elestab CPN Ultra Pure) | 7B4OP230 | Chem_NY_Cosvat (Elestab CPN Ultra Pure) | NY-Analytical | 801 | % Water | Water Content | Water Content |  | % | FDA,1 |  | 1 | J&J TMWW78017 |  |  | left_only
7B4OP284/Chem_NY_Sodium Citrate USP | 7B4OP284 | Chem_NY_Sodium Citrate USP | NY-Analytical | 396 | Odor | Odor | Odor | Passes Test (Done) |  |  |  |  | J&J TM-003188 |  |  | left_only
7B4OP284/Chem_NY_Sodium Citrate USP | 7B4OP284 | Chem_NY_Sodium Citrate USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Odor |  |  | left_only
7B4OP284/Chem_NY_Sodium Citrate USP | 7B4OP284 | Chem_NY_Sodium Citrate USP | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Colorless crystals, or as a white crystalline powder (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7B4OP284/Chem_NY_Sodium Citrate USP | 7B4OP284 | Chem_NY_Sodium Citrate USP | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7B4OP284/Chem_NY_Sodium Citrate USP | 7B4OP284 | Chem_NY_Sodium Citrate USP | NY-Analytical | 813 | Identification - USP | Identification A - Sodium | Identification A - Sodium | Passes Test (Done) |  |  |  |  | USP |  |  | left_only
7B4OP284/Chem_NY_Sodium Citrate USP | 7B4OP284 | Chem_NY_Sodium Citrate USP | NY-Analytical | 813 | Identification - USP | Identification B - Citrate | Identification B - Citrate | Passes Test (Done) |  |  |  |  | USP |  |  | left_only
7B4OP284/Chem_NY_Sodium Citrate USP | 7B4OP284 | Chem_NY_Sodium Citrate USP | NY-Analytical | 813 | Identification - USP | Identification C | Identification C | Passes Test (Done) |  |  |  |  | USP |  |  | left_only
7B4OP661/Chem_NY_Pemulen TR-2 | 7B4OP661 | Chem_NY_Pemulen TR-2 | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 |  |  | left_only
7B4OP661/Chem_NY_Pemulen TR-2 | 7B4OP661 | Chem_NY_Pemulen TR-2 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA |  |  | left_only
7B4OP661/Chem_NY_Pemulen TR-2 | 7B4OP661 | Chem_NY_Pemulen TR-2 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7B4OP661/Chem_NY_Pemulen TR-2 | 7B4OP661 | Chem_NY_Pemulen TR-2 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7B4OP661/Chem_NY_Pemulen TR-2 | 7B4OP661 | Chem_NY_Pemulen TR-2 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Slightly acetic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7B4OP661/Chem_NY_Pemulen TR-2 | 7B4OP661 | Chem_NY_Pemulen TR-2 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7B4OQ028/Chem_NY_Sorbitan Monolaurate | 7B4OQ028 | Chem_NY_Sorbitan Monolaurate | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification | 7B4OQ028/Chem_NY_Sorbitan Monolaurate | 3 | both
7B4OQ028/Chem_NY_Sorbitan Monolaurate | 7B4OQ028 | Chem_NY_Sorbitan Monolaurate | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411 | 7B4OQ028/Chem_NY_Sorbitan Monolaurate | 3 | both
7B4OQ028/Chem_NY_Sorbitan Monolaurate | 7B4OQ028 | Chem_NY_Sorbitan Monolaurate | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4OQ028/Chem_NY_Sorbitan Monolaurate | 3 | both
7B4OQ028/Chem_NY_Sorbitan Monolaurate | 7B4OQ028 | Chem_NY_Sorbitan Monolaurate | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4OQ028/Chem_NY_Sorbitan Monolaurate | 3 | both
7B4OQ028/Chem_NY_Sorbitan Monolaurate | 7B4OQ028 | Chem_NY_Sorbitan Monolaurate | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4OQ028/Chem_NY_Sorbitan Monolaurate | 3 | both
7B4OQ028/Chem_NY_Sorbitan Monolaurate | 7B4OQ028 | Chem_NY_Sorbitan Monolaurate | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4OQ028/Chem_NY_Sorbitan Monolaurate | 3 | both
7B4OQ084/Chem_60A | 7B4OQ084 | Chem_60A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection | 7B4OQ084/Chem_60A | 1 | both
7B4OQ084/Chem_60A | 7B4OQ084 | Chem_60A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Slightly Opaque Viscous Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4OQ084/Chem_60A | 1 | both
7B4OQ084/Chem_60A | 7B4OQ084 | Chem_60A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Slightly White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4OQ084/Chem_60A | 1 | both
7B4OQ084/Chem_60A | 7B4OQ084 | Chem_60A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4OQ084/Chem_60A | 1 | both
7B4OQ084/Chem_60A | 7B4OQ084 | Chem_60A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Mild (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4OQ084/Chem_60A | 1 | both
7B4OQ209/Chem_61A | 7B4OQ209 | Chem_61A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification | 7B4OQ209/Chem_61A | 3 | both
7B4OQ209/Chem_61A | 7B4OQ209 | Chem_61A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | [5 empty] | TM 411 | 7B4OQ209/Chem_61A | 3 | both
7B4OQ209/Chem_61A | 7B4OQ209 | Chem_61A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Free flowing granules or powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4OQ209/Chem_61A | 3 | both
7B4OQ209/Chem_61A | 7B4OQ209 | Chem_61A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4OQ209/Chem_61A | 3 | both
7B4OQ209/Chem_61A | 7B4OQ209 | Chem_61A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4OQ209/Chem_61A | 3 | both
7B4OQ209/Chem_61A | 7B4OQ209 | Chem_61A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4OQ209/Chem_61A | 3 | both
7B4OQ214/Chem_62A | 7B4OQ214 | Chem_62A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification | 7B4OQ214/Chem_62A | 4 | both
7B4OQ214/Chem_62A | 7B4OQ214 | Chem_62A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7B4OQ214/Chem_62A | 4 | both
7B4OQ214/Chem_62A | 7B4OQ214 | Chem_62A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Free flowing crystalline powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4OQ214/Chem_62A | 4 | both
7B4OQ214/Chem_62A | 7B4OQ214 | Chem_62A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4OQ214/Chem_62A | 4 | both
7B4OQ214/Chem_62A | 7B4OQ214 | Chem_62A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4OQ214/Chem_62A | 4 | both
7B4OQ214/Chem_62A | 7B4OQ214 | Chem_62A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4OQ214/Chem_62A | 4 | both
7B4OQ595/Chem_63A | 7B4OQ595 | Chem_63A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification | 7B4OQ595/Chem_63A | 2 | both
7B4OQ595/Chem_63A | 7B4OQ595 | Chem_63A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7B4OQ595/Chem_63A | 2 | both
7B4OQ595/Chem_63A | 7B4OQ595 | Chem_63A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Crystalline Powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4OQ595/Chem_63A | 2 | both
7B4OQ595/Chem_63A | 7B4OQ595 | Chem_63A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White to almost white (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4OQ595/Chem_63A | 2 | both
7B4OQ595/Chem_63A | 7B4OQ595 | Chem_63A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4OQ595/Chem_63A | 2 | both
7B4OQ595/Chem_63A | 7B4OQ595 | Chem_63A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Faint (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4OQ595/Chem_63A | 2 | both
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Chem_NY_Disodium EDTA-USP | NY-Analytical | 841 | Residual Solvents | Residual Solvents | Residual Solvents | Supplier Attestation Letter |  |  |  |  | Residual Solvents | 7B4OQ794/Chem_NY_Disodium EDTA-USP | 2 | both
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Chem_NY_Disodium EDTA-USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection | 7B4OQ794/Chem_NY_Disodium EDTA-USP | 2 | both
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Chem_NY_Disodium EDTA-USP | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | White crystalline powder (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | 7B4OQ794/Chem_NY_Disodium EDTA-USP | 2 | both
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Chem_NY_Disodium EDTA-USP | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | 7B4OQ794/Chem_NY_Disodium EDTA-USP | 2 | both
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Chem_NY_Disodium EDTA-USP | NY-Analytical | 398 | pH | pH (1 in 20 solution) | pH (1 in 20 solution) | 4.9 |  | FDA,1 | 4 | 6 | J&J TMWW78014 | 7B4OQ794/Chem_NY_Disodium EDTA-USP | 2 | both
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Chem_NY_Disodium EDTA-USP | NY-Analytical | 796 | Assay | Assay | Assay | 99.8 | % | FDA,1 | 99 | 101 | USP | 7B4OQ794/Chem_NY_Disodium EDTA-USP | 2 | both
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Chem_NY_Disodium EDTA-USP | NY-Analytical | 796 | Assay | Assay Final | Assay Final | 99.8 | % | FDA,1 | 99 | 101 | USP | 7B4OQ794/Chem_NY_Disodium EDTA-USP | 2 | both
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Chem_NY_Disodium EDTA-USP | NY-Analytical | 843 | Calcium | Calcium | Calcium | Passes Test (Done) |  |  |  |  | USP | 7B4OQ794/Chem_NY_Disodium EDTA-USP | 2 | both
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Chem_NY_Disodium EDTA-USP | NY-Analytical | 833 | Heavy Metals | Heavy Metals | Heavy Metals | <0.005 | % | FDA,3 |  | 0.005 | USP | 7B4OQ794/Chem_NY_Disodium EDTA-USP | 2 | both
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Chem_NY_Disodium EDTA-USP | NY-Analytical | 813 | Identification - USP | Identification A | Identification A | Passes Test (Done) |  |  |  |  | USP | 7B4OQ794/Chem_NY_Disodium EDTA-USP | 2 | both
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Chem_NY_Disodium EDTA-USP | NY-Analytical | 813 | Identification - USP | Identification B | Identification B | Passes Test (Done) |  |  |  |  | USP | 7B4OQ794/Chem_NY_Disodium EDTA-USP | 2 | both
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Chem_NY_Disodium EDTA-USP | NY-Analytical | 813 | Identification - USP | Identification C | Identification C | Passes Test (Done) |  |  |  |  | USP | 7B4OQ794/Chem_NY_Disodium EDTA-USP | 2 | both
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Chem_NY_Disodium EDTA-USP | NY-Analytical | 802 | Loss on Drying | Loss on Drying | Loss on Drying | 9.6 | % | FDA,1 | 8.7 | 11.4 | USP | 7B4OQ794/Chem_NY_Disodium EDTA-USP | 2 | both
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Chem_NY_Disodium EDTA-USP | NY-Analytical | 802 | Loss on Drying | Loss on Drying Final | Loss on Drying Final | 9.6 | % | FDA,1 | 8.7 | 11.4 | USP | 7B4OQ794/Chem_NY_Disodium EDTA-USP | 2 | both
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Chem_NY_Disodium EDTA-USP | NY-Analytical | 842 | Nitrilotriacetic Acid | Nitrilotriacetic Acid | Nitrilotriacetic Acid | 0 | % | FDA,1 |  | 0.1 | USP | 7B4OQ794/Chem_NY_Disodium EDTA-USP | 2 | both
7B4OX232/Chem_40A | 7B4OX232 | Chem_40A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Apparent proof | 7B4OX232/Chem_40A | 34 | both
7B4OX232/Chem_40A | 7B4OX232 | Chem_40A | NY-Analytical | 936 | Apparent proof | Apparent proof | Apparent proof | 189.7 |  | FDA,1 | 189.7 | 190.2 | STM 101 | 7B4OX232/Chem_40A | 34 | both
7B4OX232/Chem_40A | 7B4OX232 | Chem_40A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.8091 |  | FDA,4 | 0.8086 | 0.8098 | TM 202 | 7B4OX232/Chem_40A | 34 | both
7B4OX232/Chem_40A | 7B4OX232 | Chem_40A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Meets USP requirement (Done) |  |  |  |  | TM 411 | 7B4OX232/Chem_40A | 34 | both
7B4OX232/Chem_40A | 7B4OX232 | Chem_40A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, mobile liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4OX232/Chem_40A | 34 | both
7B4OX232/Chem_40A | 7B4OX232 | Chem_40A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4OX232/Chem_40A | 34 | both
7B4OX232/Chem_40A | 7B4OX232 | Chem_40A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4OX232/Chem_40A | 34 | both
7B4OX232/Chem_40A | 7B4OX232 | Chem_40A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Alcohol (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7B4OX232/Chem_40A | 34 | both
7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | 7C4AA580 | Chem_NY_Fragrance-Grapefruit Blossom &Orange | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C | 7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | 3 | both
7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | 7C4AA580 | Chem_NY_Fragrance-Grapefruit Blossom &Orange | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.0662 |  | FDA,4 | 1.0561 | 1.0761 | TM 202 | 7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | 3 | both
7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | 7C4AA580 | Chem_NY_Fragrance-Grapefruit Blossom &Orange | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | 3 | both
7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | 7C4AA580 | Chem_NY_Fragrance-Grapefruit Blossom &Orange | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | 3 | both
7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | 7C4AA580 | Chem_NY_Fragrance-Grapefruit Blossom &Orange | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Pale Yellow – Light Yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | 3 | both
7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | 7C4AA580 | Chem_NY_Fragrance-Grapefruit Blossom &Orange | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | 3 | both
7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | 7C4AA580 | Chem_NY_Fragrance-Grapefruit Blossom &Orange | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | 3 | both
7C4FD110/Chem_69A | 7C4FD110 | Chem_69A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection | 7C4FD110/Chem_69A | 2 | both
7C4FD110/Chem_69A | 7C4FD110 | Chem_69A | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | White or nearly white, hard, brittle material (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | 7C4FD110/Chem_69A | 2 | both
7C4FD110/Chem_69A | 7C4FD110 | Chem_69A | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | 7C4FD110/Chem_69A | 2 | both
7C4FD110/Chem_69A | 7C4FD110 | Chem_69A | NY-Analytical | 390 | Identification | Identification B | Identification B | Pass (Done) |  |  |  |  | USP | 7C4FD110/Chem_69A | 2 | both
7C4FD110/Chem_69A | 7C4FD110 | Chem_69A | NY-Analytical | 390 | Identification | Identification for Sodium (USP <191>) | Identification for Sodium | Positive response for Sodium (Done) |  |  |  |  | USP | 7C4FD110/Chem_69A | 2 | both
7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 7C4OA709 | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 961 | Peroxide Value | Peroxide Value | Peroxide Value |  |  | FDA,1 |  | 1 | AOCS Cd 8-53 | 7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 2 | both
7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 7C4OA709 | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 944 | color (Gardner) | color (Gardner) | color (Gardner) |  |  | FDA,0 |  | 3 | AOCS Td 1a-64) | 7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 2 | both
7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 7C4OA709 | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA | 7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 2 | both
7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 7C4OA709 | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 826 | Acid Value | Acid Value | Acid Value | 1.7 | mg KOH/g | FDA,2 |  | 2 | IUPAC 2.201(M) | 7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 2 | both
7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 7C4OA709 | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 826 | Acid Value | Acid Value Final | Acid Value Final | 1.7 | mg KOH/g | FDA,2 |  | 2 | IUPAC 2.201(M) | 7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 2 | both
7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 7C4OA709 | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 810 | Saponification Value | Saponification Value | Saponification Value | 22.1 | mg/g | FDA,1 | 17 | 25 | IUPAC 2.202 | 7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 2 | both
7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 7C4OA709 | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 810 | Saponification Value | Saponification Value Final | Saponification Value Final | 22.1 | mg/g | FDA,1 | 17 | 25 | IUPAC 2.202 | 7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 2 | both
7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 7C4OA709 | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 1000 | Cloudpoint | Cloudpoint | Cloudpoint | 67.4 | C | FDA,1 | 66 | 71 | TC276/WG2N30 930809C | 7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 2 | both
7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 7C4OA709 | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 959 | Water Content | Water Content | Water Content |  | % | FDA,2 |  | 1.5 | Vatten LB | 7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 2 | both
7C4OD232/Chem_NY_Methylparaben NF/USP | 7C4OD232 | Chem_NY_Methylparaben NF/USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis | 7C4OD232/Chem_NY_Methylparaben NF/USP | 7 | both
7C4OD232/Chem_NY_Propylparaben USP | 7C4OD232 | Chem_NY_Propylparaben USP | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | USP <197M> |  |  | left_only
7C4OD232/Chem_NY_Methylparaben NF/USP | 7C4OD232 | Chem_NY_Methylparaben NF/USP | NY-Analytical | 390 | Identification | IR Identification | IR Identification | [5 empty] | TM 411 | 7C4OD232/Chem_NY_Methylparaben NF/USP | 7 | both
7C4OD232/Chem_NY_Methylparaben NF/USP | 7C4OD232 | Chem_NY_Methylparaben NF/USP | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Colorless crystals or white crystalline powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7C4OD232/Chem_NY_Methylparaben NF/USP | 7 | both
7C4OD232/Chem_NY_Propylparaben USP | 7C4OD232 | Chem_NY_Propylparaben USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  |  | left_only
7C4OD232/Chem_NY_Methylparaben NF/USP | 7C4OD232 | Chem_NY_Methylparaben NF/USP | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | [5 empty] | TM208/220/230/USP/Kao LM | 7C4OD232/Chem_NY_Methylparaben NF/USP | 7 | both
7C4OD232/Chem_NY_Propylparaben USP | 7C4OD232 | Chem_NY_Propylparaben USP | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7C4OD232/Chem_NY_Methylparaben NF/USP | 7C4OD232 | Chem_NY_Methylparaben NF/USP | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | [5 empty] | TM208/220/230/USP/Kao LM | 7C4OD232/Chem_NY_Methylparaben NF/USP | 7 | both
7C4OD232/Chem_NY_Propylparaben USP | 7C4OD232 | Chem_NY_Propylparaben USP | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7C4OD232/Chem_NY_Propylparaben USP | 7C4OD232 | Chem_NY_Propylparaben USP | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Small, colorless crystals or white crystalline powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7C4OD232/Chem_NY_Propylparaben USP | 7C4OD232 | Chem_NY_Propylparaben USP | NY-Analytical | 796 | Assay | Assay | Assay |  | % | FDA,1 | 98 | 102 | USP/NF |  |  | left_only
7C4OD232/Chem_NY_Methylparaben NF/USP | 7C4OD232 | Chem_NY_Methylparaben NF/USP | NY-Analytical | 796 | Assay | Assay | Assay |  | % | FDA,1 | 98 | 102 | USP/NF | 7C4OD232/Chem_NY_Methylparaben NF/USP | 7 | both
7C4OD232/Chem_NY_Propylparaben USP | 7C4OD232 | Chem_NY_Propylparaben USP | NY-Analytical | 796 | Assay | Assay Final | Assay Final |  | % | FDA,1 | 98 | 102 | USP/NF |  |  | left_only
7C4OD232/Chem_NY_Methylparaben NF/USP | 7C4OD232 | Chem_NY_Methylparaben NF/USP | NY-Analytical | 796 | Assay | Assay Final | Assay Final |  | % | FDA,1 | 98 | 102 | USP/NF | 7C4OD232/Chem_NY_Methylparaben NF/USP | 7 | both
7C4OG169/Chem_75A | 7C4OG169 | Chem_75A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection | 7C4OG169/Chem_75A | 2 | both
7C4OG169/Chem_75A | 7C4OG169 | Chem_75A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Crystals (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7C4OG169/Chem_75A | 2 | both
7C4OG169/Chem_75A | 7C4OG169 | Chem_75A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7C4OG169/Chem_75A | 2 | both
7C4OG169/Chem_75A | 7C4OG169 | Chem_75A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7C4OG169/Chem_75A | 2 | both
7C4OG169/Chem_75A | 7C4OG169 | Chem_75A | NY-Analytical | 390 | Identification | Identification for Ammonium (USP <191>) | Identification for Ammonium | Pass (Done) |  |  |  |  | TM400/408b/421/422/USP | 7C4OG169/Chem_75A | 2 | both
7C4OG169/Chem_75A | 7C4OG169 | Chem_75A | NY-Analytical | 390 | Identification | Identification for Chloride (USP <191>) | Identification for Chloride | Pass (Done) |  |  |  |  | TM400/408b/421/422/USP | 7C4OG169/Chem_75A | 2 | both
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Chem_NY_Camellia Oleifera Leaf Extract | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 |  |  | left_only
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 509 | J&J_Pathogens | Cetrimide | Cetrimide | No Growth |  |  |  |  | TM 869 |  |  | left_only
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 509 | J&J_Pathogens | MacConkey | MacConkey | No Growth |  |  |  |  | TM 869 |  |  | left_only
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 509 | J&J_Pathogens | TSA | TSA | No Growth |  |  |  |  | TM 869 |  |  | left_only
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 509 | J&J_Pathogens | MSA or VJ | MSA or VJ | No Growth_MSA |  |  |  |  | TM 869 |  |  | left_only
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 509 | J&J_Pathogens | Biggy or SDA | Biggy or SDA | No Growth_SDA |  |  |  |  | TM 869 |  |  | left_only
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 517 | J&J_Enrichment | TSB | TSB | Not Turbid |  |  |  |  | TM 869 |  |  | left_only
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 524 | J&J Fungal | Fungal | Fungal |  | CFU/g | FDA,0 | 0 | 200 | TM 869 |  |  | left_only
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | TSA JJ |  | cfu/mL | FDA,0 | 0 | 500 | TM 869 |  |  | left_only
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 723 | J&J TSB Controls - TSA | Negative Control | Negative Control | No Growth TSB |  |  |  |  | TM 869 |  |  | left_only
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | NegativeControl | No Growth TSB |  |  |  |  | TM 869 |  |  | left_only
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Chem_NY_Camellia Oleifera Leaf Extract | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  |  | left_only
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Chem_NY_Camellia Oleifera Leaf Extract | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Light yellow to amber liquid (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Chem_NY_Camellia Oleifera Leaf Extract | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Chem_NY_Camellia Oleifera Leaf Extract | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | 7C4OP065 | Chem_NY_ Sodium Lauroamphoacetate 30% | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH | 7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | 1 | both
7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | 7C4OP065 | Chem_NY_ Sodium Lauroamphoacetate 30% | NY-Analytical | 398 | pH | pH (1 in 20 solution) | pH (1 in 20 solution) | 9.5 |  | FDA,1 | 9 | 9.5 | TM 200 | 7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | 1 | both
7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | 7C4OP065 | Chem_NY_ Sodium Lauroamphoacetate 30% | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Clean Liquid (Done) |  |  |  |  | TM 208 | 7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | 1 | both
7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | 7C4OP065 | Chem_NY_ Sodium Lauroamphoacetate 30% | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to light yellow (Done) |  |  |  |  | TM 208 | 7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | 1 | both
7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | 7C4OP065 | Chem_NY_ Sodium Lauroamphoacetate 30% | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 | 7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | 1 | both
7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | 7C4OP065 | Chem_NY_ Sodium Lauroamphoacetate 30% | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208 | 7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | 1 | both
7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | 7C4OP065 | Chem_NY_ Sodium Lauroamphoacetate 30% | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411 | 7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | 1 | both
7C4OP230/MASTER - Microbiology - J&J Chemical Testing) | 7C4OP230 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 826 | Acid Value | Acid Value | Acid Value |  |  | FDA,2 | 0 | 0.2 | J&J TMWW78011 |  |  | left_only
7C4OP230/MASTER - Microbiology - J&J Chemical Testing) | 7C4OP230 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 831 | Density @ 20° C | Density | Density |  |  | FDA,4 | 0.89 | 0.893 | J&J TMWW78012 |  |  | left_only
7C4OP284/Chem_NY_Grape Fruit Extract | 7C4OP284 | Chem_NY_Grape Fruit Extract | NY-Analytical | 390 | Identification | Identification (FTIR) | Identification (FTIR) | Matches that of the Reference Standard (Done) |  |  |  |  | J&J TMWW78013 |  |  | left_only
7C4OP284/Chem_NY_Grape Fruit Extract | 7C4OP284 | Chem_NY_Grape Fruit Extract | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  |  | left_only
7C4OP284/Chem_NY_Grape Fruit Extract | 7C4OP284 | Chem_NY_Grape Fruit Extract | NY-Analytical | 792 | Chemical Inspection | Appearance (J&J TMWW78010) | Appearance | Clear to Pale Yellow Liquid (Done) |  |  |  |  | JJTM7252 |  |  | left_only
7C4OP284/Chem_NY_Grape Fruit Extract | 7C4OP284 | Chem_NY_Grape Fruit Extract | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | JJTM7252 |  |  | left_only
7C4OP601/MASTER - Microbiology - J&J Chemical Testing) | 7C4OP601 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 748 | Visual Inspection | Appearance of Packaging | Appearance of Packaging | Meets specifications (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | 7C4OP601/MASTER - Microbiology - J&J Chemical Testing) | 1 | both
7C4OQ028/Chem_NY_Isohexadecane | 7C4OQ028 | Chem_NY_Isohexadecane | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Refractive Index at 20°C | 7C4OQ028/Chem_NY_Isohexadecane | 6 | both
7C4OQ028/Chem_NY_Isohexadecane | 7C4OQ028 | Chem_NY_Isohexadecane | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.439 |  | FDA,1 | 1.4 | 1.5 | TM 204 | 7C4OQ028/Chem_NY_Isohexadecane | 6 | both
7C4OQ028/Chem_NY_Isohexadecane | 7C4OQ028 | Chem_NY_Isohexadecane | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411 | 7C4OQ028/Chem_NY_Isohexadecane | 6 | both
7C4OQ028/Chem_NY_Isohexadecane | 7C4OQ028 | Chem_NY_Isohexadecane | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7C4OQ028/Chem_NY_Isohexadecane | 6 | both
7C4OQ028/Chem_NY_Isohexadecane | 7C4OQ028 | Chem_NY_Isohexadecane | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7C4OQ028/Chem_NY_Isohexadecane | 6 | both
7C4OQ028/Chem_NY_Isohexadecane | 7C4OQ028 | Chem_NY_Isohexadecane | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7C4OQ028/Chem_NY_Isohexadecane | 6 | both
7C4OQ028/Chem_NY_Isohexadecane | 7C4OQ028 | Chem_NY_Isohexadecane | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7C4OQ028/Chem_NY_Isohexadecane | 6 | both
7C4OQ595/Chem_81A | 7C4OQ595 | Chem_81A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection | 7C4OQ595/Chem_81A | 3 | both
7C4OQ595/Chem_81A | 7C4OQ595 | Chem_81A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Free flowing granules or powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7C4OQ595/Chem_81A | 3 | both
7C4OQ595/Chem_81A | 7C4OQ595 | Chem_81A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7C4OQ595/Chem_81A | 3 | both
7C4OQ595/Chem_81A | 7C4OQ595 | Chem_81A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7C4OQ595/Chem_81A | 3 | both
7C4OQ595/Chem_81A | 7C4OQ595 | Chem_81A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7C4OQ595/Chem_81A | 3 | both
7C4OQ595/Chem_81A | 7C4OQ595 | Chem_81A | NY-Analytical | 390 | Identification | Identification B | Identification B | Pass (Done) |  |  |  |  | TM400/408b/421/422/USP | 7C4OQ595/Chem_81A | 3 | both
7C4OQ595/Chem_81A | 7C4OQ595 | Chem_81A | NY-Analytical | 390 | Identification | Identification for Potassium (USP <191>) | Identification for Potassium | Pass (Done) |  |  |  |  | TM400/408b/421/422/USP | 7C4OQ595/Chem_81A | 3 | both
7C4OX110/Chem_84A | 7C4OX110 | Chem_84A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C | 7C4OX110/Chem_84A | 2 | both
7C4OX110/Chem_84A | 7C4OX110 | Chem_84A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.8529 |  | FDA,3 | 0.845 | 0.858 | TM 202 | 7C4OX110/Chem_84A | 2 | both
7C4OX110/Chem_84A | 7C4OX110 | Chem_84A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7C4OX110/Chem_84A | 2 | both
7C4OX110/Chem_84A | 7C4OX110 | Chem_84A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7C4OX110/Chem_84A | 2 | both
7C4OX110/Chem_84A | 7C4OX110 | Chem_84A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7C4OX110/Chem_84A | 2 | both
7C4OX110/Chem_84A | 7C4OX110 | Chem_84A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7C4OX110/Chem_84A | 2 | both
7D4AA580/Chem_NY_Monoethanolamine | 7D4AA580 | Chem_NY_Monoethanolamine | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification | 7D4AA580/Chem_NY_Monoethanolamine | 2 | both
7D4AA580/Chem_NY_Monoethanolamine | 7D4AA580 | Chem_NY_Monoethanolamine | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411 | 7D4AA580/Chem_NY_Monoethanolamine | 2 | both
7D4AA580/Chem_NY_Monoethanolamine | 7D4AA580 | Chem_NY_Monoethanolamine | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, Colorless, Viscous Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7D4AA580/Chem_NY_Monoethanolamine | 2 | both
7D4AA580/Chem_NY_Monoethanolamine | 7D4AA580 | Chem_NY_Monoethanolamine | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7D4AA580/Chem_NY_Monoethanolamine | 2 | both
7D4AA580/Chem_NY_Monoethanolamine | 7D4AA580 | Chem_NY_Monoethanolamine | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7D4AA580/Chem_NY_Monoethanolamine | 2 | both
7D4AA580/Chem_NY_Monoethanolamine | 7D4AA580 | Chem_NY_Monoethanolamine | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Ammoniacal (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7D4AA580/Chem_NY_Monoethanolamine | 2 | both
7D4OG169/Chem_88A | 7D4OG169 | Chem_88A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C | 7D4OG169/Chem_88A | 1 | both
7D4OG169/Chem_88A | 7D4OG169 | Chem_88A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.068 |  | FDA,4 | 1.0594 | 1.0794 | TM 202 | 7D4OG169/Chem_88A | 1 | both
7D4OG169/Chem_88A | 7D4OG169 | Chem_88A | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.4755 |  | FDA,4 | 1.4695 | 1.4895 | TM 204 | 7D4OG169/Chem_88A | 1 | both
7D4OG169/Chem_88A | 7D4OG169 | Chem_88A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7D4OG169/Chem_88A | 1 | both
7D4OG169/Chem_88A | 7D4OG169 | Chem_88A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7D4OG169/Chem_88A | 1 | both
7D4OG169/Chem_88A | 7D4OG169 | Chem_88A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Pale straw (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7D4OG169/Chem_88A | 1 | both
7D4OG169/Chem_88A | 7D4OG169 | Chem_88A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7D4OG169/Chem_88A | 1 | both
7D4OP065/Chem_NY_ Glydant Plus Liquid | 7D4OP065 | Chem_NY_ Glydant Plus Liquid | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH | 7D4OP065/Chem_NY_ Glydant Plus Liquid | 2 | both
7D4OP065/Chem_NY_ Glydant Plus Liquid | 7D4OP065 | Chem_NY_ Glydant Plus Liquid | NY-Analytical | 398 | pH | pH | pH | 7.66 |  | FDA,1 | 6 | 8 | TM 200 | 7D4OP065/Chem_NY_ Glydant Plus Liquid | 2 | both
7D4OP065/Chem_NY_ Glydant Plus Liquid | 7D4OP065 | Chem_NY_ Glydant Plus Liquid | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM 208 | 7D4OP065/Chem_NY_ Glydant Plus Liquid | 2 | both
7D4OP065/Chem_NY_ Glydant Plus Liquid | 7D4OP065 | Chem_NY_ Glydant Plus Liquid | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Water-White to light yellow (Done) |  |  |  |  | TM 208 | 7D4OP065/Chem_NY_ Glydant Plus Liquid | 2 | both
7D4OP065/Chem_NY_ Glydant Plus Liquid | 7D4OP065 | Chem_NY_ Glydant Plus Liquid | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 | 7D4OP065/Chem_NY_ Glydant Plus Liquid | 2 | both
7D4OP065/Chem_NY_ Glydant Plus Liquid | 7D4OP065 | Chem_NY_ Glydant Plus Liquid | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) |  |  |  |  | TM 208 | 7D4OP065/Chem_NY_ Glydant Plus Liquid | 2 | both
7D4OP065/Chem_NY_ Glydant Plus Liquid | 7D4OP065 | Chem_NY_ Glydant Plus Liquid | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411 | 7D4OP065/Chem_NY_ Glydant Plus Liquid | 2 | both
7D4OP230/Chem_NY_Zemea_01 | 7D4OP230 | Chem_NY_Zemea_01 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Appearance | 7D4OP230/Chem_NY_Zemea_01 | 1 | both
7D4OP230/Chem_NY_Zemea_01 | 7D4OP230 | Chem_NY_Zemea_01 | NY-Analytical | 397 | Appearance | Physical Properties | Physical Properties | A Clear, Colorless Liquid (Done) |  |  |  |  | TM 208 | 7D4OP230/Chem_NY_Zemea_01 | 1 | both
7D4OP230/Chem_NY_Zemea_01 | 7D4OP230 | Chem_NY_Zemea_01 | NY-Analytical | 390 | Identification | Identification A | Identification A | Passes test. (Done) |  |  |  |  | USP/NF | 7D4OP230/Chem_NY_Zemea_01 | 1 | both
7D4OP230/Chem_NY_Zemea_01 | 7D4OP230 | Chem_NY_Zemea_01 | NY-Analytical | 390 | Identification | Identification B | Identification B | Passes test. (Done) |  |  |  |  | USP/NF | 7D4OP230/Chem_NY_Zemea_01 | 1 | both
7D4OP601/Chem_91A | 7D4OP601 | Chem_91A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 |  |  | left_only
7D4OP601/Chem_91A | 7D4OP601 | Chem_91A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  |  | left_only
7D4OP601/Chem_91A | 7D4OP601 | Chem_91A | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear, yellow liquid. (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7D4OP601/Chem_91A | 7D4OP601 | Chem_91A | NY-Analytical | 792 | Chemical Inspection | Odor (J&J TM-003188) | Odor | Meets Odor Specification or of Characteristic Odor. [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7D4OP601/Chem_91A | 7D4OP601 | Chem_91A | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7D4OP601/Chem_91A | 7D4OP601 | Chem_91A | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C |  |  | FDA,3 | 1.526 | 1.538 | J&J TMWW78015 |  |  | left_only
7D4OP601/Chem_91A | 7D4OP601 | Chem_91A | NY-Analytical | 831 | Density @ 20° C | Density | Density |  |  | FDA,3 | 1.111 | 1.131 | J&J TMWW78012 |  |  | left_only
7D4OP601/Chem_91A | 7D4OP601 | Chem_91A | NY-Analytical | 844 | % Benzoic Acid Phenoxyethanol & Dehydroacetic Acid | Dehydroacetic Acid | Dehydroacetic Acid |  | % | FDA,1 | 6.5 | 7.5 | TM 882 |  |  | left_only
7D4OP601/Chem_91A | 7D4OP601 | Chem_91A | NY-Analytical | 844 | % Benzoic Acid Phenoxyethanol & Dehydroacetic Acid | Benzoic Acid | Benzoic Acid |  | % | FDA,1 | 11 | 13 | TM 882 |  |  | left_only
7D4OP601/Chem_91A | 7D4OP601 | Chem_91A | NY-Analytical | 844 | % Benzoic Acid Phenoxyethanol & Dehydroacetic Acid | Phenoxyethanol | Phenoxyethanol |  | % | FDA,1 | 71 | 77 | TM 882 |  |  | left_only
7D4OQ028/Chem_NY_Dicaprylyl Ether | 7D4OQ028 | Chem_NY_Dicaprylyl Ether | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Refractive Index at 20°C | 7D4OQ028/Chem_NY_Dicaprylyl Ether | 6 | both
7D4OQ028/Chem_NY_Dicaprylyl Ether | 7D4OQ028 | Chem_NY_Dicaprylyl Ether | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.431 |  | FDA,3 | 1.431 | 1.435 | TM 204 | 7D4OQ028/Chem_NY_Dicaprylyl Ether | 6 | both
7D4OQ028/Chem_NY_Dicaprylyl Ether | 7D4OQ028 | Chem_NY_Dicaprylyl Ether | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411 | 7D4OQ028/Chem_NY_Dicaprylyl Ether | 6 | both
7D4OQ028/Chem_NY_Dicaprylyl Ether | 7D4OQ028 | Chem_NY_Dicaprylyl Ether | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7D4OQ028/Chem_NY_Dicaprylyl Ether | 6 | both
7D4OQ028/Chem_NY_Dicaprylyl Ether | 7D4OQ028 | Chem_NY_Dicaprylyl Ether | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7D4OQ028/Chem_NY_Dicaprylyl Ether | 6 | both
7D4OQ028/Chem_NY_Dicaprylyl Ether | 7D4OQ028 | Chem_NY_Dicaprylyl Ether | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7D4OQ028/Chem_NY_Dicaprylyl Ether | 6 | both
7D4OQ028/Chem_NY_Dicaprylyl Ether | 7D4OQ028 | Chem_NY_Dicaprylyl Ether | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7D4OQ028/Chem_NY_Dicaprylyl Ether | 6 | both
7D4OQ209/Chem_92A | 7D4OQ209 | Chem_92A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 |  |  | left_only
7D4OQ209/Chem_92A | 7D4OQ209 | Chem_92A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C |  |  | FDA,4 | 1.0399 | 1.0599 | TM 202 |  |  | left_only
7D4OQ209/Chem_92A | 7D4OQ209 | Chem_92A | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C |  |  | FDA,4 | 1.507 | 1.5272 | TM 204 |  |  | left_only
7D4OQ209/Chem_92A | 7D4OQ209 | Chem_92A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C |  |  | left_only
7D4OQ209/Chem_92A | 7D4OQ209 | Chem_92A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Baby Powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7D4OQ209/Chem_92A | 7D4OQ209 | Chem_92A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7D4OQ209/Chem_92A | 7D4OQ209 | Chem_92A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Pale Yellow Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7D4OQ214/Chem_93A | 7D4OQ214 | Chem_93A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White to Off-White |  |  |  |  | Color (TM 208) | 7D4OQ214/Chem_93A | 14 | both
7D4OQ214/Chem_93A | 7D4OQ214 | Chem_93A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH | 7D4OQ214/Chem_93A | 14 | both
7D4OQ214/Chem_93A | 7D4OQ214 | Chem_93A | NY-Analytical | 398 | pH | pH (5% Solution) | pH (5% Solution) | 4.72 |  | FDA,1 | 4 | 6 | TM 200 | 7D4OQ214/Chem_93A | 14 | both
7D4OQ214/Chem_93A | 7D4OQ214 | Chem_93A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7D4OQ214/Chem_93A | 14 | both
7D4OQ214/Chem_93A | 7D4OQ214 | Chem_93A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7D4OQ214/Chem_93A | 14 | both
7D4OQ214/Chem_93A | 7D4OQ214 | Chem_93A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7D4OQ214/Chem_93A | 14 | both
7D4OQ214/Chem_93A | 7D4OQ214 | Chem_93A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7D4OQ214/Chem_93A | 14 | both
7D4OQ967/Chem_94A | 7D4OQ967 | Chem_94A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification | 7D4OQ967/Chem_94A | 1 | both
7D4OQ967/Chem_94A | 7D4OQ967 | Chem_94A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7D4OQ967/Chem_94A | 1 | both
7D4OQ967/Chem_94A | 7D4OQ967 | Chem_94A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear viscous liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7D4OQ967/Chem_94A | 1 | both
7D4OQ967/Chem_94A | 7D4OQ967 | Chem_94A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Yellow to amber (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7D4OQ967/Chem_94A | 1 | both
7D4OQ967/Chem_94A | 7D4OQ967 | Chem_94A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7D4OQ967/Chem_94A | 1 | both
7E4OD232/Chem_97A | 7E4OD232 | Chem_97A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification | 7E4OD232/Chem_97A | 5 | both
7E4OD232/Chem_97A | 7E4OD232 | Chem_97A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7E4OD232/Chem_97A | 5 | both
7E4OD232/Chem_97A | 7E4OD232 | Chem_97A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Granular Solid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7E4OD232/Chem_97A | 5 | both
7E4OD232/Chem_97A | 7E4OD232 | Chem_97A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7E4OD232/Chem_97A | 5 | both
7E4OD232/Chem_97A | 7E4OD232 | Chem_97A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7E4OD232/Chem_97A | 5 | both
7E4OD232/Chem_97A | 7E4OD232 | Chem_97A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7E4OD232/Chem_97A | 5 | both
7E4OP065/Chem_NY_Benzyl Alcohol | 7E4OP065 | Chem_NY_Benzyl Alcohol | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection | 7E4OP065/Chem_NY_Benzyl Alcohol | 6 | both
7E4OP065/Chem_NY_Benzyl Alcohol | 7E4OP065 | Chem_NY_Benzyl Alcohol | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Colorless Liquid (Done) |  |  |  |  | TM 208 | 7E4OP065/Chem_NY_Benzyl Alcohol | 6 | both
7E4OP065/Chem_NY_Benzyl Alcohol | 7E4OP065 | Chem_NY_Benzyl Alcohol | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM 208 | 7E4OP065/Chem_NY_Benzyl Alcohol | 6 | both
7E4OP065/Chem_NY_Benzyl Alcohol | 7E4OP065 | Chem_NY_Benzyl Alcohol | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 | 7E4OP065/Chem_NY_Benzyl Alcohol | 6 | both
7E4OP065/Chem_NY_Benzyl Alcohol | 7E4OP065 | Chem_NY_Benzyl Alcohol | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208 | 7E4OP065/Chem_NY_Benzyl Alcohol | 6 | both
7E4OP065/Chem_NY_Benzyl Alcohol | 7E4OP065 | Chem_NY_Benzyl Alcohol | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411 | 7E4OP065/Chem_NY_Benzyl Alcohol | 6 | both
7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | 7E4OP203 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 |  |  | left_only
7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | 7E4OP203 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Analytical | 398 | pH | pH (10% w/v aqueous solution) | pH (10% w/v aqueous solution) |  |  | FDA,1 | 6.5 | 7.5 | TM 200 |  |  | left_only
7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | 7E4OP203 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH |  |  | left_only
7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | 7E4OP203 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Analytical | 792 | Chemical Inspection | Color (TM 230) | Color |  |  | FDA,0 |  | 6 | TM208/220/230/USP/Kao LM |  |  | left_only
7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | 7E4OP203 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear viscous liquid, free of foreign matter. (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | 7E4OP203 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Similar in character and intensity to standard. Herbal Odor (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | 7E4OP203 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Microbiology | 867 | Celsis_Rapid | Bacteria Count (Obeserved) | Bacteria Count (Obeserved) |  | cfu | FDA,0 |  | 100 | TM 603 |  |  | left_only
7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | 7E4OP203 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Microbiology | 867 | Celsis_Rapid | Bacteria Count (Reported) | Bacteria Count (Reported) |  | CFU/g | FDA,0 |  | 100 | TM 603 |  |  | left_only
7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | 7E4OP203 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Microbiology | 867 | Celsis_Rapid | Mold Count (Observed) | Mold Count (Observed) |  | cfu | FDA,0 |  |  | TM 603 |  |  | left_only
7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | 7E4OP203 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Microbiology | 867 | Celsis_Rapid | Mold Count (Reported) | Mold Count (Reported) |  | CFU/g | FDA,0 |  |  | TM 603 |  |  | left_only
7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | 7E4OP203 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Microbiology | 867 | Celsis_Rapid | Celsis_ Result | Celsis Result | Negative ATP (Done) |  |  |  |  | TM 603 |  |  | left_only
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Chem_NY_Lexfeel Natural | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 |  |  | left_only
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Chem_NY_Lexfeel Natural | NY-Microbiology | 524 | J&J Fungal | Fungal | Fungal |  | CFU/g | FDA,0 | 0 | 200 | TM 869 |  |  | left_only
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Chem_NY_Lexfeel Natural | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | TSA JJ |  | cfu/mL | FDA,0 | 0 | 2000 | TM 869 |  |  | left_only
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Chem_NY_Lexfeel Natural | NY-Microbiology | 701 | J&J TSA Controls | Negative Control | Negative Control | No Growth TSA |  |  |  |  | TM 869 |  |  | left_only
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Chem_NY_Lexfeel Natural | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | NegativeControl | No Growth TSB |  |  |  |  | TM 869 |  |  | left_only
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Chem_NY_Lexfeel Natural | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  |  | left_only
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Chem_NY_Lexfeel Natural | NY-Analytical | 792 | Chemical Inspection | Colour (J&J TM7377) | Colour |  |  | FDA,0 |  | 50 | J&J TMWW78010/J&J TM7252 |  |  | left_only
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Chem_NY_Lexfeel Natural | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear colorless to pale yellow liquid. (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Chem_NY_Lexfeel Natural | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Chem_NY_Lexfeel Natural | NY-Analytical | 801 | % Water | Moisture Content | Moisture Content |  | % | FDA,1 |  | 0.1 | J&J TMWW78017 |  |  | left_only
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Chem_NY_Lexfeel Natural | NY-Analytical | 826 | Acid Value | Acid Value | Acid Value |  | mg KOH/g | FDA,1 |  | 0.1 | J&J TMWW78011 |  |  | left_only
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Chem_NY_Lexfeel Natural | NY-Analytical | 826 | Acid Value | Acid Value Final | Acid Value Final |  | mg KOH/g | FDA,1 |  | 0.1 | J&J TMWW78011 |  |  | left_only
7E4OP601/Chem_102A | 7E4OP601 | Chem_102A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 |  |  | left_only
7E4OP601/Chem_102A | 7E4OP601 | Chem_102A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Assay |  |  | left_only
7E4OP601/Chem_102A | 7E4OP601 | Chem_102A | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear, yellow liquid. (Done) |  |  |  |  | J&J TMWW78010 |  |  | left_only
7E4OP601/Chem_102A | 7E4OP601 | Chem_102A | NY-Analytical | 796 | Assay | Assay | Assay |  | % | FDA,2 | 9 | 11 | J&J TM-002528 |  |  | left_only
7E4OP601/Chem_102A | 7E4OP601 | Chem_102A | NY-Analytical | 796 | Assay | Assay Final | Assay Final |  | % | FDA,1 | 9 | 11 | J&J TM-002528 |  |  | left_only
7E4OP601/Chem_102A | 7E4OP601 | Chem_102A | NY-Analytical | 801 | % Water | Water Content | Water Content |  | % | FDA,2 |  | 1 | J&J TMWW78017 |  |  | left_only
7E4OP806/Chem_58B | 7E4OP806 | Chem_58B | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection | 7E4OP806/Chem_58B | 9 | both
7E4OP806/Chem_58B | 7E4OP806 | Chem_58B | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear colorless liquid free from suspended materials (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | 7E4OP806/Chem_58B | 9 | both
7E4OP806/Chem_58B | 7E4OP806 | Chem_58B | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | 7E4OP806/Chem_58B | 9 | both
7E4OP806/Chem_58B | 7E4OP806 | Chem_58B | NY-Analytical | 831 | Density @ 20° C | Density | Density | 0.9221 |  | FDA,3 | 0.92 | 0.923 | J&J TMWW78012 | 7E4OP806/Chem_58B | 9 | both
7E4OP806/Chem_58B | 7E4OP806 | Chem_58B | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 | 7E4OP806/Chem_58B | 9 | both
7E4OP806/Chem_58B | 7E4OP806 | Chem_58B | NY-Analytical | 801 | % Water | Moisture Content (10 grams of sample) | Moisture Content (10 grams of sample) | 0.06 | % | FDA,3 |  | 0.1 | J&J TMWW78017 | 7E4OP806/Chem_58B | 9 | both
7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | 7E4OQ028 | Chem_NY_Ethylhexyl Hydroxystearate | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C | 7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | 2 | both
7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | 7E4OQ028 | Chem_NY_Ethylhexyl Hydroxystearate | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.8964 |  | FDA,4 | 0.892 | 0.898 | TM 202 | 7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | 2 | both
7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | 7E4OQ028 | Chem_NY_Ethylhexyl Hydroxystearate | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.4555 |  | FDA,4 | 1.4555 | 1.458 | TM 204 | 7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | 2 | both
7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | 7E4OQ028 | Chem_NY_Ethylhexyl Hydroxystearate | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411 | 7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | 2 | both
7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | 7E4OQ028 | Chem_NY_Ethylhexyl Hydroxystearate | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear to slightly hazy liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | 2 | both
7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | 7E4OQ028 | Chem_NY_Ethylhexyl Hydroxystearate | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | 2 | both
7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | 7E4OQ028 | Chem_NY_Ethylhexyl Hydroxystearate | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | 2 | both
7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | 7E4OQ028 | Chem_NY_Ethylhexyl Hydroxystearate | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Slightly Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | 2 | both
7F4BP869/Chem_NY_Citric Acid Anhydrous - USP | 7F4BP869 | Chem_NY_Citric Acid Anhydrous - USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA | 7F4BP869/Chem_NY_Citric Acid Anhydrous - USP | 13 | both
7F4BP869/Chem_NY_Citric Acid Anhydrous - USP | 7F4BP869 | Chem_NY_Citric Acid Anhydrous - USP | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Crystals/crystalline powder (Done) |  |  |  |  | TM 208 | 7F4BP869/Chem_NY_Citric Acid Anhydrous - USP | 13 | both
7F4BP869/Chem_NY_Citric Acid Anhydrous - USP | 7F4BP869 | Chem_NY_Citric Acid Anhydrous - USP | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to white (Done) |  |  |  |  | TM 208 | 7F4BP869/Chem_NY_Citric Acid Anhydrous - USP | 13 | both
7F4BP869/Chem_NY_Citric Acid Anhydrous - USP | 7F4BP869 | Chem_NY_Citric Acid Anhydrous - USP | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 | 7F4BP869/Chem_NY_Citric Acid Anhydrous - USP | 13 | both
7F4BP869/Chem_NY_Citric Acid Anhydrous - USP | 7F4BP869 | Chem_NY_Citric Acid Anhydrous - USP | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless (Done) |  |  |  |  | TM 208 | 7F4BP869/Chem_NY_Citric Acid Anhydrous - USP | 13 | both
7F4BP869/Chem_NY_Citric Acid Anhydrous - USP | 7F4BP869 | Chem_NY_Citric Acid Anhydrous - USP | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7F4BP869/Chem_NY_Citric Acid Anhydrous - USP | 13 | both
7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | 7F4OA160 | Chem_NY_PEG-40 Hydrogenated Castor Oil | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH | 7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | 3 | both
7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | 7F4OA160 | Chem_NY_PEG-40 Hydrogenated Castor Oil | NY-Analytical | 398 | pH | pH (10% in water) | pH (10% in water) | 6.99 |  | FDA,1 | 6 | 7 | TM 200 | 7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | 3 | both
7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | 7F4OA160 | Chem_NY_PEG-40 Hydrogenated Castor Oil | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | 3 | both
7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | 7F4OA160 | Chem_NY_PEG-40 Hydrogenated Castor Oil | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear viscous liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | 3 | both
7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | 7F4OA160 | Chem_NY_PEG-40 Hydrogenated Castor Oil | NY-Analytical | 792 | Chemical Inspection | Color / Appearance (TM 208) | Color / Appearance | White to yellowish soft or flowing paste. Close Match to Standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | 3 | both
7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | 7F4OA160 | Chem_NY_PEG-40 Hydrogenated Castor Oil | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | 3 | both
7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | 7F4OA160 | Chem_NY_PEG-40 Hydrogenated Castor Oil | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Close match to standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | 3 | both
7F4OD110/Chem_106A | 7F4OD110 | Chem_106A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH | 7F4OD110/Chem_106A | 6 | both
7F4OD110/Chem_106A | 7F4OD110 | Chem_106A | NY-Analytical | 398 | pH | pH (1% Solution) | pH (1% Solution) | 10.6 |  | FDA,1 | 10.5 | 11.5 | TM 200 | 7F4OD110/Chem_106A | 6 | both
7F4OD110/Chem_106A | 7F4OD110 | Chem_106A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7F4OD110/Chem_106A | 6 | both
7F4OD110/Chem_106A | 7F4OD110 | Chem_106A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Crystalline Powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7F4OD110/Chem_106A | 6 | both
7F4OD110/Chem_106A | 7F4OD110 | Chem_106A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7F4OD110/Chem_106A | 6 | both
7F4OD110/Chem_106A | 7F4OD110 | Chem_106A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7F4OD110/Chem_106A | 6 | both
7F4OD110/Chem_106A | 7F4OD110 | Chem_106A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7F4OD110/Chem_106A | 6 | both
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | Chem_NY_Arlasilk Phospholipid PTC | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 |  |  | left_only
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 509 | J&J_Pathogens | Cetrimide | Cetrimide | No Growth |  |  |  |  | TM 869 |  |  | left_only
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 509 | J&J_Pathogens | MacConkey | MacConkey | No Growth |  |  |  |  | TM 869 |  |  | left_only
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 509 | J&J_Pathogens | TSA | TSA | No Growth |  |  |  |  | TM 869 |  |  | left_only
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 509 | J&J_Pathogens | MSA or VJ | MSA or VJ | No Growth_MSA |  |  |  |  | TM 869 |  |  | left_only
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 509 | J&J_Pathogens | Biggy or SDA | Biggy or SDA | No Growth_SDA |  |  |  |  | TM 869 |  |  | left_only
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 517 | J&J_Enrichment | TSB | TSB | Turbid |  |  |  |  | TM 869 |  |  | left_only
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 524 | J&J Fungal | Fungal | Fungal |  | CFU/g | FDA,0 | 0 | 200 | TM 869 |  |  | left_only
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | TSA JJ |  | cfu/mL | FDA,0 | 0 | 2000 | TM 869 |  |  | left_only
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 700 | J&J SDA Controls | Negative Control | Negative Control | No Growth SDA |  |  |  |  | TM 869 |  |  | left_only
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | NegativeControl | No Growth TSB |  |  |  |  | TM 869 |  |  | left_only
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | Chem_NY_Arlasilk Phospholipid PTC | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | color (Gardner) |  |  | left_only
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | Chem_NY_Arlasilk Phospholipid PTC | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear amber liquid. (Done) |  |  |  |  | J&J TMWW78010 |  |  | left_only
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | Chem_NY_Arlasilk Phospholipid PTC | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010 |  |  | left_only
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | Chem_NY_Arlasilk Phospholipid PTC | NY-Analytical | 792 | Chemical Inspection | Odor (J&J TM-003188) | Odor | Passes Test (Done) |  |  |  |  | J&J TMWW78010 |  |  | left_only
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | Chem_NY_Arlasilk Phospholipid PTC | NY-Analytical | 944 | color (Gardner) | color (Gardner) | color (Gardner) |  |  | FDA,0 |  | 4 | J&J TM701 |  |  | left_only
7F4OP601/Chem_109A | 7F4OP601 | Chem_109A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 |  |  | left_only
7F4OP601/MASTER - Microbiology - J&J Chemical Testing) | 7F4OP601 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 396 | Odor | Odor | Odor | Passes Test (Done) |  |  |  |  | J&J TM-003188 | 7F4OP601/MASTER - Microbiology - J&J Chemical Testing) | 4 | both
7F4OP601/Chem_109A | 7F4OP601 | Chem_109A | NY-Analytical | 396 | Odor | Odor | Odor | Passes Test (Done) |  |  |  |  | J&J TM-003188 |  |  | left_only
7F4OP601/MASTER - Microbiology - J&J Chemical Testing) | 7F4OP601 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear colorless liquid with oily consistency (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | 7F4OP601/MASTER - Microbiology - J&J Chemical Testing) | 4 | both
7F4OP601/Chem_109A | 7F4OP601 | Chem_109A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Odor |  |  | left_only
7F4OP601/Chem_109A | 7F4OP601 | Chem_109A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear colorless liquid with oily consistency (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7F4OP601/Chem_109A | 7F4OP601 | Chem_109A | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7F4OP806/Chem_NY_Phenoxyethanol_7F4OP806 | 7F4OP806 | Chem_NY_Phenoxyethanol_7F4OP806 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 |  |  | left_only
7F4OP806/Chem_NY_Phenoxyethanol_7F4OP806 | 7F4OP806 | Chem_NY_Phenoxyethanol_7F4OP806 | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C |  |  | FDA,3 | 1.105 | 1.11 | USP/NF |  |  | left_only
7F4OP806/Chem_NY_Phenoxyethanol_7F4OP806 | 7F4OP806 | Chem_NY_Phenoxyethanol_7F4OP806 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  |  | left_only
7F4OP806/Chem_NY_Phenoxyethanol_7F4OP806 | 7F4OP806 | Chem_NY_Phenoxyethanol_7F4OP806 | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear, thin colorless liquid (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7F4OP806/Chem_NY_Phenoxyethanol_7F4OP806 | 7F4OP806 | Chem_NY_Phenoxyethanol_7F4OP806 | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7F4OP806/Chem_NY_Phenoxyethanol_7F4OP806 | 7F4OP806 | Chem_NY_Phenoxyethanol_7F4OP806 | NY-Analytical | 905 | Phenol | Phenol | Phenol |  | ppm | FDA,0 |  | 25 | USP/NF |  |  | left_only
7F4OP806/Chem_NY_Phenoxyethanol_7F4OP806 | 7F4OP806 | Chem_NY_Phenoxyethanol_7F4OP806 | NY-Analytical | 986 | Phenoxyethanol Content | Phenoxyethanol Assay | Phenoxyethanol Assay |  | % | FDA,1 | 98 | 102 | USP/NF |  |  | left_only
7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | 7F4OQ028 | Chem_NY_Octyldodecyl Stearoyl Stearate | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C | 7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | 2 | both
7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | 7F4OQ028 | Chem_NY_Octyldodecyl Stearoyl Stearate | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.8747 |  | FDA,4 | 0.865 | 0.885 | TM 202 | 7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | 2 | both
7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | 7F4OQ028 | Chem_NY_Octyldodecyl Stearoyl Stearate | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.4585 |  | FDA,4 | 1.455 | 1.46 | TM 204 | 7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | 2 | both
7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | 7F4OQ028 | Chem_NY_Octyldodecyl Stearoyl Stearate | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411 | 7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | 2 | both
7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | 7F4OQ028 | Chem_NY_Octyldodecyl Stearoyl Stearate | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | 2 | both
7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | 7F4OQ028 | Chem_NY_Octyldodecyl Stearoyl Stearate | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Light Yellow Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | 2 | both
7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | 7F4OQ028 | Chem_NY_Octyldodecyl Stearoyl Stearate | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | 2 | both
7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | 7F4OQ028 | Chem_NY_Octyldodecyl Stearoyl Stearate | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Mild Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | 2 | both
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | 7G4OP027 | Chem_NY_SD Alcohol 40-B 200 Proof_ADM | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis | 7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | 2 | both
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | 7G4OP027 | Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Identity of Denatonium Benzoate | 7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | 7 | both
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | 7G4OP027 | Chem_NY_SD Alcohol 40-B 200 Proof_ADM | NY-Analytical | 824 | Identity of Denatonium Benzoate | Identification | Identification | Passes test with 5cm Quartz (Done) |  |  |  |  | J&J TM4028 | 7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | 2 | both
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | 7G4OP027 | Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | NY-Analytical | 824 | Identity of Denatonium Benzoate | Identification | Identification | [5 empty] | J&J TM4028 | 7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | 7 | both
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | 7G4OP027 | Chem_NY_SD Alcohol 40-B 200 Proof_ADM | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear colorless liquid. (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | 7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | 2 | both
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | 7G4OP027 | Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | [5 empty] | J&J TMWW78010/J&J TM7252 | 7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | 7 | both
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | 7G4OP027 | Chem_NY_SD Alcohol 40-B 200 Proof_ADM | NY-Analytical | 792 | Chemical Inspection | Odor (J&J TM-003188) | Odor | Passes Test (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | 7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | 2 | both
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | 7G4OP027 | Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | NY-Analytical | 792 | Chemical Inspection | Odor (J&J TM-003188) | Odor | [5 empty] | J&J TMWW78010/J&J TM7252 | 7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | 7 | both
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | 7G4OP027 | Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | [5 empty] | J&J TMWW78010/J&J TM7252 | 7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | 7 | both
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | 7G4OP027 | Chem_NY_SD Alcohol 40-B 200 Proof_ADM | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | 7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | 2 | both
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | 7G4OP027 | Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C |  |  | FDA,4 | 0.7905 | 0.792 | J&J TMWW78012 | 7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | 7 | both
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | 7G4OP027 | Chem_NY_SD Alcohol 40-B 200 Proof_ADM | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C |  |  | FDA,4 | 0.7905 | 0.792 | J&J TMWW78012 | 7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | 2 | both
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | 7G4OP027 | Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | NY-Analytical | 822 | Specific Gravity @ 15.56°C/ 60° F | Specific Gravity | Specific Gravity at 15.56°C | 0.7942 |  | FDA,4 |  | 0.7952 | J&J TMWW78012 | 7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | 7 | both
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | 7G4OP027 | Chem_NY_SD Alcohol 40-B 200 Proof_ADM | NY-Analytical | 822 | Specific Gravity @ 15.56°C/ 60° F | Specific Gravity | Specific Gravity at 15.56°C | 0.7941 |  | FDA,4 |  | 0.7952 | J&J TMWW78012 | 7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | 2 | both
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | 7G4OP027 | Chem_NY_SD Alcohol 40-B 200 Proof_ADM | NY-Analytical | 390 | Identification | IR Identification | IR Identification | [5 empty] | J&J TMWW78013 | 7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | 2 | both
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | 7G4OP027 | Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | NY-Analytical | 390 | Identification | IR Identification | IR Identification | [5 empty] | J&J TMWW78013 | 7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | 7 | both
7G4OP308/Chem_111A | 7G4OP308 | Chem_111A | NY-Analytical | 792 | Chemical Inspection | Odor | Odor | Matches |  |  |  |  | TM 208 | 7G4OP308/Chem_111A | 1 | both
7G4OP308/Chem_111A | 7G4OP308 | Chem_111A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection | 7G4OP308/Chem_111A | 1 | both
7G4OP308/Chem_111A | 7G4OP308 | Chem_111A | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Colorless to Light Yellow Liquid, Free of Particulate Matter (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | 7G4OP308/Chem_111A | 1 | both
7G4OP308/Chem_111A | 7G4OP308 | Chem_111A | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | 7G4OP308/Chem_111A | 1 | both
7G4OP308/Chem_111A | 7G4OP308 | Chem_111A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.9464 |  | FDA,3 | 0.942 | 0.952 | J&J TMWW78012 | 7G4OP308/Chem_111A | 1 | both
7G4OP308/Chem_111A | 7G4OP308 | Chem_111A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 | 7G4OP308/Chem_111A | 1 | both
7G4OP308/Chem_111A | 7G4OP308 | Chem_111A | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.4632 |  | FDA,3 | 1.463 | 1.473 | J&J TMWW78015 | 7G4OP308/Chem_111A | 1 | both
7G4OP601/MASTER - Microbiology - J&J Chemical Testing) | 7G4OP601 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C |  |  | FDA,4 | 1.451 | 1.453 | J&J TMWW78015 |  |  | left_only
7G4OP601/MASTER - Microbiology - J&J Chemical Testing) | 7G4OP601 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 792 | Chemical Inspection | Color (J&J TM-000839) | Color |  |  | FDA,1 |  | 1 | J&J TMWW78010/J&J TM7252 |  |  | left_only
7G4OP601/MASTER - Microbiology - J&J Chemical Testing) | 7G4OP601 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 810 | Saponification Value | Saponification Value | Saponification Value |  | mg KOH/g | FDA,1 | 100 | 115 | J&J TMWW78016 |  |  | left_only
7G4OP601/MASTER - Microbiology - J&J Chemical Testing) | 7G4OP601 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 810 | Saponification Value | Saponification Value Final | Saponification Value Final |  | mg KOH/g | FDA,1 | 100 | 115 | J&J TMWW78016 |  |  | left_only
7G4OP601/MASTER - Microbiology - J&J Chemical Testing) | 7G4OP601 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 826 | Acid Value | Acid Value Final | Acid Value Final |  | mg KOH/g | FDA,1 |  | 2 | J&J TMWW78011 |  |  | left_only
7G4OP601/MASTER - Microbiology - J&J Chemical Testing) | 7G4OP601 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 828 | Iodine Value | Iodine Value | Iodine Value |  | g I2/100g | FDA,1 |  | 2 | J&J TM-003029 |  |  | left_only
7G4OP601/MASTER - Microbiology - J&J Chemical Testing) | 7G4OP601 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 828 | Iodine Value | Iodine Value Final | Iodine Value Final |  | g I2/100g | FDA,1 |  | 2 | J&J TM-003029 |  |  | left_only
7G4OQ028/Chem_Aminomethylpropanol | 7G4OQ028 | Chem_Aminomethylpropanol | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 |  |  | left_only
7G4OQ028/Chem_Aminomethylpropanol | 7G4OQ028 | Chem_Aminomethylpropanol | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  |  | left_only
7G4OQ028/Chem_Aminomethylpropanol | 7G4OQ028 | Chem_Aminomethylpropanol | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7G4OQ028/Chem_Aminomethylpropanol | 7G4OQ028 | Chem_Aminomethylpropanol | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7G4OQ028/Chem_Aminomethylpropanol | 7G4OQ028 | Chem_Aminomethylpropanol | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7G4OQ028/Chem_Aminomethylpropanol | 7G4OQ028 | Chem_Aminomethylpropanol | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Chem_NY_Feverfew CMP-02A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 |  |  | left_only
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 509 | J&J_Pathogens | Cetrimide | Cetrimide | No Growth |  |  |  |  | TM 869 |  |  | left_only
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 509 | J&J_Pathogens | MacConkey | MacConkey | No Growth |  |  |  |  | TM 869 |  |  | left_only
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 509 | J&J_Pathogens | TSA | TSA | No Growth |  |  |  |  | TM 869 |  |  | left_only
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 509 | J&J_Pathogens | MSA or VJ | MSA or VJ | No Growth_MSA |  |  |  |  | TM 869 |  |  | left_only
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 509 | J&J_Pathogens | Biggy or SDA | Biggy or SDA | No Growth_SDA |  |  |  |  | TM 869 |  |  | left_only
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 517 | J&J_Enrichment | TSB | TSB | Not Turbid |  |  |  |  | TM 869 |  |  | left_only
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 524 | J&J Fungal | Fungal | Fungal |  | CFU/g | FDA,0 | 0 | 200 | TM 869 |  |  | left_only
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | TSA JJ |  | cfu/mL | FDA,0 | 0 | 200 | TM 869 |  |  | left_only
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 700 | J&J SDA Controls | Negative Control | Negative Control | No Growth SDA |  |  |  |  | TM 869 |  |  | left_only
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | NegativeControl | No Growth TSB |  |  |  |  | TM 869 |  |  | left_only
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Chem_NY_Feverfew CMP-02A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  |  | left_only
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Chem_NY_Feverfew CMP-02A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Chem_NY_Feverfew CMP-02A | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Chem_NY_Feverfew CMP-02A | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Red Brown Liquid (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7H4OP601/Chem_NY_DUB PTO | 7H4OP601 | Chem_NY_DUB PTO | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 |  |  | left_only
7H4OP601/Chem_NY_DUB PTO | 7H4OP601 | Chem_NY_DUB PTO | NY-Microbiology | 524 | J&J Fungal | Fungal | Fungal |  | CFU/g | FDA,0 | 0 | 200 | TM 869 |  |  | left_only
7H4OP601/J&J Chemical CHE DUB PTO | 7H4OP601 | J&J Chemical CHE DUB PTO | NY-Microbiology | 524 | J&J Fungal | Fungal | Fungal |  | CFU/g | FDA,0 | 0 | 200 | TM 869 |  |  | left_only
7H4OP601/Chem_NY_DUB PTO | 7H4OP601 | Chem_NY_DUB PTO | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | TSA JJ |  | cfu/mL | FDA,0 | 0 | 2000 | TM 869 |  |  | left_only
7H4OP601/J&J Chemical CHE DUB PTO | 7H4OP601 | J&J Chemical CHE DUB PTO | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | TSA JJ |  | cfu/mL | FDA,0 | 0 | 2000 | TM 869 |  |  | left_only
7H4OP601/Chem_NY_DUB PTO | 7H4OP601 | Chem_NY_DUB PTO | NY-Microbiology | 701 | J&J TSA Controls | Negative Control | Negative Control | [5 empty] | TM 869 |  |  | left_only
7H4OP601/J&J Chemical CHE DUB PTO | 7H4OP601 | J&J Chemical CHE DUB PTO | NY-Microbiology | 701 | J&J TSA Controls | Negative Control | Negative Control | [5 empty] | TM 869 |  |  | left_only
7H4OP601/Chem_NY_DUB PTO | 7H4OP601 | Chem_NY_DUB PTO | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | NegativeControl | No Growth TSB |  |  |  |  | TM 869 |  |  | left_only
7H4OP601/J&J Chemical CHE DUB PTO | 7H4OP601 | J&J Chemical CHE DUB PTO | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | NegativeControl | [5 empty] | TM 869 |  |  | left_only
7H4OP601/Chem_NY_DUB PTO | 7H4OP601 | Chem_NY_DUB PTO | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  |  | left_only
7H4OP601/Chem_NY_DUB PTO | 7H4OP601 | Chem_NY_DUB PTO | NY-Analytical | 792 | Chemical Inspection | Colour (J&J TM7377) | Colour |  |  | FDA,0 |  | 60 | J&J TMWW78010/J&J TM7252 |  |  | left_only
7H4OP601/Chem_NY_DUB PTO | 7H4OP601 | Chem_NY_DUB PTO | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Colorless to pale yellow liquid (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7H4OP601/Chem_NY_DUB PTO | 7H4OP601 | Chem_NY_DUB PTO | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7H4OP601/Chem_NY_DUB PTO | 7H4OP601 | Chem_NY_DUB PTO | NY-Analytical | 792 | Chemical Inspection | Odor (J&J TM-003188) | Odor | Passes Test (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7H4OP806/Chem_NY_Kelp & Sea Salt Fragrance | 7H4OP806 | Chem_NY_Kelp & Sea Salt Fragrance | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 |  |  | left_only
7H4OP806/Chem_NY_Kelp & Sea Salt Fragrance | 7H4OP806 | Chem_NY_Kelp & Sea Salt Fragrance | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C |  |  | FDA,3 | 0.957 | 0.967 | J&J TMWW78012 |  |  | left_only
7H4OP806/Chem_NY_Kelp & Sea Salt Fragrance | 7H4OP806 | Chem_NY_Kelp & Sea Salt Fragrance | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | GC Profile |  |  | left_only
7H4OP806/Chem_NY_Kelp & Sea Salt Fragrance | 7H4OP806 | Chem_NY_Kelp & Sea Salt Fragrance | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear yellow to dark yellow liquid (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7H4OP806/Chem_NY_Kelp & Sea Salt Fragrance | 7H4OP806 | Chem_NY_Kelp & Sea Salt Fragrance | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7H4OP806/Chem_NY_Kelp & Sea Salt Fragrance | 7H4OP806 | Chem_NY_Kelp & Sea Salt Fragrance | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C |  |  | FDA,4 | 1.456 | 1.466 | J&J TMWW78015 |  |  | left_only
7H4OP806/Chem_NY_Kelp & Sea Salt Fragrance | 7H4OP806 | Chem_NY_Kelp & Sea Salt Fragrance | NY-Analytical | 839 | GC Profile | GC Profile | GC Profile | To match standard. Not required when performing IR test. (Done) |  |  |  |  | J&J TM-003030 |  |  | left_only
7J4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7J4OP027 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 447 | pH as is | pH | pH |  |  | FDA,1 | 4 | 6.5 | TM WW78014 |  |  | left_only
7J4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7J4OP027 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 792 | Chemical Inspection | Solubility | Solubility | Passes Test Soluble in any proportion of water (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7J4OP485/Chem_NY_Butylene Glycol_7J4OP485 | 7J4OP485 | Chem_NY_Butylene Glycol_7J4OP485 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 |  |  | left_only
7J4OP485/Chem_NY_Butylene Glycol_7J4OP485 | 7J4OP485 | Chem_NY_Butylene Glycol_7J4OP485 | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C |  |  | FDA,3 | 1.004 | 1.006 | J&J TMWW78012 |  |  | left_only
7J4OP485/Chem_NY_Butylene Glycol_7J4OP485 | 7J4OP485 | Chem_NY_Butylene Glycol_7J4OP485 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  |  | left_only
7J4OP485/Chem_NY_Butylene Glycol_7J4OP485 | 7J4OP485 | Chem_NY_Butylene Glycol_7J4OP485 | NY-Analytical | 792 | Chemical Inspection | Colour (J&J TMP-552) | Colour |  |  | FDA,0 |  | 10 | J&J TM-003188 |  |  | left_only
7J4OP485/Chem_NY_Butylene Glycol_7J4OP485 | 7J4OP485 | Chem_NY_Butylene Glycol_7J4OP485 | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Colorless viscous liquid (Done) |  |  |  |  | J&J TM-003188 |  |  | left_only
7J4OP485/Chem_NY_Butylene Glycol_7J4OP485 | 7J4OP485 | Chem_NY_Butylene Glycol_7J4OP485 | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TM-003188 |  |  | left_only
7J4OP485/Chem_NY_Butylene Glycol_7J4OP485 | 7J4OP485 | Chem_NY_Butylene Glycol_7J4OP485 | NY-Analytical | 792 | Chemical Inspection | Odor (J&J TM-003188) | Odor | Passes Test (Done) |  |  |  |  | J&J TM-003188 |  |  | left_only
7J4OP485/Chem_NY_Butylene Glycol_7J4OP485 | 7J4OP485 | Chem_NY_Butylene Glycol_7J4OP485 | NY-Analytical | 795 | Acidity | Acidity as Acetic Acid(10ml burette 0.05ml subdiv) | Acidity as Acetic Acid(10ml burette 0.05ml subdiv) |  | % | FDA,3 |  | 0.005 | J&J TM-003589 |  |  | left_only
7J4OP485/Chem_NY_Butylene Glycol_7J4OP485 | 7J4OP485 | Chem_NY_Butylene Glycol_7J4OP485 | NY-Analytical | 796 | Assay | Assay | Assay |  | % | FDA,2 | 99.5 |  | J&J TM4075 |  |  | left_only
7J4OP485/Chem_NY_Butylene Glycol_7J4OP485 | 7J4OP485 | Chem_NY_Butylene Glycol_7J4OP485 | NY-Analytical | 796 | Assay | Assay Final | Assay Final |  | % | FDA,2 | 99.5 |  | J&J TM4075 |  |  | left_only
7J4OP485/Chem_NY_Butylene Glycol_7J4OP485 | 7J4OP485 | Chem_NY_Butylene Glycol_7J4OP485 | NY-Analytical | 801 | % Water | Water Content | Water Content |  | % | FDA,1 |  | 0.5 | J&J TMWW78017 |  |  | left_only
7K4OP027/Chem_116A | 7K4OP027 | Chem_116A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 |  |  | left_only
7K4OP027/Chem_116A | 7K4OP027 | Chem_116A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  |  | left_only
7K4OP027/Chem_116A | 7K4OP027 | Chem_116A | NY-Analytical | 792 | Chemical Inspection | Odor (CoA) | Odor (CoA) | Certified on CoA as Meets Odor Specification or of Characteristic Odor. (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7K4OP027/Chem_116A | 7K4OP027 | Chem_116A | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7K4OP027/Chem_116A | 7K4OP027 | Chem_116A | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Pale yellow to yellow liquid (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection | 7L4OP027/Chem_NY_Menthol USP | 2 | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Colorless to white needle crystal or granule (Done) |  |  |  |  | JJTMWW78010 / JJTM7252 | 7L4OP027/Chem_NY_Menthol USP | 2 | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | JJTMWW78010 / JJTM7252 | 7L4OP027/Chem_NY_Menthol USP | 2 | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 809 | Melting Range | Melting Range - End | Melting Range - End | 42.5 | C | FDA,0 | 41 | 44 | USP | 7L4OP027/Chem_NY_Menthol USP | 2 | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 809 | Melting Range | Melting Range - Start | Melting Range - Start | 41.2 | C | FDA,0 | 41 | 44 | USP | 7L4OP027/Chem_NY_Menthol USP | 2 | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 797 | Nonvolatile Residue | Nonvolatile Residue | Nonvolatile Residue | 0.02 | % w/v | FDA,2 |  | 0.05 | USP | 7L4OP027/Chem_NY_Menthol USP | 2 | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 863 | Specific Rotation | Specific Rotation | Specific Rotation | -49.9 | ° | FDA,0 | -51 | -45 | USP | 7L4OP027/Chem_NY_Menthol USP | 2 | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 796 | Assay | Assay | Assay | 100.4 | % | FDA,1 | 98 | 102 | USP/NF | 7L4OP027/Chem_NY_Menthol USP | 2 | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 796 | Assay | Assay Final | Assay Final | 100.4 | % | FDA,1 | 98 | 102 | USP/NF | 7L4OP027/Chem_NY_Menthol USP | 2 | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 455 | Identification A | Identification A | Identification A | Passes Test (Done) |  |  |  |  | USP/NF | 7L4OP027/Chem_NY_Menthol USP | 2 | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 456 | Identification B | Identification B | Identification B | Passes Test (Done) |  |  |  |  | USP/NF | 7L4OP027/Chem_NY_Menthol USP | 2 | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 964 | Individual Impurity | Iso-Menthone | Iso-Menthone | 0.3 | % | FDA,1 |  | 0.3 | USP/NF | 7L4OP027/Chem_NY_Menthol USP | 2 | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 964 | Individual Impurity | Iso-Pulegol | Iso-Pulegol | 0.14 | % | FDA,1 |  | 0.3 | USP/NF | 7L4OP027/Chem_NY_Menthol USP | 2 | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 964 | Individual Impurity | Menthone | Menthone | 0.04 | % | FDA,1 |  | 0.3 | USP/NF | 7L4OP027/Chem_NY_Menthol USP | 2 | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 964 | Individual Impurity | Neo-Menthol | Neo-Menthol | 0.18 | % | FDA,1 |  | 0.3 | USP/NF | 7L4OP027/Chem_NY_Menthol USP | 2 | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 964 | Individual Impurity | Pulegone | Pulegone | 0.08 | % | FDA,1 |  | 0.3 | USP/NF | 7L4OP027/Chem_NY_Menthol USP | 2 | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 965 | Total Impurities | Total Impurities | Total Impurities | 0.53 | % | FDA,1 |  | 2 | USP/NF | 7L4OP027/Chem_NY_Menthol USP | 2 | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 964 | Individual Impurity | Unknown impurities | Unknown impurities | 0.05 | % | FDA,1 |  | 0.1 | USP/NF | 7L4OP027/Chem_NY_Menthol USP | 2 | both
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 |  |  | left_only
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C |  |  | FDA,3 | 1.249 |  | TM 202 |  |  | left_only
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  |  | left_only
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, syrupy liquid, free of foreign matter. (Done) |  |  |  |  | See Analyte for Method # |  |  | left_only
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Similar in character and intensity to standard, practically odorless (Done) |  |  |  |  | See Analyte for Method # |  |  | left_only
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) |  |  |  |  | See Analyte for Method # |  |  | left_only
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 801 | % Water | Water | Water |  | % | FDA,1 |  | 5 | APAM 6.018 |  |  | left_only
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 825 | Limit of Ethylene Glycol and Diethylene Glycol | Diethylene Glycol | Diethylene Glycol |  | % | FDA,1 |  | 0.1 | USP |  |  | left_only
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 825 | Limit of Ethylene Glycol and Diethylene Glycol | Ethylene Glycol | Ethylene Glycol |  | % | FDA,1 |  | 0.1 | USP |  |  | left_only
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 825 | Limit of Ethylene Glycol and Diethylene Glycol | Individual Impurity | Individual Impurity |  | % | FDA,1 |  | 0.1 | USP |  |  | left_only
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 825 | Limit of Ethylene Glycol and Diethylene Glycol | Total Impurities | Total Impurities |  | % | FDA,0 |  | 1 | USP |  |  | left_only
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Chem_NY_Glycerin - RI-0420 | NY-Microbiology | 867 | Celsis_Rapid | Bacteria Count (Obeserved) | Bacteria Count (Obeserved) |  | cfu | FDA,0 |  | 100 | TM 603 |  |  | left_only
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Chem_NY_Glycerin - RI-0420 | NY-Microbiology | 867 | Celsis_Rapid | Bacteria Count (Reported) | Bacteria Count (Reported) |  | CFU/g | FDA,0 |  | 100 | TM 603 |  |  | left_only
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Chem_NY_Glycerin - RI-0420 | NY-Microbiology | 867 | Celsis_Rapid | Mold Count (Observed) | Mold Count (Observed) |  | cfu | FDA,0 |  |  | TM 603 |  |  | left_only
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Chem_NY_Glycerin - RI-0420 | NY-Microbiology | 867 | Celsis_Rapid | Mold Count (Reported) | Mold Count (Reported) |  | CFU/g | FDA,0 |  |  | TM 603 |  |  | left_only
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Chem_NY_Glycerin - RI-0420 | NY-Microbiology | 867 | Celsis_Rapid | Celsis_ Result | Celsis Result | Negative ATP (Done) |  |  |  |  | TM 603 |  |  | left_only
7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | 7L4OP485 | Chem_NY_Polysorbate 20_7L4OP485 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 |  |  | left_only
7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | 7L4OP485 | Chem_NY_Polysorbate 20_7L4OP485 | NY-Analytical | 396 | Odor | Odor | Odor | Passes Test (Done) |  |  |  |  | J&J TM-003188 |  |  | left_only
7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | 7L4OP485 | Chem_NY_Polysorbate 20_7L4OP485 | NY-Microbiology | 524 | J&J Fungal | Fungal | Fungal |  | CFU/g | FDA,0 | 0 | 200 | TM 869 |  |  | left_only
7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | 7L4OP485 | Chem_NY_Polysorbate 20_7L4OP485 | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | TSA JJ |  | CFU/g | FDA,0 | 0 | 2000 | TM 869 |  |  | left_only
7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | 7L4OP485 | Chem_NY_Polysorbate 20_7L4OP485 | NY-Microbiology | 723 | J&J TSB Controls - TSA | Negative Control | Negative Control | No Growth TSB |  |  |  |  | TM 869 |  |  | left_only
7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | 7L4OP485 | Chem_NY_Polysorbate 20_7L4OP485 | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | NegativeControl | No Growth TSB |  |  |  |  | TM 869 |  |  | left_only
7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | 7L4OP485 | Chem_NY_Polysorbate 20_7L4OP485 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Odor |  |  | left_only
7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | 7L4OP485 | Chem_NY_Polysorbate 20_7L4OP485 | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | A pale yellow to an amber colored liquid, free of foreign matter. (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | 7L4OP485 | Chem_NY_Polysorbate 20_7L4OP485 | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection | 7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 1 | both
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear colorless liquid. (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | 7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 1 | both
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | 7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 1 | both
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 796 | Assay | Assay | Assay | 48.2 | % | FDA,1 | 47.5 | 52.5 | USP/NF | 7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 1 | both
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 796 | Assay | Assay Final | Assay Final | 48.2 | % | FDA,1 | 47.5 | 52.5 | USP/NF | 7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 1 | both
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 846 | Ratio of Alkyl Compounds | C12 + C14 Content of Total | C12 + C14 Content of Total | 95.7 | % | FDA,1 | 70 |  | USP/NF | 7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 1 | both
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 846 | Ratio of Alkyl Compounds | C12 Content of Total | C12 Content of Total | 56.6 | % | FDA,1 | 40 |  | USP/NF | 7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 1 | both
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 846 | Ratio of Alkyl Compounds | C14 Content of Total | C14 Content of Total | 39.1 | % | FDA,1 | 20 |  | USP/NF | 7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 1 | both
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 390 | Identification | Identification A | Identification A | Passes test. (Done) |  |  |  |  | USP/NF | 7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 1 | both
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 390 | Identification | Identification B | Identification B | Passes test. (Done) |  |  |  |  | USP/NF | 7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 1 | both
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 390 | Identification | Identification C | Identification C | Passes test. (Done) |  |  |  |  | USP/NF | 7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 1 | both
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 390 | Identification | Identification D | Identification D | Passes test. (Done) |  |  |  |  | USP/NF | 7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 1 | both
7O4AA580/Chem_NY_Ethylene Glycol Monohexyl Ether | 7O4AA580 | Chem_NY_Ethylene Glycol Monohexyl Ether | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification | 7O4AA580/Chem_NY_Ethylene Glycol Monohexyl Ether | 3 | both
7O4AA580/Chem_NY_Ethylene Glycol Monohexyl Ether | 7O4AA580 | Chem_NY_Ethylene Glycol Monohexyl Ether | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411 | 7O4AA580/Chem_NY_Ethylene Glycol Monohexyl Ether | 3 | both
7O4AA580/Chem_NY_Ethylene Glycol Monohexyl Ether | 7O4AA580 | Chem_NY_Ethylene Glycol Monohexyl Ether | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, Colorless, Viscous Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4AA580/Chem_NY_Ethylene Glycol Monohexyl Ether | 3 | both
7O4AA580/Chem_NY_Ethylene Glycol Monohexyl Ether | 7O4AA580 | Chem_NY_Ethylene Glycol Monohexyl Ether | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4AA580/Chem_NY_Ethylene Glycol Monohexyl Ether | 3 | both
7O4AA580/Chem_NY_Ethylene Glycol Monohexyl Ether | 7O4AA580 | Chem_NY_Ethylene Glycol Monohexyl Ether | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4AA580/Chem_NY_Ethylene Glycol Monohexyl Ether | 3 | both
7O4AA580/Chem_NY_Ethylene Glycol Monohexyl Ether | 7O4AA580 | Chem_NY_Ethylene Glycol Monohexyl Ether | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Pungent (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4AA580/Chem_NY_Ethylene Glycol Monohexyl Ether | 3 | both
7O4AQ084/Chem_162A | 7O4AQ084 | Chem_162A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection | 7O4AQ084/Chem_162A | 2 | both
7O4AQ084/Chem_162A | 7O4AQ084 | Chem_162A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4AQ084/Chem_162A | 2 | both
7O4AQ084/Chem_162A | 7O4AQ084 | Chem_162A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Clear/Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4AQ084/Chem_162A | 2 | both
7O4AQ084/Chem_162A | 7O4AQ084 | Chem_162A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4AQ084/Chem_162A | 2 | both
7O4AQ084/Chem_162A | 7O4AQ084 | Chem_162A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Light Amine (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4AQ084/Chem_162A | 2 | both
7O4BJ131/Chem_NY_Acticide CBM-2 | 7O4BJ131 | Chem_NY_Acticide CBM-2 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification | 7O4BJ131/Chem_NY_Acticide CBM-2 | 4 | both
7O4BJ131/Chem_NY_Acticide CBM-2 | 7O4BJ131 | Chem_NY_Acticide CBM-2 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7O4BJ131/Chem_NY_Acticide CBM-2 | 4 | both
7O4BJ131/Chem_NY_Acticide CBM-2 | 7O4BJ131 | Chem_NY_Acticide CBM-2 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Off-white - greenish dispersion (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4BJ131/Chem_NY_Acticide CBM-2 | 4 | both
7O4BJ131/Chem_NY_Acticide CBM-2 | 7O4BJ131 | Chem_NY_Acticide CBM-2 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4BJ131/Chem_NY_Acticide CBM-2 | 4 | both
7O4BJ131/Chem_NY_Acticide CBM-2 | 7O4BJ131 | Chem_NY_Acticide CBM-2 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4BJ131/Chem_NY_Acticide CBM-2 | 4 | both
7O4BJ131/Chem_NY_Acticide CBM-2 | 7O4BJ131 | Chem_NY_Acticide CBM-2 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4BJ131/Chem_NY_Acticide CBM-2 | 4 | both
7O4BM301/Chem_126A | 7O4BM301 | Chem_126A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification | 7O4BM301/Chem_126A | 3 | both
7O4BM301/Chem_126A | 7O4BM301 | Chem_126A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7O4BM301/Chem_126A | 3 | both
7O4BM301/Chem_126A | 7O4BM301 | Chem_126A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4BM301/Chem_126A | 3 | both
7O4BM301/Chem_126A | 7O4BM301 | Chem_126A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to Amber (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4BM301/Chem_126A | 3 | both
7O4BM301/Chem_126A | 7O4BM301 | Chem_126A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4BM301/Chem_126A | 3 | both
7O4BM301/Chem_126A | 7O4BM301 | Chem_126A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4BM301/Chem_126A | 3 | both
7O4DH247/Chem_201A | 7O4DH247 | Chem_201A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C | 7O4DH247/Chem_201A | 3 | both
7O4DH247/Chem_201A | 7O4DH247 | Chem_201A | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C | 1.2883 |  | FDA,3 | 1.27 | 1.31 | TM 202 | 7O4DH247/Chem_201A | 3 | both
7O4DH247/Chem_201A | 7O4DH247 | Chem_201A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear solution, free from solids (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4DH247/Chem_201A | 3 | both
7O4DH247/Chem_201A | 7O4DH247 | Chem_201A | NY-Analytical | 792 | Chemical Inspection | Color (TM 230) | Color | 2 |  | FDA,1 |  | 2 | TM208/220/230/USP/Kao LM | 7O4DH247/Chem_201A | 3 | both
7O4DH247/Chem_201A | 7O4DH247 | Chem_201A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4DH247/Chem_201A | 3 | both
7O4DH247/Chem_201A | 7O4DH247 | Chem_201A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Bleach/chlorine (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4DH247/Chem_201A | 3 | both
7O4DX220/Chem_NY_Neutralizing Agent | 7O4DX220 | Chem_NY_Neutralizing Agent | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C | 7O4DX220/Chem_NY_Neutralizing Agent | 3 | both
7O4DX220/Chem_NY_Neutralizing Agent | 7O4DX220 | Chem_NY_Neutralizing Agent | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.0999 |  | FDA,4 | 1.088 | 1.118 | TM 202 | 7O4DX220/Chem_NY_Neutralizing Agent | 3 | both
7O4DX220/Chem_NY_Neutralizing Agent | 7O4DX220 | Chem_NY_Neutralizing Agent | NY-Analytical | 451 | Refractive Index | Refractive Index | Refractive Index at 25°C | 1.4685 |  | FDA,4 | 1.46 | 1.48 | TM 204 | 7O4DX220/Chem_NY_Neutralizing Agent | 3 | both
7O4DX220/Chem_NY_Neutralizing Agent | 7O4DX220 | Chem_NY_Neutralizing Agent | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7O4DX220/Chem_NY_Neutralizing Agent | 3 | both
7O4DX220/Chem_NY_Neutralizing Agent | 7O4DX220 | Chem_NY_Neutralizing Agent | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4DX220/Chem_NY_Neutralizing Agent | 3 | both
7O4DX220/Chem_NY_Neutralizing Agent | 7O4DX220 | Chem_NY_Neutralizing Agent | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Light Yellow Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4DX220/Chem_NY_Neutralizing Agent | 3 | both
7O4DX220/Chem_NY_Neutralizing Agent | 7O4DX220 | Chem_NY_Neutralizing Agent | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4DX220/Chem_NY_Neutralizing Agent | 3 | both
7O4DX220/Chem_NY_Neutralizing Agent | 7O4DX220 | Chem_NY_Neutralizing Agent | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic – comparable to standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4DX220/Chem_NY_Neutralizing Agent | 3 | both
7O4OA234/Chem_131A | 7O4OA234 | Chem_131A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH as is | 7O4OA234/Chem_131A | 1 | both
7O4OA234/Chem_131A | 7O4OA234 | Chem_131A | NY-Analytical | 447 | pH as is | pH | pH | 3.22 |  | FDA,2 | 3 | 3.3 | TM 200 | 7O4OA234/Chem_131A | 1 | both
7O4OA234/Chem_131A | 7O4OA234 | Chem_131A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7O4OA234/Chem_131A | 1 | both
7O4OA234/Chem_131A | 7O4OA234 | Chem_131A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Opaque, Thin Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OA234/Chem_131A | 1 | both
7O4OA234/Chem_131A | 7O4OA234 | Chem_131A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OA234/Chem_131A | 1 | both
7O4OA234/Chem_131A | 7O4OA234 | Chem_131A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OA234/Chem_131A | 1 | both
7O4OA234/Chem_131A | 7O4OA234 | Chem_131A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OA234/Chem_131A | 1 | both
7O4OA347/Chem_NY_Lauryl Glucoside | 7O4OA347 | Chem_NY_Lauryl Glucoside | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection | 7O4OA347/Chem_NY_Lauryl Glucoside | 1 | both
7O4OA347/Chem_NY_Lauryl Glucoside | 7O4OA347 | Chem_NY_Lauryl Glucoside | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Match standard (Done) |  |  |  |  | TM 208 | 7O4OA347/Chem_NY_Lauryl Glucoside | 1 | both
7O4OA347/Chem_NY_Lauryl Glucoside | 7O4OA347 | Chem_NY_Lauryl Glucoside | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Match standard (Done) |  |  |  |  | TM 208 | 7O4OA347/Chem_NY_Lauryl Glucoside | 1 | both
7O4OA347/Chem_NY_Lauryl Glucoside | 7O4OA347 | Chem_NY_Lauryl Glucoside | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 | 7O4OA347/Chem_NY_Lauryl Glucoside | 1 | both
7O4OA347/Chem_NY_Lauryl Glucoside | 7O4OA347 | Chem_NY_Lauryl Glucoside | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208 | 7O4OA347/Chem_NY_Lauryl Glucoside | 1 | both
7O4OA347/Chem_NY_Lauryl Glucoside | 7O4OA347 | Chem_NY_Lauryl Glucoside | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Comparison of scans of sample vs standard (Done) |  |  |  |  | TM 411 | 7O4OA347/Chem_NY_Lauryl Glucoside | 1 | both
7O4OA401/Chem_132A | 7O4OA401 | Chem_132A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C | 7O4OA401/Chem_132A | 1 | both
7O4OA401/Chem_132A | 7O4OA401 | Chem_132A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.0225 |  | FDA,4 | 1.0115 | 1.0315 | TM 202 | 7O4OA401/Chem_132A | 1 | both
7O4OA401/Chem_132A | 7O4OA401 | Chem_132A | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.485 |  | FDA,3 | 1.4794 | 1.4994 | TM 204 | 7O4OA401/Chem_132A | 1 | both
7O4OA401/Chem_132A | 7O4OA401 | Chem_132A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7O4OA401/Chem_132A | 1 | both
7O4OA401/Chem_132A | 7O4OA401 | Chem_132A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OA401/Chem_132A | 1 | both
7O4OA401/Chem_132A | 7O4OA401 | Chem_132A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to light yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OA401/Chem_132A | 1 | both
7O4OA401/Chem_132A | 7O4OA401 | Chem_132A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OA401/Chem_132A | 1 | both
7O4OA401/Chem_132A | 7O4OA401 | Chem_132A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Fresh Floral (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OA401/Chem_132A | 1 | both
7O4OA490/Chem_133A | 7O4OA490 | Chem_133A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C | 7O4OA490/Chem_133A | 13 | both
7O4OA490/Chem_133A | 7O4OA490 | Chem_133A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.9182 |  | FDA,3 | 0.909 | 0.929 | TM 202 | 7O4OA490/Chem_133A | 13 | both
7O4OA490/Chem_133A | 7O4OA490 | Chem_133A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7O4OA490/Chem_133A | 13 | both
7O4OA490/Chem_133A | 7O4OA490 | Chem_133A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OA490/Chem_133A | 13 | both
7O4OA490/Chem_133A | 7O4OA490 | Chem_133A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OA490/Chem_133A | 13 | both
7O4OA490/Chem_133A | 7O4OA490 | Chem_133A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OA490/Chem_133A | 13 | both
7O4OA490/Chem_133A | 7O4OA490 | Chem_133A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Ether (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OA490/Chem_133A | 13 | both
7O4OA580/Chem_NY_Glycacil SG | 7O4OA580 | Chem_NY_Glycacil SG | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 |  |  | left_only
7O4OA580/Chem_NY_Glycacil SG | 7O4OA580 | Chem_NY_Glycacil SG | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  |  | left_only
7O4OA580/Chem_NY_Glycacil SG | 7O4OA580 | Chem_NY_Glycacil SG | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic Match Standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7O4OA580/Chem_NY_Glycacil SG | 7O4OA580 | Chem_NY_Glycacil SG | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7O4OA580/Chem_NY_Glycacil SG | 7O4OA580 | Chem_NY_Glycacil SG | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to Amber (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7O4OA580/Chem_NY_Glycacil SG | 7O4OA580 | Chem_NY_Glycacil SG | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM |  |  | left_only
7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | 7O4OA634 | Chem_NY_Water Soluble Flowering Herbs Fragrance | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C | 7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | 1 | both
7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | 7O4OA634 | Chem_NY_Water Soluble Flowering Herbs Fragrance | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.0629 |  | FDA,4 | 1.0531 | 1.0831 | TM 202 | 7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | 1 | both
7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | 7O4OA634 | Chem_NY_Water Soluble Flowering Herbs Fragrance | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.47 |  | FDA,4 | 1.4655 | 1.4855 | TM 204 | 7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | 1 | both
7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | 7O4OA634 | Chem_NY_Water Soluble Flowering Herbs Fragrance | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | 1 | both
7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | 7O4OA634 | Chem_NY_Water Soluble Flowering Herbs Fragrance | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | 1 | both
7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | 7O4OA634 | Chem_NY_Water Soluble Flowering Herbs Fragrance | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Pale Yellow to Straw (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | 1 | both
7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | 7O4OA634 | Chem_NY_Water Soluble Flowering Herbs Fragrance | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | 1 | both
7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | 7O4OA634 | Chem_NY_Water Soluble Flowering Herbs Fragrance | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Floral (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | 1 | both
7O4OA709/Chem_NY_Cucumber Extract # 11-19-3 | 7O4OA709 | Chem_NY_Cucumber Extract # 11-19-3 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Specific Gravity at 25°C | 7O4OA709/Chem_NY_Cucumber Extract # 11-19-3 | 3 | both
7O4OA709/Chem_NY_Cucumber Extract # 11-19-3 | 7O4OA709 | Chem_NY_Cucumber Extract # 11-19-3 | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C |  |  | FDA,3 | 1.034 | 1.044 | TM 202 | 7O4OA709/Chem_NY_Cucumber Extract # 11-19-3 | 3 | both
7O4OA709/Chem_NY_Cucumber Extract # 11-19-3 | 7O4OA709 | Chem_NY_Cucumber Extract # 11-19-3 | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.425 |  | FDA,3 | 1.42 | 1.43 | TM 204 | 7O4OA709/Chem_NY_Cucumber Extract # 11-19-3 | 3 | both
7O4OA709/Chem_NY_Cucumber Extract # 11-19-3 | 7O4OA709 | Chem_NY_Cucumber Extract # 11-19-3 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 | 7O4OA709/Chem_NY_Cucumber Extract # 11-19-3 | 3 | both
7O4OA709/Chem_NY_Cucumber Extract # 11-19-3 | 7O4OA709 | Chem_NY_Cucumber Extract # 11-19-3 | NY-Analytical | 792 | Chemical Inspection | Description (TM 208) | Description | Clear colorless liquid having a characteristic odor (Done) |  |  |  |  | TM 208 | 7O4OA709/Chem_NY_Cucumber Extract # 11-19-3 | 3 | both
7O4OA911/Chem_NY_Zinc Sulfate Heptahydrate | 7O4OA911 | Chem_NY_Zinc Sulfate Heptahydrate | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection | 7O4OA911/Chem_NY_Zinc Sulfate Heptahydrate | 3 | both
7O4OA911/Chem_NY_Zinc Sulfate Heptahydrate | 7O4OA911 | Chem_NY_Zinc Sulfate Heptahydrate | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Crystals/crystalline powder (Done) |  |  |  |  | TM 208 | 7O4OA911/Chem_NY_Zinc Sulfate Heptahydrate | 3 | both
7O4OA911/Chem_NY_Zinc Sulfate Heptahydrate | 7O4OA911 | Chem_NY_Zinc Sulfate Heptahydrate | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to white (Done) |  |  |  |  | TM 208 | 7O4OA911/Chem_NY_Zinc Sulfate Heptahydrate | 3 | both
7O4OA911/Chem_NY_Zinc Sulfate Heptahydrate | 7O4OA911 | Chem_NY_Zinc Sulfate Heptahydrate | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 | 7O4OA911/Chem_NY_Zinc Sulfate Heptahydrate | 3 | both
7O4OA911/Chem_NY_Zinc Sulfate Heptahydrate | 7O4OA911 | Chem_NY_Zinc Sulfate Heptahydrate | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208 | 7O4OA911/Chem_NY_Zinc Sulfate Heptahydrate | 3 | both
7O4OA911/Chem_NY_Zinc Sulfate Heptahydrate | 7O4OA911 | Chem_NY_Zinc Sulfate Heptahydrate | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411 | 7O4OA911/Chem_NY_Zinc Sulfate Heptahydrate | 3 | both
7O4OD351/Chem_NY_Benzalkonium Chloride 50% | 7O4OD351 | Chem_NY_Benzalkonium Chloride 50% | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Appearance | 7O4OD351/Chem_NY_Benzalkonium Chloride 50% | 1 | both
7O4OD351/Chem_NY_Benzalkonium Chloride 50% | 7O4OD351 | Chem_NY_Benzalkonium Chloride 50% | NY-Analytical | 397 | Appearance | Appearance | Appearance | Clear liquid, free from foreign material (FFFM) (Done) |  |  |  |  | TM 208 | 7O4OD351/Chem_NY_Benzalkonium Chloride 50% | 1 | both
7O4OD351/Chem_NY_Benzalkonium Chloride 50% | 7O4OD351 | Chem_NY_Benzalkonium Chloride 50% | NY-Analytical | 938 | Ethanol | Ethanol | Ethanol |  | % v/v | FDA,2 | 5.9 | 6.5 | TM 323 | 7O4OD351/Chem_NY_Benzalkonium Chloride 50% | 1 | both
7O4OD351/Chem_NY_Benzalkonium Chloride 50% | 7O4OD351 | Chem_NY_Benzalkonium Chloride 50% | NY-Analytical | 938 | Ethanol | Ethanol Final | Ethanol Final |  | % v/v | FDA,2 | 5.9 | 6.5 | TM 323 | 7O4OD351/Chem_NY_Benzalkonium Chloride 50% | 1 | both
7O4OD351/Chem_NY_Benzalkonium Chloride 50% | 7O4OD351 | Chem_NY_Benzalkonium Chloride 50% | NY-Analytical | 455 | Identification A | Identification A - Chloride | Identification A - Chloride | [5 empty] | TM 4001 | 7O4OD351/Chem_NY_Benzalkonium Chloride 50% | 1 | both
7O4OD351/Chem_NY_Benzalkonium Chloride 50% | 7O4OD351 | Chem_NY_Benzalkonium Chloride 50% | NY-Analytical | 456 | Identification B | Identification B - Benzalkonium | Identification B - Benzalkonium | [5 empty] | TM 4001 | 7O4OD351/Chem_NY_Benzalkonium Chloride 50% | 1 | both
7O4OD351/Chem_NY_Benzalkonium Chloride 50% | 7O4OD351 | Chem_NY_Benzalkonium Chloride 50% | NY-Analytical | 1064 | Total % ADBAC | % BZK and Dual-Quat | % BZK and Dual-Quat |  | % | FDA,1 | 47.5 | 52.5 | TM 434 | 7O4OD351/Chem_NY_Benzalkonium Chloride 50% | 1 | both
7O4OD363/Chem_137A | 7O4OD363 | Chem_137A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH as is | 7O4OD363/Chem_137A | 4 | both
7O4OD363/Chem_137A | 7O4OD363 | Chem_137A | NY-Analytical | 447 | pH as is | pH | pH | 5.06 |  | FDA,1 | 4.5 | 5.5 | TM 200 | 7O4OD363/Chem_137A | 4 | both
7O4OD363/Chem_137A | 7O4OD363 | Chem_137A | NY-Analytical | 806 | Total Solids | Total Solids (2 g for 3 hr @ 105°C) | Total Solids (2 g for 3 hr @ 105°C) | 43.5 | % | FDA,1 | 41 | 45 | TM 206 | 7O4OD363/Chem_137A | 4 | both
7O4OD363/Chem_137A | 7O4OD363 | Chem_137A | NY-Analytical | 806 | Total Solids | Total Solids Final (2 g for 3 hr @ 105°C) | Total Solids Final (2 g for 3 hr @ 105°C) | 43.45 | % | FDA,1 | 41 | 45 | TM 206 | 7O4OD363/Chem_137A | 4 | both
7O4OD363/Chem_137A | 7O4OD363 | Chem_137A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7O4OD363/Chem_137A | 4 | both
7O4OD363/Chem_137A | 7O4OD363 | Chem_137A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OD363/Chem_137A | 4 | both
7O4OD363/Chem_137A | 7O4OD363 | Chem_137A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to pale yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OD363/Chem_137A | 4 | both
7O4OD419/Chem_138A | 7O4OD419 | Chem_138A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA does not meet specifications (OOS) |  |  |  |  | Identification | 7O4OD419/Chem_138A | 6 | both
7O4OD419/Chem_138A | 7O4OD419 | Chem_138A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7O4OD419/Chem_138A | 6 | both
7O4OD419/Chem_138A | 7O4OD419 | Chem_138A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Semi Solid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OD419/Chem_138A | 6 | both
7O4OD419/Chem_138A | 7O4OD419 | Chem_138A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White to light yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OD419/Chem_138A | 6 | both
7O4OD419/Chem_138A | 7O4OD419 | Chem_138A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OD419/Chem_138A | 6 | both
7O4OD419/Chem_138A | 7O4OD419 | Chem_138A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Fatty (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OD419/Chem_138A | 6 | both
7O4OD748/Chem_139A | 7O4OD748 | Chem_139A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification | 7O4OD748/Chem_139A | 2 | both
7O4OD748/Chem_139A | 7O4OD748 | Chem_139A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7O4OD748/Chem_139A | 2 | both
7O4OD748/Chem_139A | 7O4OD748 | Chem_139A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | White solid crystals (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OD748/Chem_139A | 2 | both
7O4OD748/Chem_139A | 7O4OD748 | Chem_139A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White Crystals (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OD748/Chem_139A | 2 | both
7O4OD748/Chem_139A | 7O4OD748 | Chem_139A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OD748/Chem_139A | 2 | both
7O4OD748/Chem_139A | 7O4OD748 | Chem_139A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | To match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OD748/Chem_139A | 2 | both
7O4OF575/Che_NY_Plantaren 2000N - UP | 7O4OF575 | Che_NY_Plantaren 2000N - UP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification | 7O4OF575/Che_NY_Plantaren 2000N - UP | 1 | both
7O4OF575/Che_NY_Plantaren 2000N - UP | 7O4OF575 | Che_NY_Plantaren 2000N - UP | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7O4OF575/Che_NY_Plantaren 2000N - UP | 1 | both
7O4OF575/Che_NY_Plantaren 2000N - UP | 7O4OF575 | Che_NY_Plantaren 2000N - UP | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Hazy, Pale Yellow Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OF575/Che_NY_Plantaren 2000N - UP | 1 | both
7O4OF575/Che_NY_Plantaren 2000N - UP | 7O4OF575 | Che_NY_Plantaren 2000N - UP | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OF575/Che_NY_Plantaren 2000N - UP | 1 | both
7O4OF602/Chem_141A | 7O4OF602 | Chem_141A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification | 7O4OF602/Chem_141A | 3 | both
7O4OF602/Chem_141A | 7O4OF602 | Chem_141A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7O4OF602/Chem_141A | 3 | both
7O4OF602/Chem_141A | 7O4OF602 | Chem_141A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OF602/Chem_141A | 3 | both
7O4OG169/Chem_142A | 7O4OG169 | Chem_142A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C | 7O4OG169/Chem_142A | 1 | both
7O4OG169/Chem_142A | 7O4OG169 | Chem_142A | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C | 1.08 |  | FDA,3 | 1.08 | 1.09 | TM 202 | 7O4OG169/Chem_142A | 1 | both
7O4OG169/Chem_142A | 7O4OG169 | Chem_142A | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.435 |  | FDA,3 | 1.435 | 1.437 | TM 204 | 7O4OG169/Chem_142A | 1 | both
7O4OG169/Chem_142A | 7O4OG169 | Chem_142A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7O4OG169/Chem_142A | 1 | both
7O4OG169/Chem_142A | 7O4OG169 | Chem_142A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OG169/Chem_142A | 1 | both
7O4OG169/Chem_142A | 7O4OG169 | Chem_142A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OG169/Chem_142A | 1 | both
7O4OG169/Chem_142A | 7O4OG169 | Chem_142A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OG169/Chem_142A | 1 | both
7O4OH247/Chem_143A | 7O4OH247 | Chem_143A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH | 7O4OH247/Chem_143A | 2 | both
7O4OH247/Chem_143A | 7O4OH247 | Chem_143A | NY-Analytical | 398 | pH | pH (1% w/v aqueous solution) | pH (1% w/v aqueous solution) | 2 |  | FDA,1 |  | 2 | TM 200 | 7O4OH247/Chem_143A | 2 | both
7O4OH247/Chem_143A | 7O4OH247 | Chem_143A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7O4OH247/Chem_143A | 2 | both
7O4OH247/Chem_143A | 7O4OH247 | Chem_143A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear viscous liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OH247/Chem_143A | 2 | both
7O4OH247/Chem_143A | 7O4OH247 | Chem_143A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to dark amber (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OH247/Chem_143A | 2 | both
7O4OH247/Chem_143A | 7O4OH247 | Chem_143A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OH247/Chem_143A | 2 | both
7O4OH460/MASTER - Microbiology - J&J Chemical Testing) | 7O4OH460 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 801 | % Water | Moisture Content | Moisture Content |  | % | FDA,1 |  | 1.5 | J&J TMWW78017 |  |  | left_only
7O4OJ123/Chem_NY_NET-KOB | 7O4OJ123 | Chem_NY_NET-KOB | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification | 7O4OJ123/Chem_NY_NET-KOB | 4 | both
7O4OJ123/Chem_NY_NET-KOB | 7O4OJ123 | Chem_NY_NET-KOB | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7O4OJ123/Chem_NY_NET-KOB | 4 | both
7O4OJ123/Chem_NY_NET-KOB | 7O4OJ123 | Chem_NY_NET-KOB | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Viscous Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OJ123/Chem_NY_NET-KOB | 4 | both
7O4OJ123/Chem_NY_NET-KOB | 7O4OJ123 | Chem_NY_NET-KOB | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to pale yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OJ123/Chem_NY_NET-KOB | 4 | both
7O4OJ123/Chem_NY_NET-KOB | 7O4OJ123 | Chem_NY_NET-KOB | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OJ123/Chem_NY_NET-KOB | 4 | both
7O4OJ123/Chem_NY_NET-KOB | 7O4OJ123 | Chem_NY_NET-KOB | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OJ123/Chem_NY_NET-KOB | 4 | both
7O4OM165/Chem_146A | 7O4OM165 | Chem_146A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification | 7O4OM165/Chem_146A | 6 | both
7O4OM165/Chem_146A | 7O4OM165 | Chem_146A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7O4OM165/Chem_146A | 6 | both
7O4OM165/Chem_146A | 7O4OM165 | Chem_146A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, slightly oily viscous liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OM165/Chem_146A | 6 | both
7O4OM165/Chem_146A | 7O4OM165 | Chem_146A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Light Yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OM165/Chem_146A | 6 | both
7O4OM165/Chem_146A | 7O4OM165 | Chem_146A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OM165/Chem_146A | 6 | both
7O4OM165/Chem_146A | 7O4OM165 | Chem_146A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | To match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OM165/Chem_146A | 6 | both
7O4OM666/Chem_NY_Sodium Benzoate EDF NF/FCC | 7O4OM666 | Chem_NY_Sodium Benzoate EDF NF/FCC | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection | 7O4OM666/Chem_NY_Sodium Benzoate EDF NF/FCC | 8 | both
7O4OM666/Chem_NY_Sodium Benzoate EDF NF/FCC | 7O4OM666 | Chem_NY_Sodium Benzoate EDF NF/FCC | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Extruded Pellets (Done) |  |  |  |  | TM 208 | 7O4OM666/Chem_NY_Sodium Benzoate EDF NF/FCC | 8 | both
7O4OM666/Chem_NY_Sodium Benzoate EDF NF/FCC | 7O4OM666 | Chem_NY_Sodium Benzoate EDF NF/FCC | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) |  |  |  |  | TM 208 | 7O4OM666/Chem_NY_Sodium Benzoate EDF NF/FCC | 8 | both
7O4OM666/Chem_NY_Sodium Benzoate EDF NF/FCC | 7O4OM666 | Chem_NY_Sodium Benzoate EDF NF/FCC | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 | 7O4OM666/Chem_NY_Sodium Benzoate EDF NF/FCC | 8 | both
7O4OM666/Chem_NY_Sodium Benzoate EDF NF/FCC | 7O4OM666 | Chem_NY_Sodium Benzoate EDF NF/FCC | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless (Done) |  |  |  |  | TM 208 | 7O4OM666/Chem_NY_Sodium Benzoate EDF NF/FCC | 8 | both
7O4OM666/Chem_NY_Sodium Benzoate EDF NF/FCC | 7O4OM666 | Chem_NY_Sodium Benzoate EDF NF/FCC | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7O4OM666/Chem_NY_Sodium Benzoate EDF NF/FCC | 8 | both
7O4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7O4OP027 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 827 | Hydroxyl Value | Hydroxyl Value | Hydroxyl Value | 38.5 | mg KOH/g | FDA,1 | 35 | 50 | J&J TMWW7671 | 7O4OP027/MASTER - Microbiology - J&J Chemical Testing) | 1 | both
7O4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7O4OP027 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 827 | Hydroxyl Value | Hydroxyl Value Final | Hydroxyl Value Final | 38.5 | mg KOH/g | FDA,1 | 35 | 50 | J&J TMWW7671 | 7O4OP027/MASTER - Microbiology - J&J Chemical Testing) | 1 | both
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Chem_NY_Actipone White Tea | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Density @ 20° C | 7O4OP065/Chem_NY_Actipone White Tea | 2 | both
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Chem_NY_Actipone White Tea | NY-Analytical | 831 | Density @ 20° C | Density | Density | 1.1173 |  | FDA,3 | 1.112 | 1.122 | TM 202 | 7O4OP065/Chem_NY_Actipone White Tea | 2 | both
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Chem_NY_Actipone White Tea | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.392 |  | FDA,3 | 1.388 | 1.398 | TM 204 | 7O4OP065/Chem_NY_Actipone White Tea | 2 | both
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Chem_NY_Actipone White Tea | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear to slightly cloudy liquid (Done) |  |  |  |  | TM 208 | 7O4OP065/Chem_NY_Actipone White Tea | 2 | both
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Chem_NY_Actipone White Tea | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Yellow to Dark Brown (Done) |  |  |  |  | TM 208 | 7O4OP065/Chem_NY_Actipone White Tea | 2 | both
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Chem_NY_Actipone White Tea | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 | 7O4OP065/Chem_NY_Actipone White Tea | 2 | both
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Chem_NY_Actipone White Tea | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208 | 7O4OP065/Chem_NY_Actipone White Tea | 2 | both
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Chem_NY_Actipone White Tea | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411 | 7O4OP065/Chem_NY_Actipone White Tea | 2 | both
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Chem_NY_Actipone White Tea | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram) | TAMC (Average per Gram) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | 7O4OP065/Chem_NY_Actipone White Tea | 2 | both
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Chem_NY_Actipone White Tea | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | 7O4OP065/Chem_NY_Actipone White Tea | 2 | both
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Chem_NY_Actipone White Tea | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram) | TMYC (Average per Gram) |  | CFU/g | FDA,1 |  | 100 | TM 610 | 7O4OP065/Chem_NY_Actipone White Tea | 2 | both
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Chem_NY_Actipone White Tea | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  | cfu | FDA,0 |  |  | TM 610 | 7O4OP065/Chem_NY_Actipone White Tea | 2 | both
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Chem_NY_Actipone White Tea | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Negative Rods (OOS) |  |  |  |  | WI QAS 076 | 7O4OP065/Chem_NY_Actipone White Tea | 2 | both
7O4OP075/Chem_NY_Magique Lift 02LF | 7O4OP075 | Chem_NY_Magique Lift 02LF | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA does not meet specifications (OOS) |  |  |  |  | Specific Gravity at 25°C | 7O4OP075/Chem_NY_Magique Lift 02LF | 3 | both
7O4OP075/Chem_NY_Magique Lift 02LF | 7O4OP075 | Chem_NY_Magique Lift 02LF | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.9516 |  | FDA,3 | 0.938 | 0.948 | TM 202 | 7O4OP075/Chem_NY_Magique Lift 02LF | 3 | both
7O4OP075/Chem_NY_Magique Lift 02LF | 7O4OP075 | Chem_NY_Magique Lift 02LF | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.48 |  | FDA,3 | 1.474 | 1.484 | TM 204 | 7O4OP075/Chem_NY_Magique Lift 02LF | 3 | both
7O4OP075/Chem_NY_Magique Lift 02LF | 7O4OP075 | Chem_NY_Magique Lift 02LF | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7O4OP075/Chem_NY_Magique Lift 02LF | 3 | both
7O4OP075/Chem_NY_Magique Lift 02LF | 7O4OP075 | Chem_NY_Magique Lift 02LF | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OP075/Chem_NY_Magique Lift 02LF | 3 | both
7O4OP075/Chem_NY_Magique Lift 02LF | 7O4OP075 | Chem_NY_Magique Lift 02LF | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to very light yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OP075/Chem_NY_Magique Lift 02LF | 3 | both
7O4OP075/Chem_NY_Magique Lift 02LF | 7O4OP075 | Chem_NY_Magique Lift 02LF | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OP075/Chem_NY_Magique Lift 02LF | 3 | both
7O4OP075/Chem_NY_Magique Lift 02LF | 7O4OP075 | Chem_NY_Magique Lift 02LF | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OP075/Chem_NY_Magique Lift 02LF | 3 | both
7O4OP093/Chem_NY_PEG-11 | 7O4OP093 | Chem_NY_PEG-11 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA | 7O4OP093/Chem_NY_PEG-11 | 3 | both
7O4OP093/Chem_NY_PEG-11 | 7O4OP093 | Chem_NY_PEG-11 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, Light liquid (Done) |  |  |  |  | TM 208 | 7O4OP093/Chem_NY_PEG-11 | 3 | both
7O4OP093/Chem_NY_PEG-11 | 7O4OP093 | Chem_NY_PEG-11 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Match standard (Done) |  |  |  |  | TM 208 | 7O4OP093/Chem_NY_PEG-11 | 3 | both
7O4OP093/Chem_NY_PEG-11 | 7O4OP093 | Chem_NY_PEG-11 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 | 7O4OP093/Chem_NY_PEG-11 | 3 | both
7O4OP093/Chem_NY_PEG-11 | 7O4OP093 | Chem_NY_PEG-11 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208 | 7O4OP093/Chem_NY_PEG-11 | 3 | both
7O4OP093/Chem_NY_PEG-11 | 7O4OP093 | Chem_NY_PEG-11 | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411 | 7O4OP093/Chem_NY_PEG-11 | 3 | both
7O4OP113/Chem_NY_Robertet The Matcha I | 7O4OP113 | Chem_NY_Robertet The Matcha I | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411 |  |  | left_only
7O4OP113/Chem_NY_Robertet The Matcha I | 7O4OP113 | Chem_NY_Robertet The Matcha I | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C |  |  | FDA,3 | 0.96 | 0.97 | TM 202 |  |  | left_only
7O4OP113/Chem_NY_Robertet The Matcha I | 7O4OP113 | Chem_NY_Robertet The Matcha I | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C |  |  | left_only
7O4OP113/Chem_NY_Robertet The Matcha I | 7O4OP113 | Chem_NY_Robertet The Matcha I | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208 |  |  | left_only
7O4OP113/Chem_NY_Robertet The Matcha I | 7O4OP113 | Chem_NY_Robertet The Matcha I | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 |  |  | left_only
7O4OP113/Chem_NY_Robertet The Matcha I | 7O4OP113 | Chem_NY_Robertet The Matcha I | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Oily liquid (Done) |  |  |  |  | TM 208 |  |  | left_only
7O4OP113/Chem_NY_Robertet The Matcha I | 7O4OP113 | Chem_NY_Robertet The Matcha I | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Pale Yellow – Light Yellow (Done) |  |  |  |  | TM 208 |  |  | left_only
7O4OP113/Chem_NY_Robertet The Matcha I | 7O4OP113 | Chem_NY_Robertet The Matcha I | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C |  |  | FDA,3 | 1.469 | 1.479 | TM 204 |  |  | left_only
7O4OP197/Chem_NY_Emulgade Smart/MB_01 | 7O4OP197 | Chem_NY_Emulgade Smart/MB_01 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH | 7O4OP197/Chem_NY_Emulgade Smart/MB_01 | 12 | both
7O4OP197/Chem_NY_Emulgade Smart/MB_01 | 7O4OP197 | Chem_NY_Emulgade Smart/MB_01 | NY-Analytical | 398 | pH | pH | pH | 3 |  | FDA,1 | 3 | 3.5 | TM 200 | 7O4OP197/Chem_NY_Emulgade Smart/MB_01 | 12 | both
7O4OP197/Chem_NY_Emulgade Smart/MB_01 | 7O4OP197 | Chem_NY_Emulgade Smart/MB_01 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Match standard (Done) |  |  |  |  | TM 208 | 7O4OP197/Chem_NY_Emulgade Smart/MB_01 | 12 | both
7O4OP197/Chem_NY_Emulgade Smart/MB_01 | 7O4OP197 | Chem_NY_Emulgade Smart/MB_01 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White Liquid (Done) |  |  |  |  | TM 208 | 7O4OP197/Chem_NY_Emulgade Smart/MB_01 | 12 | both
7O4OP197/Chem_NY_Emulgade Smart/MB_01 | 7O4OP197 | Chem_NY_Emulgade Smart/MB_01 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 | 7O4OP197/Chem_NY_Emulgade Smart/MB_01 | 12 | both
7O4OP197/Chem_NY_Emulgade Smart/MB_01 | 7O4OP197 | Chem_NY_Emulgade Smart/MB_01 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208 | 7O4OP197/Chem_NY_Emulgade Smart/MB_01 | 12 | both
7O4OP197/Chem_NY_Emulgade Smart/MB_01 | 7O4OP197 | Chem_NY_Emulgade Smart/MB_01 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411 | 7O4OP197/Chem_NY_Emulgade Smart/MB_01 | 12 | both
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 |  |  | left_only
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C |  |  | FDA,3 | 1.1 | 1.14 | J&J TMWW78012 |  |  | left_only
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Analytical | 447 | pH as is | pH | pH |  |  | FDA,1 | 4.5 | 5.5 | TM WW78014 |  |  | left_only
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Microbiology | 509 | J&J_Pathogens | Cetrimide | Cetrimide | No Growth |  |  |  |  | TM 869 |  |  | left_only
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Microbiology | 509 | J&J_Pathogens | MacConkey | MacConkey | No Growth |  |  |  |  | TM 869 |  |  | left_only
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Microbiology | 509 | J&J_Pathogens | TSA | TSA | No Growth |  |  |  |  | TM 869 |  |  | left_only
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Microbiology | 509 | J&J_Pathogens | MSA or VJ | MSA or VJ | No Growth_MSA |  |  |  |  | TM 869 |  |  | left_only
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Microbiology | 509 | J&J_Pathogens | Biggy or SDA | Biggy or SDA | No Growth_SDA |  |  |  |  | TM 869 |  |  | left_only
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Microbiology | 517 | J&J_Enrichment | TSB | TSB | Not Turbid |  |  |  |  | TM 869 |  |  | left_only
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Microbiology | 524 | J&J Fungal | Fungal | Fungal |  | CFU/g | FDA,0 | 0 | 200 | TM 869 |  |  | left_only
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | TSA JJ |  | cfu/mL | FDA,0 | 0 | 2000 | TM 869 |  |  | left_only
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Microbiology | 700 | J&J SDA Controls | Negative Control | Negative Control | No Growth SDA |  |  |  |  | TM 869 |  |  | left_only
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | NegativeControl | No Growth TSB |  |  |  |  | TM 869 |  |  | left_only
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  |  | left_only
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | An opalescent, pale yellow to pale brown viscous liquid. (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Analytical | 792 | Chemical Inspection | Odor (J&J TM-003188) | Odor | Passes Test (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Analytical | 916 | Reducing Sugars | Reducing Sugars | Reducing Sugars |  | % | FDA,1 | 2.5 | 3.6 | TM-002569 |  |  | left_only
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Analytical | 917 | Total Sugar | Total Sugar | Total Sugar |  | % | FDA,1 | 6.5 | 9 | TM-002568 |  |  | left_only
7O4OP253/Chem_NY_BTC 1010 | 7O4OP253 | Chem_NY_BTC 1010 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH | 7O4OP253/Chem_NY_BTC 1010 | 1 | both
7O4OP253/Chem_NY_BTC 1010 | 7O4OP253 | Chem_NY_BTC 1010 | NY-Analytical | 398 | pH | pH (10% in water) | pH (10% in water) | 7.19 |  | FDA,1 | 6 | 9 | TM 200 | 7O4OP253/Chem_NY_BTC 1010 | 1 | both
7O4OP253/Chem_NY_BTC 1010 | 7O4OP253 | Chem_NY_BTC 1010 | NY-Analytical | 796 | Assay | Assay | Assay | 50.8 | % | FDA,0 | 50 | 55 | TM 301/434 | 7O4OP253/Chem_NY_BTC 1010 | 1 | both
7O4OP253/Chem_NY_BTC 1010 | 7O4OP253 | Chem_NY_BTC 1010 | NY-Analytical | 796 | Assay | Assay Final | Assay Final | 50.85 | % | FDA,0 | 50 | 55 | TM 301/434 | 7O4OP253/Chem_NY_BTC 1010 | 1 | both
7O4OP284/Chem_155A | 7O4OP284 | Chem_155A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 |  |  | left_only
7O4OP284/Chem_155A | 7O4OP284 | Chem_155A | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C |  |  | FDA,3 | 0.963 | 0.973 | J&J TMWW78012 |  |  | left_only
7O4OP284/Chem_155A | 7O4OP284 | Chem_155A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | GC Profile |  |  | left_only
7O4OP284/Chem_155A | 7O4OP284 | Chem_155A | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear, yellow liquid. (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7O4OP284/Chem_155A | 7O4OP284 | Chem_155A | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7O4OP284/Chem_155A | 7O4OP284 | Chem_155A | NY-Analytical | 792 | Chemical Inspection | Odor (J&J TM-003188) | Odor | Supplier certifies on C of A that odor matches standard (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7O4OP284/Chem_155A | 7O4OP284 | Chem_155A | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C |  |  | FDA,4 | 1.442 | 1.452 | J&J TMWW78015 |  |  | left_only
7O4OP284/Chem_155A | 7O4OP284 | Chem_155A | NY-Analytical | 839 | GC Profile | GC Profile | GC Profile | To match standard. Not required when performing IR test. (Done) |  |  |  |  | J&J TM-001862 |  |  | left_only
7O4OP345/Chem_157A | 7O4OP345 | Chem_157A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification | 7O4OP345/Chem_157A | 2 | both
7O4OP345/Chem_157A | 7O4OP345 | Chem_157A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7O4OP345/Chem_157A | 2 | both
7O4OP345/Chem_157A | 7O4OP345 | Chem_157A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OP345/Chem_157A | 2 | both
7O4OP345/Chem_157A | 7O4OP345 | Chem_157A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OP345/Chem_157A | 2 | both
7O4OP345/Chem_157A | 7O4OP345 | Chem_157A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OP345/Chem_157A | 2 | both
7O4OP345/Chem_157A | 7O4OP345 | Chem_157A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Phenoxyethanol-like (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OP345/Chem_157A | 2 | both
7O4OP567/Chem_NY_BTC 1210-80% | 7O4OP567 | Chem_NY_BTC 1210-80% | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA | 7O4OP567/Chem_NY_BTC 1210-80% | 8 | both
7O4OP567/Chem_NY_BTC 1210-80% | 7O4OP567 | Chem_NY_BTC 1210-80% | NY-Analytical | 398 | pH | pH (10% in water) | pH (10% in water) | 7.79 |  | FDA,1 | 6 | 8.5 | TM 200 | 7O4OP567/Chem_NY_BTC 1210-80% | 8 | both
7O4OP567/Chem_NY_BTC 1210-80% | 7O4OP567 | Chem_NY_BTC 1210-80% | NY-Analytical | 397 | Appearance | Appearance | Appearance | Colorless to slight yellow (Done) |  |  |  |  | TM 208 | 7O4OP567/Chem_NY_BTC 1210-80% | 8 | both
7O4OP567/Chem_NY_BTC 1210-80% | 7O4OP567 | Chem_NY_BTC 1210-80% | NY-Analytical | 1015 | Active (BTC 1210-80%) | Active (BTC 1210-80%) | Active (BTC 1210-80%) | 80.9 | % | FDA,1 | 80 | 82 | TM 3005 | 7O4OP567/Chem_NY_BTC 1210-80% | 8 | both
7O4OP601/Chem_107A | 7O4OP601 | Chem_107A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 |  |  | left_only
7O4OP601/Chem_107A | 7O4OP601 | Chem_107A | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C |  |  | FDA,3 | 0.986 | 0.996 | J&J TMWW78012 |  |  | left_only
7O4OP601/Chem_107A | 7O4OP601 | Chem_107A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | GC Profile |  |  | left_only
7O4OP601/Chem_107A | 7O4OP601 | Chem_107A | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Colorless to pale yellow clear liquid. (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7O4OP601/Chem_107A | 7O4OP601 | Chem_107A | NY-Analytical | 792 | Chemical Inspection | Odor (J&J TM-003188) | Odor | Meets Odor Specification or of Characteristic Odor. [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7O4OP601/Chem_107A | 7O4OP601 | Chem_107A | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7O4OP601/Chem_107A | 7O4OP601 | Chem_107A | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C |  |  | FDA,4 | 1.452 | 1.462 | J&J TMWW78015 |  |  | left_only
7O4OP601/Chem_107A | 7O4OP601 | Chem_107A | NY-Analytical | 839 | GC Profile | GC Profile | GC Profile | To match standard. Not required when performing IR test. (Done) |  |  |  |  | J&J TM-003030 |  |  | left_only
7O4OP638/Chem_192A | 7O4OP638 | Chem_192A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection | 7O4OP638/Chem_192A | 4 | both
7O4OP638/Chem_192A | 7O4OP638 | Chem_192A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, colorless to light yellow liquid / White solid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OP638/Chem_192A | 4 | both
7O4OP705/Chem_NY_Cetaphil Concentrate | 7O4OP705 | Chem_NY_Cetaphil Concentrate | NY-Analytical | 1051 | Temptale Reading | Temptale Reading | Temptale Reading |  | ºC | FDA,0 | 5 | 40 | Temptale Reading | 7O4OP705/Chem_NY_Cetaphil Concentrate | 8 | both
7O4OP705/Chem_NY_Cetaphil Concentrate | 7O4OP705 | Chem_NY_Cetaphil Concentrate | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH as is | 7O4OP705/Chem_NY_Cetaphil Concentrate | 8 | both
7O4OP705/Chem_NY_Cetaphil Concentrate | 7O4OP705 | Chem_NY_Cetaphil Concentrate | NY-Analytical | 447 | pH as is | pH | pH | 7 |  | FDA,1 | 5.5 | 7 | TM 200 | 7O4OP705/Chem_NY_Cetaphil Concentrate | 8 | both
7O4OP705/Chem_NY_Cetaphil Concentrate | 7O4OP705 | Chem_NY_Cetaphil Concentrate | NY-Analytical | 747 | Viscosity | Viscosity (Spindle 3 @ 60RPM in 300 ml beaker) | Viscosity (Spindle 3 @ 60RPM in 300 ml beaker) | 915 | cps | FDA,0 | 400 | 1600 | TM 203 | 7O4OP705/Chem_NY_Cetaphil Concentrate | 8 | both
7O4OP705/Chem_NY_Cetaphil Concentrate | 7O4OP705 | Chem_NY_Cetaphil Concentrate | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 | 7O4OP705/Chem_NY_Cetaphil Concentrate | 8 | both
7O4OP705/Chem_NY_Cetaphil Concentrate | 7O4OP705 | Chem_NY_Cetaphil Concentrate | NY-Analytical | 792 | Chemical Inspection | Description (TM 208) | Description | Viscous, slightly translucent, offwhite liquid: cetyl alcohol odor (Done) |  |  |  |  | TM 208 | 7O4OP705/Chem_NY_Cetaphil Concentrate | 8 | both
7O4OP806/Chem_NY_Carbopol 2020 | 7O4OP806 | Chem_NY_Carbopol 2020 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 |  |  | left_only
7O4OP806/Chem_NY_Carbopol 2020 | 7O4OP806 | Chem_NY_Carbopol 2020 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Loss on Drying |  |  | left_only
7O4OP806/Chem_NY_Carbopol 2020 | 7O4OP806 | Chem_NY_Carbopol 2020 | NY-Analytical | 747 | Viscosity | Viscosity | Viscosity |  | cps | FDA,0 | 47000 | 77000 | J&J TM7423 |  |  | left_only
7O4OP806/Chem_NY_Carbopol 2020 | 7O4OP806 | Chem_NY_Carbopol 2020 | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Fluffy white powder free of lumps (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7O4OP806/Chem_NY_Carbopol 2020 | 7O4OP806 | Chem_NY_Carbopol 2020 | NY-Analytical | 748 | Visual Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7O4OP806/Chem_NY_Carbopol 2020 | 7O4OP806 | Chem_NY_Carbopol 2020 | NY-Analytical | 802 | Loss on Drying | Loss on Drying (80C Vacuum 1 hour) | Loss on Drying (80C Vacuum 1 hour) |  | % | FDA,2 |  | 2 | J&J TM7221 |  |  | left_only
7O4OP806/Chem_NY_Carbopol 2020 | 7O4OP806 | Chem_NY_Carbopol 2020 | NY-Analytical | 802 | Loss on Drying | Loss on Drying (80C Vacuum 1 hour) Final | Loss on Drying (80C Vacuum 1 hour) Final |  | % | FDA,2 |  | 2 | J&J TM7221 |  |  | left_only
7O4OQ028/Chem_NY_Hexylene Glycol | 7O4OQ028 | Chem_NY_Hexylene Glycol | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis | 7O4OQ028/Chem_NY_Hexylene Glycol | 4 | both
7O4OQ028/Chem_NY_Euxyl K703 | 7O4OQ028 | Chem_NY_Euxyl K703 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Density @ 20° C | 7O4OQ028/Chem_NY_Euxyl K703 | 1 | both
7O4OQ028/Chem_NY_Euxyl K703 | 7O4OQ028 | Chem_NY_Euxyl K703 | NY-Analytical | 831 | Density @ 20° C | Density | Density | 1.1282 |  | FDA,3 | 1.124 | 1.128 | TM 202 | 7O4OQ028/Chem_NY_Euxyl K703 | 1 | both
7O4OQ028/Chem_NY_Hexylene Glycol | 7O4OQ028 | Chem_NY_Hexylene Glycol | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C |  |  | FDA,4 | 0.92 | 0.923 | TM 202 | 7O4OQ028/Chem_NY_Hexylene Glycol | 4 | both
7O4OQ028/Chem_NY_Euxyl K703 | 7O4OQ028 | Chem_NY_Euxyl K703 | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.539 |  | FDA,3 | 1.539 | 1.543 | TM 204 | 7O4OQ028/Chem_NY_Euxyl K703 | 1 | both
7O4OQ028/Chem_NY_Euxyl K703 | 7O4OQ028 | Chem_NY_Euxyl K703 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7O4OQ028/Chem_NY_Euxyl K703 | 1 | both
7O4OQ028/Chem_NY_Hexylene Glycol | 7O4OQ028 | Chem_NY_Hexylene Glycol | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411 | 7O4OQ028/Chem_NY_Hexylene Glycol | 4 | both
7O4OQ028/Chem_NY_Hexylene Glycol | 7O4OQ028 | Chem_NY_Hexylene Glycol | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OQ028/Chem_NY_Hexylene Glycol | 4 | both
7O4OQ028/Chem_NY_Euxyl K703 | 7O4OQ028 | Chem_NY_Euxyl K703 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OQ028/Chem_NY_Euxyl K703 | 1 | both
7O4OQ028/Chem_NY_Hexylene Glycol | 7O4OQ028 | Chem_NY_Hexylene Glycol | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OQ028/Chem_NY_Hexylene Glycol | 4 | both
7O4OQ028/Chem_NY_Euxyl K703 | 7O4OQ028 | Chem_NY_Euxyl K703 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OQ028/Chem_NY_Euxyl K703 | 1 | both
7O4OQ028/Chem_NY_Hexylene Glycol | 7O4OQ028 | Chem_NY_Hexylene Glycol | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OQ028/Chem_NY_Hexylene Glycol | 4 | both
7O4OQ028/Chem_NY_Euxyl K703 | 7O4OQ028 | Chem_NY_Euxyl K703 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OQ028/Chem_NY_Euxyl K703 | 1 | both
7O4OQ028/Chem_NY_Hexylene Glycol | 7O4OQ028 | Chem_NY_Hexylene Glycol | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OQ028/Chem_NY_Hexylene Glycol | 4 | both
7O4OQ028/Chem_NY_Euxyl K703 | 7O4OQ028 | Chem_NY_Euxyl K703 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Phenoxyethanol-like (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OQ028/Chem_NY_Euxyl K703 | 1 | both
7O4OQ243/Chem_165A | 7O4OQ243 | Chem_165A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification | 7O4OQ243/Chem_165A | 8 | both
7O4OQ243/Chem_165A | 7O4OQ243 | Chem_165A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7O4OQ243/Chem_165A | 8 | both
7O4OQ243/Chem_165A | 7O4OQ243 | Chem_165A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Fine, free-flowing powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OQ243/Chem_165A | 8 | both
7O4OQ243/Chem_165A | 7O4OQ243 | Chem_165A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OQ243/Chem_165A | 8 | both
7O4OQ243/Chem_165A | 7O4OQ243 | Chem_165A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OQ243/Chem_165A | 8 | both
7O4OQ243/Chem_165A | 7O4OQ243 | Chem_165A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristically mild (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OQ243/Chem_165A | 8 | both
7O4OQ342/Chem_166A | 7O4OQ342 | Chem_166A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C | 7O4OQ342/Chem_166A | 1 | both
7O4OQ342/Chem_166A | 7O4OQ342 | Chem_166A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.9558 |  | FDA,4 | 0.9455 | 0.9655 | TM 202 | 7O4OQ342/Chem_166A | 1 | both
7O4OQ342/Chem_166A | 7O4OQ342 | Chem_166A | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.458 |  | FDA,4 | 1.4505 | 1.4705 | TM 204 | 7O4OQ342/Chem_166A | 1 | both
7O4OQ342/Chem_166A | 7O4OQ342 | Chem_166A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, mobile liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OQ342/Chem_166A | 1 | both
7O4OQ342/Chem_166A | 7O4OQ342 | Chem_166A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Very pale yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OQ342/Chem_166A | 1 | both
7O4OQ342/Chem_166A | 7O4OQ342 | Chem_166A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OQ342/Chem_166A | 1 | both
7O4OQ342/Chem_166A | 7O4OQ342 | Chem_166A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Citrus Fresh (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OQ342/Chem_166A | 1 | both
7O4OQ437/Chem_168A | 7O4OQ437 | Chem_168A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C | 7O4OQ437/Chem_168A | 5 | both
7O4OQ437/Chem_168A | 7O4OQ437 | Chem_168A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.944 |  | FDA,4 | 0.938 | 0.9465 | TM 202 | 7O4OQ437/Chem_168A | 5 | both
7O4OQ437/Chem_168A | 7O4OQ437 | Chem_168A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7O4OQ437/Chem_168A | 5 | both
7O4OQ437/Chem_168A | 7O4OQ437 | Chem_168A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Mobile Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OQ437/Chem_168A | 5 | both
7O4OQ437/Chem_168A | 7O4OQ437 | Chem_168A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OQ437/Chem_168A | 5 | both
7O4OQ437/Chem_168A | 7O4OQ437 | Chem_168A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OQ437/Chem_168A | 5 | both
7O4OQ437/Chem_168A | 7O4OQ437 | Chem_168A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Amine (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OQ437/Chem_168A | 5 | both
7O4OQ890/Chem_170A | 7O4OQ890 | Chem_170A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification | 7O4OQ890/Chem_170A | 10 | both
7O4OQ890/Chem_170A | 7O4OQ890 | Chem_170A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7O4OQ890/Chem_170A | 10 | both
7O4OQ890/Chem_170A | 7O4OQ890 | Chem_170A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OQ890/Chem_170A | 10 | both
7O4OQ890/Chem_170A | 7O4OQ890 | Chem_170A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OQ890/Chem_170A | 10 | both
7O4OQ890/Chem_170A | 7O4OQ890 | Chem_170A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OQ890/Chem_170A | 10 | both
7O4OQ890/Chem_170A | 7O4OQ890 | Chem_170A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Ether-like (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OQ890/Chem_170A | 10 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 792 | Chemical Inspection | Description | Description | An almost colorless or paleyellowish liquid (Done) |  |  |  |  | TM 208 | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Any individual unspecified impuritv | Any individual unspecified impuritv | 0.077 | % | FDA,1 |  | 0.1 | USP | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Ch!orhexidine oxazinone analog | Ch!orhexidine oxazinone analog | 0.042 | % | FDA,1 |  | 0.2 | USP | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Chlorhexidine | Chlorhexidine | 93.338 | % | FDA,1 |  |  | USP | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Chlorhexidine urea | Chlorhexidine urea | 0.085 | % | FDA,1 |  | 0.2 | USP | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Chlorhexidine amine | Chlorhexidine amine | 0.047 | % | FDA,1 |  | 0.3 | USP | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Chlorhexidine dimer | Chlorhexidine dimer | 0.09 | % | FDA,1 |  | 0.5 | USP | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Chlorhexidine glucityl triazine | Chlorhexidine glucityl triazine | 0.083 | % | FDA,1 |  | 0.4 | USP | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Chlorhexidine guanidine | Chlorhexidine guanidine | 0.083 | % | FDA,1 |  | 0.3 | USP | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Chlorhexidine nitrile | Chlorhexidine nitrile | 0.114 | % | FDA,1 |  | 0.4 | USP | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 390 | Identification | Identification A (USP <197k>) | Identification A | Conforms to USP (Done) |  |  |  |  | USP | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 390 | Identification | Identification B (USP <201>) | Identification B | Conforms to USP (Done) |  |  |  |  | USP | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | o-Chlorhexidine and specified unidentified impurit | o-Chlorhexidine and specified unidentified impurit | 0.164 | % | FDA,1 |  | 0.4 | USP | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Oxochlorhexidine | Oxochlorhexidine | 0.288 | % | FDA,1 |  | 0.4 | USP | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 812 | Organic Impurites | p-Chloroaniline | p-Chloroaniline | 0 | ug/mL | FDA,0 |  | 100 | USP | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 812 | Organic Impurites | p-Chloroaniline Final | p-Chloroaniline Final | 0 | ug/mL | FDA,0 |  | 100 | USP | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | p-Chlorophenvl urea | p-Chlorophenvl urea | 0 | % | FDA,1 |  | 0.2 | USP | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Any individual unspecified impuritv | RRT of Any individual unspecified impuritv | 1.13 |  | FDA,2 |  |  | USP | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Ch!orhexidine oxazinone analog | RRT of Ch!orhexidine oxazinone analog | 0.16 |  | FDA,2 |  |  | USP | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Chlorhexidine | RRT of Chlorhexidine | 1.1 |  | FDA,2 |  |  | USP | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Chlorhexidine urea | RRT of Chlorhexidine urea | 0.32 |  | FDA,2 |  |  | USP | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Chlorhexidine amine | RRT of Chlorhexidine amine | 0.21 |  | FDA,2 |  |  | USP | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Chlorhexidine dimer | RRT of Chlorhexidine dimer | 0.87 |  | FDA,2 |  |  | USP | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Chlorhexidine glucityl triazine | RRT of Chlorhexidine glucityl triazine | 0.95 |  | FDA,2 |  |  | USP | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Chlorhexidine guanidine | RRT of Chlorhexidine guanidine | 0.31 |  | FDA,2 |  |  | USP | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Chlorhexidine nitrile | RRT of Chlorhexidine nitrile | 0.52 |  | FDA,2 |  |  | USP | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of o-Chlorhexidine and specified unidentified | RRT of o-Chlorhexidine and specified unidentified | 0.9 |  | FDA,2 |  |  | USP | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Oxochlorhexidine | RRT of Oxochlorhexidine | 1.43 |  | FDA,2 |  |  | USP | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of p-Chlorophenvl urea | RRT of p-Chlorophenvl urea | 0 |  | FDA,2 |  |  | USP | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Specified unidentified impuritv 1 | RRT of Specified unidentified impuritv 1 | 0.17 |  | FDA,2 |  |  | USP | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Specified unidentified impuritv 1 | Specified unidentified impuritv 1 | 0.011 | % | FDA,1 |  | 0.2 | USP | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Total Impurities | Total Impurities | 0.983 | % | FDA,1 |  | 3 | USP | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 398 | pH | pH (Diluted 1 in 20) | pH (Diluted 1 in 20) | 5.55 |  | FDA,1 | 5.5 | 7 | USP <791> | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 796 | Assay | Assay | Assay | 20.1 | % | FDA,1 | 19 | 21 | USP-HPLC | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 796 | Assay | Assay Final | Assay Final | 20.1 | % | FDA,1 | 19 | 21 | USP-HPLC | 7O4OS407/Chem_173A | 14 | both
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.0648 |  | FDA,2 | 1.06 | 1.07 | USP/NF <841> | 7O4OS407/Chem_173A | 14 | both
7O4OU181/Chem_NY_Isopropanol 70% | 7O4OU181 | Chem_NY_Isopropanol 70% | IN-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411 |  |  | left_only
7O4OU181/Chem_NY_Isopropanol 70% | 7O4OU181 | Chem_NY_Isopropanol 70% | IN-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA |  |  | left_only
7O4OU181/Chem_NY_Isopropanol 70% | 7O4OU181 | Chem_NY_Isopropanol 70% | IN-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear (Done) |  |  |  |  | TM 208 |  |  | left_only
7O4OU181/Chem_NY_Isopropanol 70% | 7O4OU181 | Chem_NY_Isopropanol 70% | IN-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM 208 |  |  | left_only
7O4OU181/Chem_NY_Isopropanol 70% | 7O4OU181 | Chem_NY_Isopropanol 70% | IN-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208 |  |  | left_only
7O4OU181/Chem_NY_Isopropanol 70% | 7O4OU181 | Chem_NY_Isopropanol 70% | IN-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 |  |  | left_only
7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 7O4OU225 | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH | 7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 11 | both
7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 7O4OU225 | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 398 | pH | pH (10% Solution) | pH (10% Solution) | 11.8 |  | FDA,1 | 10.5 | 13 | TM 200 | 7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 11 | both
7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 7O4OU225 | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 807 | % NaOCl | % NaOCl by weight (calc) | % NaOCl by weight | 9.9316 | % | FDA,2 | 11.5 | 14 | TM 366 | 7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 11 | both
7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 7O4OU225 | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 807 | % NaOCl | % NaOCl by weight Final (calc) | % NaOCl by weight Final | 10.0533 | % | FDA,2 | 11.5 | 14 | TM 366 | 7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 11 | both
7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 7O4OU225 | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 807 | % NaOCl | Sample Weight (g) | Sample Weight | 1.0023 | g | FDA,4 |  |  | TM 366 | 7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 11 | both
7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 7O4OU225 | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 807 | % NaOCl | Titrant End Volume (mL) | Titrant End Volume | 27.5 | mL | FDA,1 |  |  | TM 366 | 7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 11 | both
7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 7O4OU225 | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 807 | % NaOCl | Titrant Start Volume (mL) | Titrant Start Volume | 0 | mL | FDA,1 |  |  | TM 366 | 7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 11 | both
7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 7O4OU225 | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 807 | % NaOCl | Total Volume of Titrant (mL) | Total Volume of Titrant | 27.5 | mL | FDA,1 |  |  | TM 366 | 7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 11 | both
7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 7O4OU225 | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear liquid with a chlorine odor (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 11 | both
7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 7O4OU225 | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | straw-yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | 7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 11 | both
7O4OU351/Chem_NY_Glycopyrronium Tosylate Monohydrate | 7O4OU351 | Chem_NY_Glycopyrronium Tosylate Monohydrate | NY-Analytical | 1051 | Temptale Reading | Shipment Temperatutre | Shipment Temperatutre | Send Temperature data to Dermira (Done) |  |  |  |  | Shipment Temperatutre | 7O4OU351/Chem_NY_Glycopyrronium Tosylate Monohydrate | 3 | both
7O4OU351/Chem_NY_Glycopyrronium Tosylate Monohydrate | 7O4OU351 | Chem_NY_Glycopyrronium Tosylate Monohydrate | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA | 7O4OU351/Chem_NY_Glycopyrronium Tosylate Monohydrate | 3 | both
7O4OU351/Chem_NY_Glycopyrronium Tosylate Monohydrate | 7O4OU351 | Chem_NY_Glycopyrronium Tosylate Monohydrate | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis (ELLI) | Certificate of Analysis | COA Meets Specification (Done) |  |  |  |  | COA | 7O4OU351/Chem_NY_Glycopyrronium Tosylate Monohydrate | 3 | both
7O4OU555/Chem_NY_Lidocaine Reduced testing | 7O4OU555 | Chem_NY_Lidocaine Reduced testing | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis | 7O4OU555/Chem_NY_Lidocaine Reduced testing | 3 | both
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Chem_NY_Lidocaine Base USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection | 7O4OU555/Chem_NY_Lidocaine Base USP | 5 | both
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Chem_NY_Lidocaine Base USP | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Match standard (Done) |  |  |  |  | TM 208 | 7O4OU555/Chem_NY_Lidocaine Base USP | 5 | both
7O4OU555/Chem_NY_Lidocaine Reduced testing | 7O4OU555 | Chem_NY_Lidocaine Reduced testing | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Match standard (Done) |  |  |  |  | TM 208 | 7O4OU555/Chem_NY_Lidocaine Reduced testing | 3 | both
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Chem_NY_Lidocaine Base USP | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Match standard (Done) |  |  |  |  | TM 208 | 7O4OU555/Chem_NY_Lidocaine Base USP | 5 | both
7O4OU555/Chem_NY_Lidocaine Reduced testing | 7O4OU555 | Chem_NY_Lidocaine Reduced testing | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Match standard (Done) |  |  |  |  | TM 208 | 7O4OU555/Chem_NY_Lidocaine Reduced testing | 3 | both
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Chem_NY_Lidocaine Base USP | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 | 7O4OU555/Chem_NY_Lidocaine Base USP | 5 | both
7O4OU555/Chem_NY_Lidocaine Reduced testing | 7O4OU555 | Chem_NY_Lidocaine Reduced testing | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208 | 7O4OU555/Chem_NY_Lidocaine Reduced testing | 3 | both
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Chem_NY_Lidocaine Base USP | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208 | 7O4OU555/Chem_NY_Lidocaine Base USP | 5 | both
7O4OU555/Chem_NY_Lidocaine Reduced testing | 7O4OU555 | Chem_NY_Lidocaine Reduced testing | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208 | 7O4OU555/Chem_NY_Lidocaine Reduced testing | 3 | both
7O4OU555/Chem_NY_Lidocaine Reduced testing | 7O4OU555 | Chem_NY_Lidocaine Reduced testing | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7O4OU555/Chem_NY_Lidocaine Reduced testing | 3 | both
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Chem_NY_Lidocaine Base USP | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | 7O4OU555/Chem_NY_Lidocaine Base USP | 5 | both
7O4OU555/Chem_NY_Lidocaine Reduced testing | 7O4OU555 | Chem_NY_Lidocaine Reduced testing | NY-Analytical | 796 | Assay | Assay | Assay | 100.8 | % | FDA,1 | 97.5 | 102.5 | USP | 7O4OU555/Chem_NY_Lidocaine Reduced testing | 3 | both
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Chem_NY_Lidocaine Base USP | NY-Analytical | 796 | Assay | Assay | Assay | 100.4 | % | FDA,1 | 97.5 | 102.5 | USP | 7O4OU555/Chem_NY_Lidocaine Base USP | 5 | both
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Chem_NY_Lidocaine Base USP | NY-Analytical | 832 | Chloride | Chloride | Chloride |  | % | FDA,4 |  | 0.0035 | USP | 7O4OU555/Chem_NY_Lidocaine Base USP | 5 | both
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Chem_NY_Lidocaine Base USP | NY-Analytical | 833 | Heavy Metals | Heavy Metals | Heavy Metals |  | ug/g | FDA,3 |  | 0.002 | USP | 7O4OU555/Chem_NY_Lidocaine Base USP | 5 | both
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Chem_NY_Lidocaine Base USP | NY-Analytical | 809 | Melting Range | Melting Range - End | Melting Range - End |  | C | FDA,1 | 66 | 69 | USP | 7O4OU555/Chem_NY_Lidocaine Base USP | 5 | both
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Chem_NY_Lidocaine Base USP | NY-Analytical | 809 | Melting Range | Melting Range - Start | Melting Range - Start |  | C | FDA,1 | 66 | 69 | USP | 7O4OU555/Chem_NY_Lidocaine Base USP | 5 | both
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Chem_NY_Lidocaine Base USP | NY-Analytical | 834 | Residue on Ignition | Residue on Ignition | Residue on Ignition |  | % | FDA,1 |  | 0.1 | USP | 7O4OU555/Chem_NY_Lidocaine Base USP | 5 | both
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Chem_NY_Lidocaine Base USP | NY-Analytical | 835 | Sulfate Content | Sulfate | Sulfate |  | % | FDA,2 |  | 0.1 | USP | 7O4OU555/Chem_NY_Lidocaine Base USP | 5 | both
7O4OX231/Chem_NY_FD&C Red 40 | 7O4OX231 | Chem_NY_FD&C Red 40 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Appearance | 7O4OX231/Chem_NY_FD&C Red 40 | 2 | both
7O4OX231/Chem_NY_FD&C Red 40 | 7O4OX231 | Chem_NY_FD&C Red 40 | NY-Analytical | 397 | Appearance | Physical Properties | Physical Properties | Brown Powder with Purple/red tint Powder (Done) |  |  |  |  | TM 208 | 7O4OX231/Chem_NY_FD&C Red 40 | 2 | both
7O4OX231/Chem_NY_FD&C Red 40 | 7O4OX231 | Chem_NY_FD&C Red 40 | NY-Analytical | 1061 | Total Color | Spectrophotometric | Spectrophotometric | 89.4 | % | FDA,0 | 85 |  | TM 3014 | 7O4OX231/Chem_NY_FD&C Red 40 | 2 | both
7R4OP485/Chem_NY_Phenoxyethanol_7R4OP485 | 7R4OP485 | Chem_NY_Phenoxyethanol_7R4OP485 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 |  |  | left_only
7R4OP485/Chem_NY_Phenoxyethanol_7R4OP485 | 7R4OP485 | Chem_NY_Phenoxyethanol_7R4OP485 | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C |  |  | FDA,3 | 1.105 | 1.11 | USP/NF |  |  | left_only
7R4OP485/Chem_NY_Phenoxyethanol_7R4OP485 | 7R4OP485 | Chem_NY_Phenoxyethanol_7R4OP485 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  |  | left_only
7R4OP485/Chem_NY_Phenoxyethanol_7R4OP485 | 7R4OP485 | Chem_NY_Phenoxyethanol_7R4OP485 | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear, thin colorless liquid (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7R4OP485/Chem_NY_Phenoxyethanol_7R4OP485 | 7R4OP485 | Chem_NY_Phenoxyethanol_7R4OP485 | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 |  |  | left_only
7R4OP485/Chem_NY_Phenoxyethanol_7R4OP485 | 7R4OP485 | Chem_NY_Phenoxyethanol_7R4OP485 | NY-Analytical | 849 | Purity | Purity | Purity |  | % | FDA,1 | 98 | 102 | USP/NF |  |  | left_only
7R4OP485/Chem_NY_Phenoxyethanol_7R4OP485 | 7R4OP485 | Chem_NY_Phenoxyethanol_7R4OP485 | NY-Analytical | 892 | Limit of Phenol | Phenol | Phenol |  | ppm | FDA,0 |  | 25 | USP/NF |  |  | left_only
8AA654SI/Label_A10 | 8AA654SI | Label_A10 | NY-Analytical | 770 | Label Inspection | Length | Length |  | in | FDA,4 | 1.3125 | 1.4375 | TM 143 / Visual |  |  | left_only
8AA654SI/Label_A10 | 8AA654SI | Label_A10 | NY-Analytical | 770 | Label Inspection | Width | Width |  | in | FDA,3 | 2.5 | 2.625 | TM 143 / Visual |  |  | left_only
8AA654SI/Label_NY_Skip lot testing | 8AA654SI | Label_NY_Skip lot testing | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | [5 empty] | TM 143 / Visual | 8AA654SI/Label_NY_Skip lot testing | 7 | both
8AA654SI/Label_A10 | 8AA654SI | Label_A10 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8AA654SI/Label_NY_Skip lot testing | 8AA654SI | Label_NY_Skip lot testing | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | 8AA654SI/Label_NY_Skip lot testing | 7 | both
8AA654SI/Label_A10 | 8AA654SI | Label_A10 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8AA654SI/Label_NY_Skip lot testing | 8AA654SI | Label_NY_Skip lot testing | NY-Analytical | 770 | Label Inspection | Length | Length | 1.375 | in | FDA,4 | 1.3125 | 1.4375 | TM 143 / Visual | 8AA654SI/Label_NY_Skip lot testing | 7 | both
8AA654SI/Label_A10 | 8AA654SI | Label_A10 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8AA654SI/Label_NY_Skip lot testing | 8AA654SI | Label_NY_Skip lot testing | NY-Analytical | 770 | Label Inspection | Odor | Odor | [5 empty] | TM 143 / Visual | 8AA654SI/Label_NY_Skip lot testing | 7 | both
8AA654SI/Label_A10 | 8AA654SI | Label_A10 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8AA654SI/Label_NY_Skip lot testing | 8AA654SI | Label_NY_Skip lot testing | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | 8AA654SI/Label_NY_Skip lot testing | 7 | both
8AA654SI/Label_A10 | 8AA654SI | Label_A10 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  |  | left_only
8AA654SI/Label_NY_Skip lot testing | 8AA654SI | Label_NY_Skip lot testing | NY-Analytical | 770 | Label Inspection | Width | Width | 2.562 | in | FDA,3 | 2.5 | 2.625 | TM 143 / Visual | 8AA654SI/Label_NY_Skip lot testing | 7 | both
8AA70900/Label_NY_NP9990_2.5-4.625 | 8AA70900 | Label_NY_NP9990_2.5-4.625 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection | 8AA70900/Label_NY_NP9990_2.5-4.625 | 1 | both
8AA70900/Label_NY_NP9990_2.5-4.625 | 8AA70900 | Label_NY_NP9990_2.5-4.625 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | 8AA70900/Label_NY_NP9990_2.5-4.625 | 1 | both
8AA70900/Label_NY_NP9990_2.5-4.625 | 8AA70900 | Label_NY_NP9990_2.5-4.625 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | 8AA70900/Label_NY_NP9990_2.5-4.625 | 1 | both
8AA70900/Label_NY_NP9990_2.5-4.625 | 8AA70900 | Label_NY_NP9990_2.5-4.625 | NY-Analytical | 770 | Label Inspection | Length | Length | 2.5 | in | FDA,4 | 2.4375 | 2.5625 | TM 143 / Visual | 8AA70900/Label_NY_NP9990_2.5-4.625 | 1 | both
8AA70900/Label_NY_NP9990_2.5-4.625 | 8AA70900 | Label_NY_NP9990_2.5-4.625 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | 8AA70900/Label_NY_NP9990_2.5-4.625 | 1 | both
8AA70900/Label_NY_NP9990_2.5-4.625 | 8AA70900 | Label_NY_NP9990_2.5-4.625 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | 8AA70900/Label_NY_NP9990_2.5-4.625 | 1 | both
8AA70900/Label_NY_NP9990_2.5-4.625 | 8AA70900 | Label_NY_NP9990_2.5-4.625 | NY-Analytical | 770 | Label Inspection | Width | Width | 4.625 | in | FDA,4 | 4.5625 | 4.6875 | TM 143 / Visual | 8AA70900/Label_NY_NP9990_2.5-4.625 | 1 | both
8AH12880/Label_NY_NP7030_9-2.625 | 8AH12880 | Label_NY_NP7030_9-2.625 | NY-Analytical | 770 | Label Inspection | Width | Width |  | in | FDA,4 | 2.5625 | 2.6875 | TM 143 / Visual |  |  | left_only
8AH12880/Label_NY_NP7030_9-2.625 | 8AH12880 | Label_NY_NP7030_9-2.625 | NY-Analytical | 770 | Label Inspection | Length | Length |  | in | FDA,4 | 8.9375 | 9.0625 | TM 143 / Visual |  |  | left_only
8AH12880/Label_NY_NP7030_9-2.625 | 8AH12880 | Label_NY_NP7030_9-2.625 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8AH12880/Label_NY_NP7030_9-2.625 | 8AH12880 | Label_NY_NP7030_9-2.625 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8AH12880/Label_NY_NP7030_9-2.625 | 8AH12880 | Label_NY_NP7030_9-2.625 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8AH12880/Label_NY_NP7030_9-2.625 | 8AH12880 | Label_NY_NP7030_9-2.625 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8AH12880/Label_NY_NP7030_9-2.625 | 8AH12880 | Label_NY_NP7030_9-2.625 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  |  | left_only
8AP02700RF/Label_NY_NP9815_73-85 | 8AP02700RF | Label_NY_NP9815_73-85 | NY-Analytical | 770 | Label Inspection | Length | Length |  | mm | FDA,4 | 71.4125 | 74.5875 | TM 143 / Visual |  |  | left_only
8AP02700RF/Label_NY_NP9815_73-85 | 8AP02700RF | Label_NY_NP9815_73-85 | NY-Analytical | 770 | Label Inspection | Width | Width |  | mm | FDA,4 | 83.4125 | 86.5875 | TM 143 / Visual |  |  | left_only
8AP02700RF/Label_NY_NP9815_73-85 | 8AP02700RF | Label_NY_NP9815_73-85 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8AP02700RF/Label_NY_NP9815_73-85 | 8AP02700RF | Label_NY_NP9815_73-85 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8AP02700RF/Label_NY_NP9815_73-85 | 8AP02700RF | Label_NY_NP9815_73-85 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8AP02700RF/Label_NY_NP9815_73-85 | 8AP02700RF | Label_NY_NP9815_73-85 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8AP02700RF/Label_NY_NP9815_73-85 | 8AP02700RF | Label_NY_NP9815_73-85 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  |  | left_only
8AP034RSL/Label_NY_NP9800_3.375-2.3437 | 8AP034RSL | Label_NY_NP9800_3.375-2.3437 | IN-Analytical | 770 | Label Inspection | Width | Width |  | in | FDA,4 | 2.2812 | 2.4062 | TM 143 / Visual |  |  | left_only
8AP034RSL/Label_NY_NP9800_3.375-2.3437 | 8AP034RSL | Label_NY_NP9800_3.375-2.3437 | IN-Analytical | 770 | Label Inspection | Length | Length |  | in | FDA,4 | 3.3125 | 3.4375 | TM 143 / Visual |  |  | left_only
8AP034RSL/Label_NY_NP9800_3.375-2.3437 | 8AP034RSL | Label_NY_NP9800_3.375-2.3437 | IN-Analytical | 770 | Label Inspection | Certificate of Conformance | Certificate of Conformance | CofC meets specifications (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8AP034RSL/Label_NY_NP9800_3.375-2.3437 | 8AP034RSL | Label_NY_NP9800_3.375-2.3437 | IN-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8AP034RSL/Label_NY_NP9800_3.375-2.3437 | 8AP034RSL | Label_NY_NP9800_3.375-2.3437 | IN-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8AP034RSL/Label_NY_NP9800_3.375-2.3437 | 8AP034RSL | Label_NY_NP9800_3.375-2.3437 | IN-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8AP034RSL/Label_NY_NP9800_3.375-2.3437 | 8AP034RSL | Label_NY_NP9800_3.375-2.3437 | IN-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8AP13800/Label_NY_NP7200_6.25-4 | 8AP13800 | Label_NY_NP7200_6.25-4 | NY-Analytical | 770 | Label Inspection | Width | Width |  | in | FDA,4 | 3.9375 | 4.0625 | TM 143 / Visual |  |  | left_only
8AP13800/Label_NY_NP7200_6.25-4 | 8AP13800 | Label_NY_NP7200_6.25-4 | NY-Analytical | 770 | Label Inspection | Length | Length |  | in | FDA,4 | 6.1875 | 6.3125 | TM 143 / Visual |  |  | left_only
8AP13800/Label_NY_NP7200_6.25-4 | 8AP13800 | Label_NY_NP7200_6.25-4 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8AP13800/Label_NY_NP7200_6.25-4 | 8AP13800 | Label_NY_NP7200_6.25-4 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8AP13800/Label_NY_NP7200_6.25-4 | 8AP13800 | Label_NY_NP7200_6.25-4 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8AP13800/Label_NY_NP7200_6.25-4 | 8AP13800 | Label_NY_NP7200_6.25-4 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8AP13800/Label_NY_NP7200_6.25-4 | 8AP13800 | Label_NY_NP7200_6.25-4 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  |  | left_only
8AP21100A/Label_NY_NP9890_85-73 | 8AP21100A | Label_NY_NP9890_85-73 | NY-Analytical | 770 | Label Inspection | Width | Width |  | mm | FDA,4 | 71.4125 | 74.5875 | TM 143 / Visual |  |  | left_only
8AP21100A/Label_NY_NP9890_85-73 | 8AP21100A | Label_NY_NP9890_85-73 | NY-Analytical | 770 | Label Inspection | Length | Length |  | mm | FDA,4 | 83.4125 | 86.5875 | TM 143 / Visual |  |  | left_only
8AP21100A/Label_NY_NP9890_85-73 | 8AP21100A | Label_NY_NP9890_85-73 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8AP21100A/Label_NY_NP9890_85-73 | 8AP21100A | Label_NY_NP9890_85-73 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8AP21100A/Label_NY_NP9890_85-73 | 8AP21100A | Label_NY_NP9890_85-73 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8AP21100A/Label_NY_NP9890_85-73 | 8AP21100A | Label_NY_NP9890_85-73 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8AP21100A/Label_NY_NP9890_85-73 | 8AP21100A | Label_NY_NP9890_85-73 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  |  | left_only
8AP22800/Label_NY_NP7200_5.8125-6.5 | 8AP22800 | Label_NY_NP7200_5.8125-6.5 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | [5 empty] | Label Inspection | 8AP22800/Label_NY_NP7200_5.8125-6.5 | 1 | both
8AP22800/Label_NY_NP7200_5.8125-6.5 | 8AP22800 | Label_NY_NP7200_5.8125-6.5 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | 8AP22800/Label_NY_NP7200_5.8125-6.5 | 1 | both
8AP22800/Label_NY_NP7200_5.8125-6.5 | 8AP22800 | Label_NY_NP7200_5.8125-6.5 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | 8AP22800/Label_NY_NP7200_5.8125-6.5 | 1 | both
8AP22800/Label_NY_NP7200_5.8125-6.5 | 8AP22800 | Label_NY_NP7200_5.8125-6.5 | NY-Analytical | 770 | Label Inspection | Length | Length | 5.75 | in | FDA,4 | 5.75 | 5.875 | TM 143 / Visual | 8AP22800/Label_NY_NP7200_5.8125-6.5 | 1 | both
8AP22800/Label_NY_NP7200_5.8125-6.5 | 8AP22800 | Label_NY_NP7200_5.8125-6.5 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | 8AP22800/Label_NY_NP7200_5.8125-6.5 | 1 | both
8AP22800/Label_NY_NP7200_5.8125-6.5 | 8AP22800 | Label_NY_NP7200_5.8125-6.5 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | 8AP22800/Label_NY_NP7200_5.8125-6.5 | 1 | both
8AP22800/Label_NY_NP7200_5.8125-6.5 | 8AP22800 | Label_NY_NP7200_5.8125-6.5 | NY-Analytical | 770 | Label Inspection | Width | Width | 6.5 | in | FDA,4 | 6.4375 | 6.5625 | TM 143 / Visual | 8AP22800/Label_NY_NP7200_5.8125-6.5 | 1 | both
8AP32800/Label_NY_NP98372_73-85 | 8AP32800 | Label_NY_NP98372_73-85 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection | 8AP32800/Label_NY_NP98372_73-85 | 1 | both
8AP32800/Label_NY_NP98372_73-85 | 8AP32800 | Label_NY_NP98372_73-85 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | 8AP32800/Label_NY_NP98372_73-85 | 1 | both
8AP32800/Label_NY_NP98372_73-85 | 8AP32800 | Label_NY_NP98372_73-85 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | 8AP32800/Label_NY_NP98372_73-85 | 1 | both
8AP32800/Label_NY_NP98372_73-85 | 8AP32800 | Label_NY_NP98372_73-85 | NY-Analytical | 770 | Label Inspection | Length | Length | 73 | mm | FDA,4 | 71.4125 | 74.5875 | TM 143 / Visual | 8AP32800/Label_NY_NP98372_73-85 | 1 | both
8AP32800/Label_NY_NP98372_73-85 | 8AP32800 | Label_NY_NP98372_73-85 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | 8AP32800/Label_NY_NP98372_73-85 | 1 | both
8AP32800/Label_NY_NP98372_73-85 | 8AP32800 | Label_NY_NP98372_73-85 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | 8AP32800/Label_NY_NP98372_73-85 | 1 | both
8AP32800/Label_NY_NP98372_73-85 | 8AP32800 | Label_NY_NP98372_73-85 | NY-Analytical | 770 | Label Inspection | Width | Width | 85 | mm | FDA,4 | 83.4125 | 86.5875 | TM 143 / Visual | 8AP32800/Label_NY_NP98372_73-85 | 1 | both
8AP60100RF/Label_NY_NP9890_73-85 | 8AP60100RF | Label_NY_NP9890_73-85 | NY-Analytical | 770 | Label Inspection | Length | Length |  | mm | FDA,4 | 71.4125 | 74.5875 | TM 143 / Visual |  |  | left_only
8AP60100RF/Label_NY_NP9890_73-85 | 8AP60100RF | Label_NY_NP9890_73-85 | NY-Analytical | 770 | Label Inspection | Width | Width |  | mm | FDA,4 | 83.4125 | 86.5875 | TM 143 / Visual |  |  | left_only
8AP60100RF/Label_NY_NP9890_73-85 | 8AP60100RF | Label_NY_NP9890_73-85 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8AP60100RF/Label_NY_NP9890_73-85 | 8AP60100RF | Label_NY_NP9890_73-85 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8AP60100RF/Label_NY_NP9890_73-85 | 8AP60100RF | Label_NY_NP9890_73-85 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8AP60100RF/Label_NY_NP9890_73-85 | 8AP60100RF | Label_NY_NP9890_73-85 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8AP60100RF/Label_NY_NP9890_73-85 | 8AP60100RF | Label_NY_NP9890_73-85 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  |  | left_only
8AU120CK/Label_NY_NP7030_6-4 | 8AU120CK | Label_NY_NP7030_6-4 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection | 8AU120CK/Label_NY_NP7030_6-4 | 1 | both
8AU120CK/Label_NY_NP7030_6-4 | 8AU120CK | Label_NY_NP7030_6-4 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | 8AU120CK/Label_NY_NP7030_6-4 | 1 | both
8AU120CK/Label_NY_NP7030_6-4 | 8AU120CK | Label_NY_NP7030_6-4 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | 8AU120CK/Label_NY_NP7030_6-4 | 1 | both
8AU120CK/Label_NY_NP7030_6-4 | 8AU120CK | Label_NY_NP7030_6-4 | NY-Analytical | 770 | Label Inspection | Length | Length | 6 | in | FDA,4 | 5.9375 | 6.0625 | TM 143 / Visual | 8AU120CK/Label_NY_NP7030_6-4 | 1 | both
8AU120CK/Label_NY_NP7030_6-4 | 8AU120CK | Label_NY_NP7030_6-4 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | 8AU120CK/Label_NY_NP7030_6-4 | 1 | both
8AU120CK/Label_NY_NP7030_6-4 | 8AU120CK | Label_NY_NP7030_6-4 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | 8AU120CK/Label_NY_NP7030_6-4 | 1 | both
8AU120CK/Label_NY_NP7030_6-4 | 8AU120CK | Label_NY_NP7030_6-4 | NY-Analytical | 770 | Label Inspection | Width | Width | 4 | in | FDA,4 | 3.9375 | 4.0625 | TM 143 / Visual | 8AU120CK/Label_NY_NP7030_6-4 | 1 | both
8BB196CK/Label_NY_NP7030_6-5 | 8BB196CK | Label_NY_NP7030_6-5 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection | 8BB196CK/Label_NY_NP7030_6-5 | 2 | both
8BB196CK/Label_NY_NP7030_6-5 | 8BB196CK | Label_NY_NP7030_6-5 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | 8BB196CK/Label_NY_NP7030_6-5 | 2 | both
8BB196CK/Label_NY_NP7030_6-5 | 8BB196CK | Label_NY_NP7030_6-5 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | 8BB196CK/Label_NY_NP7030_6-5 | 2 | both
8BB196CK/Label_NY_NP7030_6-5 | 8BB196CK | Label_NY_NP7030_6-5 | NY-Analytical | 770 | Label Inspection | Length | Length | 5.99 | in | FDA,4 | 5.9375 | 6.0625 | TM 143 / Visual | 8BB196CK/Label_NY_NP7030_6-5 | 2 | both
8BB196CK/Label_NY_NP7030_6-5 | 8BB196CK | Label_NY_NP7030_6-5 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | 8BB196CK/Label_NY_NP7030_6-5 | 2 | both
8BB196CK/Label_NY_NP7030_6-5 | 8BB196CK | Label_NY_NP7030_6-5 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | 8BB196CK/Label_NY_NP7030_6-5 | 2 | both
8BB196CK/Label_NY_NP7030_6-5 | 8BB196CK | Label_NY_NP7030_6-5 | NY-Analytical | 770 | Label Inspection | Width | Width | 5.01 | in | FDA,4 | 4.9375 | 5.0625 | TM 143 / Visual | 8BB196CK/Label_NY_NP7030_6-5 | 2 | both
8BB512CL/Label_NY_NP7030_6-4 | 8BB512CL | Label_NY_NP7030_6-4 | NY-Analytical | 770 | Label Inspection | AQL Sample Size | AQL Sample Size | [5 empty] | TM 143 / Visual |  |  | left_only
8BB512CL/Label_NY_NP7030_6-4 | 8BB512CL | Label_NY_NP7030_6-4 | NY-Analytical | 770 | Label Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Sample(s) all meet specifications (Done) | in |  |  |  | TM 143 / Visual |  |  | left_only
8BP12801/Label_NY_NP7200_4.5-4 | 8BP12801 | Label_NY_NP7200_4.5-4 | NY-Analytical | 770 | Label Inspection | Width | Width |  | in | FDA,4 | 3.9375 | 4.0625 | TM 143 / Visual |  |  | left_only
8BP12801/Label_NY_NP7200_4.5-4 | 8BP12801 | Label_NY_NP7200_4.5-4 | NY-Analytical | 770 | Label Inspection | Length | Length |  | in | FDA,4 | 4.4375 | 4.5625 | TM 143 / Visual |  |  | left_only
8BP12801/Label_NY_NP7200_4.5-4 | 8BP12801 | Label_NY_NP7200_4.5-4 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8BP12801/Label_NY_NP7200_4.5-4 | 8BP12801 | Label_NY_NP7200_4.5-4 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8BP12801/Label_NY_NP7200_4.5-4 | 8BP12801 | Label_NY_NP7200_4.5-4 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8BP12801/Label_NY_NP7200_4.5-4 | 8BP12801 | Label_NY_NP7200_4.5-4 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8BP12801/Label_NY_NP7200_4.5-4 | 8BP12801 | Label_NY_NP7200_4.5-4 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  |  | left_only
8BP24500/Label_NY_NP7200_4-4.5 | 8BP24500 | Label_NY_NP7200_4-4.5 | NY-Analytical | 770 | Label Inspection | Length | Length |  | in | FDA,4 | 3.9375 | 4.0625 | TM 143 / Visual |  |  | left_only
8BP24500/Label_NY_NP7200_4-4.5 | 8BP24500 | Label_NY_NP7200_4-4.5 | NY-Analytical | 770 | Label Inspection | Width | Width |  | in | FDA,4 | 4.4375 | 4.5625 | TM 143 / Visual |  |  | left_only
8BP24500/Label_NY_NP7200_4-4.5 | 8BP24500 | Label_NY_NP7200_4-4.5 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8BP24500/Label_NY_NP7200_4-4.5 | 8BP24500 | Label_NY_NP7200_4-4.5 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8BP24500/Label_NY_NP7200_4-4.5 | 8BP24500 | Label_NY_NP7200_4-4.5 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8BP24500/Label_NY_NP7200_4-4.5 | 8BP24500 | Label_NY_NP7200_4-4.5 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8BP24500/Label_NY_NP7200_4-4.5 | 8BP24500 | Label_NY_NP7200_4-4.5 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  |  | left_only
8BQ76500/Label_NY_NP7200_6-3.75 | 8BQ76500 | Label_NY_NP7200_6-3.75 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection | 8BQ76500/Label_NY_NP7200_6-3.75 | 1 | both
8BQ76500/Label_NY_NP7200_6-3.75 | 8BQ76500 | Label_NY_NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | 8BQ76500/Label_NY_NP7200_6-3.75 | 1 | both
8BQ76500/Label_NY_NP7200_6-3.75 | 8BQ76500 | Label_NY_NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | 8BQ76500/Label_NY_NP7200_6-3.75 | 1 | both
8BQ76500/Label_NY_NP7200_6-3.75 | 8BQ76500 | Label_NY_NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Length | Length | 6 | in | FDA,4 | 5.9375 | 6.0625 | TM 143 / Visual | 8BQ76500/Label_NY_NP7200_6-3.75 | 1 | both
8BQ76500/Label_NY_NP7200_6-3.75 | 8BQ76500 | Label_NY_NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | 8BQ76500/Label_NY_NP7200_6-3.75 | 1 | both
8BQ76500/Label_NY_NP7200_6-3.75 | 8BQ76500 | Label_NY_NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | 8BQ76500/Label_NY_NP7200_6-3.75 | 1 | both
8BQ76500/Label_NY_NP7200_6-3.75 | 8BQ76500 | Label_NY_NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Width | Width | 3.6875 | in | FDA,4 | 3.6875 | 3.8125 | TM 143 / Visual | 8BQ76500/Label_NY_NP7200_6-3.75 | 1 | both
8BQ76501/Label_NY_NP7200_3.75-6 | 8BQ76501 | Label_NY_NP7200_3.75-6 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection | 8BQ76501/Label_NY_NP7200_3.75-6 | 1 | both
8BQ76501/Label_NY_NP7200_3.75-6 | 8BQ76501 | Label_NY_NP7200_3.75-6 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | 8BQ76501/Label_NY_NP7200_3.75-6 | 1 | both
8BQ76501/Label_NY_NP7200_3.75-6 | 8BQ76501 | Label_NY_NP7200_3.75-6 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | 8BQ76501/Label_NY_NP7200_3.75-6 | 1 | both
8BQ76501/Label_NY_NP7200_3.75-6 | 8BQ76501 | Label_NY_NP7200_3.75-6 | NY-Analytical | 770 | Label Inspection | Length | Length | 3.6875 | in | FDA,4 | 3.6875 | 3.8125 | TM 143 / Visual | 8BQ76501/Label_NY_NP7200_3.75-6 | 1 | both
8BQ76501/Label_NY_NP7200_3.75-6 | 8BQ76501 | Label_NY_NP7200_3.75-6 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | 8BQ76501/Label_NY_NP7200_3.75-6 | 1 | both
8BQ76501/Label_NY_NP7200_3.75-6 | 8BQ76501 | Label_NY_NP7200_3.75-6 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | 8BQ76501/Label_NY_NP7200_3.75-6 | 1 | both
8BQ76501/Label_NY_NP7200_3.75-6 | 8BQ76501 | Label_NY_NP7200_3.75-6 | NY-Analytical | 770 | Label Inspection | Width | Width | 6 | in | FDA,4 | 5.9375 | 6.0625 | TM 143 / Visual | 8BQ76501/Label_NY_NP7200_3.75-6 | 1 | both
8CH211CK/Label_A12 | 8CH211CK | Label_A12 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection | 8CH211CK/Label_A12 | 1 | both
8CH211CK/Label_A12 | 8CH211CK | Label_A12 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | 8CH211CK/Label_A12 | 1 | both
8CH211CK/Label_A12 | 8CH211CK | Label_A12 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | 8CH211CK/Label_A12 | 1 | both
8CH211CK/Label_A12 | 8CH211CK | Label_A12 | NY-Analytical | 770 | Label Inspection | Length | Length | 9 | in | FDA,4 | 8.9375 | 9.0625 | TM 143 / Visual | 8CH211CK/Label_A12 | 1 | both
8CH211CK/Label_A12 | 8CH211CK | Label_A12 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | 8CH211CK/Label_A12 | 1 | both
8CH211CK/Label_A12 | 8CH211CK | Label_A12 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | 8CH211CK/Label_A12 | 1 | both
8CH211CK/Label_A12 | 8CH211CK | Label_A12 | NY-Analytical | 770 | Label Inspection | Width | Width | 2.625 | in | FDA,4 | 2.5625 | 2.6875 | TM 143 / Visual | 8CH211CK/Label_A12 | 1 | both
8CP13400/Label_NY_NP7100_4.5-4 | 8CP13400 | Label_NY_NP7100_4.5-4 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection | 8CP13400/Label_NY_NP7100_4.5-4 | 4 | both
8CP13400/Label_NY_NP7100_4.5-4 | 8CP13400 | Label_NY_NP7100_4.5-4 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | 8CP13400/Label_NY_NP7100_4.5-4 | 4 | both
8CP13400/Label_NY_NP7100_4.5-4 | 8CP13400 | Label_NY_NP7100_4.5-4 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | 8CP13400/Label_NY_NP7100_4.5-4 | 4 | both
8CP13400/Label_NY_NP7100_4.5-4 | 8CP13400 | Label_NY_NP7100_4.5-4 | NY-Analytical | 770 | Label Inspection | Length | Length | 4.5 | in | FDA,4 | 4.4375 | 4.5625 | TM 143 / Visual | 8CP13400/Label_NY_NP7100_4.5-4 | 4 | both
8CP13400/Label_NY_NP7100_4.5-4 | 8CP13400 | Label_NY_NP7100_4.5-4 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | 8CP13400/Label_NY_NP7100_4.5-4 | 4 | both
8CP13400/Label_NY_NP7100_4.5-4 | 8CP13400 | Label_NY_NP7100_4.5-4 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | 8CP13400/Label_NY_NP7100_4.5-4 | 4 | both
8CP13400/Label_NY_NP7100_4.5-4 | 8CP13400 | Label_NY_NP7100_4.5-4 | NY-Analytical | 770 | Label Inspection | Width | Width | 4 | in | FDA,4 | 3.9375 | 4.0625 | TM 143 / Visual | 8CP13400/Label_NY_NP7100_4.5-4 | 4 | both
8CP15900/Label_NY_NP7100/NP7200_6-3.75 | 8CP15900 | Label_NY_NP7100/NP7200_6-3.75 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection | 8CP15900/Label_NY_NP7100/NP7200_6-3.75 | 2 | both
8CP15900/Label_NY_NP7100/NP7200_6-3.75 | 8CP15900 | Label_NY_NP7100/NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | 8CP15900/Label_NY_NP7100/NP7200_6-3.75 | 2 | both
8CP15900/Label_NY_NP7100/NP7200_6-3.75 | 8CP15900 | Label_NY_NP7100/NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | 8CP15900/Label_NY_NP7100/NP7200_6-3.75 | 2 | both
8CP15900/Label_NY_NP7100/NP7200_6-3.75 | 8CP15900 | Label_NY_NP7100/NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Length | Length | 5.99 | in | FDA,4 | 5.9375 | 6.0625 | TM 143 / Visual | 8CP15900/Label_NY_NP7100/NP7200_6-3.75 | 2 | both
8CP15900/Label_NY_NP7100/NP7200_6-3.75 | 8CP15900 | Label_NY_NP7100/NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | 8CP15900/Label_NY_NP7100/NP7200_6-3.75 | 2 | both
8CP15900/Label_NY_NP7100/NP7200_6-3.75 | 8CP15900 | Label_NY_NP7100/NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | 8CP15900/Label_NY_NP7100/NP7200_6-3.75 | 2 | both
8CP15900/Label_NY_NP7100/NP7200_6-3.75 | 8CP15900 | Label_NY_NP7100/NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Width | Width | 3.705 | in | FDA,4 | 3.6875 | 3.8125 | TM 143 / Visual | 8CP15900/Label_NY_NP7100/NP7200_6-3.75 | 2 | both
8CP63801/Label_A20 | 8CP63801 | Label_A20 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection | 8CP63801/Label_A20 | 1 | both
8CP63801/Label_A20 | 8CP63801 | Label_A20 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | [5 empty] | TM 143 / Visual | 8CP63801/Label_A20 | 1 | both
8CP63801/Label_A20 | 8CP63801 | Label_A20 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | [5 empty] | TM 143 / Visual | 8CP63801/Label_A20 | 1 | both
8CP63801/Label_A20 | 8CP63801 | Label_A20 | NY-Analytical | 770 | Label Inspection | Length | Length |  | in | FDA,4 | 5.9375 | 6.0625 | TM 143 / Visual | 8CP63801/Label_A20 | 1 | both
8CP63801/Label_A20 | 8CP63801 | Label_A20 | NY-Analytical | 770 | Label Inspection | Odor | Odor | [5 empty] | TM 143 / Visual | 8CP63801/Label_A20 | 1 | both
8CP63801/Label_A20 | 8CP63801 | Label_A20 | NY-Analytical | 770 | Label Inspection | Print | Print | [5 empty] | TM 143 / Visual | 8CP63801/Label_A20 | 1 | both
8CP63801/Label_A20 | 8CP63801 | Label_A20 | NY-Analytical | 770 | Label Inspection | Width | Width |  | in | FDA,4 | 3.6875 | 3.8125 | TM 143 / Visual | 8CP63801/Label_A20 | 1 | both
8DP92000/Label_NY_NP7200_6.5-5.8125 | 8DP92000 | Label_NY_NP7200_6.5-5.8125 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | [5 empty] | Label Inspection | 8DP92000/Label_NY_NP7200_6.5-5.8125 | 2 | both
8DP92000/Label_NY_NP7200_6.5-5.8125 | 8DP92000 | Label_NY_NP7200_6.5-5.8125 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | 8DP92000/Label_NY_NP7200_6.5-5.8125 | 2 | both
8DP92000/Label_NY_NP7200_6.5-5.8125 | 8DP92000 | Label_NY_NP7200_6.5-5.8125 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | 8DP92000/Label_NY_NP7200_6.5-5.8125 | 2 | both
8DP92000/Label_NY_NP7200_6.5-5.8125 | 8DP92000 | Label_NY_NP7200_6.5-5.8125 | NY-Analytical | 770 | Label Inspection | Length | Length | 6.5 | in | FDA,4 | 6.4375 | 6.5625 | TM 143 / Visual | 8DP92000/Label_NY_NP7200_6.5-5.8125 | 2 | both
8DP92000/Label_NY_NP7200_6.5-5.8125 | 8DP92000 | Label_NY_NP7200_6.5-5.8125 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | 8DP92000/Label_NY_NP7200_6.5-5.8125 | 2 | both
8DP92000/Label_NY_NP7200_6.5-5.8125 | 8DP92000 | Label_NY_NP7200_6.5-5.8125 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | 8DP92000/Label_NY_NP7200_6.5-5.8125 | 2 | both
8DP92000/Label_NY_NP7200_6.5-5.8125 | 8DP92000 | Label_NY_NP7200_6.5-5.8125 | NY-Analytical | 770 | Label Inspection | Width | Width | 5.8125 | in | FDA,4 | 5.75 | 5.875 | TM 143 / Visual | 8DP92000/Label_NY_NP7200_6.5-5.8125 | 2 | both
8DQ85400/Label_A41 | 8DQ85400 | Label_A41 | NY-Analytical | 770 | Label Inspection | Length | Length |  | in | FDA,4 | 5.5 | 5.625 | TM 143 / Visual |  |  | left_only
8DQ85400/Label_A41 | 8DQ85400 | Label_A41 | NY-Analytical | 770 | Label Inspection | Width | Width |  | in | FDA,4 | 6.1875 | 6.3125 | TM 143 / Visual |  |  | left_only
8DQ85400/Label_A41 | 8DQ85400 | Label_A41 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8DQ85400/Label_A41 | 8DQ85400 | Label_A41 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8DQ85400/Label_A41 | 8DQ85400 | Label_A41 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8DQ85400/Label_A41 | 8DQ85400 | Label_A41 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8DQ85400/Label_A41 | 8DQ85400 | Label_A41 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  |  | left_only
8DQ85900/Label_NY_NP7200_5.5625-6.25 | 8DQ85900 | Label_NY_NP7200_5.5625-6.25 | NY-Analytical | 770 | Label Inspection | Length | Length |  | in | FDA,4 | 5.5 | 5.625 | TM 143 / Visual |  |  | left_only
8DQ85900/Label_NY_NP7200_5.5625-6.25 | 8DQ85900 | Label_NY_NP7200_5.5625-6.25 | NY-Analytical | 770 | Label Inspection | Width | Width |  | in | FDA,4 | 6.1875 | 6.3125 | TM 143 / Visual |  |  | left_only
8DQ85900/Label_NY_NP7200_5.5625-6.25 | 8DQ85900 | Label_NY_NP7200_5.5625-6.25 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8DQ85900/Label_NY_NP7200_5.5625-6.25 | 8DQ85900 | Label_NY_NP7200_5.5625-6.25 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8DQ85900/Label_NY_NP7200_5.5625-6.25 | 8DQ85900 | Label_NY_NP7200_5.5625-6.25 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8DQ85900/Label_NY_NP7200_5.5625-6.25 | 8DQ85900 | Label_NY_NP7200_5.5625-6.25 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8DQ85900/Label_NY_NP7200_5.5625-6.25 | 8DQ85900 | Label_NY_NP7200_5.5625-6.25 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  |  | left_only
8DQ914SL/Label_NY_NP7030_9-7 | 8DQ914SL | Label_NY_NP7030_9-7 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | [5 empty] | Label Inspection | 8DQ914SL/Label_NY_NP7030_9-7 | 1 | both
8DQ914SL/Label_NY_NP7030_9-7 | 8DQ914SL | Label_NY_NP7030_9-7 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | 8DQ914SL/Label_NY_NP7030_9-7 | 1 | both
8DQ914SL/Label_NY_NP7030_9-7 | 8DQ914SL | Label_NY_NP7030_9-7 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | 8DQ914SL/Label_NY_NP7030_9-7 | 1 | both
8DQ914SL/Label_NY_NP7030_9-7 | 8DQ914SL | Label_NY_NP7030_9-7 | NY-Analytical | 770 | Label Inspection | Length | Length | 9 | in | FDA,4 | 8.9375 | 9.0625 | TM 143 / Visual | 8DQ914SL/Label_NY_NP7030_9-7 | 1 | both
8DQ914SL/Label_NY_NP7030_9-7 | 8DQ914SL | Label_NY_NP7030_9-7 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | 8DQ914SL/Label_NY_NP7030_9-7 | 1 | both
8DQ914SL/Label_NY_NP7030_9-7 | 8DQ914SL | Label_NY_NP7030_9-7 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | 8DQ914SL/Label_NY_NP7030_9-7 | 1 | both
8DQ914SL/Label_NY_NP7030_9-7 | 8DQ914SL | Label_NY_NP7030_9-7 | NY-Analytical | 770 | Label Inspection | Width | Width | 7 | in | FDA,4 | 6.9375 | 7.0625 | TM 143 / Visual | 8DQ914SL/Label_NY_NP7030_9-7 | 1 | both
8GQ772SK/Label_B03 | 8GQ772SK | Label_B03 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection | 8GQ772SK/Label_B03 | 6 | both
8GQ772SK/Label_B03 | 8GQ772SK | Label_B03 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | 8GQ772SK/Label_B03 | 6 | both
8GQ772SK/Label_B03 | 8GQ772SK | Label_B03 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | 8GQ772SK/Label_B03 | 6 | both
8GQ772SK/Label_B03 | 8GQ772SK | Label_B03 | NY-Analytical | 770 | Label Inspection | Diameter | Diameter | 1.875 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | 8GQ772SK/Label_B03 | 6 | both
8GQ772SK/Label_B03 | 8GQ772SK | Label_B03 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | 8GQ772SK/Label_B03 | 6 | both
8GQ772SK/Label_B03 | 8GQ772SK | Label_B03 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | 8GQ772SK/Label_B03 | 6 | both
8OH217UP/Label_A25 | 8OH217UP | Label_A25 | NY-Analytical | 770 | Label Inspection | Width | Width |  | in | FDA,4 | 3.9375 | 4.0625 | TM 143 / Visual |  |  | left_only
8OH217UP/Label_A25 | 8OH217UP | Label_A25 | NY-Analytical | 770 | Label Inspection | Length | Length |  | in | FDA,4 | 9.9375 | 10.0625 | TM 143 / Visual |  |  | left_only
8OH217UP/Label_A25 | 8OH217UP | Label_A25 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8OH217UP/Label_A25 | 8OH217UP | Label_A25 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8OH217UP/Label_A25 | 8OH217UP | Label_A25 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8OH217UP/Label_A25 | 8OH217UP | Label_A25 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8OH217UP/Label_A25 | 8OH217UP | Label_A25 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  |  | left_only
8OJ22750/Label_A11 | 8OJ22750 | Label_A11 | NY-Analytical | 770 | Label Inspection | Width | Width |  | in | FDA,4 | 2.5625 | 2.6875 | TM 143 / Visual |  |  | left_only
8OJ22750/Label_A11 | 8OJ22750 | Label_A11 | NY-Analytical | 770 | Label Inspection | Length | Length |  | in | FDA,4 | 5.3125 | 5.4375 | TM 143 / Visual |  |  | left_only
8OJ22750/Label_A11 | 8OJ22750 | Label_A11 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8OJ22750/Label_A11 | 8OJ22750 | Label_A11 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8OJ22750/Label_A11 | 8OJ22750 | Label_A11 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8OJ22750/Label_A11 | 8OJ22750 | Label_A11 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8OJ22750/Label_A11 | 8OJ22750 | Label_A11 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  |  | left_only
8OP03700/Label_NY_NP9840_87-75 | 8OP03700 | Label_NY_NP9840_87-75 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection | 8OP03700/Label_NY_NP9840_87-75 | 2 | both
8OP03700/Label_NY_NP9840_87-75 | 8OP03700 | Label_NY_NP9840_87-75 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | 8OP03700/Label_NY_NP9840_87-75 | 2 | both
8OP03700/Label_NY_NP9840_87-75 | 8OP03700 | Label_NY_NP9840_87-75 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | 8OP03700/Label_NY_NP9840_87-75 | 2 | both
8OP03700/Label_NY_NP9840_87-75 | 8OP03700 | Label_NY_NP9840_87-75 | NY-Analytical | 770 | Label Inspection | Length | Length | 86.97 | mm | FDA,4 | 85.4125 | 88.5875 | TM 143 / Visual | 8OP03700/Label_NY_NP9840_87-75 | 2 | both
8OP03700/Label_NY_NP9840_87-75 | 8OP03700 | Label_NY_NP9840_87-75 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | 8OP03700/Label_NY_NP9840_87-75 | 2 | both
8OP03700/Label_NY_NP9840_87-75 | 8OP03700 | Label_NY_NP9840_87-75 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | 8OP03700/Label_NY_NP9840_87-75 | 2 | both
8OP03700/Label_NY_NP9840_87-75 | 8OP03700 | Label_NY_NP9840_87-75 | NY-Analytical | 770 | Label Inspection | Width | Width | 75.06 | mm | FDA,4 | 73.4125 | 76.5875 | TM 143 / Visual | 8OP03700/Label_NY_NP9840_87-75 | 2 | both
8OP065RSL/Label_NY_NP9995_2.625-3.375 | 8OP065RSL | Label_NY_NP9995_2.625-3.375 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection | 8OP065RSL/Label_NY_NP9995_2.625-3.375 | 1 | both
8OP065RSL/Label_NY_NP9995_2.625-3.375 | 8OP065RSL | Label_NY_NP9995_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | 8OP065RSL/Label_NY_NP9995_2.625-3.375 | 1 | both
8OP065RSL/Label_NY_NP9995_2.625-3.375 | 8OP065RSL | Label_NY_NP9995_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | 8OP065RSL/Label_NY_NP9995_2.625-3.375 | 1 | both
8OP065RSL/Label_NY_NP9995_2.625-3.375 | 8OP065RSL | Label_NY_NP9995_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Length | Length | 2.625 | in | FDA,4 | 2.5625 | 2.6875 | TM 143 / Visual | 8OP065RSL/Label_NY_NP9995_2.625-3.375 | 1 | both
8OP065RSL/Label_NY_NP9995_2.625-3.375 | 8OP065RSL | Label_NY_NP9995_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | 8OP065RSL/Label_NY_NP9995_2.625-3.375 | 1 | both
8OP065RSL/Label_NY_NP9995_2.625-3.375 | 8OP065RSL | Label_NY_NP9995_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | 8OP065RSL/Label_NY_NP9995_2.625-3.375 | 1 | both
8OP065RSL/Label_NY_NP9995_2.625-3.375 | 8OP065RSL | Label_NY_NP9995_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Width | Width | 3.375 | in | FDA,4 | 3.3125 | 3.4375 | TM 143 / Visual | 8OP065RSL/Label_NY_NP9995_2.625-3.375 | 1 | both
8OP08700/Label_NY_NP9800_3.375-2.625 | 8OP08700 | Label_NY_NP9800_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Width | Width |  | in | FDA,4 | 2.5625 | 2.6875 | TM 143 / Visual |  |  | left_only
8OP08700/Label_NY_NP9800_3.375-2.625 | 8OP08700 | Label_NY_NP9800_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Length | Length |  | in | FDA,4 | 3.3125 | 3.4375 | TM 143 / Visual |  |  | left_only
8OP08700/Label_NY_NP9800_3.375-2.625 | 8OP08700 | Label_NY_NP9800_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8OP08700/Label_NY_NP9800_3.375-2.625 | 8OP08700 | Label_NY_NP9800_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8OP08700/Label_NY_NP9800_3.375-2.625 | 8OP08700 | Label_NY_NP9800_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8OP08700/Label_NY_NP9800_3.375-2.625 | 8OP08700 | Label_NY_NP9800_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8OP08700/Label_NY_NP9800_3.375-2.625 | 8OP08700 | Label_NY_NP9800_3.375-2.625 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  |  | left_only
8OP09100/Label_NY_NP9991_3.375-2.625 | 8OP09100 | Label_NY_NP9991_3.375-2.625 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Label Inspection | 8OP09100/Label_NY_NP9991_3.375-2.625 | 1 | both
8OP09100/Label_NY_NP9991_3.375-2.625 | 8OP09100 | Label_NY_NP9991_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | 8OP09100/Label_NY_NP9991_3.375-2.625 | 1 | both
8OP09100/Label_NY_NP9991_3.375-2.625 | 8OP09100 | Label_NY_NP9991_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | 8OP09100/Label_NY_NP9991_3.375-2.625 | 1 | both
8OP09100/Label_NY_NP9991_3.375-2.625 | 8OP09100 | Label_NY_NP9991_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Length | Length | 3.375 | in | FDA,4 | 3.3125 | 3.4375 | TM 143 / Visual | 8OP09100/Label_NY_NP9991_3.375-2.625 | 1 | both
8OP09100/Label_NY_NP9991_3.375-2.625 | 8OP09100 | Label_NY_NP9991_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | 8OP09100/Label_NY_NP9991_3.375-2.625 | 1 | both
8OP09100/Label_NY_NP9991_3.375-2.625 | 8OP09100 | Label_NY_NP9991_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | 8OP09100/Label_NY_NP9991_3.375-2.625 | 1 | both
8OP09100/Label_NY_NP9991_3.375-2.625 | 8OP09100 | Label_NY_NP9991_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Width | Width | 2.625 | in | FDA,4 | 2.5625 | 2.6875 | TM 143 / Visual | 8OP09100/Label_NY_NP9991_3.375-2.625 | 1 | both
8OP09300/Label_NY_NP7110_2.55-3 | 8OP09300 | Label_NY_NP7110_2.55-3 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | 8OP09300/Label_NY_NP7110_2.55-3 | 3 | both
8OP09300/Label_NY_NP7110_2.55-3 | 8OP09300 | Label_NY_NP7110_2.55-3 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | 8OP09300/Label_NY_NP7110_2.55-3 | 3 | both
8OP09300/Label_NY_NP7110_2.55-3 | 8OP09300 | Label_NY_NP7110_2.55-3 | NY-Analytical | 770 | Label Inspection | Length | Length | 2.5 | in | FDA,4 | 2.4875 | 2.6125 | TM 143 / Visual | 8OP09300/Label_NY_NP7110_2.55-3 | 3 | both
8OP09300/Label_NY_NP7110_2.55-3 | 8OP09300 | Label_NY_NP7110_2.55-3 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | 8OP09300/Label_NY_NP7110_2.55-3 | 3 | both
8OP09300/Label_NY_NP7110_2.55-3 | 8OP09300 | Label_NY_NP7110_2.55-3 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | 8OP09300/Label_NY_NP7110_2.55-3 | 3 | both
8OP09300/Label_NY_NP7110_2.55-3 | 8OP09300 | Label_NY_NP7110_2.55-3 | NY-Analytical | 770 | Label Inspection | Width | Width | 3 | in | FDA,4 | 2.9375 | 3.0625 | TM 143 / Visual | 8OP09300/Label_NY_NP7110_2.55-3 | 3 | both
8OP09300/Label_NY_NP7110_2.55-3 | 8OP09300 | Label_NY_NP7110_2.55-3 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Visual | 8OP09300/Label_NY_NP7110_2.55-3 | 3 | both
8OP09301/Label_NY_NP7110_2.55-5 | 8OP09301 | Label_NY_NP7110_2.55-5 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | 8OP09301/Label_NY_NP7110_2.55-5 | 3 | both
8OP09301/Label_NY_NP7110_2.55-5 | 8OP09301 | Label_NY_NP7110_2.55-5 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | 8OP09301/Label_NY_NP7110_2.55-5 | 3 | both
8OP09301/Label_NY_NP7110_2.55-5 | 8OP09301 | Label_NY_NP7110_2.55-5 | NY-Analytical | 770 | Label Inspection | Length | Length | 2.5313 | in | FDA,4 | 2.4875 | 2.6125 | TM 143 / Visual | 8OP09301/Label_NY_NP7110_2.55-5 | 3 | both
8OP09301/Label_NY_NP7110_2.55-5 | 8OP09301 | Label_NY_NP7110_2.55-5 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | 8OP09301/Label_NY_NP7110_2.55-5 | 3 | both
8OP09301/Label_NY_NP7110_2.55-5 | 8OP09301 | Label_NY_NP7110_2.55-5 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | 8OP09301/Label_NY_NP7110_2.55-5 | 3 | both
8OP09301/Label_NY_NP7110_2.55-5 | 8OP09301 | Label_NY_NP7110_2.55-5 | NY-Analytical | 770 | Label Inspection | Width | Width | 4.9687 | in | FDA,4 | 4.9375 | 5.0625 | TM 143 / Visual | 8OP09301/Label_NY_NP7110_2.55-5 | 3 | both
8OP09301/Label_NY_NP7110_2.55-5 | 8OP09301 | Label_NY_NP7110_2.55-5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Visual | 8OP09301/Label_NY_NP7110_2.55-5 | 3 | both
8OP227RSL/Label_NY_NP9997_2.625-3.375 | 8OP227RSL | Label_NY_NP9997_2.625-3.375 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | [5 empty] | Label Inspection | 8OP227RSL/Label_NY_NP9997_2.625-3.375 | 1 | both
8OP227RSL/Label_NY_NP9997_2.625-3.375 | 8OP227RSL | Label_NY_NP9997_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | 8OP227RSL/Label_NY_NP9997_2.625-3.375 | 1 | both
8OP227RSL/Label_NY_NP9997_2.625-3.375 | 8OP227RSL | Label_NY_NP9997_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | 8OP227RSL/Label_NY_NP9997_2.625-3.375 | 1 | both
8OP227RSL/Label_NY_NP9997_2.625-3.375 | 8OP227RSL | Label_NY_NP9997_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Length | Length | 2.63 | in | FDA,4 | 2.5625 | 2.6875 | TM 143 / Visual | 8OP227RSL/Label_NY_NP9997_2.625-3.375 | 1 | both
8OP227RSL/Label_NY_NP9997_2.625-3.375 | 8OP227RSL | Label_NY_NP9997_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | 8OP227RSL/Label_NY_NP9997_2.625-3.375 | 1 | both
8OP227RSL/Label_NY_NP9997_2.625-3.375 | 8OP227RSL | Label_NY_NP9997_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | 8OP227RSL/Label_NY_NP9997_2.625-3.375 | 1 | both
8OP227RSL/Label_NY_NP9997_2.625-3.375 | 8OP227RSL | Label_NY_NP9997_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Width | Width | 3.37 | in | FDA,4 | 3.3125 | 3.4375 | TM 143 / Visual | 8OP227RSL/Label_NY_NP9997_2.625-3.375 | 1 | both
8OP22900RF/Label_NY_NP98371_78-61 | 8OP22900RF | Label_NY_NP98371_78-61 | NY-Analytical | 770 | Label Inspection | Width | Width |  | mm | FDA,4 | 59.4125 | 62.5875 | TM 143 / Visual |  |  | left_only
8OP22900RF/Label_NY_NP98371_78-61 | 8OP22900RF | Label_NY_NP98371_78-61 | NY-Analytical | 770 | Label Inspection | Length | Length |  | mm | FDA,4 | 76.4125 | 79.5875 | TM 143 / Visual |  |  | left_only
8OP22900RF/Label_NY_NP98371_78-61 | 8OP22900RF | Label_NY_NP98371_78-61 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8OP22900RF/Label_NY_NP98371_78-61 | 8OP22900RF | Label_NY_NP98371_78-61 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8OP22900RF/Label_NY_NP98371_78-61 | 8OP22900RF | Label_NY_NP98371_78-61 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8OP22900RF/Label_NY_NP98371_78-61 | 8OP22900RF | Label_NY_NP98371_78-61 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8OP22900RF/Label_NY_NP98371_78-61 | 8OP22900RF | Label_NY_NP98371_78-61 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  |  | left_only
8OP260RSL/Label_NY_NP7700_3.375-2.3437 | 8OP260RSL | Label_NY_NP7700_3.375-2.3437 | NY-Analytical | 770 | Label Inspection | Width | Width |  | in | FDA,4 | 2.2812 | 2.4062 | TM 143 / Visual |  |  | left_only
8OP260RSL/Label_NY_NP7700_3.375-2.3437 | 8OP260RSL | Label_NY_NP7700_3.375-2.3437 | NY-Analytical | 770 | Label Inspection | Length | Length |  | in | FDA,4 | 3.3125 | 3.4375 | TM 143 / Visual |  |  | left_only
8OP260RSL/Label_NY_NP7700_3.375-2.3437 | 8OP260RSL | Label_NY_NP7700_3.375-2.3437 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8OP260RSL/Label_NY_NP7700_3.375-2.3437 | 8OP260RSL | Label_NY_NP7700_3.375-2.3437 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8OP260RSL/Label_NY_NP7700_3.375-2.3437 | 8OP260RSL | Label_NY_NP7700_3.375-2.3437 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8OP260RSL/Label_NY_NP7700_3.375-2.3437 | 8OP260RSL | Label_NY_NP7700_3.375-2.3437 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8OP260RSL/Label_NY_NP7700_3.375-2.3437 | 8OP260RSL | Label_NY_NP7700_3.375-2.3437 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  |  | left_only
8OP263CK/Label_NY_NP7030_8.5-6.5 | 8OP263CK | Label_NY_NP7030_8.5-6.5 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection | 8OP263CK/Label_NY_NP7030_8.5-6.5 | 1 | both
8OP263CK/Label_NY_NP7030_8.5-6.5 | 8OP263CK | Label_NY_NP7030_8.5-6.5 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | 8OP263CK/Label_NY_NP7030_8.5-6.5 | 1 | both
8OP263CK/Label_NY_NP7030_8.5-6.5 | 8OP263CK | Label_NY_NP7030_8.5-6.5 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | 8OP263CK/Label_NY_NP7030_8.5-6.5 | 1 | both
8OP263CK/Label_NY_NP7030_8.5-6.5 | 8OP263CK | Label_NY_NP7030_8.5-6.5 | NY-Analytical | 770 | Label Inspection | Length | Length | 8.5 | in | FDA,4 | 8.4375 | 8.5625 | TM 143 / Visual | 8OP263CK/Label_NY_NP7030_8.5-6.5 | 1 | both
8OP263CK/Label_NY_NP7030_8.5-6.5 | 8OP263CK | Label_NY_NP7030_8.5-6.5 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | 8OP263CK/Label_NY_NP7030_8.5-6.5 | 1 | both
8OP263CK/Label_NY_NP7030_8.5-6.5 | 8OP263CK | Label_NY_NP7030_8.5-6.5 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | 8OP263CK/Label_NY_NP7030_8.5-6.5 | 1 | both
8OP263CK/Label_NY_NP7030_8.5-6.5 | 8OP263CK | Label_NY_NP7030_8.5-6.5 | NY-Analytical | 770 | Label Inspection | Width | Width | 6.5 | in | FDA,4 | 6.4375 | 6.5625 | TM 143 / Visual | 8OP263CK/Label_NY_NP7030_8.5-6.5 | 1 | both
8OP308SIRF/Label_NY_NP9300_2-1.375 | 8OP308SIRF | Label_NY_NP9300_2-1.375 | NY-Analytical | 770 | Label Inspection | Width | Width |  | in | FDA,4 | 1.3125 | 1.4375 | TM 143 / Visual |  |  | left_only
8OP308SIRF/Label_NY_NP9300_2-1.375 | 8OP308SIRF | Label_NY_NP9300_2-1.375 | NY-Analytical | 770 | Label Inspection | Length | Length |  | in | FDA,4 | 1.9375 | 2.0625 | TM 143 / Visual |  |  | left_only
8OP308SIRF/Label_NY_NP9300_2-1.375 | 8OP308SIRF | Label_NY_NP9300_2-1.375 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8OP308SIRF/Label_NY_NP9300_2-1.375 | 8OP308SIRF | Label_NY_NP9300_2-1.375 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8OP308SIRF/Label_NY_NP9300_2-1.375 | 8OP308SIRF | Label_NY_NP9300_2-1.375 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8OP308SIRF/Label_NY_NP9300_2-1.375 | 8OP308SIRF | Label_NY_NP9300_2-1.375 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8OP308SIRF/Label_NY_NP9300_2-1.375 | 8OP308SIRF | Label_NY_NP9300_2-1.375 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  |  | left_only
8OP585SIRF/Label_NY_NP9300_1.375-2 | 8OP585SIRF | Label_NY_NP9300_1.375-2 | NY-Analytical | 770 | Label Inspection | Length | Length |  | in | FDA,4 | 1.3125 | 1.4375 | TM 143 / Visual |  |  | left_only
8OP585SIRF/Label_NY_NP9300_1.375-2 | 8OP585SIRF | Label_NY_NP9300_1.375-2 | NY-Analytical | 770 | Label Inspection | Width | Width |  | in | FDA,4 | 1.9375 | 2.0625 | TM 143 / Visual |  |  | left_only
8OP585SIRF/Label_NY_NP9300_1.375-2 | 8OP585SIRF | Label_NY_NP9300_1.375-2 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8OP585SIRF/Label_NY_NP9300_1.375-2 | 8OP585SIRF | Label_NY_NP9300_1.375-2 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8OP585SIRF/Label_NY_NP9300_1.375-2 | 8OP585SIRF | Label_NY_NP9300_1.375-2 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8OP585SIRF/Label_NY_NP9300_1.375-2 | 8OP585SIRF | Label_NY_NP9300_1.375-2 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8OP585SIRF/Label_NY_NP9300_1.375-2 | 8OP585SIRF | Label_NY_NP9300_1.375-2 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  |  | left_only
8OP63000RF/Label_NY_NP9837_85-73 | 8OP63000RF | Label_NY_NP9837_85-73 | NY-Analytical | 770 | Label Inspection | Width | Width |  | mm | FDA,4 | 71.4125 | 74.5875 | TM 143 / Visual |  |  | left_only
8OP63000RF/Label_NY_NP9837_85-73 | 8OP63000RF | Label_NY_NP9837_85-73 | NY-Analytical | 770 | Label Inspection | Length | Length |  | mm | FDA,4 | 83.4125 | 86.5875 | TM 143 / Visual |  |  | left_only
8OP63000RF/Label_NY_NP9837_85-73 | 8OP63000RF | Label_NY_NP9837_85-73 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8OP63000RF/Label_NY_NP9837_85-73 | 8OP63000RF | Label_NY_NP9837_85-73 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8OP63000RF/Label_NY_NP9837_85-73 | 8OP63000RF | Label_NY_NP9837_85-73 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8OP63000RF/Label_NY_NP9837_85-73 | 8OP63000RF | Label_NY_NP9837_85-73 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual |  |  | left_only
8OP63000RF/Label_NY_NP9837_85-73 | 8OP63000RF | Label_NY_NP9837_85-73 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  |  | left_only
8OX351CK/Label_NY_50#W_5_4 | 8OX351CK | Label_NY_50#W_5_4 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection | 8OX351CK/Label_NY_50#W_5_4 | 1 | both
8OX351CK/Label_NY_50#W_5_4 | 8OX351CK | Label_NY_50#W_5_4 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | 8OX351CK/Label_NY_50#W_5_4 | 1 | both
8OX351CK/Label_NY_50#W_5_4 | 8OX351CK | Label_NY_50#W_5_4 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | 8OX351CK/Label_NY_50#W_5_4 | 1 | both
8OX351CK/Label_NY_50#W_5_4 | 8OX351CK | Label_NY_50#W_5_4 | NY-Analytical | 770 | Label Inspection | Length | Length | 5.01 | in | FDA,4 | 4.9375 | 5.0625 | TM 143 / Visual | 8OX351CK/Label_NY_50#W_5_4 | 1 | both
8OX351CK/Label_NY_50#W_5_4 | 8OX351CK | Label_NY_50#W_5_4 | NY-Analytical | 770 | Label Inspection | NP Structure | NP Structure | 50# White Paper Stock (Done) | in |  |  |  | TM 143 / Visual | 8OX351CK/Label_NY_50#W_5_4 | 1 | both
8OX351CK/Label_NY_50#W_5_4 | 8OX351CK | Label_NY_50#W_5_4 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | 8OX351CK/Label_NY_50#W_5_4 | 1 | both
8OX351CK/Label_NY_50#W_5_4 | 8OX351CK | Label_NY_50#W_5_4 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | 8OX351CK/Label_NY_50#W_5_4 | 1 | both
8OX351CK/Label_NY_50#W_5_4 | 8OX351CK | Label_NY_50#W_5_4 | NY-Analytical | 770 | Label Inspection | Width | Width | 3.99 | in | FDA,4 | 3.9375 | 4.0625 | TM 143 / Visual | 8OX351CK/Label_NY_50#W_5_4 | 1 | both
9OX465GL/NY_Glue_01 | 9OX465GL | NY_Glue_01 | NY-Analytical | 395 | Color | Color | Color | Very Pale Amber (Done) |  |  |  |  | Visual |  |  | left_only
9OX465GL/NY_Glue_01 | 9OX465GL | NY_Glue_01 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Visual |  |  | left_only