## Main
plan.m | Mat_MatCode | SP_TestPlan | Test_Lab | Res_TestCode | TestName | Analyte | Analyte.C | Numeric_Result | Res_Units | Rounding_Rule | Low_Release_Limit | High_Release_Limit | METHOD
2AA284SLI/Lid_01B | 2AA284SLI | Lid_01B | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection
2AA284SLI/Lid_01B | 2AA284SLI | Lid_01B | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection
2AA284SLI/Lid_01B | 2AA284SLI | Lid_01B | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual
2AA284SLI/Lid_01B | 2AA284SLI | Lid_01B | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual
2AA284SLI/Lid_01B | 2AA284SLI | Lid_01B | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.6875 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual
2AA284SLI/Lid_01B | 2AA284SLI | Lid_01B | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.875 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual
2AA284SLI/Lid_01B | 2AA284SLI | Lid_01B | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual
2AA347/Web_NP8550_10.25 | 2AA347 | Web_NP8550_10.25 | NY-Analytical | 885 | Web Inspection | AQL Sample Size | AQL Sample Size | 315 |  | FDA,0 | 0 |  | TM101/TM107/TM132/Measure
2AA347/Web_NP8550_10.25 | 2AA347 | Web_NP8550_10.25 | NY-Analytical | 885 | Web Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Sample(s) all meet specifications (Done) |  |  |  |  | TM101/TM107/TM132/Measure
2AA347/Web_NP8550_10.25 | 2AA347 | Web_NP8550_10.25 | NY-Analytical | 885 | Web Inspection | Basis_Weight | Basis Weight | 41.3 | #/Ream | FDA,1 | 37.8 | 48.2 | TM101/TM107/TM132/Measure
2AA347/Web_NP8550_10.25 | 2AA347 | Web_NP8550_10.25 | NY-Analytical | 885 | Web Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | TM101/TM107/TM132/Measure
2AA347/Web_NP8550_10.25 | 2AA347 | Web_NP8550_10.25 | NY-Analytical | 885 | Web Inspection | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 5762.95 | g/in | FDA,0 | 3500 |  | TM101/TM107/TM132/Measure
2AA347/Web_NP8550_10.25 | 2AA347 | Web_NP8550_10.25 | NY-Analytical | 885 | Web Inspection | Web Width (Inches) | Web Width | 10.25 | in | FDA,4 | 10.1875 | 10.3125 | TM101/TM107/TM132/Measure
2AA592/Web_NP8504_11.0 | 2AA592 | Web_NP8504_11.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Dyne Level
2AA592/Web_NP8504_11.0 | 2AA592 | Web_NP8504_11.0 | NY-Analytical | 836 | Dyne Level | Dyne Level (outside film) | Dyne Level (outside film) | 36 |  | FDA,0 | 36 | 38 | Dyne Test Kit
2AA592/Web_NP8504_11.0 | 2AA592 | Web_NP8504_11.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 10.968 | in | FDA,4 | 10.9375 | 11.0625 | Measurement
2AA592/Web_NP8504_11.0 | 2AA592 | Web_NP8504_11.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 57.54 | #/Ream | FDA,1 | 53.3 | 65.1 | TM 101
2AA592/Web_NP8504_11.0 | 2AA592 | Web_NP8504_11.0 | NY-Analytical | 762 | Seal Strength | Seal Strength(g/in@350F/30psi/0.5s dwell Top Jaw) | Seal Strength (@350°F/30psi/0.5s dwell Top Jaw) | 5339.63 | g/in | FDA,0 | 2700 |  | TM 107
2AA592/Web_NP8504_11.0 | 2AA592 | Web_NP8504_11.0 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 3.65 | mil | FDA,1 | 3.2 | 3.9 | TM 132
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Lid_NY_NP7600_22R1 | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Lid_NY_NP7600_22R1 | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Solo Lid Inspection
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Lid_NY_NP7600_22R1 | NY-Analytical | 818 | Solo Lid Inspection | Color | Color | Meets the specification (Done) |  |  |  |  | Solo Lid Inspection
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Lid_NY_NP7600_22R1 | NY-Analytical | 818 | Solo Lid Inspection | Functionality | Functionality | Meet specification (Done) |  |  |  |  | Solo Lid Inspection
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Lid_NY_NP7600_22R1 | NY-Analytical | 818 | Solo Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.028 | in | FDA,3 | 0.025 | 0.035 | Solo Lid Inspection
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Lid_NY_NP7600_22R1 | NY-Analytical | 818 | Solo Lid Inspection | Open Lid Overall Width | Open Lid Overall Width | 4.36 | in | FDA,3 | 4.351 | 4.401 | Solo Lid Inspection
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Lid_NY_NP7600_22R1 | NY-Analytical | 818 | Solo Lid Inspection | Overall closed lid height | Overall closed lid height | 0.37 | in | FDA,3 | 0.355 | 0.385 | Solo Lid Inspection
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Lid_NY_NP7600_22R1 | NY-Analytical | 818 | Solo Lid Inspection | Weight | Weight | 11.241 | g | FDA,1 | 10.5 | 12.5 | Measurement
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Lid_NY_NP7600_22R1 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Lid_NY_NP7600_22R1 | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Lid_NY_NP7600_22R1 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.7025 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Lid_NY_NP7600_22R1 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.8705 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Lid_NY_NP7600_22R1 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual
2AA632/Web_NP8504_11.75 | 2AA632 | Web_NP8504_11.75 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Dyne Level
2AA632/Web_NP8504_11.75 | 2AA632 | Web_NP8504_11.75 | NY-Analytical | 836 | Dyne Level | Dyne Level (outside film) | Dyne Level (outside film) | 38 |  | FDA,0 | 36 | 38 | Dyne Test Kit
2AA632/Web_NP8504_11.75 | 2AA632 | Web_NP8504_11.75 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 11.75 | in | FDA,4 | 11.6875 | 11.8125 | Measurement
2AA632/Web_NP8504_11.75 | 2AA632 | Web_NP8504_11.75 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 59.05 | #/Ream | FDA,1 | 53.3 | 65.1 | TM 101
2AA632/Web_NP8504_11.75 | 2AA632 | Web_NP8504_11.75 | NY-Analytical | 762 | Seal Strength | Seal Strength(g/in@350F/30psi/0.5s dwell Top Jaw) | Seal Strength (@350°F/30psi/0.5s dwell Top Jaw) | 5664 | g/in | FDA,0 | 2700 |  | TM 107
2AA632/Web_NP8504_11.75 | 2AA632 | Web_NP8504_11.75 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 3.7 | mil | FDA,1 | 3.2 | 3.9 | TM 132
2AA714SLI/Lid_01V | 2AA714SLI | Lid_01V | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection
2AA714SLI/Lid_01V | 2AA714SLI | Lid_01V | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection
2AA714SLI/Lid_01V | 2AA714SLI | Lid_01V | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual
2AA714SLI/Lid_01V | 2AA714SLI | Lid_01V | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual
2AA714SLI/Lid_01V | 2AA714SLI | Lid_01V | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.687 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual
2AA714SLI/Lid_01V | 2AA714SLI | Lid_01V | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.8125 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual
2AA714SLI/Lid_01V | 2AA714SLI | Lid_01V | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual
2AA922SLI/Lid_NY_Skip lot testing | 2AA922SLI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection
2AA922SLI/Lid_01V | 2AA922SLI | Lid_01V | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis
2AA922SLI/Lid_NY_Skip lot testing | 2AA922SLI | Lid_NY_Skip lot testing | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Compliance
2AA922SLI/Lid_NY_Skip lot testing | 2AA922SLI | Lid_NY_Skip lot testing | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection
2AA922SLI/Lid_NY_Skip lot testing | 2AA922SLI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual
2AA922SLI/Lid_NY_Skip lot testing | 2AA922SLI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual
2AA922SLI/Lid_NY_Skip lot testing | 2AA922SLI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.6875 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual
2AA922SLI/Lid_NY_Skip lot testing | 2AA922SLI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.8438 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual
2AA922SLI/Lid_NY_Skip lot testing | 2AA922SLI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual
2AD173/Web_NP5115_4.5 | 2AD173 | Web_NP5115_4.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2AD173/Web_NP5115_4.5 | 2AD173 | Web_NP5115_4.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 4.5 | in | FDA,4 | 4.4375 | 4.5625 | Measurement
2AD173/Web_NP5115_4.5 | 2AD173 | Web_NP5115_4.5 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Printed/Varnished) | Basis Weight | 59.5 | #/Ream | FDA,1 | 55.6 | 68 | TM 101
2AD173/Web_NP5115_4.5 | 2AD173 | Web_NP5115_4.5 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@325DegF/40PSI/0.5Sec_Dwell) | Seal Strength (@325°F/40PSI/0.5Sec Dwell) | 2057.76 | g/in | FDA,0 | 1500 |  | TM 107
2AJ123/Web_NY_Skip lot testing | 2AJ123 | Web_NY_Skip lot testing | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2AJ123/Web_NY_Skip lot testing | 2AJ123 | Web_NY_Skip lot testing | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 11 | in | FDA,4 | 10.9375 | 11.0625 | Measurement
2AJ123/Web_NY_Skip lot testing | 2AJ123 | Web_NY_Skip lot testing | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 42.33 | #/Ream | FDA,1 | 37.8 | 48.2 | TM 101
2AJ123/Web_NY_Skip lot testing | 2AJ123 | Web_NY_Skip lot testing | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 7486 | g/in | FDA,0 | 3500 |  | TM 107
2AJ143SLI/Lid_NY_NP7600_01C | 2AJ143SLI | Lid_NY_NP7600_01C | IN-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection
2AJ143SLI/Lid_NY_NP7600_01C | 2AJ143SLI | Lid_NY_NP7600_01C | IN-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual
2AJ143SLI/Lid_NY_NP7600_01C | 2AJ143SLI | Lid_NY_NP7600_01C | IN-Analytical | 765 | Canister & Lid Inspection | AQL Sample Size | AQL Sample Size | 200 |  |  | 0 |  | TM 143 / Visual
2AJ143SLI/Lid_NY_NP7600_01C | 2AJ143SLI | Lid_NY_NP7600_01C | IN-Analytical | 765 | Canister & Lid Inspection | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | TM 143 / Visual
2AJ143SLI/Lid_NY_NP7600_01C | 2AJ143SLI | Lid_NY_NP7600_01C | IN-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual
2AJ143SLI/Lid_NY_NP7600_01C | 2AJ143SLI | Lid_NY_NP7600_01C | IN-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.7088 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual
2AJ143SLI/Lid_NY_NP7600_01C | 2AJ143SLI | Lid_NY_NP7600_01C | IN-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.8438 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual
2AJ143SLI/Lid_NY_NP7600_01C | 2AJ143SLI | Lid_NY_NP7600_01C | IN-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual
2AJ170/Web_NP8530_11.0 | 2AJ170 | Web_NP8530_11.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2AJ170/Web_NP8530_11.0 | 2AJ170 | Web_NP8530_11.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 11 | in | FDA,4 | 10.9375 | 11.0625 | Measurement
2AJ170/Web_NP8530_11.0 | 2AJ170 | Web_NP8530_11.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 34.39 | #/Ream | FDA,2 | 29.4 | 37.42 | TM 101
2AJ170/Web_NP8530_11.0 | 2AJ170 | Web_NP8530_11.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 5442 | g/in | FDA,0 | 3500 |  | TM 107
2AJ170/Web_NP8530_11.0 | 2AJ170 | Web_NP8530_11.0 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 1.9 | mil | FDA,1 | 1.8 | 2.2 | TM 132
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Lid_NY_NP7600_22FA | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Lid_NY_NP7600_22FA | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Solo Lid Inspection
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Lid_NY_NP7600_22FA | NY-Analytical | 818 | Solo Lid Inspection | Color | Color | Meets the specification (Done) |  |  |  |  | Solo Lid Inspection
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Lid_NY_NP7600_22FA | NY-Analytical | 818 | Solo Lid Inspection | Functionality | Functionality | Meet specification (Done) |  |  |  |  | Solo Lid Inspection
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Lid_NY_NP7600_22FA | NY-Analytical | 818 | Solo Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.028 | in | FDA,3 | 0.025 | 0.035 | Solo Lid Inspection
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Lid_NY_NP7600_22FA | NY-Analytical | 818 | Solo Lid Inspection | Open Lid Overall Width | Open Lid Overall Width | 4.36 | in | FDA,3 | 4.351 | 4.401 | Solo Lid Inspection
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Lid_NY_NP7600_22FA | NY-Analytical | 818 | Solo Lid Inspection | Overall closed lid height | Overall closed lid height | 0.37 | in | FDA,3 | 0.355 | 0.385 | Solo Lid Inspection
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Lid_NY_NP7600_22FA | NY-Analytical | 818 | Solo Lid Inspection | Weight | Weight | 11.25 | g | FDA,1 | 10.5 | 12.5 | Measurement
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Lid_NY_NP7600_22FA | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Lid_NY_NP7600_22FA | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Lid_NY_NP7600_22FA | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.67 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Lid_NY_NP7600_22FA | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.84 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Lid_NY_NP7600_22FA | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Lid_NY_NP7600_22V | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Lid_NY_NP7600_22V | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Solo Lid Inspection
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Lid_NY_NP7600_22V | NY-Analytical | 818 | Solo Lid Inspection | Appearance | Appearance | Meet specification (Done) |  |  |  |  | Solo Lid Inspection
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Lid_NY_NP7600_22V | NY-Analytical | 818 | Solo Lid Inspection | Color | Color | Meets the specification (Done) |  |  |  |  | Solo Lid Inspection
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Lid_NY_NP7600_22V | NY-Analytical | 818 | Solo Lid Inspection | Functionality | Functionality | Meet specification (Done) |  |  |  |  | Solo Lid Inspection
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Lid_NY_NP7600_22V | NY-Analytical | 818 | Solo Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.0295 | in | FDA,3 | 0.025 | 0.035 | Solo Lid Inspection
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Lid_NY_NP7600_22V | NY-Analytical | 818 | Solo Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | Solo Lid Inspection
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Lid_NY_NP7600_22V | NY-Analytical | 818 | Solo Lid Inspection | Open Lid Overall Width | Open Lid Overall Width | 4.374 | in | FDA,3 | 4.351 | 4.401 | Solo Lid Inspection
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Lid_NY_NP7600_22V | NY-Analytical | 818 | Solo Lid Inspection | Overall closed lid height | Overall closed lid height | 0.373 | in | FDA,3 | 0.355 | 0.385 | Solo Lid Inspection
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Lid_NY_NP7600_22V | NY-Analytical | 818 | Solo Lid Inspection | Weight | Weight | 11.19 | g | FDA,1 | 10.5 | 12.5 | Measurement
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Lid_NY_NP7600_22V | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Lid_NY_NP7600_22V | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.7188 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Lid_NY_NP7600_22V | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.875 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Lid_NY_NP7600_22X | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Lid_NY_NP7600_22X | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Solo Lid Inspection
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Lid_NY_NP7600_22X | NY-Analytical | 818 | Solo Lid Inspection | Color | Color | Meets the specification (Done) |  |  |  |  | Solo Lid Inspection
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Lid_NY_NP7600_22X | NY-Analytical | 818 | Solo Lid Inspection | Functionality | Functionality | Meet specification (Done) |  |  |  |  | Solo Lid Inspection
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Lid_NY_NP7600_22X | NY-Analytical | 818 | Solo Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.0297 | in | FDA,3 | 0.025 | 0.035 | Solo Lid Inspection
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Lid_NY_NP7600_22X | NY-Analytical | 818 | Solo Lid Inspection | Open Lid Overall Width | Open Lid Overall Width | 4.371 | in | FDA,3 | 4.351 | 4.401 | Solo Lid Inspection
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Lid_NY_NP7600_22X | NY-Analytical | 818 | Solo Lid Inspection | Overall closed lid height | Overall closed lid height | 0.3705 | in | FDA,3 | 0.355 | 0.385 | Solo Lid Inspection
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Lid_NY_NP7600_22X | NY-Analytical | 818 | Solo Lid Inspection | Weight | Weight | 11.385 | g | FDA,1 | 10.5 | 12.5 | Measurement
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Lid_NY_NP7600_22X | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Lid_NY_NP7600_22X | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Lid_NY_NP7600_22X | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.7088 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Lid_NY_NP7600_22X | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.875 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Lid_NY_NP7600_22X | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual
2AP027/Web_NP8512_310 | 2AP027 | Web_NP8512_310 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2AP027/Web_NP8512_310 | 2AP027 | Web_NP8512_310 | NY-Analytical | 761 | Web Measurement | Web Width (mm) | Web Width | 310 | mm | FDA,0 | 308 | 312 | Measurement
2AP027/Web_NP8512_310 | 2AP027 | Web_NP8512_310 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/inch@130DegC/14psi/1.0 sec dwell) | Seal Strength (@130°C/14psi/1.0 sec dwell) | 8798 | g/in | FDA,0 | 2000 |  | TM 107
2AP027/Web_NP8512_310 | 2AP027 | Web_NP8512_310 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 3.5 | mil | FDA,2 | 3.06 | 3.74 | TM 132
2AP027/Web_NP8512_310 | 2AP027 | Web_NP8512_310 | NY-Analytical | 748 | Visual Inspection | Color/Copy | Color/Copy | Meets specifications (Done) |  |  |  |  | TM 208
2AP075/Web_NP8542_12.205 | 2AP075 | Web_NP8542_12.205 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2AP075/Web_NP8542_12.205 | 2AP075 | Web_NP8542_12.205 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 12.25 | in | FDA,4 | 12.1425 | 12.2675 | Measurement
2AP075/Web_NP8542_12.205 | 2AP075 | Web_NP8542_12.205 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 60.85 | #/Ream | FDA,2 | 52.53 | 64.21 | TM 101
2AP075/Web_NP8542_12.205 | 2AP075 | Web_NP8542_12.205 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/mm@130DegC/1bar/0.5sec dwell) | Seal Strength (@130°C/1bar/0.5sec dwell) | 7718 | g/mm | FDA,0 | 25 |  | TM 107
2AP087/Web_NP8550_12.205 | 2AP087 | Web_NP8550_12.205 | NY-Analytical | 885 | Web Inspection | AQL Sample Size | AQL Sample Size | 200 |  | FDA,0 | 0 |  | TM101/TM107/TM132/Measure
2AP087/Web_NP8550_12.205 | 2AP087 | Web_NP8550_12.205 | NY-Analytical | 885 | Web Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Sample(s) all meet specifications (Done) |  |  |  |  | TM101/TM107/TM132/Measure
2AP087/Web_NP8550_12.205 | 2AP087 | Web_NP8550_12.205 | NY-Analytical | 885 | Web Inspection | Basis_Weight | Basis Weight | 43.3 | #/Ream | FDA,1 | 37.8 | 48.2 | TM101/TM107/TM132/Measure
2AP087/Web_NP8550_12.205 | 2AP087 | Web_NP8550_12.205 | NY-Analytical | 885 | Web Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | TM101/TM107/TM132/Measure
2AP087/Web_NP8550_12.205 | 2AP087 | Web_NP8550_12.205 | NY-Analytical | 885 | Web Inspection | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 5039.27 | g/in | FDA,0 | 3500 |  | TM101/TM107/TM132/Measure
2AP087/Web_NP8550_12.205 | 2AP087 | Web_NP8550_12.205 | NY-Analytical | 885 | Web Inspection | UPC Verification | UPC Verification | UPC Scan is acceptable (Scan into comment section below - Done) |  |  |  |  | TM101/TM107/TM132/Measure
2AP087/Web_NP8550_12.205 | 2AP087 | Web_NP8550_12.205 | NY-Analytical | 885 | Web Inspection | Web Width (Inches) | Web Width | 12.1425 | in | FDA,4 | 12.1425 | 12.2675 | TM101/TM107/TM132/Measure
2AP087/Web_NP8550_12.205 | 2AP087 | Web_NP8550_12.205 | NY-Analytical | 885 | Web Inspection | Wind Direction | Wind Direction | Meets specification (Indicate numeric value - Done) |  |  |  |  | TM101/TM107/TM132/Measure
2AP113/Web_NP8542_11.0 | 2AP113 | Web_NP8542_11.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2AP113/Web_NP8542_11.0 | 2AP113 | Web_NP8542_11.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 11 | in | FDA,4 | 10.9375 | 11.0625 | Measurement
2AP113/Web_NP8542_11.0 | 2AP113 | Web_NP8542_11.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 61.11 | #/Ream | FDA,2 | 52.53 | 64.21 | TM 101
2AP113/Web_NP8542_11.0 | 2AP113 | Web_NP8542_11.0 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/mm@130DegC/1bar/0.5sec dwell) | Seal Strength (@130°C/1bar/0.5sec dwell) | 180.6 | g/mm | FDA,0 | 25 |  | TM 107
2AP211/Web_NP8512_12.205 | 2AP211 | Web_NP8512_12.205 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2AP211/Web_NP8512_12.205 | 2AP211 | Web_NP8512_12.205 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 12.25 | in | FDA,4 | 12.1425 | 12.2675 | Measurement
2AP211/Web_NP8512_12.205 | 2AP211 | Web_NP8512_12.205 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/inch@130DegC/14psi/1.0 sec dwell) | Seal Strength (@130°C/14psi/1.0 sec dwell) | 7698.56 | g/in | FDA,0 | 2000 |  | TM 107
2AP211/Web_NP8512_12.205 | 2AP211 | Web_NP8512_12.205 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 3.5 | mil | FDA,2 | 3.06 | 3.74 | TM 132
2AP211/Web_NP8512_12.205 | 2AP211 | Web_NP8512_12.205 | NY-Analytical | 748 | Visual Inspection | Color/Copy | Color/Copy | Meets specifications (Done) |  |  |  |  | TM 208
2AP229/Web_NP8512_9.5 | 2AP229 | Web_NP8512_9.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2AP229/Web_NP8512_9.5 | 2AP229 | Web_NP8512_9.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 9.5 | in | FDA,4 | 9.4375 | 9.5625 | Measurement
2AP229/Web_NP8512_9.5 | 2AP229 | Web_NP8512_9.5 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/inch@130DegC/14psi/1.0 sec dwell) | Seal Strength (@130°C/14psi/1.0 sec dwell) | 7419.71 | g/in | FDA,0 | 2000 |  | TM 107
2AP229/Web_NP8512_9.5 | 2AP229 | Web_NP8512_9.5 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 3.6 | mil | FDA,2 | 3.06 | 3.74 | TM 132
2AP229/Web_NP8512_9.5 | 2AP229 | Web_NP8512_9.5 | NY-Analytical | 748 | Visual Inspection | Color/Copy | Color/Copy | Meets specifications (Done) |  |  |  |  | TM 208
2AP260/Web_NP8535_12.75 | 2AP260 | Web_NP8535_12.75 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2AP260/Web_NP8535_12.75 | 2AP260 | Web_NP8535_12.75 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 12.75 | in | FDA,4 | 12.6875 | 12.8125 | Measurement
2AP260/Web_NP8535_12.75 | 2AP260 | Web_NP8535_12.75 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 41.14 | #/Ream | FDA,2 | 35.82 | 45.6 | TM 101
2AP260/Web_NP8535_12.75 | 2AP260 | Web_NP8535_12.75 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 6659.74 | g/in | FDA,0 | 3500 |  | TM 107
2AP260/Web_NP8535_12.75 | 2AP260 | Web_NP8535_12.75 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 2.4 | mil | FDA,1 | 2.2 | 2.8 | TM 132
2AP806/Web_NP8430_310 | 2AP806 | Web_NP8430_310 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2AP806/Web_NP8430_310 | 2AP806 | Web_NP8430_310 | NY-Analytical | 761 | Web Measurement | Web Width (mm) | Web Width | 311 | mm | FDA,0 | 308 | 312 | Measurement
2AP806/Web_NP8430_310 | 2AP806 | Web_NP8430_310 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 51.9 | #/Ream | FDA,2 | 44.88 | 57.12 | TM 101
2AP806/Web_NP8430_310 | 2AP806 | Web_NP8430_310 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 7699.25 | g/in | FDA,0 | 3500 |  | TM 107
2AP806/Web_NP8430_310 | 2AP806 | Web_NP8430_310 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 3.05 | mil | FDA,2 | 2.64 | 3.36 | TM 132
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Lid_04B_560_1 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Cap Thickness | Cap Thickness | 0.154 | in | FDA,3 | 0.147 | 0.157 | TM 143 / Visual
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Blue (Done) |  |  |  |  | TM 143 / Visual
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) |  |  |  |  | TM 143 / Visual
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter (as per specification) | Lid Inside Diameter | 3.7695 | in | FDA,3 | 3.76 | 3.77 | TM 143 / Visual
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.05 | in | FDA,3 | 0.04 | 0.05 | TM 143 / Visual
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.5845 | in | FDA,3 | 1.58 | 1.6 | TM 143 / Visual
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Outer Diameter (per specification) | Pop-Up Hole Outer Diameter | 0.565 | in | FDA,3 | 0.555 | 0.565 | TM 143 / Visual
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Thicknes | Pop-Up Hole Thicknes | 0.02 | in | FDA,3 | 0.015 | 0.025 | TM 143 / Visual
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 22.773 | g | FDA,1 | 20 | 24 | TM 143 / Visual
2AU191/Web_NP5115_7.5 | 2AU191 | Web_NP5115_7.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2AU191/Web_NP5115_7.5 | 2AU191 | Web_NP5115_7.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 7.4375 | in | FDA,4 | 7.4375 | 7.5625 | Measurement
2AU191/Web_NP5115_7.5 | 2AU191 | Web_NP5115_7.5 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Printed/Varnished) | Basis Weight | 60.5 | #/Ream | FDA,1 | 55.6 | 68 | TM 101
2AU191/Web_NP5115_7.5 | 2AU191 | Web_NP5115_7.5 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@325DegF/40PSI/0.5Sec_Dwell) | Seal Strength (@325°F/40PSI/0.5Sec Dwell) | 1823.34 | g/in | FDA,0 | 1500 |  | TM 107
2AU252/Web_NP3459_8.0 | 2AU252 | Web_NP3459_8.0 |  | 1044 | Basis Weight (lbs/ream) | Basis Weight | Basis Weight | 63.4 | Lb/Ream | FDA,2 | 57.56 | 70.4 | ASTM D 646
2AU252/Web_NP3459_8.0 | 2AU252 | Web_NP3459_8.0 |  | 757 | Caliper | caliper (COA) | Caliper | 3.8 | mil | FDA,3 | 3.515 | 3.515 | Caliper
2AU252/Web_3459_8 | 2AU252 | Web_3459_8 | NY-Analytical | 1044 | Basis Weight (lbs/ream) | Basis Weight | Basis Weight |  | Lb/Ream | FDA,2 | 57.56 | 70.4 | TM 101
2AU252/Web_NP3459_8.0 | 2AU252 | Web_NP3459_8.0 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/inch@290DegF/40psi/0.5 sec dwell) | Seal Strength (@290°F/40psi/0.5 sec dwell) |  | g/in | FDA,0 | 7500 |  | TM 107
2AU252/Web_3459_8 | 2AU252 | Web_3459_8 | NY-Analytical | 761 | Web Measurement | Core ID (Inches) | Core ID |  | in | FDA,3 | 2.9375 | 3.0625 | TM 143
2AU252/Web_3459_8 | 2AU252 | Web_3459_8 | NY-Analytical | 761 | Web Measurement | Roll OD | Roll OD | 15 | in | FDA,3 | 14.9375 | 15.0625 | TM 143
2AU252/Web_NP3459_8.0 | 2AU252 | Web_NP3459_8.0 |  | 761 | Web Measurement | Web Width (Inches) | Web Width | 8 | in | FDA,4 | 7.9375 | 8.0625 | TM 143
2AU252/Web_3459_8 | 2AU252 | Web_3459_8 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 8 | in | FDA,4 | 7.9375 | 8.0625 | TM 143
2AU252/Web_NP3459_8.0 | 2AU252 | Web_NP3459_8.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Visual
2AU252/Web_3459_8 | 2AU252 | Web_3459_8 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Visual
2AU252/Web_NP3459_8.0 | 2AU252 | Web_NP3459_8.0 |  | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | Visual Observation
2AU252/Web_NP3459_8.0 | 2AU252 | Web_NP3459_8.0 |  | 748 | Visual Inspection | Color/Copy | Color/Copy | Meets specifications (Done) |  |  |  |  | Visual Observation
2AU555/Web_NP3460_7.5 | 2AU555 | Web_NP3460_7.5 | NY-Analytical | 1044 | Basis Weight (lbs/ream) | Basis Weight | Basis Weight | 72.89 | Lb/Ream | FDA,1 | 66.6 | 81.4 | ASTM D 646
2AU555/Web_NP3460_7.5 | 2AU555 | Web_NP3460_7.5 | NY-Analytical | 757 | Caliper | caliper (COA) | Caliper | 4.22 | mil | FDA,4 | 3.995 | 3.995 | Caliper
2AU555/Web_NP3460_7.5 | 2AU555 | Web_NP3460_7.5 | NY-Analytical | 836 | Dyne Level | Dyne Level (outside film) | Dyne Level (outside film) | 36 |  | FDA,0 |  |  | Dyne Test Kit
2AU555/Web_NP3460_7.5 | 2AU555 | Web_NP3460_7.5 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/inch@290DegF/40psi/0.5 sec dwell) | Seal Strength (@290°F/40psi/0.5 sec dwell) | 8720.24 | g/in | FDA,0 | 5850 |  | TM 107
2AU555/Web_NP3460_7.5 | 2AU555 | Web_NP3460_7.5 | NY-Analytical | 761 | Web Measurement | Core ID (Inches) | Core ID | 3 | in | FDA,4 | 2.9375 | 3.0625 | TM 143
2AU555/Web_NP3460_7.5 | 2AU555 | Web_NP3460_7.5 | NY-Analytical | 761 | Web Measurement | Roll OD | Roll OD | 15 | in | FDA,4 | 14.9375 | 15.0625 | TM 143
2AU555/Web_NP3460_7.5 | 2AU555 | Web_NP3460_7.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 7.5 | in | FDA,4 | 7.4375 | 7.5625 | TM 143
2AU555/Web_NP3460_7.5 | 2AU555 | Web_NP3460_7.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Visual
2AU555/Web_NP3460_7.5 | 2AU555 | Web_NP3460_7.5 | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | Visual Observation
2AU555/Web_NP3460_7.5 | 2AU555 | Web_NP3460_7.5 | NY-Analytical | 748 | Visual Inspection | Color/Copy | Color/Copy | Meets specifications (Done) |  |  |  |  | Visual Observation
2BA585SLI/Lid_19AB | 2BA585SLI | Lid_19AB | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection
2BA585SLI/Lid_19AB | 2BA585SLI | Lid_19AB | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection
2BA585SLI/Lid_19AB | 2BA585SLI | Lid_19AB | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual
2BA585SLI/Lid_19AB | 2BA585SLI | Lid_19AB | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual
2BA585SLI/Lid_19AB | 2BA585SLI | Lid_19AB | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 73 | mm | FDA,0 | 71 | 75 | TM 143 / Visual
2BA585SLI/Lid_19AB | 2BA585SLI | Lid_19AB | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 54 | mm | FDA,0 | 52 | 56 | TM 143 / Visual
2BA585SLI/Lid_19AB | 2BA585SLI | Lid_19AB | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual
2BA972/Web_NP8540_13.5 | 2BA972 | Web_NP8540_13.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Web Measurement
2BA972/Web_NP8540_13.5 | 2BA972 | Web_NP8540_13.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 13.5 | in | FDA,4 | 13.4375 | 13.5625 | Measurement
2BA972/Web_NP8540_13.5 | 2BA972 | Web_NP8540_13.5 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 53.97 | #/Ream | FDA,1 | 45.4 | 57.8 | TM 101
2BA972/Web_NP8540_13.5 | 2BA972 | Web_NP8540_13.5 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) |  | g/in | FDA,0 | 3500 |  | TM 107
2BA972/Web_NP8540_13.5 | 2BA972 | Web_NP8540_13.5 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 3.1 | mil | FDA,2 | 2.86 | 3.64 | TM 132
2BC801/Web_NP3870_6.0 | 2BC801 | Web_NP3870_6.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2BC801/Web_NP3870_6.0 | 2BC801 | Web_NP3870_6.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 6 | in | FDA,4 | 5.9375 | 6.0625 | Measurement
2BC801/Web_NP3870_6.0 | 2BC801 | Web_NP3870_6.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 92.46 | #/Ream | FDA,1 | 81 | 99 | TM 101
2BC801/Web_NP3870_6.0 | 2BC801 | Web_NP3870_6.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@350DegF/40psi/1.0 sec dwell) | Seal Strength (@350°F/40psi/1.0 sec dwell) | 10097 | g/in | FDA,0 | 5000 |  | TM 107
2BH592/Web_NP3895_6.0 | 2BH592 | Web_NP3895_6.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2BH592/Web_NP3895_6.0 | 2BH592 | Web_NP3895_6.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 6 | in | FDA,4 | 5.9375 | 6.0625 | Measurement
2BH592/Web_NP3895_6.0 | 2BH592 | Web_NP3895_6.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/3000_sq_ft) | Basis Weight | 69.71 | #/3000sqft | FDA,1 | 65.3 | 79.8 | TM 101
2BH592/Web_NP3895_6.0 | 2BH592 | Web_NP3895_6.0 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/inch@250DegF/30psi/0.5 sec dwell) | Seal Strength (@250°F/30psi/0.5 sec dwell) | 5740.1 | g/in | FDA,0 | 2270 |  | TM 107
2BH592/Web_NP3895_6.0 | 2BH592 | Web_NP3895_6.0 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 4.5 | mil | FDA,1 | 3.9 | 4.7 | TM 132
2BJ123/Web_NP8550_11 | 2BJ123 | Web_NP8550_11 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2BJ123/Web_NP8550_11 | 2BJ123 | Web_NP8550_11 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 11 | in | FDA,4 | 10.9375 | 11.0625 | Measurement
2BJ123/Web_NP8550_11 | 2BJ123 | Web_NP8550_11 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 43.12 | #/Ream | FDA,1 | 37.8 | 48.2 | TM 101
2BJ123/Web_NP8550_11 | 2BJ123 | Web_NP8550_11 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 6855 | g/in | FDA,0 | 3500 |  | TM 107
2BJ143/Web_NP8530_16.25 | 2BJ143 | Web_NP8530_16.25 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2BJ143/Web_NP8530_16.25 | 2BJ143 | Web_NP8530_16.25 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 16.25 | in | FDA,4 | 16.1875 | 16.3125 | Measurement
2BJ143/Web_NP8530_16.25 | 2BJ143 | Web_NP8530_16.25 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 34.26 | #/Ream | FDA,2 | 29.4 | 37.42 | TM 101
2BJ143/Web_NP8530_16.25 | 2BJ143 | Web_NP8530_16.25 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 7619 | g/in | FDA,0 | 3500 |  | TM 107
2BJ143/Web_NP8530_16.25 | 2BJ143 | Web_NP8530_16.25 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 2.1 | mil | FDA,1 | 1.8 | 2.2 | TM 132
2BJ213/Web_NP8530_18 | 2BJ213 | Web_NP8530_18 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2BJ213/Web_NP8530_18 | 2BJ213 | Web_NP8530_18 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 18 | in | FDA,4 | 17.9375 | 18.0625 | Measurement
2BJ213/Web_NP8530_18 | 2BJ213 | Web_NP8530_18 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 33.33 | #/Ream | FDA,2 | 29.4 | 37.42 | TM 101
2BJ213/Web_NP8530_18 | 2BJ213 | Web_NP8530_18 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 6908 | g/in | FDA,0 | 3500 |  | TM 107
2BJ213/Web_NP8530_18 | 2BJ213 | Web_NP8530_18 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 1.85 | mil | FDA,1 | 1.8 | 2.2 | TM 132
2BJ227/Web_NP8530_15.75 | 2BJ227 | Web_NP8530_15.75 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2BJ227/Web_NP8530_15.75 | 2BJ227 | Web_NP8530_15.75 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 15.75 | in | FDA,4 | 15.6875 | 15.8125 | Measurement
2BJ227/Web_NP8530_15.75 | 2BJ227 | Web_NP8530_15.75 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 33.9 | #/Ream | FDA,2 | 29.4 | 37.42 | TM 101
2BJ227/Web_NP8530_15.75 | 2BJ227 | Web_NP8530_15.75 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 5068.96 | g/in | FDA,0 | 3500 |  | TM 107
2BJ227/Web_NP8530_15.75 | 2BJ227 | Web_NP8530_15.75 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 2.18 | mil | FDA,1 | 1.8 | 2.2 | TM 132
2BP308SLI/Lid_Aveeno Ultra-Calming MUR | 2BP308SLI | Lid_Aveeno Ultra-Calming MUR | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection
2BP308SLI/Lid_Aveeno Ultra-Calming MUR | 2BP308SLI | Lid_Aveeno Ultra-Calming MUR | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection
2BP308SLI/Lid_Aveeno Ultra-Calming MUR | 2BP308SLI | Lid_Aveeno Ultra-Calming MUR | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual
2BP308SLI/Lid_Aveeno Ultra-Calming MUR | 2BP308SLI | Lid_Aveeno Ultra-Calming MUR | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual
2BP308SLI/Lid_Aveeno Ultra-Calming MUR | 2BP308SLI | Lid_Aveeno Ultra-Calming MUR | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 72 | mm | FDA,0 | 71 | 75 | TM 143 / Visual
2BP308SLI/Lid_Aveeno Ultra-Calming MUR | 2BP308SLI | Lid_Aveeno Ultra-Calming MUR | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 53 | mm | FDA,0 | 52 | 56 | TM 143 / Visual
2BP308SLI/Lid_Aveeno Ultra-Calming MUR | 2BP308SLI | Lid_Aveeno Ultra-Calming MUR | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual
2BQ084LI/Lid_11A | 2BQ084LI | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid fits with a tight seal on |  |  |  |  | Canister & Lid Inspection
2BQ084LI/Lid_11A | 2BQ084LI | Lid_11A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Flip Lid Force
2BQ084LI/Lid_11A | 2BQ084LI | Lid_11A | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | Flip Lid Force | 2.32 | lb | FDA,2 | 0.75 |  | TM 143
2BQ084LI/Lid_11A | 2BQ084LI | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | TM 143 / Visual
2BQ084LI/Lid_11A | 2BQ084LI | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | Bead Diameter | 4.85 | in | FDA,3 | 4.847 | 4.867 | TM 143 / Visual
2BQ084LI/Lid_11A | 2BQ084LI | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Translucent Teal (Done) |  |  |  |  | TM 143 / Visual
2BQ084LI/Lid_11A | 2BQ084LI | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Insert Fin Gap (per specification) | Insert Fin Gap | 0.071 | in | FDA,3 | 0.065 | 0.075 | TM 143 / Visual
2BQ084LI/Lid_11A | 2BQ084LI | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter (as per specification) | Lid Inside Diameter | 4.92 | in | FDA,3 | 4.917 | 4.937 | TM 143 / Visual
2BQ084LI/Lid_11A | 2BQ084LI | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.044 | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual
2BQ084LI/Lid_11A | 2BQ084LI | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual
2BQ084LI/Lid_11A | 2BQ084LI | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness (per specification) | Orifice Flap Thickness | 0.023 | in | FDA,3 | 0.018 | 0.028 | TM 143 / Visual
2BQ084LI/Lid_11A | 2BQ084LI | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.48 | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual
2BQ084LI/Lid_11A | 2BQ084LI | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Outer Diameter (per specification) | Pop-Up Hole Outer Diameter | 0.166 | in | FDA,3 | 0.16 | 0.17 | TM 143 / Visual
2BQ084LI/Lid_11A | 2BQ084LI | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 34.11 | g | FDA,1 | 32 | 42 | TM 143 / Visual
2BQ471LI/Lid_04F_560 | 2BQ471LI | Lid_04F_560 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection
2BQ471LI/Lid_04F_560 | 2BQ471LI | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual
2BQ471LI/Lid_04F_560 | 2BQ471LI | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Cap Thickness | Cap Thickness | 0.151 | in | FDA,3 | 0.147 | 0.157 | TM 143 / Visual
2BQ471LI/Lid_04F_560 | 2BQ471LI | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Translucent Blue (Done) |  |  |  |  | TM 143 / Visual
2BQ471LI/Lid_04F_560 | 2BQ471LI | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) |  |  |  |  | TM 143 / Visual
2BQ471LI/Lid_04F_560 | 2BQ471LI | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter (as per specification) | Lid Inside Diameter | 3.7695 | in | FDA,3 | 3.76 | 3.77 | TM 143 / Visual
2BQ471LI/Lid_04F_560 | 2BQ471LI | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.046 | in | FDA,3 | 0.04 | 0.05 | TM 143 / Visual
2BQ471LI/Lid_04F_560 | 2BQ471LI | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual
2BQ471LI/Lid_04F_560 | 2BQ471LI | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.486 | in | FDA,3 | 1.473 | 1.493 | TM 143 / Visual
2BQ471LI/Lid_04F_560 | 2BQ471LI | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Outer Diameter (per specification) | Pop-Up Hole Outer Diameter | 0.555 | in | FDA,3 | 0.555 | 0.565 | TM 143 / Visual
2BQ471LI/Lid_04F_560 | 2BQ471LI | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Thicknes | Pop-Up Hole Thicknes | 0.015 | in | FDA,3 | 0.015 | 0.025 | TM 143 / Visual
2BQ471LI/Lid_04F_560 | 2BQ471LI | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 22.425 | g | FDA,1 | 20 | 24 | TM 143 / Visual
2BQ551LI/Skip Lot Release_NY_2OQ850 | 2BQ551LI | Skip Lot Release_NY_2OQ850 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA
2BQ551LI/Applicator_NY_Skip lot testing | 2BQ551LI | Applicator_NY_Skip lot testing | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Certificate of Analysis
2BQ551LI/Lid_13A | 2BQ551LI | Lid_13A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Flip Lid Force
2BQ551LI/Lid_13A | 2BQ551LI | Lid_13A | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | Flip Lid Force | 1.61 | lb | FDA,2 | 0.75 |  | TM 143
2BQ551LI/Lid_NY_Skip lot testing | 2BQ551LI | Lid_NY_Skip lot testing | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | Flip Lid Force | 1.07 | lb | FDA,2 | 0.75 |  | TM 143
2BQ551LI/Lid_13A | 2BQ551LI | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | TM 143 / Visual
2BQ551LI/Lid_13A | 2BQ551LI | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | Bead Diameter | 4.8545 | in | FDA,3 | 4.847 | 4.867 | TM 143 / Visual
2BQ551LI/Lid_NY_Skip lot testing | 2BQ551LI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | Bead Diameter | 4.85 | in | FDA,3 | 4.847 | 4.867 | TM 143 / Visual
2BQ551LI/Lid_13A | 2BQ551LI | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Center to Center Distance | Center to Center Distance | 0.45 | in | FDA,3 | 0.445 | 0.455 | TM 143 / Visual
2BQ551LI/Lid_NY_Skip lot testing | 2BQ551LI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Center to Center Distance | Center to Center Distance | 0.45 | in | FDA,3 | 0.445 | 0.455 | TM 143 / Visual
2BQ551LI/Lid_13A | 2BQ551LI | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Translucent Purple (Done) |  |  |  |  | TM 143 / Visual
2BQ551LI/Lid_NY_Skip lot testing | 2BQ551LI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Translucent Purple (Done) |  |  |  |  | TM 143 / Visual
2BQ551LI/Lid_13A | 2BQ551LI | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | End Circle Diameter | End Circle Diameter | 0.123 | in | FDA,3 | 0.12 | 0.13 | TM 143 / Visual
2BQ551LI/Lid_NY_Skip lot testing | 2BQ551LI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | End Circle Diameter | End Circle Diameter | 0.13 | in | FDA,3 | 0.12 | 0.13 | TM 143 / Visual
2BQ551LI/Lid_13A | 2BQ551LI | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) |  |  |  |  | TM 143 / Visual
2BQ551LI/Lid_NY_Skip lot testing | 2BQ551LI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) |  |  |  |  | TM 143 / Visual
2BQ551LI/Lid_13A | 2BQ551LI | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Insert Fin Gap (per specification) | Insert Fin Gap | 0.076 | in | FDA,3 | 0.07 | 0.08 | TM 143 / Visual
2BQ551LI/Lid_NY_Skip lot testing | 2BQ551LI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Insert Fin Gap (per specification) | Insert Fin Gap | 0.076 | in | FDA,3 | 0.07 | 0.08 | TM 143 / Visual
2BQ551LI/Lid_13A | 2BQ551LI | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter (as per specification) | Lid Inside Diameter | 4.923 | in | FDA,3 | 4.917 | 4.937 | TM 143 / Visual
2BQ551LI/Lid_NY_Skip lot testing | 2BQ551LI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter (as per specification) | Lid Inside Diameter | 4.92 | in | FDA,3 | 4.917 | 4.937 | TM 143 / Visual
2BQ551LI/Lid_13A | 2BQ551LI | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.04 | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual
2BQ551LI/Lid_NY_Skip lot testing | 2BQ551LI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.042 | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual
2BQ551LI/Lid_13A | 2BQ551LI | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual
2BQ551LI/Lid_13A | 2BQ551LI | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness (per specification) | Orifice Flap Thickness | 0.04 | in | FDA,3 | 0.03 | 0.04 | TM 143 / Visual
2BQ551LI/Lid_NY_Skip lot testing | 2BQ551LI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness (per specification) | Orifice Flap Thickness | 0.031 | in | FDA,3 | 0.03 | 0.04 | TM 143 / Visual
2BQ551LI/Lid_13A | 2BQ551LI | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.483 | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual
2BQ551LI/Lid_NY_Skip lot testing | 2BQ551LI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.49 | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual
2BQ551LI/Lid_13A | 2BQ551LI | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 34.73 | g | FDA,1 | 32 | 42 | TM 143 / Visual
2BQ551LI/Lid_NY_Skip lot testing | 2BQ551LI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 34.65 | g | FDA,1 | 32 | 42 | TM 143 / Visual
2BS420/Web_NP3870_11.5 | 2BS420 | Web_NP3870_11.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2BS420/Web_NP3870_11.5 | 2BS420 | Web_NP3870_11.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 11.5 | in | FDA,4 | 11.4375 | 11.5625 | Measurement
2BS420/Web_NP3870_11.5 | 2BS420 | Web_NP3870_11.5 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 91.8 | #/Ream | FDA,1 | 81 | 99 | TM 101
2BS420/Web_NP3870_11.5 | 2BS420 | Web_NP3870_11.5 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@350DegF/40psi/1.0 sec dwell) | Seal Strength (@350°F/40psi/1.0 sec dwell) | 10192.62 | g/in | FDA,0 | 5000 |  | TM 107
2BU275/Web_NP3895_8.5 | 2BU275 | Web_NP3895_8.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2BU275/Web_NP3895_8.5 | 2BU275 | Web_NP3895_8.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 8.5 | in | FDA,4 | 8.4375 | 8.5625 | Measurement
2BU275/Web_NP3895_8.5 | 2BU275 | Web_NP3895_8.5 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/3000_sq_ft) | Basis Weight | 71.69 | #/3000sqft | FDA,1 | 65.3 | 79.8 | TM 101
2BU275/Web_NP3895_8.5 | 2BU275 | Web_NP3895_8.5 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/inch@250DegF/30psi/0.5 sec dwell) | Seal Strength (@250°F/30psi/0.5 sec dwell) | 5589 | g/in | FDA,0 | 2270 |  | TM 107
2BU275/Web_NP3895_8.5 | 2BU275 | Web_NP3895_8.5 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 4.05 | mil | FDA,1 | 3.9 | 4.7 | TM 132
2CA580SLI/Lid_01R | 2CA580SLI | Lid_01R | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection
2CA580SLI/Lid_01R | 2CA580SLI | Lid_01R | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection
2CA580SLI/Lid_01R | 2CA580SLI | Lid_01R | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual
2CA580SLI/Lid_01R | 2CA580SLI | Lid_01R | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual
2CA580SLI/Lid_01R | 2CA580SLI | Lid_01R | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.73 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual
2CA580SLI/Lid_01R | 2CA580SLI | Lid_01R | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.875 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual
2CA580SLI/Lid_01R | 2CA580SLI | Lid_01R | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual
2CA735/Web_NP8810_12.125 | 2CA735 | Web_NP8810_12.125 | NY-Analytical | 1019 | Moisture Vapor Transmission rate | Report result from COA | Report result from COA | 5 | GSM | FDA,0 |  | 10 | Report result from COA
2CA735/Web_NP8810_12.125 | 2CA735 | Web_NP8810_12.125 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Dyne Level
2CA735/Web_NP8810_12.125 | 2CA735 | Web_NP8810_12.125 | NY-Analytical | 836 | Dyne Level | Dyne Level (outside film) | Dyne Level (outside film) | 36 |  | FDA,0 | 36 | 38 | Dyne Test Kit
2CA735/Web_NP8810_12.125 | 2CA735 | Web_NP8810_12.125 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 12.125 | in | FDA,4 | 12.0625 | 12.1875 | Measurement
2CA735/Web_NP8810_12.125 | 2CA735 | Web_NP8810_12.125 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 42.33 | #/Ream | FDA,1 | 39.4 | 48.2 | TM 101
2CA735/Web_NP8810_12.125 | 2CA735 | Web_NP8810_12.125 | NY-Analytical | 762 | Seal Strength | Seal Strength(g/in@350F/30psi/0.5s dwell Top Jaw) | Seal Strength (@350°F/30psi/0.5s dwell Top Jaw) | 5303 | gms/f | FDA,0 | 500 |  | TM 107
2CA735/Web_NP8810_12.125 | 2CA735 | Web_NP8810_12.125 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 2.7 | mil | FDA,2 | 2.48 | 3.03 | TM 132
2CA735/Web_NP8810_12.125 | 2CA735 | Web_NP8810_12.125 | NY-Analytical | 397 | Appearance | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208
2CB196/Web_NP5114_41.0 | 2CB196 | Web_NP5114_41.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2CB196/Web_NP5114_41.0 | 2CB196 | Web_NP5114_41.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 40.9375 | in | FDA,4 | 40.9375 | 41.0625 | Measurement
2CB196/Web_NP5114_41.0 | 2CB196 | Web_NP5114_41.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 59.79 | #/Ream | FDA,1 | 53.1 | 64.9 | TM 101
2CB196/Web_NP5114_41.0 | 2CB196 | Web_NP5114_41.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/40psi/0.5 sec dewll) | Seal Strength (@300°F/40psi/0.5 sec dewll) | 2891.81 | g/in | FDA,0 | 1300 |  | TM 107
2CB196/Web_NP5114_41.0 | 2CB196 | Web_NP5114_41.0 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 3.2 | mil | FDA,2 | 3.15 | 3.85 | TM 132
2CB512/Web_NP3870_4.5 | 2CB512 | Web_NP3870_4.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2CB512/Web_NP3870_4.5 | 2CB512 | Web_NP3870_4.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 4.5 | in | FDA,4 | 4.4375 | 4.5625 | Measurement
2CB512/Web_NP3870_4.5 | 2CB512 | Web_NP3870_4.5 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 92.6 | #/Ream | FDA,1 | 81 | 99 | TM 101
2CB512/Web_NP3870_4.5 | 2CB512 | Web_NP3870_4.5 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@350DegF/40psi/1.0 sec dwell) | Seal Strength (@350°F/40psi/1.0 sec dwell) | 9864.87 | g/in | FDA,0 | 5000 |  | TM 107
2DH211/Web_NP5116_6.0 | 2DH211 | Web_NP5116_6.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2DH211/Web_NP5116_6.0 | 2DH211 | Web_NP5116_6.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 6 | in | FDA,4 | 5.9375 | 6.0625 | Measurement
2DH211/Web_NP5116_6.0 | 2DH211 | Web_NP5116_6.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Printed/Varnished) | Basis Weight | 57.67 | #/Ream | FDA,1 | 55.8 | 68.2 | TM 101
2DH211/Web_NP5116_6.0 | 2DH211 | Web_NP5116_6.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/40psi/0.5 sec dewll) | Seal Strength (@300°F/40psi/0.5 sec dewll) | 1800.28 | g/in | FDA,0 | 1500 |  | TM 107
2ED110/Web_NP5410_4.5 | 2ED110 | Web_NP5410_4.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2ED110/Web_NP5410_4.5 | 2ED110 | Web_NP5410_4.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 4.5 | in | FDA,4 | 4.4375 | 4.5625 | Measurement
2ED110/Web_NP5410_4.5 | 2ED110 | Web_NP5410_4.5 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Printed) | Basis Weight | 54.5 | #/Ream | FDA,1 | 49.8 | 60.8 | TM 101
2ED110/Web_NP5410_4.5 | 2ED110 | Web_NP5410_4.5 | NY-Analytical | 762 | Seal Strength | Destruct Bond | Destruct Bond | Break occurs off the seal (Done) | g/in |  |  |  | TM 107
2ED110/Web_NP5410_4.5 | 2ED110 | Web_NP5410_4.5 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/40psi/0.5 sec dewll) | Seal Strength (@300°F/40psi/0.5 sec dewll) | 1814.1 | g/in | FDA,0 | 1500 |  | TM 107
2EU351/Web_NP5211_7.0 | 2EU351 | Web_NP5211_7.0 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Eyespot Repeat
2EU351/Web_NP5211_7.0 | 2EU351 | Web_NP5211_7.0 | NY-Analytical | 999 | Eyespot Repeat | Eye spot repeat over 7 impressions | Eye spot repeat over 7 impressions | 488 | mm | FDA,0 | 488 | 490 | TM 208
2EU351/Web_NP5211_7.0 | 2EU351 | Web_NP5211_7.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 7 | in | FDA,4 | 6.9375 | 7.0625 | TM 143
2EU351/Web_NP5211_7.0 | 2EU351 | Web_NP5211_7.0 | NY-Analytical | 1031 | 2D Matrix (Barcode) | 2D Matrix (Barcode) | 2D Matrix (Barcode) | Meet Specifications (Done) |  |  |  |  | Visual
2EU351/Web_NP5211_7.0 | 2EU351 | Web_NP5211_7.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Visual
2EU351/Web_NP5211_7.0 | 2EU351 | Web_NP5211_7.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis (ELLI) | Certificate of Analysis | COA Meets Specification (Done) |  |  |  |  | Visual
2EU351/Web_NP5211_7.0 | 2EU351 | Web_NP5211_7.0 | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | Visual Observation
2EU351/Web_NP5211_7.0 | 2EU351 | Web_NP5211_7.0 | NY-Analytical | 748 | Visual Inspection | Color/Copy | Color/Copy | Meets specifications (Done) |  |  |  |  | Visual Observation
2FB6038/Web_NP5410_8.0 | 2FB6038 | Web_NP5410_8.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2FB6038/Web_NP5410_8.0 | 2FB6038 | Web_NP5410_8.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width |  | in | FDA,4 | 7.9375 | 8.0625 | Measurement
2FB6038/Web_NP5410_8.0 | 2FB6038 | Web_NP5410_8.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Printed) | Basis Weight | 52.91 | #/Ream | FDA,1 | 49.8 | 60.8 | TM 101
2FB6038/Web_NP5410_8.0 | 2FB6038 | Web_NP5410_8.0 | NY-Analytical | 762 | Seal Strength | Destruct Bond | Destruct Bond |  | g/in |  |  |  | TM 107
2FB6038/Web_NP5410_8.0 | 2FB6038 | Web_NP5410_8.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/40psi/0.5 sec dewll) | Seal Strength (@300°F/40psi/0.5 sec dewll) |  | g/in | FDA,0 | 1500 |  | TM 107
2FH070/Web_NP5115_6.0 | 2FH070 | Web_NP5115_6.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2FH070/Web_NP5115_6.0 | 2FH070 | Web_NP5115_6.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 6 | in | FDA,4 | 5.9375 | 6.0625 | Measurement
2FH070/Web_NP5115_6.0 | 2FH070 | Web_NP5115_6.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Printed/Varnished) | Basis Weight | 59.66 | #/Ream | FDA,1 | 55.6 | 68 | TM 101
2FH070/Web_NP5115_6.0 | 2FH070 | Web_NP5115_6.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@325DegF/40PSI/0.5Sec_Dwell) | Seal Strength (@325°F/40PSI/0.5Sec Dwell) | 2304.16 | g/in | FDA,0 | 1500 |  | TM 107
2FJ141R/Web_NP8535_16 | 2FJ141R | Web_NP8535_16 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2FJ141R/Web_NP8535_16 | 2FJ141R | Web_NP8535_16 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 16 | in | FDA,4 | 15.9375 | 16.0625 | Measurement
2FJ141R/Web_NP8535_16 | 2FJ141R | Web_NP8535_16 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 42.06 | #/Ream | FDA,2 | 35.82 | 45.6 | TM 101
2FJ141R/Web_NP8535_16 | 2FJ141R | Web_NP8535_16 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 7567 | g/in | FDA,0 | 3500 |  | TM 107
2FJ141R/Web_NP8535_16 | 2FJ141R | Web_NP8535_16 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 2.45 | mil | FDA,1 | 2.2 | 2.8 | TM 132
2HA580/Web_NP8540_17.0 | 2HA580 | Web_NP8540_17.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2HA580/Web_NP8540_17.0 | 2HA580 | Web_NP8540_17.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 17 | in | FDA,4 | 16.9375 | 17.0625 | Measurement
2HA580/Web_NP8540_17.0 | 2HA580 | Web_NP8540_17.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 52.91 | #/Ream | FDA,1 | 45.4 | 57.8 | TM 101
2HA580/Web_NP8540_17.0 | 2HA580 | Web_NP8540_17.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 7254 | g/in | FDA,0 | 3500 |  | TM 107
2HA580/Web_NP8540_17.0 | 2HA580 | Web_NP8540_17.0 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 3 | mil | FDA,2 | 2.86 | 3.64 | TM 132
2HU120/Web_NP5115_8.5 | 2HU120 | Web_NP5115_8.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2HU120/Web_NP5115_8.5 | 2HU120 | Web_NP5115_8.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 8.5 | in | FDA,4 | 8.4375 | 8.5625 | Measurement
2HU120/Web_NP5115_8.5 | 2HU120 | Web_NP5115_8.5 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Printed/Varnished) | Basis Weight | 58.47 | #/Ream | FDA,1 | 55.6 | 68 | TM 101
2HU120/Web_NP5115_8.5 | 2HU120 | Web_NP5115_8.5 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@325DegF/40PSI/0.5Sec_Dwell) | Seal Strength (@325°F/40PSI/0.5Sec Dwell) | 2199 | g/in | FDA,0 | 1500 |  | TM 107
2JC699/Web_NP5410_6.0 | 2JC699 | Web_NP5410_6.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2JC699/Web_NP5410_6.0 | 2JC699 | Web_NP5410_6.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 6 | in | FDA,4 | 5.9375 | 6.0625 | Measurement
2JC699/Web_NP5410_6.0 | 2JC699 | Web_NP5410_6.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Printed) | Basis Weight | 55.6 | #/Ream | FDA,1 | 49.8 | 60.8 | TM 101
2JC699/Web_NP5410_6.0 | 2JC699 | Web_NP5410_6.0 | NY-Analytical | 762 | Seal Strength | Destruct Bond | Destruct Bond | Break occurs off the seal (Done) | g/in |  |  |  | TM 107
2JC699/Web_NP5410_6.0 | 2JC699 | Web_NP5410_6.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/40psi/0.5 sec dewll) | Seal Strength (@300°F/40psi/0.5 sec dewll) | 1686.57 | g/in | FDA,0 | 1500 |  | TM 107
2LB406/Web_NP3870_10.0 | 2LB406 | Web_NP3870_10.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2LB406/Web_NP3870_10.0 | 2LB406 | Web_NP3870_10.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 10 | in | FDA,4 | 9.9375 | 10.0625 | Measurement
2LB406/Web_NP3870_10.0 | 2LB406 | Web_NP3870_10.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 93.53 | #/Ream | FDA,1 | 83.7 | 96.3 | TM 101
2LB406/Web_NP3870_10.0 | 2LB406 | Web_NP3870_10.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@350DegF/40psi/1.0 sec dwell) | Seal Strength (@350°F/40psi/1.0 sec dwell) | 10196.54 | g/in | FDA,0 | 5000 |  | TM 107
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Solo Lid Inspection
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 818 | Solo Lid Inspection | Appearance | Appearance | Meet specification (Done) |  |  |  |  | Solo Lid Inspection
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 818 | Solo Lid Inspection | Color | Color | Meets the specification (Done) |  |  |  |  | Solo Lid Inspection
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 818 | Solo Lid Inspection | Functionality | Functionality | Meet specification (Done) |  |  |  |  | Solo Lid Inspection
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 818 | Solo Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.0285 | in | FDA,3 | 0.025 | 0.035 | Solo Lid Inspection
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 818 | Solo Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | Solo Lid Inspection
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 818 | Solo Lid Inspection | Open Lid Overall Width | Open Lid Overall Width | 4.367 | in | FDA,3 | 4.351 | 4.401 | Solo Lid Inspection
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 818 | Solo Lid Inspection | Overall closed lid height | Overall closed lid height | 0.3688 | in | FDA,3 | 0.355 | 0.385 | Solo Lid Inspection
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 818 | Solo Lid Inspection | Weight | Weight | 11.33 | g | FDA,1 | 10.5 | 12.5 | Measurement
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.7187 | in | FDA,4 | 3.6875 | 3.8125 | TM 143 / Visual
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.875 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Lid_NY_NP7600_22S | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Lid_NY_NP7600_22S | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Solo Lid Inspection
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Lid_NY_NP7600_22S | NY-Analytical | 818 | Solo Lid Inspection | Appearance | Appearance | Meet specification (Done) |  |  |  |  | Solo Lid Inspection
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Lid_NY_NP7600_22S | NY-Analytical | 818 | Solo Lid Inspection | Color | Color | Teal Pantone 3145C |  |  |  |  | Solo Lid Inspection
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Lid_NY_NP7600_22S | NY-Analytical | 818 | Solo Lid Inspection | Functionality | Functionality | Meet specification (Done) |  |  |  |  | Solo Lid Inspection
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Lid_NY_NP7600_22S | NY-Analytical | 818 | Solo Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.028 | in | FDA,3 | 0.025 | 0.035 | Solo Lid Inspection
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Lid_NY_NP7600_22S | NY-Analytical | 818 | Solo Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | Solo Lid Inspection
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Lid_NY_NP7600_22S | NY-Analytical | 818 | Solo Lid Inspection | Open Lid Overall Width | Open Lid Overall Width | 4.361 | in | FDA,3 | 4.351 | 4.401 | Solo Lid Inspection
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Lid_NY_NP7600_22S | NY-Analytical | 818 | Solo Lid Inspection | Overall closed lid height | Overall closed lid height | 0.3815 | in | FDA,3 | 0.355 | 0.385 | Solo Lid Inspection
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Lid_NY_NP7600_22S | NY-Analytical | 818 | Solo Lid Inspection | Weight | Weight | 11.26 | g | FDA,1 | 10.5 | 12.5 | Measurement
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Lid_NY_NP7600_22S | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Lid_NY_NP7600_22S | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.75 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Lid_NY_NP7600_22S | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.875 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Solo Lid Inspection
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 818 | Solo Lid Inspection | Color | Color | Meets the specification (Done) |  |  |  |  | Solo Lid Inspection
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 818 | Solo Lid Inspection | Functionality | Functionality | Meet specification (Done) |  |  |  |  | Solo Lid Inspection
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 818 | Solo Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.027 | in | FDA,3 | 0.025 | 0.035 | Solo Lid Inspection
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 818 | Solo Lid Inspection | Open Lid Overall Width | Open Lid Overall Width | 4.358 | in | FDA,3 | 4.351 | 4.401 | Solo Lid Inspection
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 818 | Solo Lid Inspection | Overall closed lid height | Overall closed lid height | 0.3785 | in | FDA,3 | 0.355 | 0.385 | Solo Lid Inspection
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 818 | Solo Lid Inspection | Weight | Weight | 11.235 | g | FDA,1 | 10.5 | 12.5 | Measurement
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.69 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.83 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual
2OD024/Web_NY_Skip lot testing | 2OD024 | Web_NY_Skip lot testing | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Printed/Varnished) | Basis Weight | 59.8 | #/Ream | FDA,1 | 55.6 | 68 | TM 101
2OD024/Web_NY_Skip lot testing | 2OD024 | Web_NY_Skip lot testing | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@325DegF/40PSI/0.5Sec_Dwell) | Seal Strength (@325°F/40PSI/0.5Sec Dwell) | 3353.04 | g/in | FDA,0 | 1500 |  | TM 107
2OJ124SLI/Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | 2OJ124SLI | Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection
2OJ124SLI/Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | 2OJ124SLI | Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection
2OJ124SLI/Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | 2OJ124SLI | Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual
2OJ124SLI/Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | 2OJ124SLI | Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual
2OJ124SLI/Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | 2OJ124SLI | Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.7187 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual
2OJ124SLI/Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | 2OJ124SLI | Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.875 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual
2OJ124SLI/Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | 2OJ124SLI | Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Lid_NY_NP7600_22Y | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Lid_NY_NP7600_22Y | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Solo Lid Inspection
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Lid_NY_NP7600_22Y | NY-Analytical | 818 | Solo Lid Inspection | Color | Color | Meets the specification (Done) |  |  |  |  | Solo Lid Inspection
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Lid_NY_NP7600_22Y | NY-Analytical | 818 | Solo Lid Inspection | Functionality | Functionality | Meet specification (Done) |  |  |  |  | Solo Lid Inspection
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Lid_NY_NP7600_22Y | NY-Analytical | 818 | Solo Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.0285 | in | FDA,3 | 0.025 | 0.035 | Solo Lid Inspection
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Lid_NY_NP7600_22Y | NY-Analytical | 818 | Solo Lid Inspection | Open Lid Overall Width | Open Lid Overall Width | 4.373 | in | FDA,3 | 4.351 | 4.401 | Solo Lid Inspection
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Lid_NY_NP7600_22Y | NY-Analytical | 818 | Solo Lid Inspection | Overall closed lid height | Overall closed lid height | 0.3715 | in | FDA,3 | 0.355 | 0.385 | Solo Lid Inspection
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Lid_NY_NP7600_22Y | NY-Analytical | 818 | Solo Lid Inspection | Weight | Weight | 11.4 | g | FDA,1 | 10.5 | 12.5 | Measurement
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Lid_NY_NP7600_22Y | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Lid_NY_NP7600_22Y | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Lid_NY_NP7600_22Y | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.7187 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Lid_NY_NP7600_22Y | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.875 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Lid_NY_NP7600_22Y | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual
2OJ213SLI/Lid_01A | 2OJ213SLI | Lid_01A | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection
2OJ213SLI/Lid_01A | 2OJ213SLI | Lid_01A | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection
2OJ213SLI/Lid_01A | 2OJ213SLI | Lid_01A | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual
2OJ213SLI/Lid_01A | 2OJ213SLI | Lid_01A | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual
2OJ213SLI/Lid_01A | 2OJ213SLI | Lid_01A | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 1.875 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual
2OJ213SLI/Lid_01A | 2OJ213SLI | Lid_01A | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 3.6875 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual
2OJ213SLI/Lid_01A | 2OJ213SLI | Lid_01A | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual
2OM304/Web_NP8540_16.0 | 2OM304 | Web_NP8540_16.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2OM304/Web_NP8540_16.0 | 2OM304 | Web_NP8540_16.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 16 | in | FDA,4 | 15.9375 | 16.0625 | Measurement
2OM304/Web_NP8540_16.0 | 2OM304 | Web_NP8540_16.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 53.176 | #/Ream | FDA,1 | 45.4 | 57.8 | TM 101
2OM304/Web_NP8540_16.0 | 2OM304 | Web_NP8540_16.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 7038.28 | g/in | FDA,0 | 3500 |  | TM 107
2OM304/Web_NP8540_16.0 | 2OM304 | Web_NP8540_16.0 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 3.3 | mil | FDA,2 | 2.86 | 3.64 | TM 132
2OM304SLI/Lid_No-Rinse sanitizing Solo | 2OM304SLI | Lid_No-Rinse sanitizing Solo | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection
2OM304SLI/Lid_No-Rinse sanitizing Solo | 2OM304SLI | Lid_No-Rinse sanitizing Solo | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection
2OM304SLI/Lid_No-Rinse sanitizing Solo | 2OM304SLI | Lid_No-Rinse sanitizing Solo | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual
2OM304SLI/Lid_No-Rinse sanitizing Solo | 2OM304SLI | Lid_No-Rinse sanitizing Solo | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual
2OM304SLI/Lid_No-Rinse sanitizing Solo | 2OM304SLI | Lid_No-Rinse sanitizing Solo | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.7187 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual
2OM304SLI/Lid_No-Rinse sanitizing Solo | 2OM304SLI | Lid_No-Rinse sanitizing Solo | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.875 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual
2OM304SLI/Lid_No-Rinse sanitizing Solo | 2OM304SLI | Lid_No-Rinse sanitizing Solo | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual
2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | 2OM701SLI | Lid_No-Rinse sanitizing Solo_01 | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection
2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | 2OM701SLI | Lid_No-Rinse sanitizing Solo_01 | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection
2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | 2OM701SLI | Lid_No-Rinse sanitizing Solo_01 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual
2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | 2OM701SLI | Lid_No-Rinse sanitizing Solo_01 | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual
2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | 2OM701SLI | Lid_No-Rinse sanitizing Solo_01 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.687 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual
2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | 2OM701SLI | Lid_No-Rinse sanitizing Solo_01 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.8125 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual
2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | 2OM701SLI | Lid_No-Rinse sanitizing Solo_01 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual
2OP037/Web_NP8511_245 | 2OP037 | Web_NP8511_245 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2OP037/Web_NP8511_245 | 2OP037 | Web_NP8511_245 | NY-Analytical | 761 | Web Measurement | Web Width (mm) | Web Width | 245 | mm | FDA,0 | 243 | 247 | Measurement
2OP037/Web_NP8511_245 | 2OP037 | Web_NP8511_245 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 47.88 | Lbs/Ream | FDA,1 | 43.1 | 52.7 | TM 101
2OP037/Web_NP8511_245 | 2OP037 | Web_NP8511_245 | NY-Analytical | 1048 | Seal strength (NP8511) | Seal strength 2500g/in@140DegC/1sec/14psi | Seal Strength (@140°C/1sec/14psi | 6296.05 | g/in | FDA,0 | 2500 |  | TM 107
2OP037/Web_NP8511_245 | 2OP037 | Web_NP8511_245 | NY-Analytical | 763 | Web Thickness | Thickness | Thickness | 71 | micron | FDA,1 | 64.8 | 79.2 | TM 132
2OP037/Web_NP8511_245 | 2OP037 | Web_NP8511_245 | NY-Analytical | 885 | Web Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Sample(s) all meet specifications (Done) |  |  |  |  | Visual Observation
2OP063/Web_NP8511_320 | 2OP063 | Web_NP8511_320 | NY-Analytical | 836 | Dyne Level | Dyne Level (outside film) | Dyne Level (outside film) | 36 |  | FDA,0 |  |  | Dyne Level
2OP063/Web_NP8511_320 | 2OP063 | Web_NP8511_320 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2OP063/Web_NP8511_320 | 2OP063 | Web_NP8511_320 | NY-Analytical | 761 | Web Measurement | Web Width (mm) | Web Width | 320 | mm | FDA,0 | 318 | 322 | Measurement
2OP063/Web_NP8511_320 | 2OP063 | Web_NP8511_320 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 47.8 | Lbs/Ream | FDA,1 | 43.1 | 52.7 | TM 101
2OP063/Web_NP8511_320 | 2OP063 | Web_NP8511_320 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/inch@140C/14psi/1 s dwellTopJaw) | Seal Strength (@140°C/14psi/1 s dwellTopJaw) | 7207.05 | g/in | FDA,0 | 2500 |  | TM 107
2OP063/Web_NP8511_320 | 2OP063 | Web_NP8511_320 | NY-Analytical | 1048 | Seal strength (NP8511) | Seal strength 2500g/in@140DegC/1sec/14psi | Seal Strength (@140°C/1sec/14psi | 5948 | g/in | FDA,0 | 2500 |  | TM 107
2OP063/Web_NP8511_320 | 2OP063 | Web_NP8511_320 | NY-Analytical | 763 | Web Thickness | Thickness | Thickness | 74 | micron | FDA,1 | 64.8 | 79.2 | TM 132
2OP065/Web_NP8536_11.875 | 2OP065 | Web_NP8536_11.875 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2OP065/Web_NP8536_11.875 | 2OP065 | Web_NP8536_11.875 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 11.875 | in | FDA,4 | 11.8125 | 11.9375 | Measurement
2OP065/Web_NP8536_11.875 | 2OP065 | Web_NP8536_11.875 | NY-Analytical | 1044 | Basis Weight (lbs/ream) | Basis Weight | Basis Weight | 41.4 | Lb/Ream | FDA,2 | 40.1 | 49.02 | TM 101
2OP065/Web_NP8536_11.875 | 2OP065 | Web_NP8536_11.875 | NY-Analytical | 1049 | Seal strength (NP8536) | Seal Strength (g/inch@320DegF/40psi/0.5 sec dwell) | Seal Strength (@320°F/40psi/0.5 sec dwell) | 8624 | g/in | FDA,0 | 2000 |  | TM 107
2OP065/Web_NP8536_11.875 | 2OP065 | Web_NP8536_11.875 | NY-Analytical | 763 | Web Thickness | Thickness | Thickness | 2.4 | mil | FDA,3 | 2.367 | 2.893 | TM 132
2OP065/Web_NP8536_11.875 | 2OP065 | Web_NP8536_11.875 | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | Visual
2OP065/Web_NP8536_11.875 | 2OP065 | Web_NP8536_11.875 | NY-Analytical | 748 | Visual Inspection | Barcode Legibility/Quality | Barcode Legibility/Quality | Barcode should scan and code should be matched under the barcode (Done) |  |  |  |  | Visual
2OP065/Web_NP8536_11.875 | 2OP065 | Web_NP8536_11.875 | NY-Analytical | 748 | Visual Inspection | Color/Copy | Color/Copy | Meets specifications (Done) |  |  |  |  | Visual
2OP067/Web_NY_Skip lot testing | 2OP067 | Web_NY_Skip lot testing | NY-Analytical | 1044 | Basis Weight (lbs/ream) | Basis Weight | Basis Weight | 42.86 | Lb/Ream | FDA,2 | 40.1 | 49.02 | TM 101
2OP067/Web_NY_Skip lot testing | 2OP067 | Web_NY_Skip lot testing | NY-Analytical | 1049 | Seal strength (NP8536) | Seal Strength (g/inch@320DegF/40psi/0.5 sec dwell) | Seal Strength (@320°F/40psi/0.5 sec dwell) | 9089 | g/in | FDA,0 | 2000 |  | TM 107
2OP067/Web_NY_Skip lot testing | 2OP067 | Web_NY_Skip lot testing | NY-Analytical | 763 | Web Thickness | Thickness | Thickness | 2.6 | mil | FDA,3 | 2.367 | 2.893 | TM 132
2OP067/Web_NY_Skip lot testing | 2OP067 | Web_NY_Skip lot testing | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | Visual
2OP067/Web_NY_Skip lot testing | 2OP067 | Web_NY_Skip lot testing | NY-Analytical | 748 | Visual Inspection | Barcode Legibility/Quality | Barcode Legibility/Quality | Barcode should scan and code should be matched under the barcode (Done) |  |  |  |  | Visual
2OP067/Web_NY_Skip lot testing | 2OP067 | Web_NY_Skip lot testing | NY-Analytical | 748 | Visual Inspection | Color/Copy | Color/Copy | Meets specifications (Done) |  |  |  |  | Visual
2OP072/Web_NP8511_245 | 2OP072 | Web_NP8511_245 | NY-Analytical | 836 | Dyne Level | Dyne Level (outside film) | Dyne Level (outside film) | 34 |  | FDA,0 |  |  | Dyne Level
2OP073/Web_NP8536_11.875 | 2OP073 | Web_NP8536_11.875 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight |  | lbs/rm | FDA,2 |  |  | TM 101
2OP073/Web_NP8536_11.875 | 2OP073 | Web_NP8536_11.875 | NY-Analytical | 762 | Seal Strength | Seal Strength(g/in@350DegF/40psi/0.5sec dwell) | Seal Strength (@350°F/40psi/0.5sec dwell) |  | g/in | FDA,0 | 2500 |  | TM 107
2OP075/Web_NP8541_12.205 | 2OP075 | Web_NP8541_12.205 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2OP075/Web_NP8541_12.205 | 2OP075 | Web_NP8541_12.205 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 12.187 | in | FDA,4 | 12.1425 | 12.2675 | Measurement
2OP075/Web_NP8541_12.205 | 2OP075 | Web_NP8541_12.205 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 86.11 | GSM | FDA,1 | 74.3 | 90.9 | TM 101
2OP075/Web_NP8541_12.205 | 2OP075 | Web_NP8541_12.205 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/inch@130DegC/0.5 sec/1 bar) | Seal Strength (@130°C/0.5 sec/1 bar) | 8916 | g/in | FDA,0 | 1000 |  | TM 107
2OP075/Web_NP8541_12.205 | 2OP075 | Web_NP8541_12.205 | NY-Analytical | 763 | Web Thickness | Thickness | Thickness | 0.084 | mm | FDA,3 | 0.074 | 0.094 | TM 132
2OP075/Web_NP8541_12.205 | 2OP075 | Web_NP8541_12.205 | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | Visual
2OP075/Web_NP8541_12.205 | 2OP075 | Web_NP8541_12.205 | NY-Analytical | 748 | Visual Inspection | Barcode Legibility/Quality | Barcode Legibility/Quality | Barcode should scan and code should be matched under the barcode (Done) |  |  |  |  | Visual
2OP075/Web_NP8541_12.205 | 2OP075 | Web_NP8541_12.205 | NY-Analytical | 748 | Visual Inspection | Color/Copy | Color/Copy | Meets specifications (Done) |  |  |  |  | Visual
2OP092/Web_NP8536_10 | 2OP092 | Web_NP8536_10 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2OP092/Web_NP8536_10 | 2OP092 | Web_NP8536_10 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 10 | in | FDA,4 | 9.9375 | 10.0625 | Measurement
2OP092/Web_NP8536_10 | 2OP092 | Web_NP8536_10 | NY-Analytical | 1044 | Basis Weight (lbs/ream) | Basis Weight | Basis Weight | 42.3 | Lb/Ream | FDA,2 | 40.1 | 49.02 | TM 101
2OP092/Web_NP8536_10 | 2OP092 | Web_NP8536_10 | NY-Analytical | 1049 | Seal strength (NP8536) | Seal Strength (g/inch@320DegF/40psi/0.5 sec dwell) | Seal Strength (@320°F/40psi/0.5 sec dwell) | 8116.37 | g/in | FDA,0 | 2000 |  | TM 107
2OP092/Web_NP8536_10 | 2OP092 | Web_NP8536_10 | NY-Analytical | 763 | Web Thickness | Thickness | Thickness | 2.8 | mil | FDA,3 | 2.367 | 2.893 | TM 132
2OP092/Web_NP8536_10 | 2OP092 | Web_NP8536_10 | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | Visual
2OP092/Web_NP8536_10 | 2OP092 | Web_NP8536_10 | NY-Analytical | 748 | Visual Inspection | Barcode Legibility/Quality | Barcode Legibility/Quality | Barcode should scan and code should be matched under the barcode (Done) |  |  |  |  | Visual
2OP092/Web_NP8536_10 | 2OP092 | Web_NP8536_10 | NY-Analytical | 748 | Visual Inspection | Color/Copy | Color/Copy | Meets specifications (Done) |  |  |  |  | Visual
2OP093/Skip Lot Release_NY_2OP093 | 2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis
2OP093/Skip Lot Release_NY_2OP093 | 2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection
2OP093/Skip Lot Release_NY_2OP093 | 2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Matches Specification (Done) |  |  |  |  | Canister & Lid Inspection
2OP093/Skip Lot Release_NY_2OP093 | 2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Matches Specification (Done) |  |  |  |  | Canister & Lid Inspection
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | Diameter | Diameter | 2.746 | in | FDA,3 | 2.709 | 2.803 | Canister & Lid Inspection
2OP093/Skip Lot Release_NY_2OP093 | 2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 765 | Canister & Lid Inspection | Diameter | Diameter | 2.75 | in | FDA,3 | 2.709 | 2.803 | Canister & Lid Inspection
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | E Diameter | E Diameter | 2.51 | in | FDA,3 | 2.509 | 2.539 | Canister & Lid Inspection
2OP093/Skip Lot Release_NY_2OP093 | 2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 765 | Canister & Lid Inspection | E Diameter | E Diameter | 2.52 | in | FDA,3 | 2.509 | 2.539 | Canister & Lid Inspection
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) |  |  |  |  | Canister & Lid Inspection
2OP093/Skip Lot Release_NY_2OP093 | 2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) |  |  |  |  | Canister & Lid Inspection
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | H Dimension | H Dimension | 0.458 | in | FDA,3 | 0.457 | 0.487 | Canister & Lid Inspection
2OP093/Skip Lot Release_NY_2OP093 | 2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 765 | Canister & Lid Inspection | H Dimension | H Dimension | 0.47 | in | FDA,3 | 0.457 | 0.487 | Canister & Lid Inspection
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | Height | Height | 3.52 | in | FDA,3 | 3.45 | 3.544 | Canister & Lid Inspection
2OP093/Skip Lot Release_NY_2OP093 | 2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 765 | Canister & Lid Inspection | Height | Height | 3.5 | in | FDA,3 | 3.45 | 3.544 | Canister & Lid Inspection
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | Canister & Lid Inspection
2OP093/Skip Lot Release_NY_2OP093 | 2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | Canister & Lid Inspection
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | S Dimension | S Dimension | 0.053 | in | FDA,3 | 0.035 | 0.065 | Canister & Lid Inspection
2OP093/Skip Lot Release_NY_2OP093 | 2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 765 | Canister & Lid Inspection | S Dimension | S Dimension | 0.042 | in | FDA,3 | 0.035 | 0.065 | Canister & Lid Inspection
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | T Diameter | T Diameter | 2.6 | in | FDA,3 | 2.593 | 2.623 | Canister & Lid Inspection
2OP093/Skip Lot Release_NY_2OP093 | 2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 765 | Canister & Lid Inspection | T Diameter | T Diameter | 2.61 | in | FDA,3 | 2.593 | 2.623 | Canister & Lid Inspection
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight | 26.19 | g | FDA,2 | 24.5 | 27.5 | Measurement
2OP093/Skip Lot Release_NY_2OP093 | 2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight | 26.32 | g | FDA,2 | 24.5 | 27.5 | Measurement
2OP093LI/Lid_NY_Threaded Blue Cap | 2OP093LI | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid fits with a tight seal on |  |  |  |  | Canister & Lid Inspection
2OP093LI/Lid_NY_Threaded Blue Cap | 2OP093LI | Lid_NY_Threaded Blue Cap | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection
2OP093LI/Lid_NY_Threaded Blue Cap | 2OP093LI | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | TM 143 / Visual
2OP093LI/Lid_NY_Threaded Blue Cap | 2OP093LI | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Matches Specification (Done) |  |  |  |  | TM 143 / Visual
2OP093LI/Lid_NY_Threaded Blue Cap | 2OP093LI | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | Diameter | Diameter | 2.7515 | in | FDA,3 | 2.748 | 2.764 | TM 143 / Visual
2OP093LI/Lid_NY_Threaded Blue Cap | 2OP093LI | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | E Diameter | E Diameter | 2.5495 | in | FDA,3 | 2.548 | 2.562 | TM 143 / Visual
2OP093LI/Lid_NY_Threaded Blue Cap | 2OP093LI | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | Height | Height | 0.5625 | in | FDA,3 | 0.56 | 0.58 | TM 143 / Visual
2OP093LI/Lid_NY_Threaded Blue Cap | 2OP093LI | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual
2OP093LI/Lid_NY_Threaded Blue Cap | 2OP093LI | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | T Diameter | T Diameter | 2.619 | in | FDA,3 | 2.619 | 2.639 | TM 143 / Visual
2OP093LI/Lid_NY_Threaded Blue Cap | 2OP093LI | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 10.465 | g | FDA,2 |  |  | TM 143 / Visual
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Cap Closure | Functionality - Cap Closure | Cap of lid |  |  |  |  | Canister & Lid Inspection
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid fits with a tight seal on |  |  |  |  | Canister & Lid Inspection
2OP134/Skip Lot Release_NY_2OP134 | 2OP134 | Skip Lot Release_NY_2OP134 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | White (Done) |  |  |  |  | Canister & Lid Inspection
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Bead OD | Bead OD | 3.77 | in | FDA,3 | 3.768 | 3.792 | Canister & Lid Inspection
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Canister Height OD | Canister Height OD | 6.66 | in | FDA,3 | 6.608 | 6.732 | Canister & Lid Inspection
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Lip (width) | Lip Width | 0.138 | in | FDA,3 | 0.127 | 0.143 | Canister & Lid Inspection
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Neck OD | Neck OD | 3.68 | in | FDA,3 | 3.673 | 3.697 | Canister & Lid Inspection
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | HDPE Resin Identification Marker | Meets the Specification (Done) |  |  |  |  | Canister & Lid Inspection
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | Canister & Lid Inspection
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight | 53.94 | g | FDA,1 | 52 | 56 | Measurement
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 766 | Lid Pull Force | Lid Pull-Off Force | Lid Pull-Off Force | 2997 | g | FDA,0 | 2000 | 3500 | TM 148
2OP138/Skip Lot Relrase_NY_2OP138C | 2OP138 | Skip Lot Relrase_NY_2OP138C | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA
2OP138/Skip Lot Relrase_NY_2OP638 | 2OP138 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | [5 empty] | Canister & Lid Inspection
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | [5 empty] | Canister & Lid Inspection
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Bead OD | Bead OD |  | in | FDA,3 | 4.922 | 4.958 | Canister & Lid Inspection
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Canister OD | Canister OD |  | in | FDA,3 | 4.956 | 5.076 | Canister & Lid Inspection
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Neck OD | Neck OD |  | in | FDA,3 | 4.824 | 4.86 | Canister & Lid Inspection
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Total Height OD | Total Height OD |  | in | FDA,3 | 6.745 | 6.865 | Canister & Lid Inspection
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | [5 empty] | Canister & Lid Inspection
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | HDPE Resin Identification Marker | [5 empty] | Canister & Lid Inspection
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Lip Width (for reference only) | Lip Width (for reference only) |  | in | FDA,3 | 0.14 |  | Canister & Lid Inspection
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | [5 empty] | Canister & Lid Inspection
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Wall Thickness (Bottom base area) | Wall Thickness (Bottom base area) |  | in | FDA,3 | 0.01 |  | Canister & Lid Inspection
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Wall Thickness (Top area) | Wall Thickness (Top area) |  | in | FDA,3 | 0.02 |  | Canister & Lid Inspection
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight |  | g | FDA,1 | 66.5 | 73.5 | Measurement
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 1052 | Top Load Compression Strength | Top load Compression Strength (0.5"/mSR,0.450"TD) | Top load Compression Strength (0.5"/mSR,0.450"TD) | 150.3 | lb | FDA,0 | 84 |  | NPTM 1019
2OP138/Skip Lot Relrase_NY_2OP638 | 2OP138 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | [5 empty] | TM 143 / Visual
2OP138/Skip Lot Relrase_NY_2OP638 | 2OP138 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | White (Done) |  |  |  |  | TM 143 / Visual
2OP138/Skip Lot Relrase_NY_2OP638 | 2OP138 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | [5 empty] | TM 143 / Visual
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 1053 | Cavity | cavity Inserts | cavity Inserts | [5 empty] | Visual
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | [5 empty] | Canister & Lid Inspection
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | [5 empty] | Canister & Lid Inspection
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Bead OD | Bead OD |  | in | FDA,3 | 4.922 | 4.958 | Canister & Lid Inspection
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Canister OD | Canister OD |  | in | FDA,3 | 4.956 | 5.076 | Canister & Lid Inspection
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Neck OD | Neck OD |  | in | FDA,3 | 4.824 | 4.86 | Canister & Lid Inspection
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Total Height OD | Total Height OD | 6.785 | in | FDA,3 | 6.745 | 6.865 | Canister & Lid Inspection
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | [5 empty] | Canister & Lid Inspection
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | HDPE Resin Identification Marker | [5 empty] | Canister & Lid Inspection
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | [5 empty] | Canister & Lid Inspection
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Wall Thickness (Bottom base area) | Wall Thickness (Bottom base area) |  | in | FDA,3 | 0.01 |  | Canister & Lid Inspection
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Wall Thickness (Top area) | Wall Thickness (Top area) |  | in | FDA,3 | 0.02 |  | Canister & Lid Inspection
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight |  | g | FDA,1 | 66.5 | 73.5 | Measurement
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 1052 | Top Load Compression Strength | Top load Compression Strength (0.5"/mSR,0.450"TD) | Top load Compression Strength (0.5"/mSR,0.450"TD) |  | lb | FDA,0 | 84 |  | NPTM 1019
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 1053 | Cavity | cavity Inserts | cavity Inserts | [5 empty] | Visual
2OP159LI/Lid_03B | 2OP159LI | Lid_03B | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection
2OP159LI/Lid_03B | 2OP159LI | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual
2OP159LI/Lid_03B | 2OP159LI | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | Bead Diameter | 4.856 | in | FDA,3 | 4.847 | 4.867 | TM 143 / Visual
2OP159LI/Lid_03B | 2OP159LI | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Cap Thickness | Cap Thickness | 0.049 | in | FDA,3 | 0.045 | 0.055 | TM 143 / Visual
2OP159LI/Lid_03B | 2OP159LI | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Blue (Done) |  |  |  |  | TM 143 / Visual
2OP159LI/Lid_03B | 2OP159LI | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) |  |  |  |  | TM 143 / Visual
2OP159LI/Lid_03B | 2OP159LI | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Insert Fin Gap (per specification) | Insert Fin Gap | 0.07 | in | FDA,3 | 0.065 | 0.075 | TM 143 / Visual
2OP159LI/Lid_03B | 2OP159LI | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter (as per specification) | Lid Inside Diameter | 4.929 | in | FDA,3 | 4.922 | 4.932 | TM 143 / Visual
2OP159LI/Lid_03B | 2OP159LI | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.042 | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual
2OP159LI/Lid_03B | 2OP159LI | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual
2OP159LI/Lid_03B | 2OP159LI | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.486 | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual
2OP159LI/Lid_03B | 2OP159LI | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Outer Diameter (per specification) | Pop-Up Hole Outer Diameter | 0.163 | in | FDA,3 | 0.16 | 0.17 | TM 143 / Visual
2OP159LI/Lid_03B | 2OP159LI | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Thicknes | Pop-Up Hole Thicknes | 0.02 | in | FDA,3 | 0.012 | 0.022 | TM 143 / Visual
2OP159LI/Lid_03B | 2OP159LI | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 34.335 | g | FDA,1 | 32 | 42 | TM 143 / Visual
2OP213/Web_NP8541_11 | 2OP213 | Web_NP8541_11 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2OP213/Web_NP8541_11 | 2OP213 | Web_NP8541_11 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 11 | in | FDA,4 | 10.9375 | 11.0625 | Measurement
2OP213/Web_NP8541_11 | 2OP213 | Web_NP8541_11 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 86.76 | GSM | FDA,1 | 74.3 | 90.9 | TM 101
2OP213/Web_NP8541_11 | 2OP213 | Web_NP8541_11 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/inch@130DegC/0.5 sec/1 bar) | Seal Strength (@130°C/0.5 sec/1 bar) | 8561 | g/in | FDA,0 | 1000 |  | TM 107
2OP213/Web_NP8541_11 | 2OP213 | Web_NP8541_11 | NY-Analytical | 763 | Web Thickness | Thickness | Thickness | 0.088 | mm | FDA,3 | 0.074 | 0.094 | TM 132
2OP213/Web_NP8541_11 | 2OP213 | Web_NP8541_11 | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | Visual
2OP213/Web_NP8541_11 | 2OP213 | Web_NP8541_11 | NY-Analytical | 748 | Visual Inspection | Barcode Legibility/Quality | Barcode Legibility/Quality | Barcode should scan and code should be matched under the barcode (Done) |  |  |  |  | Visual
2OP213/Web_NP8541_11 | 2OP213 | Web_NP8541_11 | NY-Analytical | 748 | Visual Inspection | Color/Copy | Color/Copy | Meets specifications (Done) |  |  |  |  | Visual
2OP223/Web_NP8511_320 | 2OP223 | Web_NP8511_320 | NY-Analytical | 885 | Web Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Sample(s) all meet specifications (Done) |  |  |  |  | Visual Observation
2OP242LI/Lid_17C | 2OP242LI | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Cap Closure | Functionality - Cap Closure | Cap of lid |  |  |  |  | Canister & Lid Inspection
2OP242LI/Lid_17C | 2OP242LI | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid fits with a tight seal on |  |  |  |  | Canister & Lid Inspection
2OP242LI/Lid_17C | 2OP242LI | Lid_17C | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Flip Lid Force
2OP242LI/Lid_17C | 2OP242LI | Lid_17C | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | Flip Lid Force | 1.11 | lb | FDA,2 | 0.75 |  | TM 143
2OP242LI/Lid_17C | 2OP242LI | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | TM 143 / Visual
2OP242LI/Lid_17C | 2OP242LI | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | Bead Diameter | 4.835 | in | FDA,3 | 4.831 | 4.851 | TM 143 / Visual
2OP242LI/Lid_17C | 2OP242LI | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Translucent Plum (Done) |  |  |  |  | TM 143 / Visual
2OP242LI/Lid_17C | 2OP242LI | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.045 | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual
2OP242LI/Lid_17C | 2OP242LI | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual
2OP242LI/Lid_17C | 2OP242LI | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Diameter | Orifice Diameter | 0.62 | in | FDA,3 | 0.62 | 0.63 | TM 143 / Visual
2OP242LI/Lid_17C | 2OP242LI | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness (per specification) | Orifice Flap Thickness | 0.021 | in | FDA,3 | 0.017 | 0.023 | TM 143 / Visual
2OP242LI/Lid_17C | 2OP242LI | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.483 | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual
2OP242LI/Lid_17C | 2OP242LI | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 38.41 | g | FDA,1 | 32.5 | 41.5 | TM 143 / Visual
2OP280LI/Lid_17A | 2OP280LI | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Cap Closure | Functionality - Cap Closure | Cap of lid |  |  |  |  | Canister & Lid Inspection
2OP280LI/Lid_17A | 2OP280LI | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid fits with a tight seal on |  |  |  |  | Canister & Lid Inspection
2OP280LI/Lid_17A | 2OP280LI | Lid_17A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Flip Lid Force
2OP280LI/Lid_17A | 2OP280LI | Lid_17A | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | Flip Lid Force | 2.88 | lb | FDA,2 | 0.75 |  | TM 143
2OP280LI/Lid_17A | 2OP280LI | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | TM 143 / Visual
2OP280LI/Lid_17A | 2OP280LI | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | Bead Diameter | 4.851 | in | FDA,3 | 4.831 | 4.851 | TM 143 / Visual
2OP280LI/Lid_17A | 2OP280LI | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Translucent Orange (Done) |  |  |  |  | TM 143 / Visual
2OP280LI/Lid_NY_Skip lot testing | 2OP280LI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Cap Closure | Functionality - Cap Closure | [5 empty] | TM 143 / Visual
2OP280LI/Lid_NY_Skip lot testing | 2OP280LI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | [5 empty] | TM 143 / Visual
2OP280LI/Lid_17A | 2OP280LI | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.035 | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual
2OP280LI/Lid_17A | 2OP280LI | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual
2OP280LI/Lid_17A | 2OP280LI | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Diameter | Orifice Diameter | 1.468 | in | FDA,3 | 1.468 | 1.478 | TM 143 / Visual
2OP280LI/Lid_NY_Skip lot testing | 2OP280LI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Diameter | Orifice Diameter |  | in | FDA,3 | 1.468 | 1.478 | TM 143 / Visual
2OP280LI/Lid_17A | 2OP280LI | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness (per specification) | Orifice Flap Thickness | 0.035 | in | FDA,3 | 0.03 | 0.04 | TM 143 / Visual
2OP280LI/Lid_17A | 2OP280LI | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.484 | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual
2OP280LI/Lid_17A | 2OP280LI | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 37.555 | g | FDA,1 | 36 | 42 | TM 143 / Visual
2OP300LI/Lid_13E | 2OP300LI | Lid_13E | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection
2OP300LI/Lid_13E | 2OP300LI | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual
2OP300LI/Lid_13E | 2OP300LI | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | Bead Diameter | 4.852 | in | FDA,3 | 4.847 | 4.867 | TM 143 / Visual
2OP300LI/Lid_13E | 2OP300LI | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Center to Center Distance | Center to Center Distance | 0.455 | in | FDA,3 | 0.445 | 0.455 | TM 143 / Visual
2OP300LI/Lid_13E | 2OP300LI | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Translucent Blue (Done) |  |  |  |  | TM 143 / Visual
2OP300LI/Lid_13E | 2OP300LI | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | End Circle Diameter | End Circle Diameter | 0.128 | in | FDA,3 | 0.12 | 0.13 | TM 143 / Visual
2OP300LI/Lid_13E | 2OP300LI | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) |  |  |  |  | TM 143 / Visual
2OP300LI/Lid_13E | 2OP300LI | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Insert Fin Gap (per specification) | Insert Fin Gap | 0.08 | in | FDA,3 | 0.07 | 0.08 | TM 143 / Visual
2OP300LI/Lid_13E | 2OP300LI | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.04 | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual
2OP300LI/Lid_13E | 2OP300LI | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual
2OP300LI/Lid_13E | 2OP300LI | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness (per specification) | Orifice Flap Thickness | 0.031 | in | FDA,3 | 0.03 | 0.04 | TM 143 / Visual
2OP300LI/Lid_13E | 2OP300LI | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.4885 | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual
2OP300LI/Lid_13E | 2OP300LI | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 35.535 | g | FDA,1 | 32 | 42 | TM 143 / Visual
2OP300LI/Lid_13E | 2OP300LI | Lid_13E | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | Flip Lid Force | 1.05 | lb | FDA,2 | 0.75 |  | TM 148
2OP435LI/Lid_04C_560 | 2OP435LI | Lid_04C_560 | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection
2OP435LI/Lid_04C_560 | 2OP435LI | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual
2OP435LI/Lid_04C_560 | 2OP435LI | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Cap Thickness | Cap Thickness | 0.157 | in | FDA,3 | 0.147 | 0.157 | TM 143 / Visual
2OP435LI/Lid_04C_560 | 2OP435LI | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Translucent Purple (Done) |  |  |  |  | TM 143 / Visual
2OP435LI/Lid_04C_560 | 2OP435LI | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) |  |  |  |  | TM 143 / Visual
2OP435LI/Lid_04C_560 | 2OP435LI | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter (as per specification) | Lid Inside Diameter | 3.765 | in | FDA,3 | 3.76 | 3.77 | TM 143 / Visual
2OP435LI/Lid_04C_560 | 2OP435LI | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.04 | in | FDA,3 | 0.04 | 0.05 | TM 143 / Visual
2OP435LI/Lid_04C_560 | 2OP435LI | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual
2OP435LI/Lid_04C_560 | 2OP435LI | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.492 | in | FDA,3 | 1.473 | 1.493 | TM 143 / Visual
2OP435LI/Lid_04C_560 | 2OP435LI | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Outer Diameter (per specification) | Pop-Up Hole Outer Diameter | 0.565 | in | FDA,3 | 0.555 | 0.565 | TM 143 / Visual
2OP435LI/Lid_04C_560 | 2OP435LI | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Thicknes | Pop-Up Hole Thicknes | 0.02 | in | FDA,3 | 0.015 | 0.025 | TM 143 / Visual
2OP435LI/Lid_04C_560 | 2OP435LI | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 21.1 | g | FDA,1 | 20 | 24 | TM 143 / Visual
2OP630/Web_NP8440_310 | 2OP630 | Web_NP8440_310 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2OP630/Web_NP8440_310 | 2OP630 | Web_NP8440_310 | NY-Analytical | 761 | Web Measurement | Web Width (mm) | Web Width | 310 | mm | FDA,0 | 308 | 312 | Measurement
2OP630/Web_NP8440_310 | 2OP630 | Web_NP8440_310 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 50.66 | #/Ream | FDA,2 | 44.88 | 57.12 | TM 101
2OP630/Web_NP8440_310 | 2OP630 | Web_NP8440_310 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 6841.04 | g/in | FDA,0 | 3500 |  | TM 107
2OP630/Web_NP8440_310 | 2OP630 | Web_NP8440_310 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 3.2 | mil | FDA,0 | 2.64 | 3.36 | TM 132
2OP638/Skip Lot Relrase_NY_2OP638 | 2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid has a tight friction fit |  |  |  |  | Canister & Lid Inspection
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA
2OP638/Skip Lot Relrase_NY_2OP638 | 2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection
2OP638/Skip Lot Relrase_NY_2OP638 | 2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Bead OD | Bead OD | 4.9245 | in | FDA,3 | 4.922 | 4.958 | Canister & Lid Inspection
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Bead OD | Bead OD | 4.9545 | in | FDA,3 | 4.922 | 4.958 | Canister & Lid Inspection
2OP638/Skip Lot Relrase_NY_2OP638 | 2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Canister OD | Canister OD | 5.03 | in | FDA,3 | 4.956 | 5.076 | Canister & Lid Inspection
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Canister OD | Canister OD | 5.031 | in | FDA,3 | 4.956 | 5.076 | Canister & Lid Inspection
2OP638/Skip Lot Relrase_NY_2OP638 | 2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Neck OD | Neck OD | 4.8545 | in | FDA,3 | 4.824 | 4.86 | Canister & Lid Inspection
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Neck OD | Neck OD | 4.8425 | in | FDA,3 | 4.824 | 4.86 | Canister & Lid Inspection
2OP638/Skip Lot Relrase_NY_2OP638 | 2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Total Height OD | Total Height OD | 8.75 | in | FDA,3 | 8.688 | 8.808 | Canister & Lid Inspection
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Total Height OD | Total Height OD | 8.75 | in | FDA,3 | 8.688 | 8.808 | Canister & Lid Inspection
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | [5 empty] | Canister & Lid Inspection
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | HDPE Resin Identification Marker | [5 empty] | Canister & Lid Inspection
2OP638/Skip Lot Relrase_NY_2OP638 | 2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | HDPE Resin Identification Marker | Meets the Specification (Done) |  |  |  |  | Canister & Lid Inspection
2OP638/Skip Lot Relrase_NY_2OP638 | 2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Lip Width (for reference only) | Lip Width (for reference only) | 0.154 | in | FDA,3 | 0.14 |  | Canister & Lid Inspection
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Lip Width (for reference only) | Lip Width (for reference only) | 0.15 | in | FDA,3 | 0.14 |  | Canister & Lid Inspection
2OP638/Skip Lot Relrase_NY_2OP638 | 2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Wall Thickness (Bottom base area) | Wall Thickness (Bottom base area) | 0.0325 | in | FDA,3 | 0.01 |  | Canister & Lid Inspection
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Wall Thickness (Bottom base area) | Wall Thickness (Bottom base area) |  | in | FDA,3 | 0.01 |  | Canister & Lid Inspection
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Wall Thickness (Top area) | Wall Thickness (Top area) |  | in | FDA,3 | 0.02 |  | Canister & Lid Inspection
2OP638/Skip Lot Relrase_NY_2OP638 | 2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Wall Thickness (Top area) | Wall Thickness (Top area) | 0.043 | in | FDA,3 | 0.02 |  | Canister & Lid Inspection
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight | 90.2 | g | FDA,1 | 85.5 | 94.5 | Measurement
2OP638/Skip Lot Relrase_NY_2OP638 | 2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight | 89.75 | g | FDA,1 | 85.5 | 94.5 | Measurement
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 1052 | Top Load Compression Strength | Top load Compression Strength (0.5"/mSR,0.450"TD) | Top load Compression Strength (0.5"/mSR,0.450"TD) |  | lb | FDA,0 | 97 |  | NPTM 1019
2OP638/Skip Lot Relrase_NY_2OP638 | 2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 1052 | Top Load Compression Strength | Top load Compression Strength (0.5"/mSR,0.450"TD) | Top load Compression Strength (0.5"/mSR,0.450"TD) | 138.8 | lb | FDA,0 | 97 |  | NPTM 1019
2OP638/Bulk_IN_4ABBY12 - 20190218 | 2OP638 | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 902 | Bulk Liquid Inspection | Appearance (TM 208) | Appearance | [5 empty] | See Analyte for Method #
2OP638/Bulk_IN_4ABBY12 - 20190218 | 2OP638 | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 902 | Bulk Liquid Inspection | Color (TM 208) | Color | [5 empty] | See Analyte for Method #
2OP638/Bulk_IN_4ABBY12 - 20190218 | 2OP638 | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 902 | Bulk Liquid Inspection | Contamination (TM 208) | Contamination | [5 empty] | See Analyte for Method #
2OP638/Bulk_IN_4ABBY12 - 20190218 | 2OP638 | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 902 | Bulk Liquid Inspection | Odor (TM 208) | Odor | [5 empty] | See Analyte for Method #
2OP638/Bulk_IN_4ABBY12 - 20190218 | 2OP638 | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 902 | Bulk Liquid Inspection | pH (TM 200) | pH |  |  | FDA,2 | 4.5 | 4.9 | See Analyte for Method #
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | TM 143 / Visual
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | White (Done) |  |  |  |  | TM 143 / Visual
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual
2OP638/Bulk_IN_4ABBY12 - 20190218 | 2OP638 | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 943 | %Phenoxyethanol & Sodium Benzoate | Phenoxyethanol | Phenoxyethanol |  | % | FDA,3 | 0.63 | 0.77 | TM 395
2OP638/Bulk_IN_4ABBY12 - 20190218 | 2OP638 | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 943 | %Phenoxyethanol & Sodium Benzoate | Phenoxyethanol_Final | Phenoxyethanol Final |  | % | FDA,3 | 0.63 | 0.77 | TM 395
2OP638/Bulk_IN_4ABBY12 - 20190218 | 2OP638 | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 943 | %Phenoxyethanol & Sodium Benzoate | Sodium_Benzoate | Sodium Benzoate |  | % | FDA,3 | 0.423 | 0.517 | TM 395
2OP638/Bulk_IN_4ABBY12 - 20190218 | 2OP638 | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 943 | %Phenoxyethanol & Sodium Benzoate | Sodium_Benzoate_Final | Sodium Benzoate Final |  | % | FDA,3 | 0.423 | 0.517 | TM 395
2OP638/Skip Lot Relrase_NY_2OP638 | 2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 1053 | Cavity | cavity Inserts | cavity Inserts | Meets specification (Done) |  |  |  |  | Visual
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 1053 | Cavity | cavity Inserts | cavity Inserts | [5 empty] | Visual
2OP638LI/Lid_13D | 2OP638LI | Lid_13D | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection
2OP638LI/Lid_13D | 2OP638LI | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual
2OP638LI/Lid_13D | 2OP638LI | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | Bead Diameter | 4.8615 | in | FDA,3 | 4.847 | 4.867 | TM 143 / Visual
2OP638LI/Lid_13D | 2OP638LI | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Center to Center Distance | Center to Center Distance | 0.4535 | in | FDA,3 | 0.445 | 0.455 | TM 143 / Visual
2OP638LI/Lid_13D | 2OP638LI | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Translucent Gray (Done) |  |  |  |  | TM 143 / Visual
2OP638LI/Lid_13D | 2OP638LI | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | End Circle Diameter | End Circle Diameter | 0.122 | in | FDA,3 | 0.12 | 0.13 | TM 143 / Visual
2OP638LI/Lid_13D | 2OP638LI | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) |  |  |  |  | TM 143 / Visual
2OP638LI/Lid_13D | 2OP638LI | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Insert Fin Gap (per specification) | Insert Fin Gap | 0.0745 | in | FDA,3 | 0.07 | 0.08 | TM 143 / Visual
2OP638LI/Lid_13D | 2OP638LI | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter (as per specification) | Lid Inside Diameter | 4.9218 | in | FDA,3 | 4.917 | 4.937 | TM 143 / Visual
2OP638LI/Lid_13D | 2OP638LI | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.041 | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual
2OP638LI/Lid_13D | 2OP638LI | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual
2OP638LI/Lid_13D | 2OP638LI | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness (per specification) | Orifice Flap Thickness | 0.0345 | in | FDA,3 | 0.03 | 0.04 | TM 143 / Visual
2OP638LI/Lid_13D | 2OP638LI | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.487 | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual
2OP638LI/Lid_13D | 2OP638LI | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 33.92 | g | FDA,1 | 32 | 42 | TM 143 / Visual
2OP638LI/Lid_13D | 2OP638LI | Lid_13D | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | Flip Lid Force | 1.79 | lb | FDA,2 | 0.75 |  | TM 148
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid has a tight friction fit |  |  |  |  | Canister & Lid Inspection
2OQ753/Skip Lot Release_NY_2OQ753 | 2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid has a tight friction fit |  |  |  |  | Canister & Lid Inspection
2OQ753/Skip Lot Release_NY_2OQ753 | 2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis
2OQ753/Skip Lot Release_NY_2OQ753 | 2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection
2OQ753/Skip Lot Release_NY_2OQ753 | 2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | White (Done) |  |  |  |  | Canister & Lid Inspection
2OQ753/Skip Lot Release_NY_2OQ753 | 2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | White (Done) |  |  |  |  | Canister & Lid Inspection
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Bead OD | Bead OD | 4.9355 | in | FDA,3 | 4.908 | 4.944 | Canister & Lid Inspection
2OQ753/Skip Lot Release_NY_2OQ753 | 2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Bead OD | Bead OD | 4.911 | in | FDA,3 | 4.908 | 4.944 | Canister & Lid Inspection
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Canister OD | Canister OD | 5.0145 | in | FDA,3 | 4.998 | 5.034 | Canister & Lid Inspection
2OQ753/Skip Lot Release_NY_2OQ753 | 2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Canister OD | Canister OD | 4.998 | in | FDA,3 | 4.998 | 5.034 | Canister & Lid Inspection
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Lip (width) | Lip Width | 0.17 | in | FDA,3 | 0.142 | 0.178 | Canister & Lid Inspection
2OQ753/Skip Lot Release_NY_2OQ753 | 2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Lip (width) | Lip Width | 0.15 | in | FDA,3 | 0.142 | 0.178 | Canister & Lid Inspection
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Total Height OD | Total Height OD | 6.748 | in | FDA,3 | 6.706 | 6.826 | Canister & Lid Inspection
2OQ753/Skip Lot Release_NY_2OQ753 | 2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Total Height OD | Total Height OD | 6.825 | in | FDA,3 | 6.706 | 6.826 | Canister & Lid Inspection
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | HDPE Resin Identification Marker | Meets the Specification (Done) |  |  |  |  | Canister & Lid Inspection
2OQ753/Skip Lot Release_NY_2OQ753 | 2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | HDPE Resin Identification Marker | Meets the Specification (Done) |  |  |  |  | Canister & Lid Inspection
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | Canister & Lid Inspection
2OQ753/Skip Lot Release_NY_2OQ753 | 2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | Canister & Lid Inspection
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight | 69.21 | g | FDA,1 | 67.5 | 72.5 | Measurement
2OQ753/Skip Lot Release_NY_2OQ753 | 2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight | 68.86 | g | FDA,1 | 67.5 | 72.5 | Measurement
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 766 | Lid Pull Force | Lid Pull-Off Force | Lid Pull-Off Force | 2859.39 | g | FDA,0 | 2000 | 3500 | TM 148
2OQ753/Skip Lot Release_NY_2OQ753 | 2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 766 | Lid Pull Force | Lid Pull-Off Force | Lid Pull-Off Force | 2979.67 | g | FDA,0 | 2000 | 3500 | TM 148
2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 2OQ753X | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid has a tight friction fit |  |  |  |  | Canister & Lid Inspection
2OQ753X/Skip Lot Relrase_NY_2OQ753X | 2OQ753X | Skip Lot Relrase_NY_2OQ753X | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA
2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 2OQ753X | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA
2OQ753X/Skip Lot Release_NY_2OQ850X | 2OQ753X | Skip Lot Release_NY_2OQ850X | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA
2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 2OQ753X | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Canister & Lid Inspection
2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 2OQ753X | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection
2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 2OQ753X | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | White (Done) |  |  |  |  | Canister & Lid Inspection
2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 2OQ753X | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | HDPE Resin Identification Marker | Meets the Specification (Done) |  |  |  |  | Canister & Lid Inspection
2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 2OQ753X | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | Canister & Lid Inspection
2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 2OQ753X | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight | 70.655 | g | FDA,1 | 66.5 | 73.5 | Measurement
2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 2OQ753X | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 1052 | Top Load Compression Strength | Top load Compression Strength (0.5"/mSR,0.450"TD) | Top load Compression Strength (0.5"/mSR,0.450"TD) |  | lb | FDA,0 | 84 |  | NPTM 1019
2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 2OQ753X | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 1053 | Cavity | cavity Inserts | cavity Inserts | Meets specification (Done) |  |  |  |  | Visual
2OQ850/Reduced Testing_NY_Can 2OQ850X | 2OQ850 | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid has a tight friction fit |  |  |  |  | Canister & Lid Inspection
2OQ850/Reduced Testing_NY_Can 2OQ850 | 2OQ850 | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid has a tight friction fit |  |  |  |  | Canister & Lid Inspection
2OQ850/Skip Lot Release_NY_2OQ850 | 2OQ850 | Skip Lot Release_NY_2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid has a tight friction fit |  |  |  |  | Canister & Lid Inspection
2OQ850/Reduced Testing_NY_Can 2OQ850 | 2OQ850 | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA
2OQ850/Skip Lot Release_NY_2OQ850 | 2OQ850 | Skip Lot Release_NY_2OQ850 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis
2OQ850/Reduced Testing_NY_Can 2OQ850 | 2OQ850 | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis
2OQ850/Reduced Testing_NY_Can 2OQ850X | 2OQ850 | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection
2OQ850/Reduced Testing_NY_Can 2OQ850X | 2OQ850 | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection
2OQ850/Reduced Testing_NY_Can 2OQ850 | 2OQ850 | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection
2OQ850/Skip Lot Release_NY_2OQ850 | 2OQ850 | Skip Lot Release_NY_2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection
2OQ850/Reduced Testing_NY_Can 2OQ850X | 2OQ850 | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | White (Done) |  |  |  |  | Canister & Lid Inspection
2OQ850/Reduced Testing_NY_Can 2OQ850 | 2OQ850 | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | White (Done) |  |  |  |  | Canister & Lid Inspection
2OQ850/Skip Lot Release_NY_2OQ850 | 2OQ850 | Skip Lot Release_NY_2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | White (Done) |  |  |  |  | Canister & Lid Inspection
2OQ850/Reduced Testing_NY_Can 2OQ850X | 2OQ850 | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | HDPE Resin Identification Marker | Meets the Specification (Done) |  |  |  |  | Canister & Lid Inspection
2OQ850/Reduced Testing_NY_Can 2OQ850 | 2OQ850 | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | HDPE Resin Identification Marker | Meets the Specification (Done) |  |  |  |  | Canister & Lid Inspection
2OQ850/Skip Lot Release_NY_2OQ850 | 2OQ850 | Skip Lot Release_NY_2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | HDPE Resin Identification Marker | Meets the Specification (Done) |  |  |  |  | Canister & Lid Inspection
2OQ850/Reduced Testing_NY_Can 2OQ850X | 2OQ850 | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | Canister & Lid Inspection
2OQ850/Reduced Testing_NY_Can 2OQ850 | 2OQ850 | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | Canister & Lid Inspection
2OQ850/Skip Lot Release_NY_2OQ850 | 2OQ850 | Skip Lot Release_NY_2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | Canister & Lid Inspection
2OQ850/Reduced Testing_NY_Can 2OQ850X | 2OQ850 | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight | 88.41 | g | FDA,1 | 85.5 | 94.5 | Measurement
2OQ850/Reduced Testing_NY_Can 2OQ850 | 2OQ850 | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight | 88.41 | g | FDA,1 | 87 | 93 | Measurement
2OQ850/Skip Lot Release_NY_2OQ850 | 2OQ850 | Skip Lot Release_NY_2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight | 89.99 | g | FDA,1 | 87 | 93 | Measurement
2OQ850/Reduced Testing_NY_Can 2OQ850X | 2OQ850 | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 1052 | Top Load Compression Strength | Top load Compression Strength (0.5"/mSR,0.450"TD) | Top load Compression Strength (0.5"/mSR,0.450"TD) |  | lb | FDA,0 | 97 |  | NPTM 1019
2OQ850/Reduced Testing_NY_Can 2OQ850 | 2OQ850 | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 766 | Lid Pull Force | Lid Pull-Off Force | Lid Pull-Off Force |  | g | FDA,0 | 2000 | 3500 | TM 148
2OQ850/Skip Lot Release_NY_2OQ850 | 2OQ850 | Skip Lot Release_NY_2OQ850 | NY-Analytical | 766 | Lid Pull Force | Lid Pull-Off Force | Lid Pull-Off Force | 2719.98 | g | FDA,0 | 2000 | 3500 | TM 148
2OQ850/Reduced Testing_NY_Can 2OQ850X | 2OQ850 | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 1053 | Cavity | cavity Inserts | cavity Inserts | [5 empty] | Visual
2OQ850LI/Lid_11B | 2OQ850LI | Lid_11B | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Flip Lid Force
2OQ850LI/Lid_11B | 2OQ850LI | Lid_11B | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | Flip Lid Force | 2.41 | lb | FDA,2 | 0.75 |  | TM 143
2OQ850LI/Lid_11B | 2OQ850LI | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual
2OQ850LI/Lid_11B | 2OQ850LI | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | Bead Diameter | 4.861 | in | FDA,3 | 4.847 | 4.867 | TM 143 / Visual
2OQ850LI/Lid_11B | 2OQ850LI | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Translucent Red (Done) |  |  |  |  | TM 143 / Visual
2OQ850LI/Lid_11B | 2OQ850LI | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) |  |  |  |  | TM 143 / Visual
2OQ850LI/Lid_11B | 2OQ850LI | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Insert Fin Gap (per specification) | Insert Fin Gap | 0.073 | in | FDA,3 | 0.065 | 0.075 | TM 143 / Visual
2OQ850LI/Lid_11B | 2OQ850LI | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter (as per specification) | Lid Inside Diameter | 4.919 | in | FDA,3 | 4.917 | 4.937 | TM 143 / Visual
2OQ850LI/Lid_11B | 2OQ850LI | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.0425 | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual
2OQ850LI/Lid_11B | 2OQ850LI | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual
2OQ850LI/Lid_11B | 2OQ850LI | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness (per specification) | Orifice Flap Thickness | 0.024 | in | FDA,3 | 0.018 | 0.028 | TM 143 / Visual
2OQ850LI/Lid_11B | 2OQ850LI | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.488 | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual
2OQ850LI/Lid_11B | 2OQ850LI | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Outer Diameter (per specification) | Pop-Up Hole Outer Diameter | 0.165 | in | FDA,3 | 0.16 | 0.17 | TM 143 / Visual
2OQ850LI/Lid_NY_Skip lot testing | 2OQ850LI | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Outer Diameter (per specification) | Pop-Up Hole Outer Diameter | 0.163 | in | FDA,3 | 0.16 | 0.17 | TM 143 / Visual
2OQ850LI/Lid_11B | 2OQ850LI | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 34.04 | g | FDA,1 | 33 | 42 | TM 143 / Visual
2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | 2OQ850X | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 |  |  | TM 141 / Visual
2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | 2OQ850X | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | [5 empty] | TM 141 / Visual
2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | 2OQ850X | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | 2OQ850X | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | [5 empty] | TM 141 / Visual
2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | 2OQ850X | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height |  | in | FDA,4 | 4.0625 | 4.3125 | TM 141 / Visual
2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | 2OQ850X | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  | in | FDA,4 | 12.3125 | 12.5625 | TM 141 / Visual
2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | 2OQ850X | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | [5 empty] | TM 141 / Visual
2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | 2OQ850X | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | [5 empty] | TM 141 / Visual
2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | 2OQ850X | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  | in | FDA,4 | 5.3125 | 5.5625 | TM 141 / Visual
2OQ882LI/Lid_12A | 2OQ882LI | Lid_12A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Flip Lid Force
2OQ882LI/Lid_12A | 2OQ882LI | Lid_12A | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | Flip Lid Force | 1.61 | lb | FDA,2 | 0.75 |  | TM 143
2OQ882LI/Lid_12A | 2OQ882LI | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual
2OQ882LI/Lid_12A | 2OQ882LI | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | Bead Diameter | 4.854 | in | FDA,3 | 4.852 | 4.862 | TM 143 / Visual
2OQ882LI/Lid_12A | 2OQ882LI | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Cap Thickness | Cap Thickness | 0.048 | in | FDA,3 | 0.045 | 0.055 | TM 143 / Visual
2OQ882LI/Lid_12A | 2OQ882LI | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Dark Green (Done) |  |  |  |  | TM 143 / Visual
2OQ882LI/Lid_12A | 2OQ882LI | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) |  |  |  |  | TM 143 / Visual
2OQ882LI/Lid_12A | 2OQ882LI | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Insert Fin Gap (per specification) | Insert Fin Gap | 0.074 | in | FDA,3 | 0.065 | 0.075 | TM 143 / Visual
2OQ882LI/Lid_12A | 2OQ882LI | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.045 | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual
2OQ882LI/Lid_12A | 2OQ882LI | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual
2OQ882LI/Lid_12A | 2OQ882LI | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness (per specification) | Orifice Flap Thickness | 0.021 | in | FDA,3 | 0.018 | 0.028 | TM 143 / Visual
2OQ882LI/Lid_12A | 2OQ882LI | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.488 | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual
2OQ882LI/Lid_12A | 2OQ882LI | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Outer Diameter (per specification) | Pop-Up Hole Outer Diameter | 0.168 | in | FDA,3 | 0.16 | 0.17 | TM 143 / Visual
2OQ882LI/Lid_12A | 2OQ882LI | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 35.675 | g | FDA,1 | 32 | 42 | TM 143 / Visual
2OU130/Web_NP5115_9.0 | 2OU130 | Web_NP5115_9.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2OU130/Web_NP5115_9.0 | 2OU130 | Web_NP5115_9.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 8.9375 | in | FDA,4 | 8.9375 | 9.0625 | Measurement
2OU130/Web_NP5115_9.0 | 2OU130 | Web_NP5115_9.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Printed/Varnished) | Basis Weight | 60.8 | #/Ream | FDA,1 | 55.6 | 68 | TM 101
2OU130/Web_NP5115_9.0 | 2OU130 | Web_NP5115_9.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@325DegF/40PSI/0.5Sec_Dwell) | Seal Strength (@325°F/40PSI/0.5Sec Dwell) | 2198.05 | g/in | FDA,0 | 1500 |  | TM 107
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | NY-Analytical | 765 | Canister & Lid Inspection | Bottle Internal Diameter at Flange | Bottle Internal Diameter at Flange | 0.722 | in | FDA,3 | 0.709 | 0.729 | Canister & Lid Inspection
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Natural (Done) |  |  |  |  | Canister & Lid Inspection
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | NY-Analytical | 765 | Canister & Lid Inspection | Overall Length | Overall Length | 4.855 | in | FDA,3 | 4.825 | 4.875 | Canister & Lid Inspection
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | NY-Analytical | 765 | Canister & Lid Inspection | Seal Flange Diameter | Seal Flange Diameter | 0.8735 | in | FDA,3 | 0.864 | 0.874 | Canister & Lid Inspection
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 2 | g | FDA,2 |  |  | TM 610
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer |  |  |  |  | TM 610
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 1 | cfu/device | FDA,1 |  |  | TM 610
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0.0028852 | CFU/g | FDA,1 |  | 15 | TM 610
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Negative Rods (OOS) |  |  |  |  | WI QAS 076
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Mold (OOS) |  |  |  |  | WI QAS 076
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Certificate of Conformance
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 765 | Canister & Lid Inspection | Bottle Internal Diameter at Flange | Bottle Internal Diameter at Flange | 0.4425 | in | FDA,3 | 0.442 | 0.452 | Canister & Lid Inspection
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | NY-Analytical | 765 | Canister & Lid Inspection | Bottle Internal Diameter at Flange | Bottle Internal Diameter at Flange | 0.443 | in | FDA,3 | 0.442 | 0.452 | Canister & Lid Inspection
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Natural (Done) |  |  |  |  | Canister & Lid Inspection
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Natural (Done) |  |  |  |  | Canister & Lid Inspection
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 765 | Canister & Lid Inspection | Overall Length | Overall Length | 2.8805 | in | FDA,3 | 2.848 | 2.888 | Canister & Lid Inspection
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | NY-Analytical | 765 | Canister & Lid Inspection | Overall Length | Overall Length | 2.8737 | in | FDA,3 | 2.848 | 2.888 | Canister & Lid Inspection
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 765 | Canister & Lid Inspection | Seal Flange Diameter | Seal Flange Diameter | 0.55 | in | FDA,3 | 0.544 | 0.554 | Canister & Lid Inspection
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | NY-Analytical | 765 | Canister & Lid Inspection | Seal Flange Diameter | Seal Flange Diameter | 0.55 | in | FDA,3 | 0.544 | 0.554 | Canister & Lid Inspection
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight |  | g | FDA,2 |  |  | TM 610
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 4.5 | g | FDA,2 |  |  | TM 610
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 1.1 | cfu/device | FDA,1 |  |  | TM 610
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 100 | TM 610
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0.24444 | CFU/g | FDA,1 |  | 100 | TM 610
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 1.1 | cfu/device | FDA,1 |  |  | TM 610
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 1 | cfu/device | FDA,1 |  |  | TM 610
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0.015027 | CFU/g | FDA,1 |  | 15 | TM 610
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0.022222 | CFU/g | FDA,1 |  | 15 | TM 610
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 1 | cfu | FDA,0 |  |  | TM 610
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Positive Rods (Done) |  |  |  |  | WI QAS 076
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Positive Rods (Done) |  |  |  |  | WI QAS 076
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Mold (OOS) |  |  |  |  | WI QAS 076
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Mold (OOS) |  |  |  |  | WI QAS 076
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil seal 5.1 ml | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil seal 5.1 ml | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 1.78 | g | FDA,2 |  |  | TM 610
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil seal 5.1 ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil seal 5.1 ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil seal 5.1 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 11.6 | cfu/device | FDA,1 |  |  | TM 610
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil seal 5.1 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 6.5169 | CFU/g | FDA,1 |  | 100 | TM 610
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil seal 5.1 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 11 | cfu | FDA,0 |  |  | TM 610
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil seal 5.1 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil seal 5.1 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil seal 5.1 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil seal 5.1 ml | NY-Analytical | 1056 | Functionality | Foil diameter to fit Bottle(Item#2OX510B)opening | Foil diameter to fit Bottle(Item#2OX510B)opening | Pass |  |  |  |  | Visual
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil seal 5.1 ml | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Negative Rods (OOS) |  |  |  |  | WI QAS 076
2PD351/Web_NP5116_4.5 | 2PD351 | Web_NP5116_4.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2PD351/Web_NP5116_4.5 | 2PD351 | Web_NP5116_4.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 4.5 | in | FDA,4 | 4.4375 | 4.5625 | Measurement
2PD351/Web_NP5116_4.5 | 2PD351 | Web_NP5116_4.5 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Printed/Varnished) | Basis Weight | 59.82 | #/Ream | FDA,1 | 55.8 | 68.2 | TM 101
2PD351/Web_NP5116_4.5 | 2PD351 | Web_NP5116_4.5 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/40psi/0.5 sec dewll) | Seal Strength (@300°F/40psi/0.5 sec dewll) | 2163 | g/in | FDA,0 | 1500 |  | TM 107
2QB114X/Web_NP3860_5.0 | 2QB114X | Web_NP3860_5.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2QB114X/Web_NP3860_5.0 | 2QB114X | Web_NP3860_5.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 5 | in | FDA,4 | 4.9375 | 5.0625 | Measurement
2QB114X/Web_NP3860_5.0 | 2QB114X | Web_NP3860_5.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 89.29 | #/Ream | FDA,1 | 78.9 | 96.5 | TM 101
2QB114X/Web_NP3860_5.0 | 2QB114X | Web_NP3860_5.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@350DegF/40psi/1.0 sec dwell) | Seal Strength (@350°F/40psi/1.0 sec dwell) | 9928 | g/in | FDA,0 | 7250 |  | TM 107
2QB114X/Web_NP3860_5.0 | 2QB114X | Web_NP3860_5.0 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 5 | mil | FDA,1 | 4.7 | 5.7 | TM 132
2RQ890/Web_NP9508_18.125 | 2RQ890 | Web_NP9508_18.125 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2RQ890/Web_NP9508_18.125 | 2RQ890 | Web_NP9508_18.125 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 18.125 | in | FDA,4 | 18.0625 | 18.1875 | Measurement
2RQ890/Web_NP9508_18.125 | 2RQ890 | Web_NP9508_18.125 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Unprinted) | Basis Weight |  | #/Ream | FDA,1 | 37.8 | 46.2 | TM 101
2RQ890/Web_NP9508_18.125 | 2RQ890 | Web_NP9508_18.125 | NY-Analytical | 762 | Seal Strength | Seal Strength (@300DegF/40psi/0.5 sec dewll) | Seal Strength (@300°F/40psi/0.5 sec dewll) | Sample peeled apart and showed stress whitening when sealed (Done) |  |  |  |  | TM 107
2RQ890/Web_NP9508_18.125 | 2RQ890 | Web_NP9508_18.125 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 2.6 | mil | FDA,1 | 2.3 | 2.8 | TM 132
2RS324/Web_NP3860_12.5 | 2RS324 | Web_NP3860_12.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2RS324/Web_NP3860_12.5 | 2RS324 | Web_NP3860_12.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 12.5 | in | FDA,4 | 12.4375 | 12.5625 | Measurement
2RS324/Web_NP3860_12.5 | 2RS324 | Web_NP3860_12.5 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 89 | #/Ream | FDA,1 | 78.9 | 96.5 | TM 101
2RS324/Web_NP3860_12.5 | 2RS324 | Web_NP3860_12.5 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@350DegF/40psi/1.0 sec dwell) | Seal Strength (@350°F/40psi/1.0 sec dwell) | 9622 | g/in | FDA,0 | 7250 |  | TM 107
2RS324/Web_NP3860_12.5 | 2RS324 | Web_NP3860_12.5 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 4.91 | mil | FDA,1 | 4.7 | 5.7 | TM 132
2RS407/Web_NP3885_13.0 | 2RS407 | Web_NP3885_13.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2RS407/Web_NP3885_13.0 | 2RS407 | Web_NP3885_13.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 12.9375 | in | FDA,4 | 12.9375 | 13.0625 | Measurement
2RS407/Web_NP3885_13.0 | 2RS407 | Web_NP3885_13.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 80.69 | #/Ream | FDA,0 | 74 | 90 | TM 101
2RS407/Web_NP3885_13.0 | 2RS407 | Web_NP3885_13.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@350DegF/50psi/1.0 sec dwell) | Seal Strength (@350°F/50psi/1.0 sec dwell) | 1941 | g/in | FDA,0 | 700 |  | TM 107
2RS419/Web_NP3885_16 | 2RS419 | Web_NP3885_16 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2RS419/Web_NP3885_16 | 2RS419 | Web_NP3885_16 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 16 | in | FDA,4 | 15.9375 | 16.0625 | Measurement
2RS419/Web_NP3885_16 | 2RS419 | Web_NP3885_16 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 81.85 | #/Ream | FDA,0 | 74 | 90 | TM 101
2RS419/Web_NP3885_16 | 2RS419 | Web_NP3885_16 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@350DegF/50psi/1.0 sec dwell) | Seal Strength (@350°F/50psi/1.0 sec dwell) | 2030.5 | g/in | FDA,0 | 700 |  | TM 107
2SS273/Web_NP3860_15.5 | 2SS273 | Web_NP3860_15.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement
2SS273/Web_NP3860_15.5 | 2SS273 | Web_NP3860_15.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 15.5 | in | FDA,4 | 15.4375 | 15.5625 | Measurement
2SS273/Web_NP3860_15.5 | 2SS273 | Web_NP3860_15.5 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 87.7 | #/Ream | FDA,1 | 78.9 | 96.5 | TM 101
2SS273/Web_NP3860_15.5 | 2SS273 | Web_NP3860_15.5 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@350DegF/40psi/1.0 sec dwell) | Seal Strength (@350°F/40psi/1.0 sec dwell) | 10111.34 | g/in | FDA,0 | 7250 |  | TM 107
2SS273/Web_NP3860_15.5 | 2SS273 | Web_NP3860_15.5 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 5.2 | mil | FDA,1 | 4.7 | 5.7 | TM 132
2SS428/Web_NP3860_8.0 | 2SS428 | Web_NP3860_8.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Basis Weight
2SS428/Web_NP3860_8.0 | 2SS428 | Web_NP3860_8.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 87.36 | #/Ream | FDA,1 | 78.9 | 96.5 | TM 101
2SS428/Web_NP3860_8.0 | 2SS428 | Web_NP3860_8.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@350DegF/40psi/1.0 sec dwell) | Seal Strength (@350°F/40psi/1.0 sec dwell) | 9981.29 | g/in | FDA,0 | 7250 |  | TM 107
2SS428/Web_NP3860_8.0 | 2SS428 | Web_NP3860_8.0 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 5.2 | mil | FDA,1 | 4.7 | 5.7 | TM 132
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 44 | in | FDA,0 | 44 | 46 | Applicator Measurements
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 37.5 | in | FDA,1 | 37.3 | 37.7 | Applicator Measurements
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 55.54 | GSM | FDA,1 | 49.5 | 60.5 | TM 101
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 1147.63 | g/in | FDA,0 | 692 | 1438 | TM 103
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 4131.08 | g/in | FDA,0 | 2785 | 5785 | TM 103
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 748 | Visual Inspection | Color | Color | White to Off-White (Done) |  |  |  |  | TM 208
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 509 | J&J_Pathogens | Biggy or SDA | Biggy or SDA | No Growth_SDA |  |  |  |  | TM 869
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 509 | J&J_Pathogens | Cetrimide | Cetrimide | No Growth |  |  |  |  | TM 869
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 524 | J&J Fungal | JJ Fungal_RM | JJ Fungal RM | 0 | CFU/g |  |  | 200 | TM 869
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 509 | J&J_Pathogens | MacConkey | MacConkey | No Growth |  |  |  |  | TM 869
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 509 | J&J_Pathogens | MSA or VJ | MSA or VJ | No Growth_MSA |  |  |  |  | TM 869
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 700 | J&J SDA Controls | Negative Control | Negative Control | No Growth SDA |  |  |  |  | TM 869
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | NegativeControl | No Growth TSB |  |  |  |  | TM 869
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 509 | J&J_Pathogens | TSA | TSA | No Growth |  |  |  |  | TM 869
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ RM | TSA JJ RM | 0 | CFU/g |  |  | 200 | TM 869
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 517 | J&J_Enrichment | TSB | TSB | Not Turbid |  |  |  |  | TM 869
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 573 | GramStain_Candida | Gram Stain_J&JCandida | Gram Stain J&JCandida | Non-Subculturable Growth (Done) |  |  |  |  | WI QAS 076
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 545 | GramStain_J&J_TSA | GramStain_J&J_TSA | GramStain J&J TSA | Gram Positive Rods (Done) |  |  |  |  | WI QAS 076
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 3AQ470-A | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading |  |  |  |  | Color (TM 208)
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 3AQ470-A | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Free of hard wrinkles. Formation uniform (Done) |  |  |  |  | See Analyte for Method #
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 3AQ470-A | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | See Analyte for Method #
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 3AQ470-A | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 49.51 | GSM | FDA,1 | 45 | 55 | See Analyte for Method #
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 3AQ470-A | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (mm) (CofA) | Caliper | 0.59 | mm | FDA,2 | 0.51 | 0.71 | See Analyte for Method #
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 3AQ470-A | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method #
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 3AQ470-A | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | May have slight odor when wet (Done) |  |  |  |  | See Analyte for Method #
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 3AQ470-A | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter | 50 | in | FDA,0 | 45 | 60 | See Analyte for Method #
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 3AQ470-A | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Width (Inches) (Measurement) | Roll Width | 37.75 | in | FDA,2 | 37.5 | 38 | See Analyte for Method #
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 3AQ470-A | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Wet CD (g/in) (TM 103) | Wet CD | 1634 | g/in | FDA,0 | 1100 | 2300 | See Analyte for Method #
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 3AQ470-A | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Wet MD (g/in) (TM 103) | Wet MD | 6147 | g/in | FDA,0 | 3000 | 7000 | See Analyte for Method #
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 30 | in | FDA,0 | 28 | 32 | Applicator Measurements
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 7.5 | in | FDA,3 | 7.375 | 7.625 | Applicator Measurements
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 50.29 | GSM | FDA,1 | 49.5 | 60.5 | TM 101
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper | 0.52 | mm | FDA,2 | 0.5 | 0.64 | TM 102
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 1925.71 | g/in | FDA,0 | 1443 | 2463 | TM 103
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 6181.64 | g/in | FDA,0 | 4883 | 7967 | TM 103
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 1.03 | g | FDA,2 |  |  | TM 610
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer |  |  |  |  | TM 610
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610
3AU228/Applicator_06I | 3AU228 | Applicator_06I | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | REDUCED TESTING_NY_3BB406 | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading |  |  |  |  | Color (TM 208)
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | REDUCED TESTING_NY_3BB406 | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | May have residual sulfur odor when wet |  |  |  |  | Odor (TM 208)
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | REDUCED TESTING_NY_3BB406 | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Meets specifications (Done) |  |  |  |  | See Analyte for Method #
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | REDUCED TESTING_NY_3BB406 | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 200 |  | FDA,0 |  |  | See Analyte for Method #
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | REDUCED TESTING_NY_3BB406 | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 52.74 | GSM | FDA,1 | 48.9 | 56.3 | See Analyte for Method #
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | REDUCED TESTING_NY_3BB406 | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method #
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight |  | g | FDA,2 |  |  | TM 610
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer |  |  |  |  | TM 610
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 100 | TM 610
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 1 | cfu/device | FDA,1 |  |  | TM 610
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 15 | TM 610
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  | cfu | FDA,0 |  |  | TM 610
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Mold (OOS) |  |  |  |  | WI QAS 076
3BJ141/Applicator_07K | 3BJ141 | Applicator_07K | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements
3BJ141/Applicator_07K | 3BJ141 | Applicator_07K | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 11.625 | in | FDA,2 | 11.5 | 11.75 | Applicator Measurements
3BJ141/Applicator_07K | 3BJ141 | Applicator_07K | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 41.33 | GSM | FDA,1 | 38.7 | 47.3 | TM 101
3BJ141/Applicator_07K | 3BJ141 | Applicator_07K | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(inch) | Caliper | 0.0225 | in | FDA,3 | 0.022 | 0.03 | TM 102
3BJ141/Applicator_07K | 3BJ141 | Applicator_07K | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 1064.23 | g/in | FDA,0 | 921 | 1535 | TM 103
3BJ141/Applicator_07K | 3BJ141 | Applicator_07K | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 5060.98 | g/in | FDA,0 | 3623 | 5650 | TM 103
3BJ141/Applicator_07K | 3BJ141 | Applicator_07K | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208
3BJ141/Applicator_07K | 3BJ141 | Applicator_07K | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208
3BJ141/Applicator_07K | 3BJ141 | Applicator_07K | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208
3DQ595/Applicator_10D | 3DQ595 | Applicator_10D | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements
3DQ595/Applicator_10D | 3DQ595 | Applicator_10D | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 46 | in | FDA,0 | 44 | 47 | Applicator Measurements
3DQ595/Applicator_10D | 3DQ595 | Applicator_10D | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 30 | in | FDA,3 | 29.875 | 30.125 | Applicator Measurements
3DQ595/Applicator_10D | 3DQ595 | Applicator_10D | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 39.18 | GSM | FDA,1 | 38 | 42 | TM 101
3DQ595/Applicator_10D | 3DQ595 | Applicator_10D | NY-Analytical | 757 | Caliper | Caliper-1_Ply_(mm) | Caliper | 0.36 | mm | FDA,2 | 0.35 | 0.45 | TM 102
3DQ595/Applicator_10D | 3DQ595 | Applicator_10D | NY-Analytical | 758 | Tensile Strength | Dry_CD_(g/in) | Dry CD | 667.15 | g/in | FDA,0 | 513 | 1130 | TM 103
3DQ595/Applicator_10D | 3DQ595 | Applicator_10D | NY-Analytical | 758 | Tensile Strength | Dry_MD_(g/in) | Dry CD | 2695.76 | g/in | FDA,0 | 2054 | 4108 | TM 103
3DQ595/Applicator_10D | 3DQ595 | Applicator_10D | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208
3DQ595/Applicator_10D | 3DQ595 | Applicator_10D | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208
3DQ595/Applicator_10D | 3DQ595 | Applicator_10D | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208
3DQ595N/Applicator_10I | 3DQ595N | Applicator_10I | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements
3DQ595N/Applicator_10I | 3DQ595N | Applicator_10I | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 46 | in | FDA,0 | 44 | 47 | Applicator Measurements
3DQ595N/Applicator_10I | 3DQ595N | Applicator_10I | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 30 | in | FDA,3 | 29.875 | 30.125 | Applicator Measurements
3DQ595N/Applicator_10I | 3DQ595N | Applicator_10I | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 39.83 | GSM | FDA,0 | 38 | 42 | TM 101
3DQ595N/Applicator_10I | 3DQ595N | Applicator_10I | NY-Analytical | 757 | Caliper | Caliper-1_Ply_(mm) | Caliper | 0.41 | mm | FDA,2 | 0.35 | 0.45 | TM 102
3DQ595N/Applicator_10I | 3DQ595N | Applicator_10I | NY-Analytical | 758 | Tensile Strength | Dry_CD_(g/in) | Dry CD | 573.11 | g/in | FDA,0 | 513 | 1130 | TM 103
3DQ595N/Applicator_10I | 3DQ595N | Applicator_10I | NY-Analytical | 758 | Tensile Strength | Dry_MD_(g/in) | Dry CD | 3431.99 | g/in | FDA,0 | 2054 | 4108 | TM 103
3DQ595N/Applicator_10I | 3DQ595N | Applicator_10I | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208
3DQ595N/Applicator_10I | 3DQ595N | Applicator_10I | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208
3DQ595N/Applicator_10I | 3DQ595N | Applicator_10I | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading |  |  |  |  | Color (TM 208)
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading |  |  |  |  | Color (TM 208)
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | May have residual sulfur odor when wet |  |  |  |  | Odor (TM 208)
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | May have residual sulfur odor when wet |  |  |  |  | Odor (TM 208)
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator_Inspection
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Smooth free of hard wrinkles, formation uniform matches embossing pattern (Done) |  |  |  |  | See Analyte for Method #
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Smooth free of hard wrinkles, formation uniform matches embossing pattern (Done) |  |  |  |  | See Analyte for Method #
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | See Analyte for Method #
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | See Analyte for Method #
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 61.79 | GSM | FDA,1 | 60 | 70 | See Analyte for Method #
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 65.66 | GSM | FDA,1 | 60 | 70 | See Analyte for Method #
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (mm) (TM 102) | Caliper | 0.64 | mm | FDA,2 | 0.54 | 0.66 | See Analyte for Method #
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (mm) (TM 102) | Caliper | 0.66 | mm | FDA,2 | 0.54 | 0.66 | See Analyte for Method #
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Ratio: Wet MD / Wet CD (TM 103) | Ratio: Wet MD / Wet CD | 0 |  | FDA,1 |  | 4.5 | See Analyte for Method #
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Ratio: Wet MD / Wet CD (TM 103) | Ratio: Wet MD / Wet CD | 0 |  | FDA,1 |  | 4.5 | See Analyte for Method #
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter | 45 | in | FDA,0 | 42 | 48 | See Analyte for Method #
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter | 43 | in | FDA,0 | 42 | 48 | See Analyte for Method #
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Width (mm) (Measurement) | Roll Width | 6.062 | in | FDA,2 | 6 | 6.125 | See Analyte for Method #
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Roll Width (mm) (Measurement) | Roll Width | 6 | in | FDA,2 | 6 | 6.125 | See Analyte for Method #
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Wet CD (g/in) (TM 103) | Wet CD | 416 | g/in | FDA,0 | 375 | 575 | See Analyte for Method #
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Wet CD (g/in) (TM 103) | Wet CD | 426 | g/in | FDA,0 | 375 | 575 | See Analyte for Method #
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Wet MD (g/in) (TM 103) | Wet MD | 692 | g/in | FDA,0 | 485 | 915 | See Analyte for Method #
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Wet MD (g/in) (TM 103) | Wet MD | 769 | g/in | FDA,0 | 485 | 915 | See Analyte for Method #
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 67.38 | GSM | FDA,1 | 60 | 70 | TM 101
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight |  | GSM | FDA,1 | 60 | 70 | TM 101
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper | 0.64 | mm | FDA,2 | 0.54 | 0.66 | TM 102
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper | 0.63 | mm | FDA,2 | 0.54 | 0.66 | TM 102
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 758 | Tensile Strength | Elongation % Wet CD | Elongation % Wet CD | 43.05 | % | FDA,0 | 15 | 65 | TM 103
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 758 | Tensile Strength | Elongation % Wet CD | Elongation % Wet CD | 25.08 | % | FDA,0 | 15 | 65 | TM 103
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 758 | Tensile Strength | Elongation % Wet MD | Elongation % Wet MD | 30.3 | % | FDA,0 | 2 | 50 | TM 103
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 758 | Tensile Strength | Elongation % Wet MD | Elongation % Wet MD | 17.02 | % | FDA,0 | 2 | 50 | TM 103
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 561.7 | g/in | FDA,0 | 375 | 575 | TM 103
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 495 | g/in | FDA,0 | 375 | 575 | TM 103
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 694.79 | g/in | FDA,0 | 485 | 915 | TM 103
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 648 | g/in | FDA,0 | 485 | 915 | TM 103
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 2.7 | g | FDA,2 |  |  | TM 610
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 1 | cfu/device | FDA,1 |  |  | TM 610
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0.96774 | CFU/g | FDA,1 |  | 15 | TM 610
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Positive Rods (Done) |  |  |  |  | WI QAS 076
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Mold (OOS) |  |  |  |  | WI QAS 076
3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | 3OA426 | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading |  |  |  |  | Color (TM 208)
3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | 3OA426 | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Free of hard wrinkles, formation uniform, large dot embossing pattern. (Done) |  |  |  |  | See Analyte for Method #
3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | 3OA426 | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 200 |  | FDA,0 |  |  | See Analyte for Method #
3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | 3OA426 | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 52.01 | GSM | FDA,0 | 47 | 57 | See Analyte for Method #
3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | 3OA426 | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (mm) (CofA) | Caliper | 0.79 | mm | FDA,2 | 0.74 | 0.83 | See Analyte for Method #
3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | 3OA426 | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method #
3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | 3OA426 | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | May have slight odor when wet (Done) |  |  |  |  | See Analyte for Method #
3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | 3OA426 | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter | 45 | in | FDA,0 | 43 | 48 | See Analyte for Method #
3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | 3OA426 | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Width (Inches) (Measurement) | Roll Width | 8 | in | FDA,3 | 7.875 | 8.125 | See Analyte for Method #
3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | 3OA426 | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Wet CD (g/in) (TM 103) | Wet CD | 1337.35 | g/in | FDA,0 | 1200 | 2100 | See Analyte for Method #
3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | 3OA426 | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Wet MD (g/in) (TM 103) | Wet MD | 5490.37 | g/in | FDA,0 | 3800 | 6270 | See Analyte for Method #
3OA580/Applicator_38A | 3OA580 | Applicator_38A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements
3OA580/Applicator_38A | 3OA580 | Applicator_38A | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 45 | in | FDA,0 | 43 | 47 | Applicator Measurements
3OA580/Applicator_38A | 3OA580 | Applicator_38A | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 11.5 | in | FDA,3 | 11.375 | 11.625 | Applicator Measurements
3OA580/Applicator_38A | 3OA580 | Applicator_38A | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 44.35 | GSM | FDA,0 | 40 | 50 | TM 101
3OA580/Applicator_38A | 3OA580 | Applicator_38A | NY-Analytical | 757 | Caliper | Caliper_at_Pt_of_Mfr_(micron) | Caliper | 500 | um | FDA,0 | 416 | 534 | TM 102
3OA580/Applicator_38A | 3OA580 | Applicator_38A | NY-Analytical | 758 | Tensile Strength | Dry_CD_CoA_(g/in) | Dry CD (CoA) | 1326 | g/in | FDA,0 | 610 |  | TM 103
3OA580/Applicator_38A | 3OA580 | Applicator_38A | NY-Analytical | 758 | Tensile Strength | Dry_MD_CoA_(g/in) | Dry MD (CoA) | 3086 | g/in | FDA,0 | 1825 |  | TM 103
3OA580/Applicator_38A | 3OA580 | Applicator_38A | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208
3OA580/Applicator_38A | 3OA580 | Applicator_38A | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208
3OA580/Applicator_38A | 3OA580 | Applicator_38A | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208
3OA580/Applicator_NY_Skip lot testing | 3OA580 | Applicator_NY_Skip lot testing | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 2.05 | g | FDA,2 |  |  | TM 610
3OA580/Applicator_NY_Skip lot testing | 3OA580 | Applicator_NY_Skip lot testing | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer |  |  |  |  | TM 610
3OA580/Applicator_NY_Skip lot testing | 3OA580 | Applicator_NY_Skip lot testing | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610
3OA580/Applicator_NY_Skip lot testing | 3OA580 | Applicator_NY_Skip lot testing | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OA580/Applicator_NY_Skip lot testing | 3OA580 | Applicator_NY_Skip lot testing | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610
3OA580/Applicator_NY_Skip lot testing | 3OA580 | Applicator_NY_Skip lot testing | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OA580/Applicator_NY_Skip lot testing | 3OA580 | Applicator_NY_Skip lot testing | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OA580/Applicator_NY_Skip lot testing | 3OA580 | Applicator_NY_Skip lot testing | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610
3OA580/Applicator_NY_Skip lot testing | 3OA580 | Applicator_NY_Skip lot testing | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OA580/Applicator_NY_Skip lot testing | 3OA580 | Applicator_NY_Skip lot testing | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Mold (OOS) |  |  |  |  | WI QAS 076
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 45 | in | FDA,0 | 45 | 48 | Applicator Measurements
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 6 | in | FDA,3 | 6 | 6.125 | Applicator Measurements
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 69.3 | GSM | FDA,0 | 65 | 75 | TM 101
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper | 0.6 | mm | FDA,2 | 0.54 | 0.66 | TM 102
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 387.69 | g/in | FDA,0 | 300 | 500 | TM 103
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 793.49 | g/in | FDA,0 | 610 | 1040 | TM 103
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 2.76 | g | FDA,2 |  |  | TM 610
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer |  |  |  |  | TM 610
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 7.9 | in | FDA,3 | 7.9 | 8.025 | Applicator Measurements
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 54.9 | GSM | FDA,0 | 50 | 60 | TM 101
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_CoA_(mm) | Caliper | 0.72 | mm | FDA,2 | 0.63 | 0.87 | TM 102
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 758 | Tensile Strength | Dry_CD_(g/in) | Dry CD | 1382.58 | g/in | FDA,0 | 1036 |  | TM 103
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 758 | Tensile Strength | Dry_MD_(g/in) | Dry CD | 6608.28 | g/in | FDA,0 | 4404 |  | TM 103
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 1.05 | g | FDA,2 |  |  | TM 610
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | [5 empty] | TM 610
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | [5 empty] | TM 610
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OA911/Applicator_NY_Skip lot testing | 3OA911 | Applicator_NY_Skip lot testing | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Non-Subculturable Growth (Done) |  |  |  |  | WI QAS 076
3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | 3OA972 | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements
3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | 3OA972 | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 44 | in | FDA,0 | 43 | 47 | Applicator Measurements
3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | 3OA972 | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 9 | in | FDA,3 | 8.875 | 9.125 | Applicator Measurements
3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | 3OA972 | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 40.9 | GSM | FDA,1 | 35.2 | 44.8 | TM 101
3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | 3OA972 | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 757 | Caliper | Caliper_CoA_(uM) | Caliper | 437.8 | micron | FDA,0 | 340 | 540 | TM 102
3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | 3OA972 | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 758 | Tensile Strength | Wet_CD_CoA_(g/in) | Wet CD CoA | 1145 | g/in | FDA,0 | 400 | 1800 | TM 103
3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | 3OA972 | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 758 | Tensile Strength | Wet_MD_CoA_(g/in) | Wet MD CoA | 2050 | g/in | FDA,0 | 1000 | 3400 | TM 103
3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | 3OA972 | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208
3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | 3OA972 | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208
3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | 3OA972 | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | REDUCED TESTING_NY_3OB108 | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading |  |  |  |  | Color (TM 208)
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | REDUCED TESTING_NY_3OB108 | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Meets specifications (Done) |  |  |  |  | See Analyte for Method #
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | REDUCED TESTING_NY_3OB108 | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | See Analyte for Method #
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | REDUCED TESTING_NY_3OB108 | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 54.9 | GSM | FDA,1 | 52.3 | 59.6 | See Analyte for Method #
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | REDUCED TESTING_NY_3OB108 | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method #
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | REDUCED TESTING_NY_3OB108 | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | Meets specifications (Done) |  |  |  |  | See Analyte for Method #
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 1.25 | g | FDA,2 |  |  | TM 610
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading |  |  |  |  | Color (TM 208)
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | May have residual sulfur odor when wet |  |  |  |  | Odor (TM 208)
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Free of hard wrinkles, and formation uniform, with micro crepe texture (Done) |  |  |  |  | See Analyte for Method #
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | See Analyte for Method #
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 54.04 | GSM | FDA,0 | 49.5 | 60.5 | See Analyte for Method #
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (mm) (TM 102) | Caliper | 0.55 | mm | FDA,2 | 0.52 | 0.74 | See Analyte for Method #
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method #
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter | 32 | in | FDA,0 | 28 | 36 | See Analyte for Method #
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Roll Width (Inches) (Measurement) | Roll Width | 6 | in | FDA,0 | 5.875 | 6.125 | See Analyte for Method #
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Wet CD (g/in) (TM 103) | Wet CD | 1004.84 | g/in | FDA,0 | 992 | 2208 | See Analyte for Method #
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Wet MD (g/in) (TM 103) | Wet MD | 4435.12 | g/in | FDA,0 | 4392 | 7808 | See Analyte for Method #
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 1.5 | g | FDA,2 |  |  | TM 610
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading |  |  |  |  | Color (TM 208)
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | May have residual sulfur odor when wet |  |  |  |  | Odor (TM 208)
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Meets specifications (Done) |  |  |  |  | See Analyte for Method #
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 200 |  | FDA,0 |  |  | See Analyte for Method #
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (lbs/ream) (TM 101) | Basis Weight | 23.81 | #/Ream | FDA,1 | 21.5 | 25 | See Analyte for Method #
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Caliper at point of manufacture (mils) (TM 102) | Caliper | 4.9 | mil | FDA,1 | 4.2 | 5.6 | See Analyte for Method #
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method #
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter | 32 | in | FDA,0 | 31 | 33 | See Analyte for Method #
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Roll Width (Inches) (Measurement) | Roll Width | 7.75 | in | FDA,4 | 7.6875 | 7.8125 | See Analyte for Method #
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Wet CD (g/in) (TM 103) | Wet CD | 1261.89 | g/in | FDA,0 | 1000 | 2500 | See Analyte for Method #
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Wet MD (g/in) (TM 103) | Wet MD | 2429.95 | g/in | FDA,0 | 1500 | 3850 | See Analyte for Method #
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 1.1 | g | FDA,2 |  |  | TM 610
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading |  |  |  |  | Color (TM 208)
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | REDUCED TESTING_NY_3OD438 | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading |  |  |  |  | Color (TM 208)
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | May have residual sulfur odor when wet |  |  |  |  | Odor (TM 208)
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | REDUCED TESTING_NY_3OD438 | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | May have residual sulfur odor when wet |  |  |  |  | Odor (TM 208)
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Smooth free of hard wrinkles, formation uniform matches embossing pattern (Done) |  |  |  |  | See Analyte for Method #
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | REDUCED TESTING_NY_3OD438 | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Smooth free of hard wrinkles, formation uniform matches embossing pattern (Done) |  |  |  |  | See Analyte for Method #
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | See Analyte for Method #
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | REDUCED TESTING_NY_3OD438 | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 315 |  | FDA,0 |  |  | See Analyte for Method #
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 34.66 | GSM | FDA,1 | 31.5 | 38.5 | See Analyte for Method #
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | REDUCED TESTING_NY_3OD438 | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 35.09 | GSM | FDA,1 | 31.5 | 38.5 | See Analyte for Method #
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (inches) (TM 102) | Caliper | 0.0077 | in | FDA,4 | 0.006 | 0.0087 | See Analyte for Method #
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method #
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | REDUCED TESTING_NY_3OD438 | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method #
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Dry CD (g/in) (TM103) | Dry CD | 1868.87 | g/in | FDA,0 | 900 | 2200 | See Analyte for Method #
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Dry MD (g/in) (TM 103) | Dry CD | 4655.69 | g/in | FDA,0 | 2000 | 5000 | See Analyte for Method #
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter | 32 | in | FDA,0 | 30 | 34 | See Analyte for Method #
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Roll Width (Inches) (Measurement) | Roll Width | 7.875 | in | FDA,2 | 7.75 | 8 | See Analyte for Method #
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Wet CD (g/in) (TM 103) | Wet CD | 789.96 | g/in | FDA,0 | 250 | 1250 | See Analyte for Method #
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Wet MD (g/in) (TM 103) | Wet MD | 2141.89 | g/in | FDA,0 | 800 | 2600 | See Analyte for Method #
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 1.1 | g | FDA,2 |  |  | TM 610
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 0.97 | g | FDA,2 |  |  | TM 610
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 1 | cfu/device | FDA,1 |  |  | TM 610
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0.18182 | CFU/g | FDA,1 |  | 100 | TM 610
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Positive Rods (Done) |  |  |  |  | WI QAS 076
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading |  |  |  |  | Color (TM 208)
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | May have residual sulfur odor when wet |  |  |  |  | Odor (TM 208)
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Free of hard wrinkles and formation uniform (Done) |  |  |  |  | See Analyte for Method #
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 315 |  | FDA,0 |  |  | See Analyte for Method #
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 46.72 | GSM | FDA,1 | 40.5 | 49.5 | See Analyte for Method #
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (mm) (TM 102) | Caliper | 0.64 | mm | FDA,2 | 0.55 | 0.67 | See Analyte for Method #
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method #
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter | 32 | in | FDA,0 | 31 | 33 | See Analyte for Method #
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Roll Width (Inches) (Measurement) | Roll Width | 7 | in | FDA,0 | 6.875 | 7.125 | See Analyte for Method #
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Wet CD (g/in) (TM 103) | Wet CD | 2598 | g/in | FDA,0 | 1131 | 1593 | See Analyte for Method #
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Wet MD (g/in) (TM 103) | Wet MD | 3711 | g/in | FDA,0 | 3906 | 4934 | See Analyte for Method #
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 1.75 | g | FDA,2 |  |  | TM 610
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading |  |  |  |  | Color (TM 208)
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Smooth free of hard wrinkles, formation uniform matches embossing pattern (Done) |  |  |  |  | See Analyte for Method #
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 200 |  | FDA,0 |  |  | See Analyte for Method #
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 52.74 | GSM | FDA,1 | 45 | 55 | See Analyte for Method #
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Caliper 1-ply (mm) (TM 102) | Caliper | 0.55 | mm | FDA,2 | 0.51 | 0.63 | See Analyte for Method #
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method #
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Dry CD (g/in) (TM103) | Dry CD | 2389.39 | g/in | FDA,0 | 1801 | 2821 | See Analyte for Method #
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Dry MD (g/in) (TM 103) | Dry CD | 8511.74 | g/in | FDA,0 | 6675 | 9447 | See Analyte for Method #
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | No objectionable odor detectable wet or dry (Done) |  |  |  |  | See Analyte for Method #
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter | 30 | in | FDA,0 | 30 | 32 | See Analyte for Method #
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Roll Width (Inches) (Measurement) | Roll Width | 7 | in | FDA,3 | 6.875 | 7.125 | See Analyte for Method #
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 1.07 | g | FDA,2 |  |  | TM 610
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer |  |  |  |  | TM 610
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OH592/Applicator_36A | 3OH592 | Applicator_36A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements
3OH592/Applicator_36A | 3OH592 | Applicator_36A | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 32 | in | FDA,0 | 30 | 34 | Applicator Measurements
3OH592/Applicator_36A | 3OH592 | Applicator_36A | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 7.875 | in | FDA,3 | 7.75 | 8 | Applicator Measurements
3OH592/Applicator_36A | 3OH592 | Applicator_36A | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 39.181 | GSM | FDA,0 | 35 | 45 | TM 101
3OH592/Applicator_36A | 3OH592 | Applicator_36A | NY-Analytical | 757 | Caliper | Caliper_CoA_(uM) | Caliper | 424 | um | FDA,0 | 340 | 540 | TM 102
3OH592/Applicator_36A | 3OH592 | Applicator_36A | NY-Analytical | 758 | Tensile Strength | Wet_CD_CoA_(g/in) | Wet CD CoA | 1235 | g/in | FDA,0 | 400 |  | TM 103
3OH592/Applicator_36A | 3OH592 | Applicator_36A | NY-Analytical | 758 | Tensile Strength | Wet_MD_CoA_(g/in) | Wet MD CoA | 2903 | g/in | FDA,0 | 1000 |  | TM 103
3OH592/Applicator_36A | 3OH592 | Applicator_36A | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208
3OH592/Applicator_36A | 3OH592 | Applicator_36A | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208
3OH592/Applicator_36A | 3OH592 | Applicator_36A | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Applicator Measurements
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width |  | in | FDA,3 | 11.875 | 12.125 | Applicator Measurements
3OJ213/Applicator_NY_Skip lot testing | 3OJ213 | Applicator_NY_Skip lot testing | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 12 | in | FDA,3 | 11.875 | 12.125 | Applicator Measurements
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 41.76 | GSM | FDA,0 | 36 | 44 | TM 101
3OJ213/Applicator_NY_Skip lot testing | 3OJ213 | Applicator_NY_Skip lot testing | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 40.9 | GSM | FDA,0 | 36 | 44 | TM 101
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_CoA_(mm) | Caliper | 0.54 | mm | FDA,2 | 0.48 | 0.61 | TM 102
3OJ213/Applicator_NY_Skip lot testing | 3OJ213 | Applicator_NY_Skip lot testing | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_CoA_(mm) | Caliper | 0.54 | mm | FDA,2 | 0.48 | 0.61 | TM 102
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 1096.49 | g/in | FDA,0 | 1030 | 1546 | TM 103
3OJ213/Applicator_NY_Skip lot testing | 3OJ213 | Applicator_NY_Skip lot testing | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 865 | g/in | FDA,0 | 1030 | 1546 | TM 103
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 4699.94 | g/in | FDA,0 | 3343 | 4350 | TM 103
3OJ213/Applicator_NY_Skip lot testing | 3OJ213 | Applicator_NY_Skip lot testing | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 3407 | g/in | FDA,0 | 3343 | 4350 | TM 103
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208
3OJ213/Applicator_NY_Skip lot testing | 3OJ213 | Applicator_NY_Skip lot testing | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208
3OJ213/Applicator_NY_Skip lot testing | 3OJ213 | Applicator_NY_Skip lot testing | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208
3OJ213/Applicator_NY_Skip lot testing | 3OJ213 | Applicator_NY_Skip lot testing | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight |  | g | FDA,2 |  |  | TM 610
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | [5 empty] | TM 610
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | [5 empty] | TM 610
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 100 | TM 610
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) |  | cfu | FDA,0 |  |  | TM 610
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 15 | TM 610
3OJ213/Applicator_08A | 3OJ213 | Applicator_08A | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  | cfu | FDA,0 |  |  | TM 610
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Meets specifications (Done) |  |  |  |  | See Analyte for Method #
3OM304/Applicator_NY_Skip lot testing | 3OM304 | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Meets specifications (Done) |  |  |  |  | See Analyte for Method #
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 315 |  | FDA,0 |  |  | See Analyte for Method #
3OM304/Applicator_NY_Skip lot testing | 3OM304 | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 315 |  | FDA,0 |  |  | See Analyte for Method #
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 47.36 | GSM | FDA,0 | 45 | 55 | See Analyte for Method #
3OM304/Applicator_NY_Skip lot testing | 3OM304 | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 50.38 | GSM | FDA,0 | 45 | 55 | See Analyte for Method #
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Caliper 1-ply (mm) (TM 102) | Caliper | 0.46 | mm | FDA,2 | 0.51 | 0.63 | See Analyte for Method #
3OM304/Applicator_NY_Skip lot testing | 3OM304 | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Caliper 1-ply (mm) (TM 102) | Caliper | 0.53 | mm | FDA,2 | 0.51 | 0.63 | See Analyte for Method #
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method #
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | Meets specifications (Done) |  |  |  |  | See Analyte for Method #
3OM304/Applicator_NY_Skip lot testing | 3OM304 | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | Meets specifications (Done) |  |  |  |  | See Analyte for Method #
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Dry CD (g/in) (TM103) | Dry CD | 2142.67 | g/in | FDA,0 | 1801 | 2821 | See Analyte for Method #
3OM304/Applicator_NY_Skip lot testing | 3OM304 | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Dry CD (g/in) (TM103) | Dry CD | 2395.04 | g/in | FDA,0 | 1801 | 2821 | See Analyte for Method #
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Dry MD (g/in) (TM 103) | Dry CD | 7892 | g/in | FDA,0 | 6675 | 9447 | See Analyte for Method #
3OM304/Applicator_NY_Skip lot testing | 3OM304 | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Dry MD (g/in) (TM 103) | Dry CD | 8931.7 | g/in | FDA,0 | 6675 | 9447 | See Analyte for Method #
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | Meets specifications (Done) |  |  |  |  | See Analyte for Method #
3OM304/Applicator_NY_Skip lot testing | 3OM304 | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | Meets specifications (Done) |  |  |  |  | See Analyte for Method #
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (COA) | Roll Diameter (COA) | 43 |  | FDA,0 | 41 | 45 | See Analyte for Method #
3OM304/Applicator_NY_Skip lot testing | 3OM304 | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (COA) | Roll Diameter (COA) | 43 |  | FDA,0 | 41 | 45 | See Analyte for Method #
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Roll width (COA) | Roll width (COA) | 9 | in | FDA,2 | 8.75 | 9.25 | See Analyte for Method #
3OM304/Applicator_NY_Skip lot testing | 3OM304 | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Roll width (COA) | Roll width (COA) | 9 | in | FDA,2 | 8.75 | 9.25 | See Analyte for Method #
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 0.89 | g | FDA,2 |  |  | TM 610
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | 3OP065 | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading |  |  |  |  | Color (TM 208)
3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | 3OP065 | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Free of hard wrinkles. Formation uniform (Done) |  |  |  |  | See Analyte for Method #
3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | 3OP065 | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | See Analyte for Method #
3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | 3OP065 | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 57.26 | GSM | FDA,1 | 54 | 66 | See Analyte for Method #
3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | 3OP065 | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (mm) (CofA) | Caliper | 0.8 | mm | FDA,2 | 0.67 | 0.81 | See Analyte for Method #
3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | 3OP065 | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method #
3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | 3OP065 | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | Meets specifications (Done) |  |  |  |  | See Analyte for Method #
3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | 3OP065 | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter | 46 | in | FDA,0 | 45 | 47 | See Analyte for Method #
3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | 3OP065 | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Width (Inches) (Measurement) | Roll Width | 30 | in | FDA,2 | 29.75 | 30.25 | See Analyte for Method #
3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | 3OP065 | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Wet CD (g/in) (TM 103) | Wet CD | 1584 | g/in | FDA,0 | 1300 | 2000 | See Analyte for Method #
3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | 3OP065 | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Wet MD (g/in) (TM 103) | Wet MD | 5195.31 | g/in | FDA,0 | 4000 | 5500 | See Analyte for Method #
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter |  | in | FDA,0 | 45 | 47 | Applicator Measurements
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width |  | in | FDA,2 | 29.75 | 30.25 | Applicator Measurements
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 57.7 | GSM | FDA,1 | 49.5 | 60.5 | TM 101
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper | 0.51 | mm | FDA,2 | 0.4 | 0.6 | TM 102
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 758 | Tensile Strength | Elongation % Wet CD | Elongation % Wet CD | 103.07 | % | FDA,0 | 80 | 130 | TM 103
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 758 | Tensile Strength | Elongation % Wet MD | Elongation % Wet MD | 24.68 | % | FDA,0 | 20 | 35 | TM 103
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 893.28 | g/in | FDA,0 | 650 | 1300 | TM 103
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 2898.19 | g/in | FDA,0 | 2400 | 3400 | TM 103
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | [5 empty] | TM 208
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 748 | Visual Inspection | Color | Color | [5 empty] | TM 208
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 748 | Visual Inspection | Odor | Odor | [5 empty] | TM 208
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 0.84 | g | FDA,2 |  |  | TM 610
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer |  |  |  |  | TM 610
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Applicator_NY_Die Cut Pads | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Applicator_NY_Die Cut Pads | NY-Analytical | 759 | Applicator Measurements | Pad Diameter | Pad Diameter | 2.2188 | in | FDA,2 | 2.19 | 2.31 | Applicator Measurements
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Applicator_NY_Die Cut Pads | NY-Analytical | 759 | Applicator Measurements | Stack Count | Stack Count | 90 | Units | FDA,0 | 90 | 90 | Applicator Measurements
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Applicator_NY_Die Cut Pads | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Applicator_NY_Die Cut Pads | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Applicator_NY_Die Cut Pads | NY-Analytical | 748 | Visual Inspection | Odor | Odor | No objectionable odor Wet or Dry [Done] |  |  |  |  | TM 208
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Applicator_NY_Die Cut Pads | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 2.07 | g | FDA,2 |  |  | TM 610
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Applicator_NY_Die Cut Pads | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Applicator_NY_Die Cut Pads | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Applicator_NY_Die Cut Pads | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 1 | cfu/device | FDA,1 |  |  | TM 610
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Applicator_NY_Die Cut Pads | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0.042735 | CFU/g | FDA,1 |  | 100 | TM 610
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Applicator_NY_Die Cut Pads | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Applicator_NY_Die Cut Pads | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Applicator_NY_Die Cut Pads | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Applicator_NY_Die Cut Pads | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Applicator_NY_Die Cut Pads | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Positive Cocci (OOS) |  |  |  |  | WI QAS 076
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Applicator_NY_Die Cut Pads | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Non-Subculturable Growth (Done) |  |  |  |  | WI QAS 076
3OP113/Applicator_NY_Skip lot testing | 3OP113 | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (mm) (CofA) | Caliper | 0.62 | mm | FDA,2 | 0.5 | 0.62 | See Analyte for Method #
3OP113/Applicator_NY_Skip lot testing | 3OP113 | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter | 40 | in | FDA,0 | 40 | 46 | See Analyte for Method #
3OP113/Applicator_NY_Skip lot testing | 3OP113 | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Roll Width (Inches) (Measurement) | Roll Width | 30 | in | FDA,3 | 29.875 | 30.125 | See Analyte for Method #
3OP113/Applicator_NY_Skip lot testing | 3OP113 | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Wet CD (g/in) (TM 103) | Wet CD | 1469.66 | g/in | FDA,0 | 1142 | 1714 | See Analyte for Method #
3OP113/Applicator_NY_Skip lot testing | 3OP113 | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Wet MD (g/in) (TM 103) | Wet MD | 4083.47 | g/in | FDA,0 | 3265 | 4897 | See Analyte for Method #
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 40 | in | FDA,0 | 39 | 40 | Applicator Measurements
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 30 | in | FDA,3 | 29.875 | 30.125 | Applicator Measurements
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 34.01 | GSM | FDA,1 | 31.5 | 38.5 | TM 101
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(inch) | Caliper | 0.008 | in | FDA,4 | 0.006 | 0.009 | TM 102
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_CoA_(mm) | Caliper | 0.0079 | in |  |  |  | TM 102
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 812.99 | g/in | FDA,0 | 250 | 1250 | TM 103
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 1821.41 | g/in | FDA,0 | 800 | 2600 | TM 103
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 748 | Visual Inspection | Color | Color | White with no off colors or shading [Done] |  |  |  |  | TM 208
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 0.97 | g | FDA,2 |  |  | TM 610
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | [5 empty] | WI QAS 076
3OP138/Applicator_36D | 3OP138 | Applicator_36D | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Applicator Measurements
3OP138/Applicator_36D | 3OP138 | Applicator_36D | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter |  | in | FDA,0 | 43 | 47 | Applicator Measurements
3OP138/Applicator_36D | 3OP138 | Applicator_36D | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 30 | in | FDA,3 | 29.875 | 30.125 | Applicator Measurements
3OP138/Applicator_36D | 3OP138 | Applicator_36D | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 40.26 | GSM | FDA,0 | 35 | 45 | TM 101
3OP138/Applicator_36D | 3OP138 | Applicator_36D | NY-Analytical | 757 | Caliper | Caliper_CoA_(uM) | Caliper |  | um | FDA,0 | 340 | 540 | TM 102
3OP138/Applicator_36D | 3OP138 | Applicator_36D | NY-Analytical | 758 | Tensile Strength | Wet_CD_CoA_(g/in) | Wet CD CoA | 1228 | g/in | FDA,0 | 400 |  | TM 103
3OP138/Applicator_36D | 3OP138 | Applicator_36D | NY-Analytical | 758 | Tensile Strength | Wet_MD_CoA_(g/in) | Wet MD CoA | 2545 | g/in | FDA,0 | 1000 |  | TM 103
3OP138/Applicator_36D | 3OP138 | Applicator_36D | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208
3OP138/Applicator_36D | 3OP138 | Applicator_36D | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208
3OP138/Applicator_36D | 3OP138 | Applicator_36D | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208
3OP197A/Applicator_NY_Skip lot testing | 3OP197A | Applicator_NY_Skip lot testing | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Bioburden_TMYC
3OP197A/Applicator_NY_Skip lot testing | 3OP197A | Applicator_NY_Skip lot testing | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 0.92 | g | FDA,2 |  |  | TM 610
3OP197A/Applicator_NY_Skip lot testing | 3OP197A | Applicator_NY_Skip lot testing | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer |  |  |  |  | TM 610
3OP197A/Applicator_NY_Skip lot testing | 3OP197A | Applicator_NY_Skip lot testing | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610
3OP197A/Applicator_NY_Skip lot testing | 3OP197A | Applicator_NY_Skip lot testing | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OP197A/Applicator_NY_Skip lot testing | 3OP197A | Applicator_NY_Skip lot testing | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610
3OP197A/Applicator_NY_Skip lot testing | 3OP197A | Applicator_NY_Skip lot testing | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OP197A/Applicator_NY_Skip lot testing | 3OP197A | Applicator_NY_Skip lot testing | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 1 | cfu/device | FDA,1 |  |  | TM 610
3OP197A/Applicator_NY_Skip lot testing | 3OP197A | Applicator_NY_Skip lot testing | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0.76087 | CFU/g | FDA,1 |  | 15 | TM 610
3OP197A/Applicator_NY_Skip lot testing | 3OP197A | Applicator_NY_Skip lot testing | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 1 | cfu | FDA,0 |  |  | TM 610
3OP197A/Applicator_NY_Skip lot testing | 3OP197A | Applicator_NY_Skip lot testing | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Non-Subculturable Growth (Done) |  |  |  |  | WI QAS 076
3OP197A-SU/Label_NY_NP7200_6.25-4.5 | 3OP197A-SU | Label_NY_NP7200_6.25-4.5 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection
3OP197A-SU/Label_NY_NP7200_6.25-4.5 | 3OP197A-SU | Label_NY_NP7200_6.25-4.5 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
3OP197A-SU/Label_NY_NP7200_6.25-4.5 | 3OP197A-SU | Label_NY_NP7200_6.25-4.5 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
3OP197A-SU/Label_NY_NP7200_6.25-4.5 | 3OP197A-SU | Label_NY_NP7200_6.25-4.5 | NY-Analytical | 770 | Label Inspection | Length | Length | 6.25 | in | FDA,4 | 6.1875 | 6.3125 | TM 143 / Visual
3OP197A-SU/Label_NY_NP7200_6.25-4.5 | 3OP197A-SU | Label_NY_NP7200_6.25-4.5 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
3OP197A-SU/Label_NY_NP7200_6.25-4.5 | 3OP197A-SU | Label_NY_NP7200_6.25-4.5 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
3OP197A-SU/Label_NY_NP7200_6.25-4.5 | 3OP197A-SU | Label_NY_NP7200_6.25-4.5 | NY-Analytical | 770 | Label Inspection | Width | Width | 4.5 | in | FDA,4 | 4.4375 | 4.5625 | TM 143 / Visual
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 37.5 | in | FDA,3 | 37.375 | 37.625 | Applicator Measurements
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Basis Weight
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 35.95 | GSM | FDA,1 | 31.5 | 38.5 | TM 101
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper | 0.34 | mm | FDA,2 | 0.27 | 0.41 | TM 102
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 758 | Tensile Strength | Wet_CD_CoA_(g/in) | Wet CD CoA | 520.88 | g/in | FDA,0 | 300 |  | TM 103
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 758 | Tensile Strength | Wet_MD_CoA_(g/in) | Wet MD CoA | 3556.57 | g/in | FDA,0 | 2000 | 4400 | TM 103
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 748 | Visual Inspection | Color | Color | White with no off colors or shading [Done] |  |  |  |  | TM 208
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 0.95 | g | FDA,2 |  |  | TM 610
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 1.5 | cfu/device | FDA,1 |  |  | TM 610
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 1.5789 | CFU/g | FDA,1 |  | 100 | TM 610
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 1 | cfu | FDA,0 |  |  | TM 610
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | 3OP242 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Non-Subculturable Growth (Done) |  |  |  |  | WI QAS 076
3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | 3OP257 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (COA) | Roll Diameter (COA) | 43 | in | FDA,0 | 41 | 45 | Applicator Measurements
3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | 3OP257 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 759 | Applicator Measurements | Roll width (COA) | Roll width (COA) | 37.5 | in | FDA,2 | 37.25 | 37.75 | Applicator Measurements
3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | 3OP257 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Basis Weight
3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | 3OP257 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 51.02 | GSM | FDA,0 | 45 | 55 | TM 101
3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | 3OP257 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper | 0.52 | mm | FDA,2 | 0.51 | 0.63 | TM 102
3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | 3OP257 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 758 | Tensile Strength | Dry_CD_(g/in) | Dry CD | 2193.6 | g/in | FDA,0 | 1801 | 2821 | TM 103
3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | 3OP257 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 758 | Tensile Strength | Dry_MD_(g/in) | Dry CD | 7519.45 | g/in | FDA,0 | 6675 | 9447 | TM 103
3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | 3OP257 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208
3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | 3OP257 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208
3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | 3OP257 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 748 | Visual Inspection | Odor | Odor | No objectionable odor Wet or Dry [Done] |  |  |  |  | TM 208
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading |  |  |  |  | Color (TM 208)
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 759 | Applicator Measurements | Roll Width (mm) | Roll Width | 953 | mm | FDA,0 | 950 | 956 | Applicator Measurements
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Free of hard wrinkles, formation uniform/quilted big dot emboss pattern (Done) |  |  |  |  | See Analyte for Method #
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 315 |  | FDA,0 |  |  | See Analyte for Method #
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 32.08 | GSM | FDA,1 | 31.5 | 38.5 | See Analyte for Method #
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (mm) (TM 102) | Caliper | 0.34 | mm | FDA,2 | 0.26 | 0.4 | See Analyte for Method #
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method #
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Dry CD (g/in) (TM103) | Dry CD | 1254.21 | g/in | FDA,0 | 1100 | 1700 | See Analyte for Method #
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Dry MD (g/in) (TM 103) | Dry CD | 2070.52 | g/in | FDA,0 | 2000 | 3000 | See Analyte for Method #
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | May have slight odor when wet (Done) |  |  |  |  | See Analyte for Method #
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter | 40 | in | FDA,0 | 40 |  | See Analyte for Method #
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Roll Width (Inches) (Measurement) | Roll Width | 37.5 | in | FDA,2 | 37.25 | 37.75 | See Analyte for Method #
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 34.44 | GSM | FDA,1 | 31.5 | 38.5 | TM 101
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper | 0.29 | mm | FDA,2 | 0.26 | 0.4 | TM 102
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 758 | Tensile Strength | Dry_CD_(g/in) | Dry CD | 1143.94 | g/in | FDA,0 | 1100 | 1700 | TM 103
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 758 | Tensile Strength | Dry_MD_(g/in) | Dry CD | 2220.05 | g/in | FDA,0 | 2000 | 3000 | TM 103
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 0.95 | g | FDA,2 |  |  | TM 610
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer |  |  |  |  | TM 610
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 1 | cfu/device | FDA,1 |  |  | TM 610
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0.21053 | CFU/g | FDA,1 |  | 15 | TM 610
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ222 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Certificate of Analysis
3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Applicator Measurements
3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ222 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter |  | in | FDA,0 | 42 | 46 | Applicator Measurements
3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 44 | in | FDA,0 | 42 | 46 | Applicator Measurements
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter |  | in | FDA,0 | 42 | 46 | Applicator Measurements
3OQ222/Applicator_NY_Skip lot testing | 3OQ222 | Applicator_NY_Skip lot testing | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter |  | in | FDA,0 | 42 | 46 | Applicator Measurements
3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ222 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width |  | in | FDA,3 | 37.375 | 37.625 | Applicator Measurements
3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 36 | in | FDA,3 | 35.875 | 36.125 | Applicator Measurements
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 30 | in | FDA,3 | 29.875 | 30.125 | Applicator Measurements
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 41.12 | GSM | FDA,1 | 35.5 | 44.5 | TM 101
3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 39.61 | GSM | FDA,1 | 35.5 | 44.5 | TM 101
3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ222 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight |  | GSM | FDA,1 | 35.5 | 44.5 | TM 101
3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper | 0.46 | mm | FDA,2 | 0.4 | 0.48 | TM 102
3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ222 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_CoA_(mm) | Caliper |  | mm | FDA,2 | 0.4 | 0.48 | TM 102
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_CoA_(mm) | Caliper | 0.45 | mm | FDA,2 | 0.4 | 0.48 | TM 102
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 740.59 | g/in | FDA,0 | 600 | 1200 | TM 103
3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 521.87 | g/in | FDA,0 | 600 | 1200 | TM 103
3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ222 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD |  | g/in | FDA,0 | 600 | 1200 | TM 103
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD |  | g/in | FDA,0 | 2200 | 3400 | TM 103
3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 1610.35 | g/in | FDA,0 | 2200 | 3400 | TM 103
3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ222 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD |  | g/in | FDA,0 | 2200 | 3400 | TM 103
3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | [5 empty] | TM 208
3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ222 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | [5 empty] | TM 208
3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 748 | Visual Inspection | Color | Color | White with no off colors or shading [Done] |  |  |  |  | TM 208
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 748 | Visual Inspection | Color | Color | [5 empty] | TM 208
3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ222 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 748 | Visual Inspection | Color | Color | [5 empty] | TM 208
3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 748 | Visual Inspection | Odor | Odor | [5 empty] | TM 208
3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ222 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 748 | Visual Inspection | Odor | Odor | [5 empty] | TM 208
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight |  | g | FDA,2 |  |  | TM 610
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | [5 empty] | TM 610
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | [5 empty] | TM 610
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 100 | TM 610
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) |  | cfu | FDA,0 |  |  | TM 610
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 15 | TM 610
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  | cfu | FDA,0 |  |  | TM 610
3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ333 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 1.32 | g | FDA,2 |  |  | TM 610
3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ333 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610
3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ333 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610
3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ333 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ333 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610
3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ333 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ333 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ333 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610
3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ333 | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OQ444/Applicator_NY_Skip lot testing | 3OQ444 | Applicator_NY_Skip lot testing | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper | 0.41 | mm | FDA,2 | 0.4 | 0.48 | TM 102
3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ444 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight |  | g | FDA,2 |  |  | TM 610
3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ444 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | [5 empty] | TM 610
3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ444 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | [5 empty] | TM 610
3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ444 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610
3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ444 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 100 | TM 610
3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ444 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) |  | cfu | FDA,0 |  |  | TM 610
3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ444 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610
3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ444 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 15 | TM 610
3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ444 | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  | cfu | FDA,0 |  |  | TM 610
3OQ551/Applicator_15E | 3OQ551 | Applicator_15E | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (COA) | Roll Diameter (COA) | 45 | in | FDA,0 | 43 | 47 | Applicator Measurements
3OQ551/Applicator_15E | 3OQ551 | Applicator_15E | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 42 | in | FDA,3 | 41.875 | 42.125 | Applicator Measurements
3OQ551/Applicator_15E | 3OQ551 | Applicator_15E | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Basis Weight
3OQ551/Applicator_15E | 3OQ551 | Applicator_15E | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight |  | GSM | FDA,0 | 35 | 45 | TM 101
3OQ551/Applicator_15E | 3OQ551 | Applicator_15E | NY-Analytical | 757 | Caliper | Caliper_at_Pt_of_Mfr_(micron) | Caliper | 431.8 | um | FDA,0 | 340 | 540 | TM 102
3OQ551/Applicator_15E | 3OQ551 | Applicator_15E | NY-Analytical | 758 | Tensile Strength | Wet_CD_CoA_(g/25mm) | Wet CD CoA | 1073 | g/25mm | FDA,0 | 400 |  | TM 103
3OQ551/Applicator_15E | 3OQ551 | Applicator_15E | NY-Analytical | 758 | Tensile Strength | Wet_MD_CoA_(g/25mm) | Wet MD CoA | 2506 | g/25mm | FDA,0 | 1000 |  | TM 103
3OQ551/Applicator_15E | 3OQ551 | Applicator_15E | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208
3OQ551/Applicator_15E | 3OQ551 | Applicator_15E | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208
3OQ551/Applicator_15E | 3OQ551 | Applicator_15E | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208
3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | 3OQ850N | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading |  |  |  |  | Color (TM 208)
3OQ850N/Applicator_01E | 3OQ850N | Applicator_01E | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Applicator Measurements
3OQ850N/Applicator_01E | 3OQ850N | Applicator_01E | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter |  | in | FDA,0 | 44 | 47 | Applicator Measurements
3OQ850N/Applicator_01E | 3OQ850N | Applicator_01E | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width |  | in | FDA,3 | 35.875 | 36.125 | Applicator Measurements
3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | 3OQ850N | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Smooth, free of hard wrinkles and formation uniform (Done) |  |  |  |  | See Analyte for Method #
3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | 3OQ850N | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 315 |  | FDA,0 |  |  | See Analyte for Method #
3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | 3OQ850N | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 40.69 | GSM | FDA,1 | 38 | 42 | See Analyte for Method #
3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | 3OQ850N | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (mm) (TM 102) | Caliper | 0.41 | mm |  |  |  | See Analyte for Method #
3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | 3OQ850N | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method #
3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | 3OQ850N | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Dry CD (g/in) (TM103) | Dry CD | 1098.41 | g/in | FDA,0 | 513 | 1130 | See Analyte for Method #
3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | 3OQ850N | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Dry MD (g/in) (TM 103) | Dry CD | 2976.77 | g/in | FDA,0 | 2054 | 4108 | See Analyte for Method #
3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | 3OQ850N | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | No objectionable odor detectable wet or dry (Done) |  |  |  |  | See Analyte for Method #
3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | 3OQ850N | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter | 46 | in | FDA,0 | 44 | 47 | See Analyte for Method #
3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | 3OQ850N | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Roll Width (Inches) (Measurement) | Roll Width | 37.5 | in | FDA,3 | 37.375 | 37.625 | See Analyte for Method #
3OQ850N/Applicator_01E | 3OQ850N | Applicator_01E | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight |  | GSM | FDA,0 | 38 | 42 | TM 101
3OQ850N/Applicator_01E | 3OQ850N | Applicator_01E | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper |  | mm | FDA,2 | 0.35 | 0.45 | TM 102
3OQ850N/Applicator_01E | 3OQ850N | Applicator_01E | NY-Analytical | 758 | Tensile Strength | Dry_CD_(g/in) | Dry CD |  | g/in | FDA,0 | 513 | 1130 | TM 103
3OQ850N/Applicator_01E | 3OQ850N | Applicator_01E | NY-Analytical | 758 | Tensile Strength | Dry_MD_(g/in) | Dry CD |  | g/in | FDA,0 | 2054 | 4108 | TM 103
3OQ850N/Applicator_01E | 3OQ850N | Applicator_01E | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | [5 empty] | TM 208
3OQ850N/Applicator_01E | 3OQ850N | Applicator_01E | NY-Analytical | 748 | Visual Inspection | Color | Color | [5 empty] | TM 208
3OQ850N/Applicator_01E | 3OQ850N | Applicator_01E | NY-Analytical | 748 | Visual Inspection | Odor | Odor | [5 empty] | TM 208
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Swab Measurements
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Analytical | 760 | Swab Measurements | Mitt_Length_(Inches) | Mitt Length | 0.99 | in | FDA,3 | 0.97 | 1.03 | TM 109
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Analytical | 760 | Swab Measurements | Mitt_Thickness_(Inches) | Mitt Thickness | 0.314 | in | FDA,3 | 0.265 | 0.385 | TM 109
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Analytical | 760 | Swab Measurements | Mitt_Width_(Inches) | Mitt Width | 0.49 | in | FDA,2 | 0.47 | 0.53 | TM 109
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Analytical | 760 | Swab Measurements | Overall_Swab_Length_(Inches) | Overall Swab Length | 4.77 | in | FDA,3 | 4.69 | 4.81 | TM 109
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Analytical | 760 | Swab Measurements | Scrub_Test | Scrub Test | Meets specifications (Done) |  |  |  |  | TM 109
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Analytical | 760 | Swab Measurements | Tuft_Adhesion_(Lb.) | Tuft Adhesion | 2.59 | lb | FDA,1 | 1.4 |  | TM 109
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Analytical | 760 | Swab Measurements | Use_Test-Solvents | Use Test-Solvents | Meets specifications (Done) |  |  |  |  | TM 109
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 1.2 | g | FDA,2 |  |  | TM 610
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) |  | cfu | FDA,0 |  |  | TM 610
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 15 | TM 610
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  | cfu | FDA,0 |  |  | TM 610
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Positive Rods (Done) |  |  |  |  | WI QAS 076
3OS407/Applicator_25B | 3OS407 | Applicator_25B | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Non-Subculturable Growth (Done) |  |  |  |  | WI QAS 076
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Swab Measurements
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Analytical | 760 | Swab Measurements | Mitt_Length_(Inches) | Mitt Length | 1.22 | in | FDA,2 | 1.16 | 1.36 | TM 109
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Analytical | 760 | Swab Measurements | Mitt_Thickness_(Inches) | Mitt Thickness | 0.47 | in | FDA,3 | 0.41 | 0.53 | TM 109
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Analytical | 760 | Swab Measurements | Mitt_Width_(Inches) | Mitt Width | 1.19 | in | FDA,2 | 1.1 | 1.3 | TM 109
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Analytical | 760 | Swab Measurements | Overall_Swab_Length_(Inches) | Overall Swab Length | 6.03 | in | FDA,2 | 5.85 | 6.15 | TM 109
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Analytical | 760 | Swab Measurements | Scrub_Test | Scrub Test | Meets specifications (Done) |  |  |  |  | TM 109
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Analytical | 760 | Swab Measurements | Tuft_Adhesion_(Lb.) | Tuft Adhesion | 5.12 | lb | FDA,1 | 1.4 |  | TM 109
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Analytical | 760 | Swab Measurements | Use_Test-Solvents | Use Test-Solvents | Meets specifications (Done) |  |  |  |  | TM 109
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 1.26 | g | FDA,2 |  |  | TM 610
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | [5 empty] | TM 610
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | [5 empty] | TM 610
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 100 | TM 610
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) |  | cfu | FDA,0 |  |  | TM 610
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OS419/Applicator_25C | 3OS419 | Applicator_25C | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Positive Rods (Done) |  |  |  |  | WI QAS 076
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | REDUCED TESTING_NY_3OS428 | NY-Analytical | 760 | Swab Measurements | AQL Sample Size | AQL Sample Size | 315 | in |  |  |  | TM 109
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | REDUCED TESTING_NY_3OS428 | NY-Analytical | 760 | Swab Measurements | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | TM 109
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | REDUCED TESTING_NY_3OS428 | NY-Analytical | 760 | Swab Measurements | Scrub_Test | Scrub Test | Pass (Done) |  |  |  |  | TM 109
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | REDUCED TESTING_NY_3OS428 | NY-Analytical | 760 | Swab Measurements | Tuft_Adhesion_(Lb.) | Tuft Adhesion | 2.4 | lb | FDA,1 | 1.4 |  | TM 109
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 2.5 | g | FDA,2 |  |  | TM 610
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OS484/Applicator_26A | 3OS484 | Applicator_26A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Swab Measurements
3OS484/Applicator_26A | 3OS484 | Applicator_26A | NY-Analytical | 760 | Swab Measurements | Overall_Swab_Length_(Inches) | Overall Swab Length | 7.875 | in | FDA,2 | 7.82 | 7.94 | TM 109
3OS484/Applicator_26A | 3OS484 | Applicator_26A | NY-Analytical | 760 | Swab Measurements | Scrub_Test | Scrub Test | Meets specifications (Done) |  |  |  |  | TM 109
3OS484/Applicator_26A | 3OS484 | Applicator_26A | NY-Analytical | 760 | Swab Measurements | Tuft_Diameter_(Inches) | Tuft Diameter | 0.48 | in | FDA,2 | 0.42 | 0.53 | TM 109
3OS484/Applicator_26A | 3OS484 | Applicator_26A | NY-Analytical | 760 | Swab Measurements | Use_Test-Solvents | Use Test-Solvents | Meets specifications (Done) |  |  |  |  | TM 109
3OS484/Applicator_26A | 3OS484 | Applicator_26A | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 1.01 | g | FDA,2 |  |  | TM 610
3OS484/Applicator_26A | 3OS484 | Applicator_26A | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610
3OS484/Applicator_26A | 3OS484 | Applicator_26A | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610
3OS484/Applicator_26A | 3OS484 | Applicator_26A | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 1 | cfu/device | FDA,1 |  |  | TM 610
3OS484/Applicator_26A | 3OS484 | Applicator_26A | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0.20833 | CFU/g | FDA,1 |  | 100 | TM 610
3OS484/Applicator_26A | 3OS484 | Applicator_26A | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OS484/Applicator_26A | 3OS484 | Applicator_26A | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OS484/Applicator_26A | 3OS484 | Applicator_26A | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610
3OS484/Applicator_26A | 3OS484 | Applicator_26A | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OS484/Applicator_26A | 3OS484 | Applicator_26A | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Negative Rods (OOS) |  |  |  |  | WI QAS 076
3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | 3OU252 | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | [5 empty] | See Analyte for Method #
3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | 3OU252 | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 |  |  | See Analyte for Method #
3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | 3OU252 | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight |  | GSM | FDA,1 | 64.4 | 71.2 | See Analyte for Method #
3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | 3OU252 | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (mm) (TM 102) | Caliper |  | mm | FDA,2 | 0.28 | 0.42 | See Analyte for Method #
3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | 3OU252 | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | [5 empty] | See Analyte for Method #
3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | 3OU252 | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | [5 empty] | See Analyte for Method #
3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | 3OU252 | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | [5 empty] | See Analyte for Method #
3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | 3OU252 | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (COA) | Roll Diameter (COA) |  |  | FDA,0 | 41 | 45 | See Analyte for Method #
3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | 3OU252 | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Roll width (COA) | Roll width (COA) |  | in | FDA,2 | 8.94 | 9.13 | See Analyte for Method #
3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | 3OU252 | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Wet CD (g/in) (TM 103) | Wet CD |  | g/in | FDA,0 | 1500 | 2500 | See Analyte for Method #
3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | 3OU252 | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Wet MD (g/in) (TM 103) | Wet MD |  | g/in | FDA,0 | 4500 | 6500 | See Analyte for Method #
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 44 | in | FDA,0 | 41 | 45 | Applicator Measurements
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 11.5 | in | FDA,2 | 11.25 | 11.75 | Applicator Measurements
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 50.6 | GSM | FDA,1 | 45 | 55 | TM 101
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Analytical | 757 | Caliper | Caliper-1_Ply_(mm) | Caliper | 0.52 | mm | FDA,2 | 0.51 | 0.63 | TM 102
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Analytical | 758 | Tensile Strength | Dry_CD_(g/in) | Dry CD | 2302.98 | g/in | FDA,0 | 1801 | 2821 | TM 103
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Analytical | 758 | Tensile Strength | Dry_MD_(g/in) | Dry CD | 8886.78 | g/in | FDA,0 | 6675 | 9447 | TM 103
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 1.15 | g | FDA,2 |  |  | TM 610
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer |  |  |  |  | TM 610
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610
3OU265/Applicator_10H | 3OU265 | Applicator_10H | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OU275/Applicator_36C | 3OU275 | Applicator_36C | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements
3OU275/Applicator_36C | 3OU275 | Applicator_36C | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 32 | in | FDA,0 | 30 | 34 | Applicator Measurements
3OU275/Applicator_36C | 3OU275 | Applicator_36C | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 11.75 | in | FDA,3 | 11.625 | 11.875 | Applicator Measurements
3OU275/Applicator_36C | 3OU275 | Applicator_36C | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 37.67 | GSM | FDA,0 | 35 | 45 | TM 101
3OU275/Applicator_36C | 3OU275 | Applicator_36C | NY-Analytical | 757 | Caliper | Caliper_CoA_(uM) | Caliper | 458 | um | FDA,0 | 340 | 540 | TM 102
3OU275/Applicator_36C | 3OU275 | Applicator_36C | NY-Analytical | 758 | Tensile Strength | Wet_CD_CoA_(g/in) | Wet CD CoA | 1256.91 | g/in | FDA,0 | 400 |  | TM 103
3OU275/Applicator_36C | 3OU275 | Applicator_36C | NY-Analytical | 758 | Tensile Strength | Wet_MD_CoA_(g/in) | Wet MD CoA | 3006.85 | g/in | FDA,0 | 1000 |  | TM 103
3OU275/Applicator_36C | 3OU275 | Applicator_36C | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208
3OU275/Applicator_36C | 3OU275 | Applicator_36C | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208
3OU275/Applicator_36C | 3OU275 | Applicator_36C | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208
3OU351/Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | 3OU351 | Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA
3OU351/Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | 3OU351 | Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis (ELLI) | Certificate of Analysis | COA Meets Specification (Done) |  |  |  |  | COA
3OU351/Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | 3OU351 | Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 5.968 | in | FDA,3 | 5.875 | 6.125 | TM 143
3OU351/Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | 3OU351 | Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | NY-Analytical | 397 | Appearance | Appearance/Matches Embossing Pattern | Appearance/Matches Embossing Pattern | Free of Hard wrinkles. Formation uniform (Done) |  |  |  |  | TM 208
3OU351/Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | 3OU351 | Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | NY-Analytical | 395 | Color | Color | Color | White with no off colors or shading (Done) |  |  |  |  | TM 208
3OU351/Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | 3OU351 | Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | NY-Analytical | 1038 | KC Part Number | Verify KC Part Number | Verify KC Part Number | SPEC-08468 (Done)
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Applicatpr_NY_Foam 10.5ml | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Conformance
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Applicatpr_NY_Foam 10.5ml | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Bioburden_TAMC
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 0.91 | g | FDA,2 |  |  | TM 610
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 1 | cfu/device | FDA,1 |  |  | TM 610
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0.32967 | CFU/g | FDA,1 |  | 15 | TM 610
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Applicatpr_NY_Foam 10.5ml | NY-Analytical | 397 | Appearance | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | Visual
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Mold (OOS) |  |  |  |  | WI QAS 076
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Natural (Done) |  |  |  |  | Canister & Lid Inspection
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | NY-Analytical | 765 | Canister & Lid Inspection | Handle Internal Diameter | Handle Internal Diameter | 0.9575 | in | FDA,3 | 0.951 | 0.957 | Canister & Lid Inspection
3OX465H/Microbiologu - Bioburden_Applicator_NY | 3OX465H | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 765 | Canister & Lid Inspection | Handle Internal Diameter | Handle Internal Diameter | 0.9545 | in | FDA,3 | 0.951 | 0.957 | Canister & Lid Inspection
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | NY-Analytical | 765 | Canister & Lid Inspection | Overall Length | Overall Length | 5.23 | in | FDA,3 | 5.2 | 5.25 | Canister & Lid Inspection
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 87.86 | g | FDA,2 |  |  | TM 610
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 3 | cfu/device | FDA,1 |  |  | TM 610
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0.034145 | CFU/g | FDA,1 |  | 100 | TM 610
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 13 | cfu | FDA,0 |  |  | TM 610
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Non-Subculturable Growth (Done) |  |  |  |  | WI QAS 076
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Gram Positive Rods (Done) |  |  |  |  | WI QAS 076
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swab_NY_Swabstick, Cotton-Tipped | NY-Analytical | 760 | Swab Measurements | Mitt_Length_(Inches) | Mitt Length |  | in | FDA,3 | 0.593 | 0.781 | Swab Measurements
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swab_NY_Swabstick, Cotton-Tipped | NY-Analytical | 760 | Swab Measurements | Mitt_Width_(Inches) | Mitt Width |  | in | FDA,3 | 0.188 | 0.28 | Swab Measurements
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swab_NY_Swabstick, Cotton-Tipped | NY-Analytical | 760 | Swab Measurements | Overall_Swab_Length_(Inches) | Overall Swab Length |  | in | FDA,3 | 5.875 | 5.999 | Swab Measurements
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swab_NY_Swabstick, Cotton-Tipped | NY-Analytical | 760 | Swab Measurements | Stick Diameter | Stick Diameter |  | in | FDA,3 | 0.093 | 0.107 | Swab Measurements
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swab_NY_Swabstick, Cotton-Tipped | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Certificate of Conformance
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swab_NY_Swabstick, Cotton-Tipped | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | [5 empty] | Bioburden_TAMC
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 0.58 | g | FDA,2 |  |  | TM 610
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram) | TAMC (Average per Gram) | 0 | CFU/g |  |  |  | TM 610
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu |  |  |  | TM 610
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Natural (Done) |  |  |  |  | Canister & Lid Inspection
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | NY-Analytical | 765 | Canister & Lid Inspection | Handle Internal Diameter | Handle Internal Diameter | 0.6405 | in | FDA,3 | 0.639 | 0.645 | Canister & Lid Inspection
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | NY-Analytical | 765 | Canister & Lid Inspection | Overall Length | Overall Length | 3.156 | in | FDA,3 | 3.149 | 3.189 | Canister & Lid Inspection
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 1.5 | g | FDA,2 |  |  | TM 610
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer |  |  |  |  | TM 610
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 39 | in | FDA,0 | 39 | 40 | Applicator Measurements
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 36 | in | FDA,3 | 35.875 | 36.125 | Applicator Measurements
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (inches) (TM 102) | Caliper | 0.0082 | in | FDA,3 | 0.006 | 0.009 | See Analyte for Method #
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 34.44 | GSM | FDA,1 | 31.5 | 38.5 | TM 101
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 842.87 | g/in | FDA,0 | 250 | 1250 | TM 103
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 1755.59 | g/in | FDA,0 | 800 | 2600 | TM 103
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 748 | Visual Inspection | Color | Color | White with no off colors or shading [Done] |  |  |  |  | TM 208
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 1.55 | g | FDA,2 |  |  | TM 610
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 1 | cfu/device | FDA,1 |  |  | TM 610
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0.12903 | CFU/g | FDA,1 |  | 100 | TM 610
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 2 | cfu | FDA,0 |  |  | TM 610
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Negative Rods (OOS) |  |  |  |  | WI QAS 076
3XP138/Applicator_36B | 3XP138 | Applicator_36B | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements
3XP138/Applicator_36B | 3XP138 | Applicator_36B | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 45 | in | FDA,0 | 43 | 47 | Applicator Measurements
3XP138/Applicator_36B | 3XP138 | Applicator_36B | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 37.5 | in | FDA,3 | 37.375 | 37.625 | Applicator Measurements
3XP138/Applicator_36B | 3XP138 | Applicator_36B | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 40.04 | GSM | FDA,0 | 35 | 45 | TM 101
3XP138/Applicator_36B | 3XP138 | Applicator_36B | NY-Analytical | 757 | Caliper | Caliper_CoA_(uM) | Caliper | 419.8 | um | FDA,0 | 340 | 540 | TM 102
3XP138/Applicator_36B | 3XP138 | Applicator_36B | NY-Analytical | 758 | Tensile Strength | Wet_CD_CoA_(g/in) | Wet CD CoA | 1069 | g/in | FDA,0 | 400 |  | TM 103
3XP138/Applicator_36B | 3XP138 | Applicator_36B | NY-Analytical | 758 | Tensile Strength | Wet_MD_CoA_(g/in) | Wet MD CoA | 2099 | g/in | FDA,0 | 1000 |  | TM 103
3XP138/Applicator_36B | 3XP138 | Applicator_36B | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208
3XP138/Applicator_36B | 3XP138 | Applicator_36B | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208
3XP138/Applicator_36B | 3XP138 | Applicator_36B | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 30 | in | FDA,3 | 29.875 | 30.125 | Applicator Measurements
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 37.03 | GSM | FDA,1 | 31.5 | 38.5 | TM 101
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_CoA_(mm) | Caliper | 0.33 | mm | FDA,2 | 0.27 | 0.41 | TM 102
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 599.1 | g/in | FDA,0 | 300 |  | TM 103
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 3682.03 | g/in | FDA,0 | 2000 | 4400 | TM 103
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 748 | Visual Inspection | Color | Color | White with no off colors or shading [Done] |  |  |  |  | TM 208
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) |  |  |  |  | TM 208
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 0.98 | g | FDA,2 |  |  | TM 610
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 1 | cfu/device | FDA,1 |  |  | TM 610
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0.91837 | CFU/g | FDA,1 |  | 100 | TM 610
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 1 | cfu | FDA,0 |  |  | TM 610
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | 3XP253 | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Positive Rods (Done) |  |  |  |  | WI QAS 076
3XP540/Applicator_NY_Skip lot testing | 3XP540 | Applicator_NY_Skip lot testing | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (COA) | Roll Diameter (COA) |  | in | FDA,0 | 41 | 45 | Applicator Measurements
3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 3XP540 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (COA) | Roll Diameter (COA) | 43 | in | FDA,0 | 41 | 45 | Applicator Measurements
3XP540/Applicator_NY_Skip lot testing | 3XP540 | Applicator_NY_Skip lot testing | NY-Analytical | 759 | Applicator Measurements | Roll width (COA) | Roll width (COA) |  | in | FDA,2 | 41.75 | 42.25 | Applicator Measurements
3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 3XP540 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 759 | Applicator Measurements | Roll width (COA) | Roll width (COA) | 42 | in | FDA,2 | 41.75 | 42.25 | Applicator Measurements
3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 3XP540 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Basis Weight
3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 3XP540 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 52.31 | GSM | FDA,0 | 45 | 55 | TM 101
3XP540/Applicator_NY_Skip lot testing | 3XP540 | Applicator_NY_Skip lot testing | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 49.551 | GSM | FDA,0 | 45 | 55 | TM 101
3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 3XP540 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper | 0.42 | mm | FDA,2 | 0.51 | 0.63 | TM 102
3XP540/Applicator_NY_Skip lot testing | 3XP540 | Applicator_NY_Skip lot testing | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper | 0.46 | mm | FDA,2 | 0.51 | 0.63 | TM 102
3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 3XP540 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 758 | Tensile Strength | Dry_CD_(g/in) | Dry CD | 2372 | g/in | FDA,0 | 1801 | 2821 | TM 103
3XP540/Applicator_NY_Skip lot testing | 3XP540 | Applicator_NY_Skip lot testing | NY-Analytical | 758 | Tensile Strength | Dry_CD_(g/in) | Dry CD |  | g/in | FDA,0 | 1801 | 2821 | TM 103
3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 3XP540 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 758 | Tensile Strength | Dry_MD_(g/in) | Dry CD | 9752 | g/in | FDA,0 | 6675 | 9447 | TM 103
3XP540/Applicator_NY_Skip lot testing | 3XP540 | Applicator_NY_Skip lot testing | NY-Analytical | 758 | Tensile Strength | Dry_MD_(g/in) | Dry CD |  | g/in | FDA,0 | 6675 | 9447 | TM 103
3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 3XP540 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208
3XP540/Applicator_NY_Skip lot testing | 3XP540 | Applicator_NY_Skip lot testing | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208
3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 3XP540 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208
3XP540/Applicator_NY_Skip lot testing | 3XP540 | Applicator_NY_Skip lot testing | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208
3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 3XP540 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 748 | Visual Inspection | Odor | Odor | No objectionable odor Wet or Dry [Done] |  |  |  |  | TM 208
3XP540/Applicator_NY_Skip lot testing | 3XP540 | Applicator_NY_Skip lot testing | NY-Analytical | 748 | Visual Inspection | Odor | Odor | No objectionable odor Wet or Dry [Done] |  |  |  |  | TM 208
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 43 | in | FDA,0 | 41 | 45 | Applicator Measurements
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 38.75 | in | FDA,2 | 38.5 | 39 | Applicator Measurements
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Basis Weight
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 49.94 | GSM | FDA,1 | 45 | 55 | TM 101
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_CoA_(mm) | Caliper | 0.56 | mm | FDA,2 | 0.51 | 0.63 | TM 102
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 758 | Tensile Strength | Dry_CD_(g/in) | Dry CD | 2635 | g/in | FDA,0 | 1801 | 2821 | TM 103
3XP667/Applicator_NY_Skip lot testing | 3XP667 | Applicator_NY_Skip lot testing | NY-Analytical | 758 | Tensile Strength | Dry_CD_(g/in) | Dry CD | 1805 | g/in | FDA,0 | 1801 | 2821 | TM 103
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 758 | Tensile Strength | Dry_MD_(g/in) | Dry CD | 8881 | g/in | FDA,0 | 6675 | 9447 | TM 103
3XP667/Applicator_NY_Skip lot testing | 3XP667 | Applicator_NY_Skip lot testing | NY-Analytical | 758 | Tensile Strength | Dry_MD_(g/in) | Dry CD | 8039 | g/in | FDA,0 | 6675 | 9447 | TM 103
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) |  |  |  |  | TM 208
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 748 | Visual Inspection | Odor | Odor | No objectionable odor Wet or Dry [Done] |  |  |  |  | TM 208
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 0.94 | g | FDA,2 |  |  | TM 610
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer |  |  |  |  | TM 610
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
4YBBY12/4YBBY12 | 4YBBY12 | 4YBBY12 | JB-Analytical | 398 | pH | pH | pH |  |  | FDA,1 | 10.8 | 11.3 | TM200
4YBBY12/4YBBY12 | 4YBBY12 | 4YBBY12 | JB-Analytical | 902 | Bulk Liquid Inspection | Appearance (TM 208) | Appearance | [5 empty] | See Analyte for Method #
4YBBY12/4YBBY12 | 4YBBY12 | 4YBBY12 | JB-Analytical | 902 | Bulk Liquid Inspection | Color (TM 208) | Color | [5 empty] | See Analyte for Method #
4YBBY12/4YBBY12 | 4YBBY12 | 4YBBY12 | JB-Analytical | 902 | Bulk Liquid Inspection | Contamination (TM 208) | Contamination | [5 empty] | See Analyte for Method #
4YBBY12/4YBBY12 | 4YBBY12 | 4YBBY12 | JB-Analytical | 902 | Bulk Liquid Inspection | Odor (TM 208) | Odor | [5 empty] | See Analyte for Method #
4YBBY12/BL_JB 20X Sodium Benz Base Blend | 4YBBY12 | BL_JB 20X Sodium Benz Base Blend | JB-Analytical | 447 | pH as is | pH | pH | 11.86 |  | FDA,2 | 11.2 | 12.6 | TM 200
4YBBY12/BL_JB 20X Sodium Benz Base Blend | 4YBBY12 | BL_JB 20X Sodium Benz Base Blend | JB-Analytical | 748 | Visual Inspection | Appearance | Appearance | Slighlty Hazy (Done) |  |  |  |  | TM 208
4YBBY12/BL_JB 20X Sodium Benz Base Blend | 4YBBY12 | BL_JB 20X Sodium Benz Base Blend | JB-Analytical | 748 | Visual Inspection | Color | Color | Colorless to Slight Yellow (Done) |  |  |  |  | TM 208
4YBBY12/BL_JB 20X Sodium Benz Base Blend | 4YBBY12 | BL_JB 20X Sodium Benz Base Blend | JB-Analytical | 748 | Visual Inspection | Contamination | Contamination | None (Done) |  |  |  |  | TM 208
4YBBY12/BL_JB 20X Sodium Benz Base Blend | 4YBBY12 | BL_JB 20X Sodium Benz Base Blend | JB-Analytical | 748 | Visual Inspection | Odor | Odor | Unscented (Done) |  |  |  |  | TM 208
4YBBY12/BL_JB 20X Sodium Benz Base Blend | 4YBBY12 | BL_JB 20X Sodium Benz Base Blend | JB-Analytical | 425 | % Sodium Benzoate | Sodium Benzoate | Sodium Benzoate | 20.47 | % | FDA,2 | 17.86 | 20.68 | TM 395
4YBBY12/4YBBY12 | 4YBBY12 | 4YBBY12 | JB-Analytical | 425 | % Sodium Benzoate | Sodium Benzoate | Sodium Benzoate |  | % | FDA,2 | 8.46 | 10.34 | TM 395
4YBBY12/BL_JB 20X Sodium Benz Base Blend | 4YBBY12 | BL_JB 20X Sodium Benz Base Blend | JB-Analytical | 425 | % Sodium Benzoate | Sodium Benzoate Final | Sodium Benzoate Final | 20.41 | % | FDA,2 | 17.86 | 20.68 | TM 395
4YBBY12/4YBBY12 | 4YBBY12 | 4YBBY12 | JB-Analytical | 425 | % Sodium Benzoate | Sodium Benzoate Final | Sodium Benzoate Final |  | % | FDA,2 | 8.46 | 10.34 | TM 395
4YBBY23/4YBBY23 | 4YBBY23 | 4YBBY23 | JB-Analytical | 902 | Bulk Liquid Inspection | Appearance (TM 208) | Appearance | [5 empty] | See Analyte for Method #
4YBBY23/BL_JB 40X Natural Baby W/Vitamin E | 4YBBY23 | BL_JB 40X Natural Baby W/Vitamin E | JB-Analytical | 902 | Bulk Liquid Inspection | Appearance (TM 208) | Appearance | Slightly Hazy Liquid (Done) |  |  |  |  | See Analyte for Method #
4YBBY23/4YBBY23 | 4YBBY23 | 4YBBY23 | JB-Analytical | 902 | Bulk Liquid Inspection | Color (TM 208) | Color | [5 empty] | See Analyte for Method #
4YBBY23/BL_JB 40X Natural Baby W/Vitamin E | 4YBBY23 | BL_JB 40X Natural Baby W/Vitamin E | JB-Analytical | 902 | Bulk Liquid Inspection | Color (TM 208) | Color | Slightly Yellow Liquid (Done) |  |  |  |  | See Analyte for Method #
4YBBY23/4YBBY23 | 4YBBY23 | 4YBBY23 | JB-Analytical | 902 | Bulk Liquid Inspection | Contamination (TM 208) | Contamination | [5 empty] | See Analyte for Method #
4YBBY23/BL_JB 40X Natural Baby W/Vitamin E | 4YBBY23 | BL_JB 40X Natural Baby W/Vitamin E | JB-Analytical | 902 | Bulk Liquid Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | See Analyte for Method #
4YBBY23/4YBBY23 | 4YBBY23 | 4YBBY23 | JB-Analytical | 902 | Bulk Liquid Inspection | Odor (TM 208) | Odor | [5 empty] | See Analyte for Method #
4YBBY23/BL_JB 40X Natural Baby W/Vitamin E | 4YBBY23 | BL_JB 40X Natural Baby W/Vitamin E | JB-Analytical | 902 | Bulk Liquid Inspection | Odor (TM 208) | Odor | Unscented (Done) |  |  |  |  | See Analyte for Method #
4YBBY23/BL_JB 40X Natural Baby W/Vitamin E | 4YBBY23 | BL_JB 40X Natural Baby W/Vitamin E | JB-Analytical | 902 | Bulk Liquid Inspection | pH (TM 200) | pH | 11.88 |  | FDA,2 | 11.2 | 12.6 | See Analyte for Method #
4YBBY23/4YBBY23 | 4YBBY23 | 4YBBY23 | JB-Analytical | 398 | pH | pH | pH |  |  | FDA,1 | 10.8 | 11.3 | TM 200
4YBBY23/4YBBY23 | 4YBBY23 | 4YBBY23 | JB-Analytical | 425 | % Sodium Benzoate | Sodium Benzoate | Sodium Benzoate |  | % | FDA,2 | 8.46 | 10.34 | TM 395
4YBBY23/BL_JB 40X Natural Baby W/Vitamin E | 4YBBY23 | BL_JB 40X Natural Baby W/Vitamin E | JB-Analytical | 425 | % Sodium Benzoate | Sodium Benzoate | Sodium Benzoate | 20.11 | % | FDA,2 | 17.86 | 20.68 | TM 395
4YBBY23/4YBBY23 | 4YBBY23 | 4YBBY23 | JB-Analytical | 425 | % Sodium Benzoate | Sodium Benzoate Final | Sodium Benzoate Final | 0 | % | FDA,2 | 8.46 | 10.34 | TM 395
4YBBY23/BL_JB 40X Natural Baby W/Vitamin E | 4YBBY23 | BL_JB 40X Natural Baby W/Vitamin E | JB-Analytical | 425 | % Sodium Benzoate | Sodium Benzoate Final | Sodium Benzoate Final | 20.13 | % | FDA,2 | 17.86 | 20.68 | TM 395
5AA347BT/Box_NY_0.020SBS_300_3.375-1.313-2.625 | 5AA347BT | Box_NY_0.020SBS_300_3.375-1.313-2.625 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 315 |  | FDA,0 |  |  | Box Inspection
5AA347BT/Box_NY_0.020SBS_300_3.375-1.313-2.625 | 5AA347BT | Box_NY_0.020SBS_300_3.375-1.313-2.625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5AA347BT/Box_NY_0.020SBS_300_3.375-1.313-2.625 | 5AA347BT | Box_NY_0.020SBS_300_3.375-1.313-2.625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.019 | in | FDA,3 | 0.019 | 0.021 | Box Inspection
5AA347BT/Box_NY_0.020SBS_300_3.375-1.313-2.625 | 5AA347BT | Box_NY_0.020SBS_300_3.375-1.313-2.625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection
5AA347BT/Box_NY_0.020SBS_300_3.375-1.313-2.625 | 5AA347BT | Box_NY_0.020SBS_300_3.375-1.313-2.625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Open Top/Auto Bottom Tray (Done) |  |  |  |  | Box Inspection
5AA347BT/Box_NY_Skip lot testing | 5AA347BT | Box_NY_Skip lot testing | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Box Inspection
5AA347BT/Box_NY_0.020SBS_300_3.375-1.313-2.625 | 5AA347BT | Box_NY_0.020SBS_300_3.375-1.313-2.625 | NY-Analytical | 772 | Box Inspection | Height | Height | 2.625 | in | FDA,5 | 2.5938 | 2.6562 | Box Inspection
5AA347BT/Box_NY_0.020SBS_300_3.375-1.313-2.625 | 5AA347BT | Box_NY_0.020SBS_300_3.375-1.313-2.625 | NY-Analytical | 772 | Box Inspection | Length | Length | 3.375 | in | FDA,5 | 3.3438 | 3.4062 | Box Inspection
5AA347BT/Box_NY_0.020SBS_300_3.375-1.313-2.625 | 5AA347BT | Box_NY_0.020SBS_300_3.375-1.313-2.625 | NY-Analytical | 772 | Box Inspection | Width | Width | 1.3125 | in | FDA,5 | 1.2817 | 1.3442 | Box Inspection
5AA404S48/Box_NY_0.018SBS_111_6-5.875-3.625 | 5AA404S48 | Box_NY_0.018SBS_111_6-5.875-3.625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy |  | in |  |  |  | Box Inspection
5AA404S48/Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | 5AA404S48 | Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5AA404S48/Box_NY_0.018SBS_111_6-5.875-3.625 | 5AA404S48 | Box_NY_0.018SBS_111_6-5.875-3.625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper |  | in | FDA,3 | 0.017 | 0.019 | Box Inspection
5AA404S48/Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | 5AA404S48 | Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.018 | in | FDA,3 | 0.017 | 0.019 | Box Inspection
5AA404S48/Box_NY_0.018SBS_111_6-5.875-3.625 | 5AA404S48 | Box_NY_0.018SBS_111_6-5.875-3.625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | [5 empty] | Box Inspection
5AA404S48/Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | 5AA404S48 | Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection
5AA404S48/Box_NY_0.018SBS_111_6-5.875-3.625 | 5AA404S48 | Box_NY_0.018SBS_111_6-5.875-3.625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | [5 empty] | Box Inspection
5AA404S48/Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | 5AA404S48 | Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection
5AA404S48/Box_NY_0.018SBS_111_6-5.875-3.625 | 5AA404S48 | Box_NY_0.018SBS_111_6-5.875-3.625 | NY-Analytical | 772 | Box Inspection | Height | Height |  | in | FDA,5 | 3.5938 | 3.6562 | Box Inspection
5AA404S48/Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | 5AA404S48 | Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | NY-Analytical | 772 | Box Inspection | Height | Height | 6.0625 | in | FDA,5 | 6.0312 | 6.0938 | Box Inspection
5AA404S48/Box_NY_0.018SBS_111_6-5.875-3.625 | 5AA404S48 | Box_NY_0.018SBS_111_6-5.875-3.625 | NY-Analytical | 772 | Box Inspection | Length | Length |  | in | FDA,5 | 5.9688 | 6.0312 | Box Inspection
5AA404S48/Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | 5AA404S48 | Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | NY-Analytical | 772 | Box Inspection | Length | Length | 7.1875 | in | FDA,5 | 7.1562 | 7.2188 | Box Inspection
5AA404S48/Box_NY_0.018SBS_111_6-5.875-3.625 | 5AA404S48 | Box_NY_0.018SBS_111_6-5.875-3.625 | NY-Analytical | 772 | Box Inspection | Width | Width |  | in | FDA,5 | 5.8438 | 5.9062 | Box Inspection
5AA404S48/Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | 5AA404S48 | Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.6848 | in | FDA,5 | 3.6562 | 3.7188 | Box Inspection
5AA617S96/Box_NY_0.018SBS_47B_3.625-3.875-6 | 5AA617S96 | Box_NY_0.018SBS_47B_3.625-3.875-6 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5AA617S96/Box_NY_0.018SBS_47B_3.625-3.875-6 | 5AA617S96 | Box_NY_0.018SBS_47B_3.625-3.875-6 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.019 | in | FDA,3 | 0.017 | 0.019 | Box Inspection
5AA617S96/Box_NY_0.018SBS_47B_3.625-3.875-6 | 5AA617S96 | Box_NY_0.018SBS_47B_3.625-3.875-6 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection
5AA617S96/Box_NY_0.018SBS_47B_3.625-3.875-6 | 5AA617S96 | Box_NY_0.018SBS_47B_3.625-3.875-6 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Straight Tuck End - French Style (Done) |  |  |  |  | Box Inspection
5AA617S96/Box_NY_0.018SBS_47B_3.625-3.875-6 | 5AA617S96 | Box_NY_0.018SBS_47B_3.625-3.875-6 | NY-Analytical | 772 | Box Inspection | Height | Height | 6 | in | FDA,5 | 5.9688 | 6.0312 | Box Inspection
5AA617S96/Box_NY_0.018SBS_47B_3.625-3.875-6 | 5AA617S96 | Box_NY_0.018SBS_47B_3.625-3.875-6 | NY-Analytical | 772 | Box Inspection | Length | Length | 3.625 | in | FDA,5 | 3.5938 | 3.6562 | Box Inspection
5AA617S96/Box_NY_0.018SBS_47B_3.625-3.875-6 | 5AA617S96 | Box_NY_0.018SBS_47B_3.625-3.875-6 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.875 | in | FDA,5 | 3.8438 | 3.9062 | Box Inspection
5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | 5AA632S60 | Box_NY_0.020SBS_101_6-3.625-7 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 3 |  |  |  |  | Box Inspection
5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | 5AA632S60 | Box_NY_0.020SBS_101_6-3.625-7 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | 5AA632S60 | Box_NY_0.020SBS_101_6-3.625-7 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.021 | in | FDA,3 | 0.019 | 0.021 | Box Inspection
5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | 5AA632S60 | Box_NY_0.020SBS_101_6-3.625-7 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection
5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | 5AA632S60 | Box_NY_0.020SBS_101_6-3.625-7 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom - French Style (Done) |  |  |  |  | Box Inspection
5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | 5AA632S60 | Box_NY_0.020SBS_101_6-3.625-7 | NY-Analytical | 772 | Box Inspection | Height | Height | 7 | in | FDA,5 | 6.9688 | 7.0312 | Box Inspection
5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | 5AA632S60 | Box_NY_0.020SBS_101_6-3.625-7 | NY-Analytical | 772 | Box Inspection | Length | Length | 6 | in | FDA,5 | 5.9688 | 6.0312 | Box Inspection
5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | 5AA632S60 | Box_NY_0.020SBS_101_6-3.625-7 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.625 | in | FDA,5 | 3.5938 | 3.6562 | Box Inspection
5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | 5AA911TY | Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom w/perf |  |  |  |  | Box Inspection
5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | 5AA911TY | Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | Box Inspection
5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | 5AA911TY | Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | 5AA911TY | Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0195 | in | FDA,3 | 0.019 | 0.021 | Box Inspection
5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | 5AA911TY | Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection
5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | 5AA911TY | Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | NY-Analytical | 772 | Box Inspection | Height | Height | 7.94 | in | FDA,5 | 7.9062 | 7.9688 | Box Inspection
5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | 5AA911TY | Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | NY-Analytical | 772 | Box Inspection | Length | Length | 4.51 | in | FDA,5 | 4.4688 | 4.5312 | Box Inspection
5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | 5AA911TY | Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | NY-Analytical | 772 | Box Inspection | Width | Width | 7.94 | in | FDA,5 | 7.9062 | 7.9688 | Certificate of Conformance
5AB123CP/Label_NY_Skip lot testing | 5AB123CP | Label_NY_Skip lot testing | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Certificate of Conformance
5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | 5AB123CP | Label_NY_5AB123CP_1.9375-1.875 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | [5 empty] | Label Inspection
5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | 5AB123CP | Label_NY_5AB123CP_1.9375-1.875 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | 5AB123CP | Label_NY_5AB123CP_1.9375-1.875 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | 5AB123CP | Label_NY_5AB123CP_1.9375-1.875 | NY-Analytical | 770 | Label Inspection | Length | Length | 1.875 | in | FDA,4 | 1.875 | 2 | TM 143 / Visual
5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | 5AB123CP | Label_NY_5AB123CP_1.9375-1.875 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | 5AB123CP | Label_NY_5AB123CP_1.9375-1.875 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | 5AB123CP | Label_NY_5AB123CP_1.9375-1.875 | NY-Analytical | 770 | Label Inspection | Width | Width | 1.9375 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual
5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | 5AB123ST | Box_NY_0.018SBS_107_10.25-2.375-2.75 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | 5AB123ST | Box_NY_0.018SBS_107_10.25-2.375-2.75 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.018 | in | FDA,3 | 0.017 | 0.019 | Box Inspection
5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | 5AB123ST | Box_NY_0.018SBS_107_10.25-2.375-2.75 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection
5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | 5AB123ST | Box_NY_0.018SBS_107_10.25-2.375-2.75 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | TT/AB w/friction lock (Done) |  |  |  |  | Box Inspection
5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | 5AB123ST | Box_NY_0.018SBS_107_10.25-2.375-2.75 | NY-Analytical | 772 | Box Inspection | Height | Height | 2.7515 | in | FDA,5 | 2.7188 | 2.7812 | Box Inspection
5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | 5AB123ST | Box_NY_0.018SBS_107_10.25-2.375-2.75 | NY-Analytical | 772 | Box Inspection | Length | Length | 10.22 | in | FDA,5 | 10.2188 | 10.2812 | Box Inspection
5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | 5AB123ST | Box_NY_0.018SBS_107_10.25-2.375-2.75 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.351 | in | FDA,5 | 2.3438 | 2.4062 | Box Inspection
5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | 5AD02424 | Box_NY_0.018SBS_107_4-2.03125-4.5 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top Auto Bottom w/ Tab Lock |  |  |  |  | Box Inspection
5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | 5AD02424 | Box_NY_0.018SBS_107_4-2.03125-4.5 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | 5AD02424 | Box_NY_0.018SBS_107_4-2.03125-4.5 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0175 | in | FDA,3 | 0.017 | 0.019 | Box Inspection
5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | 5AD02424 | Box_NY_0.018SBS_107_4-2.03125-4.5 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection
5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | 5AD02424 | Box_NY_0.018SBS_107_4-2.03125-4.5 | NY-Analytical | 772 | Box Inspection | Height | Height | 4.5 | in | FDA,5 | 4.4688 | 4.5312 | Box Inspection
5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | 5AD02424 | Box_NY_0.018SBS_107_4-2.03125-4.5 | NY-Analytical | 772 | Box Inspection | Length | Length | 4 | in | FDA,5 | 3.9688 | 4.0312 | Box Inspection
5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | 5AD02424 | Box_NY_0.018SBS_107_4-2.03125-4.5 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.0312 | in | FDA,5 | 2 | 2.0625 | Box Inspection
5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | 5AH45016 | Box_NY_0.018SBS_101_3.875-2-3.5 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 200 |  |  |  |  | Box Inspection
5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | 5AH45016 | Box_NY_0.018SBS_101_3.875-2-3.5 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | 5AH45016 | Box_NY_0.018SBS_101_3.875-2-3.5 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.018 | in | FDA,3 | 0.017 | 0.019 | Box Inspection
5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | 5AH45016 | Box_NY_0.018SBS_101_3.875-2-3.5 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection
5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | 5AH45016 | Box_NY_0.018SBS_101_3.875-2-3.5 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | French Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection
5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | 5AH45016 | Box_NY_0.018SBS_101_3.875-2-3.5 | NY-Analytical | 772 | Box Inspection | Height | Height | 3.5 | in | FDA,5 | 3.4688 | 3.5312 | Box Inspection
5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | 5AH45016 | Box_NY_0.018SBS_101_3.875-2-3.5 | NY-Analytical | 772 | Box Inspection | Length | Length | 3.875 | in | FDA,5 | 3.8438 | 3.9062 | Box Inspection
5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | 5AH45016 | Box_NY_0.018SBS_101_3.875-2-3.5 | NY-Analytical | 772 | Box Inspection | Width | Width | 2 | in | FDA,5 | 1.9688 | 2.0312 | Box Inspection
5AH58195/Box_NY_0.020SBS_127_6.25-2.5625-4.625 | 5AH58195 | Box_NY_0.020SBS_127_6.25-2.5625-4.625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Straight Tuck w/Partition Dispenser Box |  |  |  |  | Box Inspection
5AH58195/Box_NY_0.020SBS_127_6.25-2.5625-4.625 | 5AH58195 | Box_NY_0.020SBS_127_6.25-2.5625-4.625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5AH58195/Box_NY_0.020SBS_127_6.25-2.5625-4.625 | 5AH58195 | Box_NY_0.020SBS_127_6.25-2.5625-4.625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0195 | in | FDA,3 | 0.019 | 0.021 | Box Inspection
5AH58195/Box_NY_0.020SBS_127_6.25-2.5625-4.625 | 5AH58195 | Box_NY_0.020SBS_127_6.25-2.5625-4.625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection
5AH58195/Box_NY_0.020SBS_127_6.25-2.5625-4.625 | 5AH58195 | Box_NY_0.020SBS_127_6.25-2.5625-4.625 | NY-Analytical | 772 | Box Inspection | Height | Height | 4.625 | in | FDA,5 | 4.5938 | 4.6562 | Box Inspection
5AH58195/Box_NY_0.020SBS_127_6.25-2.5625-4.625 | 5AH58195 | Box_NY_0.020SBS_127_6.25-2.5625-4.625 | NY-Analytical | 772 | Box Inspection | Length | Length | 6.25 | in | FDA,5 | 6.2188 | 6.2812 | Box Inspection
5AH58195/Box_NY_0.020SBS_127_6.25-2.5625-4.625 | 5AH58195 | Box_NY_0.020SBS_127_6.25-2.5625-4.625 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.5625 | in | FDA,5 | 2.5312 | 2.5938 | Box Inspection
5AJ130BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | 5AJ130BT | Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size |  |  | FDA,0 |  |  | Box Inspection
5AJ130BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | 5AJ130BT | Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | [5 empty] | Box Inspection
5AJ130BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | 5AJ130BT | Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper |  | in | FDA,3 | 0.019 | 0.021 | Box Inspection
5AJ130BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | 5AJ130BT | Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | [5 empty] | Box Inspection
5AJ130BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | 5AJ130BT | Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | [5 empty] | Box Inspection
5AJ130BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | 5AJ130BT | Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | NY-Analytical | 772 | Box Inspection | Height | Height |  | in | FDA,5 | 2.5938 | 2.6562 | Box Inspection
5AJ130BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | 5AJ130BT | Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | NY-Analytical | 772 | Box Inspection | Length | Length |  | in | FDA,5 | 3.2812 | 3.3438 | Box Inspection
5AJ130BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | 5AJ130BT | Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | NY-Analytical | 772 | Box Inspection | Width | Width |  | in | FDA,5 | 1.4062 | 1.4688 | Box Inspection
5AP027TY/Box_57A | 5AP027TY | Box_57A | NY-Analytical | 772 | Box Inspection | Appearance | Appearance | Meets specification (Done) |  |  |  |  | Box Inspection
5AP027TY/Box_57A | 5AP027TY | Box_57A | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Box Inspection
5AP027TY/Box_57A | 5AP027TY | Box_57A | NY-Analytical | 772 | Box Inspection | Color | Color | Clear (Done) |  |  |  |  | Box Inspection
5AP027TY/Box_57A | 5AP027TY | Box_57A | NY-Analytical | 772 | Box Inspection | Functionality | Functionality | Meets specification (Done) |  |  |  |  | Box Inspection
5AP027TY/Box_57A | 5AP027TY | Box_57A | NY-Analytical | 772 | Box Inspection | Length | Length | 10.75 | in | FDA,3 | 10.66 | 10.84 | Box Inspection
5AP027TY/Box_57A | 5AP027TY | Box_57A | NY-Analytical | 772 | Box Inspection | Odor | Odor | None Detectable (Done) |  |  |  |  | Box Inspection
5AP027TY/Box_57A | 5AP027TY | Box_57A | NY-Analytical | 772 | Box Inspection | Width | Width | 5.625 | in | FDA,3 | 5.535 | 5.715 | Box Inspection
5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | 5AP07530 | Box_NY_0.020BC_110_6.375-3.4375-4.75 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | Box Inspection
5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | 5AP07530 | Box_NY_0.020BC_110_6.375-3.4375-4.75 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | 5AP07530 | Box_NY_0.020BC_110_6.375-3.4375-4.75 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.02 | in | FDA,3 | 0.019 | 0.021 | Box Inspection
5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | 5AP07530 | Box_NY_0.020BC_110_6.375-3.4375-4.75 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | BC (Done) |  |  |  |  | Box Inspection
5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | 5AP07530 | Box_NY_0.020BC_110_6.375-3.4375-4.75 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | TT/AB w/finger holes (Done) |  |  |  |  | Box Inspection
5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | 5AP07530 | Box_NY_0.020BC_110_6.375-3.4375-4.75 | NY-Analytical | 772 | Box Inspection | Height | Height | 4.771 | in | FDA,5 | 4.7188 | 4.7812 | Box Inspection
5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | 5AP07530 | Box_NY_0.020BC_110_6.375-3.4375-4.75 | NY-Analytical | 772 | Box Inspection | Length | Length | 6.395 | in | FDA,5 | 6.3438 | 6.4062 | Box Inspection
5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | 5AP07530 | Box_NY_0.020BC_110_6.375-3.4375-4.75 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.4225 | in | FDA,5 | 3.4062 | 3.4688 | Box Inspection
5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | 5AP209SL | Box_NY_0.020BC_220_5.09375-3.6875-4.625 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 30 |  | FDA,0 |  |  | Box Inspection
5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | 5AP209SL | Box_NY_0.020BC_220_5.09375-3.6875-4.625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | 5AP209SL | Box_NY_0.020BC_220_5.09375-3.6875-4.625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0195 | in | FDA,3 | 0.019 | 0.021 | Box Inspection
5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | 5AP209SL | Box_NY_0.020BC_220_5.09375-3.6875-4.625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | BC (Done) |  |  |  |  | Box Inspection
5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | 5AP209SL | Box_NY_0.020BC_220_5.09375-3.6875-4.625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Diecut Glued Sleeve w/ 2-1" fingerholes (Done) |  |  |  |  | Box Inspection
5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | 5AP209SL | Box_NY_0.020BC_220_5.09375-3.6875-4.625 | NY-Analytical | 772 | Box Inspection | Height | Height | 4.61 | in | FDA,5 | 4.5938 | 4.6562 | Box Inspection
5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | 5AP209SL | Box_NY_0.020BC_220_5.09375-3.6875-4.625 | NY-Analytical | 772 | Box Inspection | Length | Length | 5.07 | in | FDA,5 | 5.0625 | 5.125 | Box Inspection
5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | 5AP209SL | Box_NY_0.020BC_220_5.09375-3.6875-4.625 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.7 | in | FDA,5 | 3.6562 | 3.7188 | Box Inspection
5AP21195/Box_01M | 5AP21195 | Box_01M | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | Bending Chip |  |  |  |  | Box Inspection
5AP21195/Box_01M | 5AP21195 | Box_01M | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5AP21195/Box_01M | 5AP21195 | Box_01M | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0235 | in | FDA,3 | 0.023 | 0.025 | Box Inspection
5AP21195/Box_01M | 5AP21195 | Box_01M | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection
5AP21195/Box_01M | 5AP21195 | Box_01M | NY-Analytical | 772 | Box Inspection | Height | Height | 5.875 | in | FDA,5 | 5.8438 | 5.9062 | Box Inspection
5AP21195/Box_01M | 5AP21195 | Box_01M | NY-Analytical | 772 | Box Inspection | Length | Length | 4.625 | in | FDA,5 | 4.5938 | 4.6562 | Box Inspection
5AP21195/Box_01M | 5AP21195 | Box_01M | NY-Analytical | 772 | Box Inspection | Width | Width | 4.5 | in | FDA,5 | 4.4688 | 4.5312 | Box Inspection
5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | 5AP227TY | Box_NY_0.020SBS_300_4.625-5-3.625 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | Box Inspection
5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | 5AP227TY | Box_NY_0.020SBS_300_4.625-5-3.625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | 5AP227TY | Box_NY_0.020SBS_300_4.625-5-3.625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.02 | in | FDA,3 | 0.019 | 0.021 | Box Inspection
5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | 5AP227TY | Box_NY_0.020SBS_300_4.625-5-3.625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection
5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | 5AP227TY | Box_NY_0.020SBS_300_4.625-5-3.625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Open Top/Auto Bottom Tray (Done) |  |  |  |  | Box Inspection
5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | 5AP227TY | Box_NY_0.020SBS_300_4.625-5-3.625 | NY-Analytical | 772 | Box Inspection | Height | Height | 3.61 | in | FDA,5 | 3.5938 | 3.6562 | Box Inspection
5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | 5AP227TY | Box_NY_0.020SBS_300_4.625-5-3.625 | NY-Analytical | 772 | Box Inspection | Length | Length | 4.6 | in | FDA,5 | 4.5938 | 4.6562 | Box Inspection
5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | 5AP227TY | Box_NY_0.020SBS_300_4.625-5-3.625 | NY-Analytical | 772 | Box Inspection | Width | Width | 5.01 | in | FDA,5 | 4.9688 | 5.0312 | Box Inspection
5AP229TY/Box_NY_0.024SBS_303_8.228-3.661-6.024 | 5AP229TY | Box_NY_0.024SBS_303_8.228-3.661-6.024 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5AP229TY/Box_NY_0.024SBS_303_8.228-3.661-6.024 | 5AP229TY | Box_NY_0.024SBS_303_8.228-3.661-6.024 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.023 | in | FDA,3 | 0.023 | 0.025 | Box Inspection
5AP229TY/Box_NY_0.024SBS_303_8.228-3.661-6.024 | 5AP229TY | Box_NY_0.024SBS_303_8.228-3.661-6.024 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection
5AP229TY/Box_NY_0.024SBS_303_8.228-3.661-6.024 | 5AP229TY | Box_NY_0.024SBS_303_8.228-3.661-6.024 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | A/B tray w/tear-off lid (Done) |  |  |  |  | Box Inspection
5AP229TY/Box_NY_0.024SBS_303_8.228-3.661-6.024 | 5AP229TY | Box_NY_0.024SBS_303_8.228-3.661-6.024 | NY-Analytical | 772 | Box Inspection | Height | Height | 6.03 | in | FDA,5 | 5.9928 | 6.0553 | Box Inspection
5AP229TY/Box_NY_0.024SBS_303_8.228-3.661-6.024 | 5AP229TY | Box_NY_0.024SBS_303_8.228-3.661-6.024 | NY-Analytical | 772 | Box Inspection | Length | Length | 8.22 | in | FDA,5 | 8.1967 | 8.2592 | Box Inspection
5AP229TY/Box_NY_0.024SBS_303_8.228-3.661-6.024 | 5AP229TY | Box_NY_0.024SBS_303_8.228-3.661-6.024 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.64 | in | FDA,5 | 3.6298 | 3.6923 | Box Inspection
5AP308TY/Box_63A | 5AP308TY | Box_63A | NY-Analytical | 772 | Box Inspection | Appearance | Appearance | Meets specification (Done) |  |  |  |  | Box Inspection
5AP308TY/Box_63A | 5AP308TY | Box_63A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Box Inspection
5AP308TY/Box_63A | 5AP308TY | Box_63A | NY-Analytical | 772 | Box Inspection | Color | Color | Clear (Done) |  |  |  |  | Box Inspection
5AP308TY/Box_63A | 5AP308TY | Box_63A | NY-Analytical | 772 | Box Inspection | Flange - Sides | Flange - Sides | 0.075 | in | FDA,3 | 0.045 | 0.135 | Box Inspection
5AP308TY/Box_63A | 5AP308TY | Box_63A | NY-Analytical | 772 | Box Inspection | Flange - Top & Bottom | Flange - Top & Bottom | 0.099 | in | FDA,3 | 0.045 | 0.135 | Box Inspection
5AP308TY/Box_63A | 5AP308TY | Box_63A | NY-Analytical | 772 | Box Inspection | Length | Length | 7.2 | in | FDA,3 | 7.156 | 7.22 | Box Inspection
5AP308TY/Box_63A | 5AP308TY | Box_63A | NY-Analytical | 772 | Box Inspection | Odor | Odor | None Detectable (Done) |  |  |  |  | Box Inspection
5AP308TY/Box_63A | 5AP308TY | Box_63A | NY-Analytical | 772 | Box Inspection | Total Height | Total Height | 2.95 | in | FDA,3 | 2.899 | 2.963 | Box Inspection
5AP308TY/Box_63A | 5AP308TY | Box_63A | NY-Analytical | 772 | Box Inspection | Width | Width | 6.1 | in | FDA,3 | 6.093 | 6.157 | Box Inspection
5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | 5AU25206 | Box_NY_0.018SBS_100_101.6-44.45-101.6 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 200 |  | FDA,0 |  |  | Box Inspection
5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | 5AU25206 | Box_NY_0.018SBS_100_101.6-44.45-101.6 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | 5AU25206 | Box_NY_0.018SBS_100_101.6-44.45-101.6 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.018 | in | FDA,3 | 0.017 | 0.019 | Box Inspection
5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | 5AU25206 | Box_NY_0.018SBS_100_101.6-44.45-101.6 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection
5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | 5AU25206 | Box_NY_0.018SBS_100_101.6-44.45-101.6 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection
5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | 5AU25206 | Box_NY_0.018SBS_100_101.6-44.45-101.6 | NY-Analytical | 772 | Box Inspection | Height | Height | 102 | mm | FDA,1 | 100.8 | 102.4 | Box Inspection
5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | 5AU25206 | Box_NY_0.018SBS_100_101.6-44.45-101.6 | NY-Analytical | 772 | Box Inspection | Length | Length | 102 | mm | FDA,1 | 100.8 | 102.4 | Box Inspection
5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | 5AU25206 | Box_NY_0.018SBS_100_101.6-44.45-101.6 | NY-Analytical | 772 | Box Inspection | Width | Width | 45 | mm | FDA,1 | 43.7 | 45.2 | Box Inspection
5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | 5AU252IN | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | Box Inspection
5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | 5AU252IN | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Does not meet specifications (OOS) | in |  |  |  | Box Inspection
5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | 5AU252IN | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.019 | in | FDA,3 | 0.019 | 0.021 | Box Inspection
5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | 5AU252IN | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | TBC (Done) |  |  |  |  | Box Inspection
5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | 5AU252IN | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | RT (Done) |  |  |  |  | Box Inspection
5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | 5AU252IN | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Box Inspection
5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | 5AU252IN | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 772 | Box Inspection | Height | Height | 4.2812 | in | FDA,5 | 4.2812 | 4.3438 | Box Inspection
5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | 5AU252IN | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 772 | Box Inspection | Length | Length | 4.125 | in | FDA,5 | 4.0625 | 4.125 | Box Inspection
5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | 5AU252IN | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.625 | in | FDA,5 | 3.5938 | 3.6562 | Box Inspection
5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | 5AU31450 | Box_NY_0.020SBS_169_6-3.125-5.9375 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | 5AU31450 | Box_NY_0.020SBS_169_6-3.125-5.9375 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.019 | in | FDA,3 | 0.019 | 0.021 | Box Inspection
5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | 5AU31450 | Box_NY_0.020SBS_169_6-3.125-5.9375 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection
5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | 5AU31450 | Box_NY_0.020SBS_169_6-3.125-5.9375 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Dispenser (Done) |  |  |  |  | Box Inspection
5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | 5AU31450 | Box_NY_0.020SBS_169_6-3.125-5.9375 | NY-Analytical | 772 | Box Inspection | Height | Height | 5.9375 | in | FDA,5 | 5.9062 | 5.9688 | Box Inspection
5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | 5AU31450 | Box_NY_0.020SBS_169_6-3.125-5.9375 | NY-Analytical | 772 | Box Inspection | Length | Length | 6 | in | FDA,5 | 5.9688 | 6.0312 | Box Inspection
5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | 5AU31450 | Box_NY_0.020SBS_169_6-3.125-5.9375 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.125 | in | FDA,5 | 3.0938 | 3.1562 | Box Inspection
5AU44410/Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | 5AU44410 | Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5AU44410/Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | 5AU44410 | Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0155 | in | FDA,3 | 0.015 | 0.017 | Box Inspection
5AU44410/Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | 5AU44410 | Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection
5AU44410/Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | 5AU44410 | Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Straight Tuck End w/1" glued double header (Done) |  |  |  |  | Box Inspection
5AU44410/Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | 5AU44410 | Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Height | Height | 3.8125 | in | FDA,5 | 3.7812 | 3.8438 | Box Inspection
5AU44410/Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | 5AU44410 | Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Length | Length | 2.9375 | in | FDA,5 | 2.9062 | 2.9688 | Box Inspection
5AU44410/Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | 5AU44410 | Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Width | Width | 1.875 | in | FDA,5 | 1.8438 | 1.9062 | Box Inspection
5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | 5AU57314 | Box_NY_0.016CCSBS_100_3-2.5-4 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 200 |  | FDA,0 |  |  | Box Inspection
5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | 5AU57314 | Box_NY_0.016CCSBS_100_3-2.5-4 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | 5AU57314 | Box_NY_0.016CCSBS_100_3-2.5-4 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0155 | in | FDA,3 | 0.015 | 0.017 | Box Inspection
5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | 5AU57314 | Box_NY_0.016CCSBS_100_3-2.5-4 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | Clay Coated SBS (Done) |  |  |  |  | Box Inspection
5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | 5AU57314 | Box_NY_0.016CCSBS_100_3-2.5-4 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection
5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | 5AU57314 | Box_NY_0.016CCSBS_100_3-2.5-4 | NY-Analytical | 772 | Box Inspection | Height | Height | 4 | in | FDA,5 | 3.9688 | 4.0312 | Box Inspection
5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | 5AU57314 | Box_NY_0.016CCSBS_100_3-2.5-4 | NY-Analytical | 772 | Box Inspection | Length | Length | 3 | in | FDA,5 | 2.9688 | 3.0312 | Box Inspection
5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | 5AU57314 | Box_NY_0.016CCSBS_100_3-2.5-4 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.5 | in | FDA,5 | 2.4688 | 2.5312 | Box Inspection
5AU66522/Box_NY_0.016BMPB_122_3.625-3-3.8125 | 5AU66522 | Box_NY_0.016BMPB_122_3.625-3-3.8125 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5AU66522/Box_NY_0.016BMPB_122_3.625-3-3.8125 | 5AU66522 | Box_NY_0.016BMPB_122_3.625-3-3.8125 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.017 | in | FDA,3 | 0.015 | 0.017 | Box Inspection
5AU66522/Box_NY_0.016BMPB_122_3.625-3-3.8125 | 5AU66522 | Box_NY_0.016BMPB_122_3.625-3-3.8125 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | Bright Metalized Poly Board (Done) |  |  |  |  | Box Inspection
5AU66522/Box_NY_0.016BMPB_122_3.625-3-3.8125 | 5AU66522 | Box_NY_0.016BMPB_122_3.625-3-3.8125 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Straight Tuck End (Done) |  |  |  |  | Box Inspection
5AU66522/Box_NY_0.016BMPB_122_3.625-3-3.8125 | 5AU66522 | Box_NY_0.016BMPB_122_3.625-3-3.8125 | NY-Analytical | 772 | Box Inspection | Height | Height | 3.812 | in | FDA,5 | 3.7812 | 3.8438 | Box Inspection
5AU66522/Box_NY_0.016BMPB_122_3.625-3-3.8125 | 5AU66522 | Box_NY_0.016BMPB_122_3.625-3-3.8125 | NY-Analytical | 772 | Box Inspection | Length | Length | 3.625 | in | FDA,5 | 3.5938 | 3.6562 | Box Inspection
5AU66522/Box_NY_0.016BMPB_122_3.625-3-3.8125 | 5AU66522 | Box_NY_0.016BMPB_122_3.625-3-3.8125 | NY-Analytical | 772 | Box Inspection | Width | Width | 3 | in | FDA,5 | 2.9688 | 3.0312 | Box Inspection
5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | 5BU19101 | Box_NY_0.020SBS_129_5.75-3-3.875 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | 5BU19101 | Box_NY_0.020SBS_129_5.75-3-3.875 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0207 | in | FDA,3 | 0.019 | 0.021 | Box Inspection
5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | 5BU19101 | Box_NY_0.020SBS_129_5.75-3-3.875 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection
5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | 5BU19101 | Box_NY_0.020SBS_129_5.75-3-3.875 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Straight Tuck End French Style Dipsenser w/False Bottom |  |  |  |  | Box Inspection
5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | 5BU19101 | Box_NY_0.020SBS_129_5.75-3-3.875 | NY-Analytical | 772 | Box Inspection | Height | Height | 3.875 | in | FDA,5 | 3.8438 | 3.9062 | Box Inspection
5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | 5BU19101 | Box_NY_0.020SBS_129_5.75-3-3.875 | NY-Analytical | 772 | Box Inspection | Length | Length | 5.75 | in | FDA,5 | 5.7188 | 5.7812 | Box Inspection
5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | 5BU19101 | Box_NY_0.020SBS_129_5.75-3-3.875 | NY-Analytical | 772 | Box Inspection | Width | Width | 3 | in | FDA,5 | 2.9688 | 3.0312 | Box Inspection
5BU26595/Box_NY_0.024SBS_102_4.390625-4.390625-12 | 5BU26595 | Box_NY_0.024SBS_102_4.390625-4.390625-12 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom Dispenser Box |  |  |  |  | Box Inspection
5BU26595/Box_NY_0.024SBS_102_4.390625-4.390625-12 | 5BU26595 | Box_NY_0.024SBS_102_4.390625-4.390625-12 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5BU26595/Box_NY_0.024SBS_102_4.390625-4.390625-12 | 5BU26595 | Box_NY_0.024SBS_102_4.390625-4.390625-12 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.024 | in | FDA,3 | 0.023 | 0.025 | Box Inspection
5BU26595/Box_NY_0.024SBS_102_4.390625-4.390625-12 | 5BU26595 | Box_NY_0.024SBS_102_4.390625-4.390625-12 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection
5BU26595/Box_NY_0.024SBS_102_4.390625-4.390625-12 | 5BU26595 | Box_NY_0.024SBS_102_4.390625-4.390625-12 | NY-Analytical | 772 | Box Inspection | Height | Height | 12 | in | FDA,5 | 11.9688 | 12.0312 | Box Inspection
5BU26595/Box_NY_0.024SBS_102_4.390625-4.390625-12 | 5BU26595 | Box_NY_0.024SBS_102_4.390625-4.390625-12 | NY-Analytical | 772 | Box Inspection | Length | Length | 4.41 | in | FDA,5 | 4.3594 | 4.4219 | Box Inspection
5BU26595/Box_NY_0.024SBS_102_4.390625-4.390625-12 | 5BU26595 | Box_NY_0.024SBS_102_4.390625-4.390625-12 | NY-Analytical | 772 | Box Inspection | Width | Width | 4.39 | in | FDA,5 | 4.3594 | 4.4219 | Box Inspection
5BU27500/Box_NY_0.024SBS_102_4.3906-4.3906-12 | 5BU27500 | Box_NY_0.024SBS_102_4.3906-4.3906-12 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom Dispenser Box |  |  |  |  | Box Inspection
5BU27500/Box_NY_0.024SBS_102_4.3906-4.3906-12 | 5BU27500 | Box_NY_0.024SBS_102_4.3906-4.3906-12 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5BU27500/Box_NY_0.024SBS_102_4.3906-4.3906-12 | 5BU27500 | Box_NY_0.024SBS_102_4.3906-4.3906-12 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.024 | in | FDA,3 | 0.023 | 0.025 | Box Inspection
5BU27500/Box_NY_0.024SBS_102_4.3906-4.3906-12 | 5BU27500 | Box_NY_0.024SBS_102_4.3906-4.3906-12 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection
5BU27500/Box_NY_0.024SBS_102_4.3906-4.3906-12 | 5BU27500 | Box_NY_0.024SBS_102_4.3906-4.3906-12 | NY-Analytical | 772 | Box Inspection | Height | Height | 12 | in | FDA,5 | 11.9688 | 12.0312 | Box Inspection
5BU27500/Box_NY_0.024SBS_102_4.3906-4.3906-12 | 5BU27500 | Box_NY_0.024SBS_102_4.3906-4.3906-12 | NY-Analytical | 772 | Box Inspection | Length | Length | 4.4 | in | FDA,5 | 4.3594 | 4.4219 | Box Inspection
5BU27500/Box_NY_0.024SBS_102_4.3906-4.3906-12 | 5BU27500 | Box_NY_0.024SBS_102_4.3906-4.3906-12 | NY-Analytical | 772 | Box Inspection | Width | Width | 4.41 | in | FDA,5 | 4.3594 | 4.4219 | Box Inspection
5BU37230/Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | 5BU37230 | Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 125 |  |  |  |  | Box Inspection
5BU37230/Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | 5BU37230 | Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5BU37230/Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | 5BU37230 | Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.017 | in | FDA,3 | 0.017 | 0.019 | Box Inspection
5BU37230/Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | 5BU37230 | Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection
5BU37230/Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | 5BU37230 | Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom - French Style (Done) |  |  |  |  | Box Inspection
5BU37230/Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | 5BU37230 | Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | NY-Analytical | 772 | Box Inspection | Height | Height | 2.93 | in | FDA,5 | 2.9062 | 2.9688 | Box Inspection
5BU37230/Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | 5BU37230 | Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | NY-Analytical | 772 | Box Inspection | Length | Length | 3.84 | in | FDA,5 | 3.7812 | 3.8438 | Box Inspection
5BU37230/Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | 5BU37230 | Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.38 | in | FDA,5 | 3.3438 | 3.4062 | Box Inspection
5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | 5BU55510 | Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | 5BU55510 | Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0155 | in | FDA,3 | 0.015 | 0.017 | Box Inspection
5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | 5BU55510 | Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | Bright Metalized Poly Board (Done) |  |  |  |  | Box Inspection
5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | 5BU55510 | Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Straight Tuck End w/glued header (Done) |  |  |  |  | Box Inspection
5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | 5BU55510 | Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Height | Height | 3.8125 | in | FDA,5 | 3.7812 | 3.8438 | Box Inspection
5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | 5BU55510 | Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Length | Length | 2.9375 | in | FDA,5 | 2.9062 | 2.9688 | Box Inspection
5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | 5BU55510 | Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Width | Width | 1.875 | in | FDA,5 | 1.8438 | 1.9062 | Box Inspection
5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | 5BU71230 | Box_NY_0.018SBS_122_4.6875-3-3.8125 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | 5BU71230 | Box_NY_0.018SBS_122_4.6875-3-3.8125 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.018 | in | FDA,3 | 0.017 | 0.019 | Box Inspection
5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | 5BU71230 | Box_NY_0.018SBS_122_4.6875-3-3.8125 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection
5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | 5BU71230 | Box_NY_0.018SBS_122_4.6875-3-3.8125 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Straight Tuck End (Done) |  |  |  |  | Box Inspection
5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | 5BU71230 | Box_NY_0.018SBS_122_4.6875-3-3.8125 | NY-Analytical | 772 | Box Inspection | Height | Height | 3.82 | in | FDA,5 | 3.7812 | 3.8438 | Box Inspection
5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | 5BU71230 | Box_NY_0.018SBS_122_4.6875-3-3.8125 | NY-Analytical | 772 | Box Inspection | Length | Length | 4.7 | in | FDA,5 | 4.6562 | 4.7188 | Box Inspection
5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | 5BU71230 | Box_NY_0.018SBS_122_4.6875-3-3.8125 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.01 | in | FDA,5 | 2.9688 | 3.0312 | Box Inspection
5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | 5CA284TR | Box_NY_0.018SBS_111_6-5.3125-3.6875 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | 5CA284TR | Box_NY_0.018SBS_111_6-5.3125-3.6875 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0175 | in | FDA,3 | 0.017 | 0.019 | Box Inspection
5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | 5CA284TR | Box_NY_0.018SBS_111_6-5.3125-3.6875 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection
5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | 5CA284TR | Box_NY_0.018SBS_111_6-5.3125-3.6875 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom French Style w/Perfs (Done) |  |  |  |  | Box Inspection
5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | 5CA284TR | Box_NY_0.018SBS_111_6-5.3125-3.6875 | NY-Analytical | 772 | Box Inspection | Height | Height | 3.6875 | in | FDA,5 | 3.6562 | 3.7188 | Box Inspection
5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | 5CA284TR | Box_NY_0.018SBS_111_6-5.3125-3.6875 | NY-Analytical | 772 | Box Inspection | Length | Length | 6 | in | FDA,5 | 5.9688 | 6.0312 | Box Inspection
5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | 5CA284TR | Box_NY_0.018SBS_111_6-5.3125-3.6875 | NY-Analytical | 772 | Box Inspection | Width | Width | 5.3125 | in | FDA,5 | 5.2812 | 5.3438 | Box Inspection
5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 5CB19600 | Box_NY_0.018SBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | Box Inspection
5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 5CB19600 | Box_NY_0.018SBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 5CB19600 | Box_NY_0.018SBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.018 | in | FDA,3 | 0.017 | 0.019 | Box Inspection
5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 5CB19600 | Box_NY_0.018SBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection
5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 5CB19600 | Box_NY_0.018SBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection
5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 5CB19600 | Box_NY_0.018SBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Height | Height | 2.0938 | in | FDA,5 | 2.0625 | 2.125 | Box Inspection
5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 5CB19600 | Box_NY_0.018SBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Length | Length | 5.125 | in | FDA,5 | 5.0938 | 5.1562 | Box Inspection
5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 5CB19600 | Box_NY_0.018SBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.25 | in | FDA,5 | 2.2188 | 2.2812 | Box Inspection
5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | 5CB51200 | Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | UCSBS |  |  |  |  | Box Inspection
5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | 5CB51200 | Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | 5CB51200 | Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0155 | in | FDA,3 | 0.015 | 0.017 | Box Inspection
5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | 5CB51200 | Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection
5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | 5CB51200 | Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | NY-Analytical | 772 | Box Inspection | Height | Height | 2.375 | in | FDA,5 | 2.3438 | 2.4062 | Box Inspection
5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | 5CB51200 | Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | NY-Analytical | 772 | Box Inspection | Length | Length | 6.125 | in | FDA,5 | 6.0938 | 6.1562 | Box Inspection
5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | 5CB51200 | Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.75 | in | FDA,5 | 2.7188 | 2.7812 | Box Inspection
5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | 5CC80100 | Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | UCSBS |  |  |  |  | Box Inspection
5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | 5CC80100 | Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | Box Inspection
5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | 5CC80100 | Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | 5CC80100 | Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.015 | in | FDA,3 | 0.015 | 0.017 | Box Inspection
5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | 5CC80100 | Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection
5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | 5CC80100 | Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Height | Height | 3.07 | in | FDA,5 | 3.0312 | 3.0938 | Box Inspection
5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | 5CC80100 | Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Length | Length | 5.07 | in | FDA,5 | 5.0312 | 5.0938 | Box Inspection
5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | 5CC80100 | Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Width | Width | 1.73 | in | FDA,5 | 1.7188 | 1.7812 | Box Inspection
5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 5CD35185 | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | UCSBS |  |  |  |  | Box Inspection
5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 5CD35185 | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | Box Inspection
5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 5CD35185 | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 5CD35185 | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.016 | in | FDA,3 | 0.015 | 0.017 | Box Inspection
5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 5CD35185 | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection
5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 5CD35185 | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 772 | Box Inspection | Height | Height | 6.078 | in | FDA,5 | 6.0312 | 6.0938 | Box Inspection
5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 5CD35185 | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 772 | Box Inspection | Length | Length | 6.5 | in | FDA,5 | 6.4688 | 6.5312 | Box Inspection
5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 5CD35185 | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.312 | in | FDA,5 | 2.2812 | 2.3438 | Box Inspection
5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 5CS42050 | Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | UCSBS |  |  |  |  | Box Inspection
5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 5CS42050 | Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | Box Inspection
5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 5CS42050 | Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 5CS42050 | Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.016 | in | FDA,3 | 0.015 | 0.017 | Box Inspection
5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 5CS42050 | Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection
5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 5CS42050 | Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | NY-Analytical | 772 | Box Inspection | Height | Height | 5.81 | in | FDA,5 | 5.7812 | 5.8438 | Box Inspection
5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 5CS42050 | Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | NY-Analytical | 772 | Box Inspection | Length | Length | 6.51 | in | FDA,5 | 6.4688 | 6.5312 | Box Inspection
5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 5CS42050 | Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.32 | in | FDA,5 | 3.2812 | 3.3438 | Box Inspection
5DA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | 5DA941TY | Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5DA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | 5DA941TY | Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | NY-Analytical | 772 | Box Inspection | Back Height | Back Height | 8.625 | in | FDA,5 | 8.5938 | 8.6562 | Box Inspection
5DA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | 5DA941TY | Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.024 | in | FDA,3 | 0.023 | 0.025 | Box Inspection
5DA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | 5DA941TY | Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | WCK (Done) |  |  |  |  | Box Inspection
5DA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | 5DA941TY | Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom French Style w/Perfs (Done) |  |  |  |  | Box Inspection
5DA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | 5DA941TY | Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | NY-Analytical | 772 | Box Inspection | Front Height | Front Height | 3.5 | in | FDA,5 | 3.4688 | 3.5312 | Box Inspection
5DA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | 5DA941TY | Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | NY-Analytical | 772 | Box Inspection | Length | Length | 5.24 | in | FDA,5 | 5.2188 | 5.2812 | Box Inspection
5DA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | 5DA941TY | Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | NY-Analytical | 772 | Box Inspection | Width | Width | 6.26 | in | FDA,5 | 6.2188 | 6.2812 | Box Inspection
5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | 5DH12824 | Box_NY_0.018SBS_100_4-2.125-6.125 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 30 |  | FDA,0 |  |  | Box Inspection
5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | 5DH12824 | Box_NY_0.018SBS_100_4-2.125-6.125 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | 5DH12824 | Box_NY_0.018SBS_100_4-2.125-6.125 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.017 | in | FDA,3 | 0.017 | 0.019 | Box Inspection
5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | 5DH12824 | Box_NY_0.018SBS_100_4-2.125-6.125 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection
5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | 5DH12824 | Box_NY_0.018SBS_100_4-2.125-6.125 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection
5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | 5DH12824 | Box_NY_0.018SBS_100_4-2.125-6.125 | NY-Analytical | 772 | Box Inspection | Height | Height | 6.125 | in | FDA,5 | 6.0938 | 6.1562 | Box Inspection
5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | 5DH12824 | Box_NY_0.018SBS_100_4-2.125-6.125 | NY-Analytical | 772 | Box Inspection | Length | Length | 4 | in | FDA,5 | 3.9688 | 4.0312 | Box Inspection
5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | 5DH12824 | Box_NY_0.018SBS_100_4-2.125-6.125 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.125 | in | FDA,5 | 2.0938 | 2.1562 | Box Inspection
5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | 5EH21130 | Box_NY_0.016SBS_100_4.75-2.375-3.25 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | Box Inspection
5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | 5EH21130 | Box_NY_0.016SBS_100_4.75-2.375-3.25 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | 5EH21130 | Box_NY_0.016SBS_100_4.75-2.375-3.25 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.015 | in | FDA,3 | 0.015 | 0.017 | Box Inspection
5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | 5EH21130 | Box_NY_0.016SBS_100_4.75-2.375-3.25 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection
5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | 5EH21130 | Box_NY_0.016SBS_100_4.75-2.375-3.25 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection
5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | 5EH21130 | Box_NY_0.016SBS_100_4.75-2.375-3.25 | NY-Analytical | 772 | Box Inspection | Height | Height | 3.26 | in | FDA,5 | 3.2188 | 3.2812 | Box Inspection
5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | 5EH21130 | Box_NY_0.016SBS_100_4.75-2.375-3.25 | NY-Analytical | 772 | Box Inspection | Length | Length | 4.75 | in | FDA,5 | 4.7188 | 4.7812 | Box Inspection
5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | 5EH21130 | Box_NY_0.016SBS_100_4.75-2.375-3.25 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.36 | in | FDA,5 | 2.3438 | 2.4062 | Certificate of Analysis
5FD41900/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 5FD41900 | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Box Inspection
5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | 5FH07050 | Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | 5FH07050 | Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.02 | in | FDA,3 | 0.019 | 0.021 | Box Inspection
5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | 5FH07050 | Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection
5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | 5FH07050 | Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Straight Tuck End w/Glued Internal Dividers&15/16" Dummy Platform Bottom (Done) |  |  |  |  | Box Inspection
5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | 5FH07050 | Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | NY-Analytical | 772 | Box Inspection | Height | Height | 6 | in | FDA,5 | 5.9844 | 6.0469 | Box Inspection
5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | 5FH07050 | Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | NY-Analytical | 772 | Box Inspection | Length | Length | 6.14 | in | FDA,5 | 6.0938 | 6.1562 | Box Inspection
5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | 5FH07050 | Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.44 | in | FDA,5 | 2.4062 | 2.4688 | Box Inspection
5GB10300/Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | 5GB10300 | Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | UCSBS |  |  |  |  | Box Inspection
5GB10300/Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | 5GB10300 | Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5GB10300/Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | 5GB10300 | Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.015 | in | FDA,3 | 0.015 | 0.017 | Box Inspection
5GB10300/Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | 5GB10300 | Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection
5GB10300/Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | 5GB10300 | Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | NY-Analytical | 772 | Box Inspection | Height | Height | 2.375 | in | FDA,5 | 2.3438 | 2.4062 | Box Inspection
5GB10300/Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | 5GB10300 | Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | NY-Analytical | 772 | Box Inspection | Length | Length | 6.125 | in | FDA,5 | 6.0938 | 6.1562 | Box Inspection
5GB10300/Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | 5GB10300 | Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.75 | in | FDA,5 | 2.7188 | 2.7812 | Box Inspection
5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 5HB59800 | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | Box Inspection
5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 5HB59800 | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 5HB59800 | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | Back Flip | Back Flip | 5.158 | in | FDA,5 | 4.6875 | 5.3125 | Box Inspection
5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 5HB59800 | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.001 | in | FDA,4 | 0.0008 | 0.0028 | Box Inspection
5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 5HB59800 | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | LDPE (Done) |  |  |  |  | Box Inspection
5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 5HB59800 | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Poly Bag w/fold over top on 2” wicketts (Done) |  |  |  |  | Box Inspection
5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 5HB59800 | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | Gussett | Gussett | 3.8 | in | FDA,5 | 3.6875 | 4.3125 | Box Inspection
5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 5HB59800 | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | Height | Height | 5.76 | in | FDA,5 | 5.1875 | 5.8125 | Box Inspection
5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 5HB59800 | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | Width | Width | 11.812 | in | FDA,5 | 11.6875 | 12.3125 | Box Inspection
5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | 5HC12400 | Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | UCSBS |  |  |  |  | Box Inspection
5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | 5HC12400 | Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 20 |  | FDA,0 |  |  | Box Inspection
5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | 5HC12400 | Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | 5HC12400 | Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0169 | in | FDA,3 | 0.015 | 0.017 | Box Inspection
5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | 5HC12400 | Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection
5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | 5HC12400 | Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | NY-Analytical | 772 | Box Inspection | Height | Height | 3.0625 | in | FDA,5 | 3.0312 | 3.0938 | Box Inspection
5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | 5HC12400 | Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | NY-Analytical | 772 | Box Inspection | Length | Length | 5.0625 | in | FDA,5 | 5.0312 | 5.0938 | Box Inspection
5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | 5HC12400 | Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | NY-Analytical | 772 | Box Inspection | Width | Width | 4.25 | in | FDA,5 | 4.2188 | 4.2812 | Box Inspection
5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | 5HC22324 | Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | 5HC22324 | Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0198 | in | FDA,3 | 0.019 | 0.021 | Box Inspection
5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | 5HC22324 | Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | WCCN (Done) |  |  |  |  | Box Inspection
5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | 5HC22324 | Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection
5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | 5HC22324 | Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Height | Height | 3.0625 | in | FDA,5 | 3.0312 | 3.0938 | Box Inspection
5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | 5HC22324 | Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Length | Length | 5.0625 | in | FDA,5 | 5.0312 | 5.0938 | Box Inspection
5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | 5HC22324 | Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Width | Width | 1.75 | in | FDA,5 | 1.7188 | 1.7812 | Box Inspection
5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | 5JB16400 | Box_NY_0.016UCSBS_100_4-3.125-2.375 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | UCSBS |  |  |  |  | Box Inspection
5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | 5JB16400 | Box_NY_0.016UCSBS_100_4-3.125-2.375 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | 5JB16400 | Box_NY_0.016UCSBS_100_4-3.125-2.375 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.015 | in | FDA,3 | 0.015 | 0.017 | Box Inspection
5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | 5JB16400 | Box_NY_0.016UCSBS_100_4-3.125-2.375 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection
5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | 5JB16400 | Box_NY_0.016UCSBS_100_4-3.125-2.375 | NY-Analytical | 772 | Box Inspection | Height | Height | 2.375 | in | FDA,5 | 2.3438 | 2.4062 | Box Inspection
5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | 5JB16400 | Box_NY_0.016UCSBS_100_4-3.125-2.375 | NY-Analytical | 772 | Box Inspection | Length | Length | 4 | in | FDA,5 | 3.9688 | 4.0312 | Box Inspection
5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | 5JB16400 | Box_NY_0.016UCSBS_100_4-3.125-2.375 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.125 | in | FDA,5 | 3.0938 | 3.1562 | Box Inspection
5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | 5LB60302 | Box_NY_0.016UCSBS_100_6.125-4-2.375 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | UCSBS |  |  |  |  | Box Inspection
5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | 5LB60302 | Box_NY_0.016UCSBS_100_6.125-4-2.375 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | Box Inspection
5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | 5LB60302 | Box_NY_0.016UCSBS_100_6.125-4-2.375 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | 5LB60302 | Box_NY_0.016UCSBS_100_6.125-4-2.375 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.016 | in | FDA,3 | 0.015 | 0.017 | Box Inspection
5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | 5LB60302 | Box_NY_0.016UCSBS_100_6.125-4-2.375 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection
5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | 5LB60302 | Box_NY_0.016UCSBS_100_6.125-4-2.375 | NY-Analytical | 772 | Box Inspection | Height | Height | 2.375 | in | FDA,5 | 2.3438 | 2.4062 | Box Inspection
5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | 5LB60302 | Box_NY_0.016UCSBS_100_6.125-4-2.375 | NY-Analytical | 772 | Box Inspection | Length | Length | 6.0938 | in | FDA,5 | 6.0938 | 6.1562 | Box Inspection
5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | 5LB60302 | Box_NY_0.016UCSBS_100_6.125-4-2.375 | NY-Analytical | 772 | Box Inspection | Width | Width | 4 | in | FDA,5 | 3.9688 | 4.0312 | Box Inspection
5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | 5MS48050 | Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | UCSBS |  |  |  |  | Box Inspection
5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | 5MS48050 | Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | Box Inspection
5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | 5MS48050 | Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | 5MS48050 | Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0155 | in | FDA,3 | 0.015 | 0.017 | Box Inspection
5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | 5MS48050 | Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection
5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | 5MS48050 | Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | NY-Analytical | 772 | Box Inspection | Height | Height | 5.8125 | in | FDA,5 | 5.7812 | 5.8438 | Box Inspection
5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | 5MS48050 | Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | NY-Analytical | 772 | Box Inspection | Length | Length | 6.5 | in | FDA,5 | 6.4688 | 6.5312 | Box Inspection
5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | 5MS48050 | Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.3125 | in | FDA,5 | 3.2812 | 3.3438 | Box Inspection
5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | 5NB40600 | Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | UCSBS |  |  |  |  | Box Inspection
5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | 5NB40600 | Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | Box Inspection
5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | 5NB40600 | Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | 5NB40600 | Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0155 | in | FDA,3 | 0.015 | 0.017 | Box Inspection
5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | 5NB40600 | Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection
5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | 5NB40600 | Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Height | Height | 2.07 | in | FDA,5 | 2.0625 | 2.125 | Box Inspection
5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | 5NB40600 | Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Length | Length | 5.13 | in | FDA,5 | 5.0938 | 5.1562 | Box Inspection
5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | 5NB40600 | Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.24 | in | FDA,5 | 2.2188 | 2.2812 | Box Inspection
5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | 5NU87250 | Box_NY_0.024SBS_127_4.375-4.375-12 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom Dispenser Box |  |  |  |  | Box Inspection
5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | 5NU87250 | Box_NY_0.024SBS_127_4.375-4.375-12 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | 5NU87250 | Box_NY_0.024SBS_127_4.375-4.375-12 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.024 | in | FDA,3 | 0.023 | 0.025 | Box Inspection
5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | 5NU87250 | Box_NY_0.024SBS_127_4.375-4.375-12 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection
5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | 5NU87250 | Box_NY_0.024SBS_127_4.375-4.375-12 | NY-Analytical | 772 | Box Inspection | Height | Height | 12 | in | FDA,5 | 11.9688 | 12.0312 | Box Inspection
5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | 5NU87250 | Box_NY_0.024SBS_127_4.375-4.375-12 | NY-Analytical | 772 | Box Inspection | Length | Length | 4.39 | in | FDA,5 | 4.3438 | 4.4062 | Box Inspection
5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | 5NU87250 | Box_NY_0.024SBS_127_4.375-4.375-12 | NY-Analytical | 772 | Box Inspection | Width | Width | 4.39 | in | FDA,5 | 4.3438 | 4.4062 | Box Inspection
5OD125PB/Box_80A | 5OD125PB | Box_80A | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5OD125PB/Box_80A | 5OD125PB | Box_80A | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0005 | in | FDA,4 | 0.0005 | 0.0025 | Box Inspection
5OD125PB/Box_80A | 5OD125PB | Box_80A | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | LDPE (Done) |  |  |  |  | Box Inspection
5OD125PB/Box_80A | 5OD125PB | Box_80A | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Poly Bag w/fold over top on 4” wicketts (Done) |  |  |  |  | Box Inspection
5OD125PB/Box_80A | 5OD125PB | Box_80A | NY-Analytical | 772 | Box Inspection | Length | Length | 6.5 | in | FDA,5 | 6.0625 | 6.5 | Box Inspection
5OD125PB/Box_80A | 5OD125PB | Box_80A | NY-Analytical | 772 | Box Inspection | Outside Flap | Outside Flap | 2.25 | in | FDA,5 | 2.2188 | 2.2812 | Box Inspection
5OD125PB/Box_80A | 5OD125PB | Box_80A | NY-Analytical | 772 | Box Inspection | Width | Width | 8.25 | in | FDA,5 | 8.25 | 8.5 | Box Inspection
5OD239PD/Box_51A | 5OD239PD | Box_51A | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5OD239PD/Box_51A | 5OD239PD | Box_51A | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0175 | in | FDA,3 |  |  | Box Inspection
5OD239PD/Box_51A | 5OD239PD | Box_51A | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | Chip Board (Done) |  |  |  |  | Box Inspection
5OD239PD/Box_51A | 5OD239PD | Box_51A | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Layer Pad (Done) |  |  |  |  | Box Inspection
5OD239PD/Box_51A | 5OD239PD | Box_51A | NY-Analytical | 772 | Box Inspection | Length | Length | 15.375 | in | FDA,5 | 15.3438 | 15.4062 | Box Inspection
5OD239PD/Box_51A | 5OD239PD | Box_51A | NY-Analytical | 772 | Box Inspection | Width | Width | 5.5 | in | FDA,5 | 5.4688 | 5.5312 | Box Inspection
5OH382PD/Box_50D | 5OH382PD | Box_50D | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5OH382PD/Box_50D | 5OH382PD | Box_50D | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0018 | in | FDA,3 |  |  | Box Inspection
5OH382PD/Box_50D | 5OH382PD | Box_50D | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | Any Board (Done) |  |  |  |  | Box Inspection
5OH382PD/Box_50D | 5OH382PD | Box_50D | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Cut Pad (Done) |  |  |  |  | Box Inspection
5OH382PD/Box_50D | 5OH382PD | Box_50D | NY-Analytical | 772 | Box Inspection | Length | Length | 11.5938 | in | FDA,5 | 11.5938 | 11.6562 | Box Inspection
5OH382PD/Box_50D | 5OH382PD | Box_50D | NY-Analytical | 772 | Box Inspection | Width | Width | 9.4688 | in | FDA,5 | 9.4688 | 9.5312 | Box Inspection
5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | 5OP067BT | Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 50 |  | FDA,0 |  |  | Box Inspection
5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | 5OP067BT | Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | 5OP067BT | Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.016 | in | FDA,3 | 0.015 | 0.017 | Box Inspection
5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | 5OP067BT | Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS C1S (Done) |  |  |  |  | Box Inspection
5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | 5OP067BT | Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Open Top/Tuck Bottom (Done) |  |  |  |  | Box Inspection
5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | 5OP067BT | Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Height | Height | 2.44 | in | FDA,5 | 2.4062 | 2.4688 | Box Inspection
5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | 5OP067BT | Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Length | Length | 3.2 | in | FDA,5 | 3.1488 | 3.2113 | Box Inspection
5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | 5OP067BT | Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.19 | in | FDA,5 | 2.1562 | 2.2188 | Box Inspection
5OP073BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | 5OP073BT | Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | Box Inspection
5OP073BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | 5OP073BT | Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5OP073BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | 5OP073BT | Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.016 | in | FDA,3 | 0.015 | 0.017 | Box Inspection
5OP073BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | 5OP073BT | Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS C1S (Done) |  |  |  |  | Box Inspection
5OP073BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | 5OP073BT | Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Open Top/Tuck Bottom (Done) |  |  |  |  | Box Inspection
5OP073BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | 5OP073BT | Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Height | Height | 2.42 | in | FDA,5 | 2.4062 | 2.4688 | Box Inspection
5OP073BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | 5OP073BT | Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Length | Length | 3.2 | in | FDA,5 | 3.1488 | 3.2113 | Box Inspection
5OP073BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | 5OP073BT | Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.17 | in | FDA,5 | 2.1562 | 2.2188 | Box Inspection
5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | 5OP092BT | Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 315 |  | FDA,0 |  |  | Box Inspection
5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | 5OP092BT | Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | 5OP092BT | Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.016 | in | FDA,3 | 0.015 | 0.017 | Box Inspection
5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | 5OP092BT | Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS C1S (Done) |  |  |  |  | Box Inspection
5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | 5OP092BT | Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Open Top/Tuck Bottom (Done) |  |  |  |  | Box Inspection
5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | 5OP092BT | Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | NY-Analytical | 772 | Box Inspection | Height | Height | 2.6875 | in | FDA,5 | 2.6562 | 2.7188 | Box Inspection
5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | 5OP092BT | Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | NY-Analytical | 772 | Box Inspection | Length | Length | 3.1875 | in | FDA,5 | 3.1488 | 3.2113 | Box Inspection
5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | 5OP092BT | Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | NY-Analytical | 772 | Box Inspection | Width | Width | 1.4375 | in | FDA,5 | 1.4062 | 1.4688 | Box Inspection
5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | 5OP092PB | Box_NY_LDPE_407_13.5-6.5-6-4 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | Box Inspection
5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | 5OP092PB | Box_NY_LDPE_407_13.5-6.5-6-4 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | 5OP092PB | Box_NY_LDPE_407_13.5-6.5-6-4 | NY-Analytical | 772 | Box Inspection | Back Flip | Back Flip | 5.75 | in | FDA,5 | 5.6875 | 6.3125 | Box Inspection
5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | 5OP092PB | Box_NY_LDPE_407_13.5-6.5-6-4 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | LDPE (Done) |  |  |  |  | Box Inspection
5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | 5OP092PB | Box_NY_LDPE_407_13.5-6.5-6-4 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Poly Bag w/fold over top on 2” wicketts (Done) |  |  |  |  | Box Inspection
5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | 5OP092PB | Box_NY_LDPE_407_13.5-6.5-6-4 | NY-Analytical | 772 | Box Inspection | Gussett | Gussett | 4 | in | FDA,5 | 3.6875 | 4.3125 | Box Inspection
5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | 5OP092PB | Box_NY_LDPE_407_13.5-6.5-6-4 | NY-Analytical | 772 | Box Inspection | Height | Height | 6.5 | in | FDA,5 | 6.1875 | 6.8125 | Box Inspection
5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | 5OP092PB | Box_NY_LDPE_407_13.5-6.5-6-4 | NY-Analytical | 772 | Box Inspection | Width | Width | 13.1875 | in | FDA,5 | 13.1875 | 13.8125 | Box Inspection
5OP308PD/Box_50C | 5OP308PD | Box_50C | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5OP308PD/Box_50C | 5OP308PD | Box_50C | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.02 | in | FDA,3 | 0.018 | 0.02 | Box Inspection
5OP308PD/Box_50C | 5OP308PD | Box_50C | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | Chip Board (Done) |  |  |  |  | Box Inspection
5OP308PD/Box_50C | 5OP308PD | Box_50C | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Cut Pad (Done) |  |  |  |  | Box Inspection
5OP308PD/Box_50C | 5OP308PD | Box_50C | NY-Analytical | 772 | Box Inspection | Length | Length | 7 | in | FDA,5 | 6.9688 | 7.0312 | Box Inspection
5OP308PD/Box_50C | 5OP308PD | Box_50C | NY-Analytical | 772 | Box Inspection | Width | Width | 12.28 | in | FDA,5 | 12.2188 | 12.2812 | Box Inspection
5PB114IS/Box_NY_NA50#_402_2.5-3.25 | 5PB114IS | Box_NY_NA50#_402_2.5-3.25 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 500 |  | FDA,0 |  |  | Box Inspection
5PB114IS/Box_NY_NA50#_402_2.5-3.25 | 5PB114IS | Box_NY_NA50#_402_2.5-3.25 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5PB114IS/Box_NY_NA50#_402_2.5-3.25 | 5PB114IS | Box_NY_NA50#_402_2.5-3.25 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | 50# White Paper Stock (Done) |  |  |  |  | Box Inspection
5PB114IS/Box_NY_NA50#_402_2.5-3.25 | 5PB114IS | Box_NY_NA50#_402_2.5-3.25 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Insert (Done) |  |  |  |  | Box Inspection
5PB114IS/Box_NY_NA50#_402_2.5-3.25 | 5PB114IS | Box_NY_NA50#_402_2.5-3.25 | NY-Analytical | 772 | Box Inspection | Length | Length | 2.5 | in | FDA,5 | 2.4688 | 2.5312 | Box Inspection
5PB114IS/Box_NY_NA50#_402_2.5-3.25 | 5PB114IS | Box_NY_NA50#_402_2.5-3.25 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.25 | in | FDA,5 | 3.2188 | 3.2812 | Box Inspection
5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | 5PS40750 | Box_NY_0.018SBS_100_7.125-3-6.625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | 5PS40750 | Box_NY_0.018SBS_100_7.125-3-6.625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0175 | in | FDA,3 | 0.017 | 0.019 | Box Inspection
5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | 5PS40750 | Box_NY_0.018SBS_100_7.125-3-6.625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection
5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | 5PS40750 | Box_NY_0.018SBS_100_7.125-3-6.625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection
5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | 5PS40750 | Box_NY_0.018SBS_100_7.125-3-6.625 | NY-Analytical | 772 | Box Inspection | Height | Height | 6.625 | in | FDA,5 | 6.5938 | 6.6562 | Box Inspection
5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | 5PS40750 | Box_NY_0.018SBS_100_7.125-3-6.625 | NY-Analytical | 772 | Box Inspection | Length | Length | 7.125 | in | FDA,5 | 7.0938 | 7.1562 | Box Inspection
5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | 5PS40750 | Box_NY_0.018SBS_100_7.125-3-6.625 | NY-Analytical | 772 | Box Inspection | Width | Width | 3 | in | FDA,5 | 2.9688 | 3.0312 | Box Inspection
5QS273IS/Box_NY_NA50#_402_5-3.25 | 5QS273IS | Box_NY_NA50#_402_5-3.25 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 500 |  | FDA,0 |  |  | Box Inspection
5QS273IS/Box_NY_NA50#_402_5-3.25 | 5QS273IS | Box_NY_NA50#_402_5-3.25 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5QS273IS/Box_NY_NA50#_402_5-3.25 | 5QS273IS | Box_NY_NA50#_402_5-3.25 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | 50# White Paper Stock (Done) |  |  |  |  | Box Inspection
5QS273IS/Box_NY_NA50#_402_5-3.25 | 5QS273IS | Box_NY_NA50#_402_5-3.25 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Insert (Done) |  |  |  |  | Box Inspection
5QS273IS/Box_NY_NA50#_402_5-3.25 | 5QS273IS | Box_NY_NA50#_402_5-3.25 | NY-Analytical | 772 | Box Inspection | Length | Length | 5 | in | FDA,5 | 4.9688 | 5.0312 | Box Inspection
5QS273IS/Box_NY_NA50#_402_5-3.25 | 5QS273IS | Box_NY_NA50#_402_5-3.25 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.2188 | in | FDA,5 | 3.2188 | 3.2812 | Box Inspection
5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | 5QS41950 | Box_NY_0.018SBS_100_6.625-4.125-8.125 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection
5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | 5QS41950 | Box_NY_0.018SBS_100_6.625-4.125-8.125 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.017 | in | FDA,3 | 0.017 | 0.019 | Box Inspection
5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | 5QS41950 | Box_NY_0.018SBS_100_6.625-4.125-8.125 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) |  |  |  |  | Box Inspection
5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | 5QS41950 | Box_NY_0.018SBS_100_6.625-4.125-8.125 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection
5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | 5QS41950 | Box_NY_0.018SBS_100_6.625-4.125-8.125 | NY-Analytical | 772 | Box Inspection | Height | Height | 8.125 | in | FDA,5 | 8.0938 | 8.1562 | Box Inspection
5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | 5QS41950 | Box_NY_0.018SBS_100_6.625-4.125-8.125 | NY-Analytical | 772 | Box Inspection | Length | Length | 6.625 | in | FDA,5 | 6.5938 | 6.6562 | Box Inspection
5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | 5QS41950 | Box_NY_0.018SBS_100_6.625-4.125-8.125 | NY-Analytical | 772 | Box Inspection | Width | Width | 4.125 | in | FDA,5 | 4.0938 | 4.1562 | Box Inspection
6AA34724/Shipper_NY_Skip lot testing | 6AA34724 | Shipper_NY_Skip lot testing | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Shipper Inspection
6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | 6AA34724 | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 20 |  | FDA,0 |  |  | TM 141 / Visual
6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | 6AA34724 | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | 6AA34724 | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | 6AA34724 | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | 6AA34724 | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6.1875 | in | FDA,4 | 6.0625 | 6.4375 | TM 141 / Visual
6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | 6AA34724 | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 8.4375 | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual
6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | 6AA34724 | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | 6AA34724 | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | 6AA34724 | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.1875 | in | FDA,4 | 6.8125 | 7.1875 | TM 141 / Visual
6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | 6AA617S96 | SHIP_NY_44K_100_12.75-11.625-7.875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | 6AA617S96 | SHIP_NY_44K_100_12.75-11.625-7.875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual
6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | 6AA617S96 | SHIP_NY_44K_100_12.75-11.625-7.875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | 6AA617S96 | SHIP_NY_44K_100_12.75-11.625-7.875 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7.75 | in | FDA,4 | 7.6875 | 8.0625 | TM 141 / Visual
6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | 6AA617S96 | SHIP_NY_44K_100_12.75-11.625-7.875 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 12.75 | in | FDA,4 | 12.5625 | 12.9375 | TM 141 / Visual
6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | 6AA617S96 | SHIP_NY_44K_100_12.75-11.625-7.875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | 6AA617S96 | SHIP_NY_44K_100_12.75-11.625-7.875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | 6AA617S96 | SHIP_NY_44K_100_12.75-11.625-7.875 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 11.5 | in | FDA,4 | 11.4375 | 11.8125 | TM 141 / Visual
6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | 6AA632TY | SHIP_NY_44K_231_18.5-7.375-14 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | 6AA632TY | SHIP_NY_44K_231_18.5-7.375-14 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual
6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | 6AA632TY | SHIP_NY_44K_231_18.5-7.375-14 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | 6AA632TY | SHIP_NY_44K_231_18.5-7.375-14 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 14 | in | FDA,4 | 13.875 | 14.125 | TM 141 / Visual
6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | 6AA632TY | SHIP_NY_44K_231_18.5-7.375-14 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 18.5 | in | FDA,4 | 18.375 | 18.625 | TM 141 / Visual
6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | 6AA632TY | SHIP_NY_44K_231_18.5-7.375-14 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | 6AA632TY | SHIP_NY_44K_231_18.5-7.375-14 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | D/C (Done) |  |  |  |  | TM 141 / Visual
6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | 6AA632TY | SHIP_NY_44K_231_18.5-7.375-14 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.5 | in | FDA,4 | 7.25 | 7.5 | TM 141 / Visual
6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 6AA709S48 | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 200 |  | FDA,0 |  |  | TM 141 / Visual
6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 6AA709S48 | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 6AA709S48 | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 40 | Lbs./In. | FDA,0 | 40 |  | TM 141 / Visual
6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 6AA709S48 | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 6AA709S48 | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6.125 | in | FDA,4 | 5.9375 | 6.3125 | TM 141 / Visual
6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 6AA709S48 | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15 | in | FDA,4 | 14.8125 | 15.1875 | TM 141 / Visual
6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 6AA709S48 | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 6AA709S48 | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 6AA709S48 | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 11 | in | FDA,4 | 10.8125 | 11.1875 | TM 141 / Visual
6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 6AA73530 | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual
6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 6AA73530 | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 6AA73530 | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 6AA73530 | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 6AA73530 | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 8.25 | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual
6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 6AA73530 | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 10 | in | FDA,4 | 9.8125 | 10.1875 | TM 141 / Visual
6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 6AA73530 | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 6AA73530 | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 6AA73530 | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 8.375 | in | FDA,4 | 8.1875 | 8.5625 | TM 141 / Visual
6AA76930/SHIP_NY_44K_100_9-9-4.75 | 6AA76930 | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual
6AA76930/SHIP_NY_44K_100_9-9-4.75 | 6AA76930 | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AA76930/SHIP_NY_44K_100_9-9-4.75 | 6AA76930 | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual
6AA76930/SHIP_NY_44K_100_9-9-4.75 | 6AA76930 | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AA76930/SHIP_NY_44K_100_9-9-4.75 | 6AA76930 | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 4.75 | in | FDA,4 | 4.625 | 4.875 | TM 141 / Visual
6AA76930/SHIP_NY_44K_100_9-9-4.75 | 6AA76930 | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 9 | in | FDA,4 | 8.875 | 9.125 | TM 141 / Visual
6AA76930/SHIP_NY_44K_100_9-9-4.75 | 6AA76930 | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AA76930/SHIP_NY_44K_100_9-9-4.75 | 6AA76930 | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AA76930/SHIP_NY_44K_100_9-9-4.75 | 6AA76930 | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 9 | in | FDA,4 | 8.875 | 9.125 | TM 141 / Visual
6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 6AB16400 | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual
6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 6AB16400 | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 6AB16400 | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 6AB16400 | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 6AB16400 | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 5.25 | in | FDA,4 | 4.9375 | 5.3125 | TM 141 / Visual
6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 6AB16400 | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 16.125 | in | FDA,4 | 15.8125 | 16.1875 | TM 141 / Visual
6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 6AB16400 | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 6AB16400 | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 6AB16400 | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 4.25 | in | FDA,4 | 4.0625 | 4.4375 | TM 141 / Visual
6AD11055/SHIP_NY_32K_100_10.625-6.375-11.25 | 6AD11055 | SHIP_NY_32K_100_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AD11055/SHIP_NY_32K_100_10.625-6.375-11.25 | 6AD11055 | SHIP_NY_32K_100_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6AD11055/SHIP_NY_32K_100_10.625-6.375-11.25 | 6AD11055 | SHIP_NY_32K_100_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AD11055/SHIP_NY_32K_100_10.625-6.375-11.25 | 6AD11055 | SHIP_NY_32K_100_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 11.375 | in | FDA,4 | 11.0625 | 11.4375 | TM 141 / Visual
6AD11055/SHIP_NY_32K_100_10.625-6.375-11.25 | 6AD11055 | SHIP_NY_32K_100_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 10.625 | in | FDA,4 | 10.4375 | 10.8125 | TM 141 / Visual
6AD11055/SHIP_NY_32K_100_10.625-6.375-11.25 | 6AD11055 | SHIP_NY_32K_100_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AD11055/SHIP_NY_32K_100_10.625-6.375-11.25 | 6AD11055 | SHIP_NY_32K_100_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AD11055/SHIP_NY_32K_100_10.625-6.375-11.25 | 6AD11055 | SHIP_NY_32K_100_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.375 | in | FDA,4 | 6.1875 | 6.5625 | TM 141 / Visual
6AD41900/SHIP_NY_32K_100_13.25-12-6.125 | 6AD41900 | SHIP_NY_32K_100_13.25-12-6.125 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AD41900/SHIP_NY_32K_100_13.25-12-6.125 | 6AD41900 | SHIP_NY_32K_100_13.25-12-6.125 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6AD41900/SHIP_NY_32K_100_13.25-12-6.125 | 6AD41900 | SHIP_NY_32K_100_13.25-12-6.125 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AD41900/SHIP_NY_32K_100_13.25-12-6.125 | 6AD41900 | SHIP_NY_32K_100_13.25-12-6.125 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6.25 | in | FDA,4 | 5.9375 | 6.3125 | TM 141 / Visual
6AD41900/SHIP_NY_32K_100_13.25-12-6.125 | 6AD41900 | SHIP_NY_32K_100_13.25-12-6.125 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 13.25 | in | FDA,4 | 13.0625 | 13.4375 | TM 141 / Visual
6AD41900/SHIP_NY_32K_100_13.25-12-6.125 | 6AD41900 | SHIP_NY_32K_100_13.25-12-6.125 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AD41900/SHIP_NY_32K_100_13.25-12-6.125 | 6AD41900 | SHIP_NY_32K_100_13.25-12-6.125 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AD41900/SHIP_NY_32K_100_13.25-12-6.125 | 6AD41900 | SHIP_NY_32K_100_13.25-12-6.125 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 11.875 | in | FDA,4 | 11.8125 | 12.1875 | TM 141 / Visual
6AH58195/SHIP_NY_32K_100_13.4375-6.25-9.6875 | 6AH58195 | SHIP_NY_32K_100_13.4375-6.25-9.6875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AH58195/SHIP_NY_32K_100_13.4375-6.25-9.6875 | 6AH58195 | SHIP_NY_32K_100_13.4375-6.25-9.6875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6AH58195/SHIP_NY_32K_100_13.4375-6.25-9.6875 | 6AH58195 | SHIP_NY_32K_100_13.4375-6.25-9.6875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AH58195/SHIP_NY_32K_100_13.4375-6.25-9.6875 | 6AH58195 | SHIP_NY_32K_100_13.4375-6.25-9.6875 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 9.75 | in | FDA,4 | 9.5 | 9.875 | TM 141 / Visual
6AH58195/SHIP_NY_32K_100_13.4375-6.25-9.6875 | 6AH58195 | SHIP_NY_32K_100_13.4375-6.25-9.6875 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 13.4375 | in | FDA,4 | 13.25 | 13.625 | TM 141 / Visual
6AH58195/SHIP_NY_32K_100_13.4375-6.25-9.6875 | 6AH58195 | SHIP_NY_32K_100_13.4375-6.25-9.6875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AH58195/SHIP_NY_32K_100_13.4375-6.25-9.6875 | 6AH58195 | SHIP_NY_32K_100_13.4375-6.25-9.6875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AH58195/SHIP_NY_32K_100_13.4375-6.25-9.6875 | 6AH58195 | SHIP_NY_32K_100_13.4375-6.25-9.6875 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.25 | in | FDA,4 | 6.0625 | 6.4375 | TM 141 / Visual
6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 6AJ19025 | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual
6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 6AJ19025 | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 6AJ19025 | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 6AJ19025 | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 6AJ19025 | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6.75 | in | FDA,4 | 6.5625 | 6.9375 | TM 141 / Visual
6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 6AJ19025 | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 12 | in | FDA,4 | 11.8125 | 12.1875 | TM 141 / Visual
6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 6AJ19025 | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 6AJ19025 | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 6AJ19025 | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.25 | in | FDA,4 | 7.0625 | 7.4375 | TM 141 / Visual
6AM30472/SHIP_NY_44K_100_15.5-9-12 | 6AM30472 | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual
6AM30472/SHIP_NY_44K_100_15.5-9-12 | 6AM30472 | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AM30472/SHIP_NY_44K_100_15.5-9-12 | 6AM30472 | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual
6AM30472/SHIP_NY_44K_100_15.5-9-12 | 6AM30472 | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AM30472/SHIP_NY_44K_100_15.5-9-12 | 6AM30472 | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 12.125 | in | FDA,4 | 11.8125 | 12.1875 | TM 141 / Visual
6AM30472/SHIP_NY_44K_100_15.5-9-12 | 6AM30472 | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.375 | in | FDA,4 | 15.3125 | 15.6875 | TM 141 / Visual
6AM30472/SHIP_NY_44K_100_15.5-9-12 | 6AM30472 | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AM30472/SHIP_NY_44K_100_15.5-9-12 | 6AM30472 | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AM30472/SHIP_NY_44K_100_15.5-9-12 | 6AM30472 | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 8.875 | in | FDA,4 | 8.8125 | 9.1875 | TM 141 / Visual
6AP03707/SHIP_NY_32K_100_13-7.3125-4 | 6AP03707 | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual
6AP03707/SHIP_NY_32K_100_13-7.3125-4 | 6AP03707 | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AP03707/SHIP_NY_32K_100_13-7.3125-4 | 6AP03707 | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6AP03707/SHIP_NY_32K_100_13-7.3125-4 | 6AP03707 | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AP03707/SHIP_NY_32K_100_13-7.3125-4 | 6AP03707 | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 4.1875 | in | FDA,4 | 3.8125 | 4.1875 | TM 141 / Visual
6AP03707/SHIP_NY_32K_100_13-7.3125-4 | 6AP03707 | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 13.062 | in | FDA,4 | 12.8125 | 13.1875 | TM 141 / Visual
6AP03707/SHIP_NY_32K_100_13-7.3125-4 | 6AP03707 | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AP03707/SHIP_NY_32K_100_13-7.3125-4 | 6AP03707 | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AP03707/SHIP_NY_32K_100_13-7.3125-4 | 6AP03707 | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.5 | in | FDA,4 | 7.125 | 7.5 | TM 141 / Visual
6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | 6AP06528 | SHIP_NY_32K_100_13.25-7.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | 6AP06528 | SHIP_NY_32K_100_13.25-7.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | 6AP06528 | SHIP_NY_32K_100_13.25-7.75-5.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | 6AP06528 | SHIP_NY_32K_100_13.25-7.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 5.9375 | in | FDA,4 | 5.5625 | 5.9375 | TM 141 / Visual
6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | 6AP06528 | SHIP_NY_32K_100_13.25-7.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 13.25 | in | FDA,4 | 13.0625 | 13.4375 | TM 141 / Visual
6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | 6AP06528 | SHIP_NY_32K_100_13.25-7.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | 6AP06528 | SHIP_NY_32K_100_13.25-7.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | 6AP06528 | SHIP_NY_32K_100_13.25-7.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.75 | in | FDA,4 | 7.5625 | 7.9375 | TM 141 / Visual
6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | 6AP06550 | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 20 |  | FDA,0 |  |  | TM 141 / Visual
6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | 6AP06550 | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | 6AP06550 | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | 6AP06550 | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | 6AP06550 | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 10.375 | in | FDA,4 | 10.375 | 10.625 | TM 141 / Visual
6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | 6AP06550 | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 17.5 | in | FDA,4 | 17.25 | 17.5 | TM 141 / Visual
6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | 6AP06550 | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | 6AP06550 | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | 6AP06550 | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 11.9375 | in | FDA,4 | 11.9375 | 12.1875 | TM 141 / Visual
6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 6AP06728 | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual
6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 6AP06728 | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 6AP06728 | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 6AP06728 | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 6AP06728 | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6 | in | FDA,4 | 5.875 | 6.25 | TM 141 / Visual
6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 6AP06728 | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 7 | in | FDA,4 | 6.9375 | 7.3125 | TM 141 / Visual
6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 6AP06728 | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 6AP06728 | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 6AP06728 | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.125 | in | FDA,4 | 6.8125 | 7.1875 | TM 141 / Visual
6AP07207/Simple Kind To Skin facial 7ct Shipper Canada | 6AP07207 | Simple Kind To Skin facial 7ct Shipper Canada | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AP07207/Simple Kind To Skin facial 7ct Shipper Canada | 6AP07207 | Simple Kind To Skin facial 7ct Shipper Canada | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6AP07207/Simple Kind To Skin facial 7ct Shipper Canada | 6AP07207 | Simple Kind To Skin facial 7ct Shipper Canada | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AP07207/Simple Kind To Skin facial 7ct Shipper Canada | 6AP07207 | Simple Kind To Skin facial 7ct Shipper Canada | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 4 | in | FDA,4 | 3.9375 | 4.0625 | TM 141 / Visual
6AP07207/Simple Kind To Skin facial 7ct Shipper Canada | 6AP07207 | Simple Kind To Skin facial 7ct Shipper Canada | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 13 | in | FDA,4 | 12.9375 | 13.0625 | TM 141 / Visual
6AP07207/Simple Kind To Skin facial 7ct Shipper Canada | 6AP07207 | Simple Kind To Skin facial 7ct Shipper Canada | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AP07207/Simple Kind To Skin facial 7ct Shipper Canada | 6AP07207 | Simple Kind To Skin facial 7ct Shipper Canada | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AP07207/Simple Kind To Skin facial 7ct Shipper Canada | 6AP07207 | Simple Kind To Skin facial 7ct Shipper Canada | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.375 | in | FDA,4 | 7.25 | 7.375 | TM 141 / Visual
6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 6AP09215 | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual
6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 6AP09215 | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 6AP09215 | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 6AP09215 | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 6AP09215 | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 5.75 | in | FDA,4 | 5.5625 | 5.9375 | TM 141 / Visual
6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 6AP09215 | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 13.5 | in | FDA,4 | 13.25 | 13.625 | TM 141 / Visual
6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 6AP09215 | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 6AP09215 | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 6AP09215 | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 9.25 | in | FDA,4 | 9.0625 | 9.4375 | TM 141 / Visual
6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 6AP09384 | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual
6AP09384/Shipper_NY_Skip lot testing | 6AP09384 | Shipper_NY_Skip lot testing | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual
6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 6AP09384 | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AP09384/Shipper_NY_Skip lot testing | 6AP09384 | Shipper_NY_Skip lot testing | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 6AP09384 | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6AP09384/Shipper_NY_Skip lot testing | 6AP09384 | Shipper_NY_Skip lot testing | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 6AP09384 | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AP09384/Shipper_NY_Skip lot testing | 6AP09384 | Shipper_NY_Skip lot testing | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 6AP09384 | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 3.8125 | in | FDA,4 | 3.5625 | 3.8125 | TM 141 / Visual
6AP09384/Shipper_NY_Skip lot testing | 6AP09384 | Shipper_NY_Skip lot testing | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 3.75 | in | FDA,4 | 3.5625 | 3.8125 | TM 141 / Visual
6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 6AP09384 | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 8.5 | in | FDA,4 | 8.375 | 8.625 | TM 141 / Visual
6AP09384/Shipper_NY_Skip lot testing | 6AP09384 | Shipper_NY_Skip lot testing | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 8.5 | in | FDA,4 | 8.375 | 8.625 | TM 141 / Visual
6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 6AP09384 | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AP09384/Shipper_NY_Skip lot testing | 6AP09384 | Shipper_NY_Skip lot testing | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 6AP09384 | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AP09384/Shipper_NY_Skip lot testing | 6AP09384 | Shipper_NY_Skip lot testing | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 6AP09384 | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 5.75 | in | FDA,4 | 5.625 | 5.875 | TM 141 / Visual
6AP09384/Shipper_NY_Skip lot testing | 6AP09384 | Shipper_NY_Skip lot testing | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 5.75 | in | FDA,4 | 5.625 | 5.875 | TM 141 / Visual
6AP11325/SHIP_NY_44K_100_10.625-9.75-7.75 | 6AP11325 | SHIP_NY_44K_100_10.625-9.75-7.75 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AP11325/SHIP_NY_44K_100_10.625-9.75-7.75 | 6AP11325 | SHIP_NY_44K_100_10.625-9.75-7.75 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual
6AP11325/SHIP_NY_44K_100_10.625-9.75-7.75 | 6AP11325 | SHIP_NY_44K_100_10.625-9.75-7.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AP11325/SHIP_NY_44K_100_10.625-9.75-7.75 | 6AP11325 | SHIP_NY_44K_100_10.625-9.75-7.75 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7.75 | in | FDA,4 | 7.5625 | 7.9375 | TM 141 / Visual
6AP11325/SHIP_NY_44K_100_10.625-9.75-7.75 | 6AP11325 | SHIP_NY_44K_100_10.625-9.75-7.75 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 10.625 | in | FDA,4 | 10.4375 | 10.8125 | TM 141 / Visual
6AP11325/SHIP_NY_44K_100_10.625-9.75-7.75 | 6AP11325 | SHIP_NY_44K_100_10.625-9.75-7.75 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AP11325/SHIP_NY_44K_100_10.625-9.75-7.75 | 6AP11325 | SHIP_NY_44K_100_10.625-9.75-7.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AP11325/SHIP_NY_44K_100_10.625-9.75-7.75 | 6AP11325 | SHIP_NY_44K_100_10.625-9.75-7.75 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 9.75 | in | FDA,4 | 9.5625 | 9.9375 | TM 141 / Visual
6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | 6AP13472 | SHIP_NY_32K_100_16.3125-12.25-7.5625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | 6AP13472 | SHIP_NY_32K_100_16.3125-12.25-7.5625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | 6AP13472 | SHIP_NY_32K_100_16.3125-12.25-7.5625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | 6AP13472 | SHIP_NY_32K_100_16.3125-12.25-7.5625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7.75 | in | FDA,4 | 7.375 | 7.75 | TM 141 / Visual
6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | 6AP13472 | SHIP_NY_32K_100_16.3125-12.25-7.5625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 16.25 | in | FDA,4 | 16.125 | 16.5 | TM 141 / Visual
6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | 6AP13472 | SHIP_NY_32K_100_16.3125-12.25-7.5625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | 6AP13472 | SHIP_NY_32K_100_16.3125-12.25-7.5625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | 6AP13472 | SHIP_NY_32K_100_16.3125-12.25-7.5625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 12.25 | in | FDA,4 | 12.0625 | 12.4375 | TM 141 / Visual
6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | 6AP15984 | SHIP_NY_44K_100_15.25-10.375-7.6875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | 6AP15984 | SHIP_NY_44K_100_15.25-10.375-7.6875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual
6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | 6AP15984 | SHIP_NY_44K_100_15.25-10.375-7.6875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | 6AP15984 | SHIP_NY_44K_100_15.25-10.375-7.6875 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7.75 | in | FDA,4 | 7.5 | 7.875 | TM 141 / Visual
6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | 6AP15984 | SHIP_NY_44K_100_15.25-10.375-7.6875 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.125 | in | FDA,4 | 15.0625 | 15.4375 | TM 141 / Visual
6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | 6AP15984 | SHIP_NY_44K_100_15.25-10.375-7.6875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | 6AP15984 | SHIP_NY_44K_100_15.25-10.375-7.6875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | 6AP15984 | SHIP_NY_44K_100_15.25-10.375-7.6875 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 10.5 | in | FDA,4 | 10.1875 | 10.5625 | TM 141 / Visual
6AP22325/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | 6AP22325 | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Shipper Inspection
6AP22907/SHIP_NY_32K_100_8.5625-7.8125-6.4375 | 6AP22907 | SHIP_NY_32K_100_8.5625-7.8125-6.4375 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual
6AP22907/SHIP_NY_32K_100_8.5625-7.8125-6.4375 | 6AP22907 | SHIP_NY_32K_100_8.5625-7.8125-6.4375 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AP22907/SHIP_NY_32K_100_8.5625-7.8125-6.4375 | 6AP22907 | SHIP_NY_32K_100_8.5625-7.8125-6.4375 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6AP22907/SHIP_NY_32K_100_8.5625-7.8125-6.4375 | 6AP22907 | SHIP_NY_32K_100_8.5625-7.8125-6.4375 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AP22907/SHIP_NY_32K_100_8.5625-7.8125-6.4375 | 6AP22907 | SHIP_NY_32K_100_8.5625-7.8125-6.4375 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6.5 | in | FDA,4 | 6.25 | 6.625 | TM 141 / Visual
6AP22907/SHIP_NY_32K_100_8.5625-7.8125-6.4375 | 6AP22907 | SHIP_NY_32K_100_8.5625-7.8125-6.4375 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 8.5 | in | FDA,4 | 8.375 | 8.75 | TM 141 / Visual
6AP22907/SHIP_NY_32K_100_8.5625-7.8125-6.4375 | 6AP22907 | SHIP_NY_32K_100_8.5625-7.8125-6.4375 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AP22907/SHIP_NY_32K_100_8.5625-7.8125-6.4375 | 6AP22907 | SHIP_NY_32K_100_8.5625-7.8125-6.4375 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AP22907/SHIP_NY_32K_100_8.5625-7.8125-6.4375 | 6AP22907 | SHIP_NY_32K_100_8.5625-7.8125-6.4375 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.75 | in | FDA,4 | 7.625 | 8 | TM 141 / Visual
6AP23025/SHIP_NY_32K_100_11.25-5.8125-4.3125 | 6AP23025 | SHIP_NY_32K_100_11.25-5.8125-4.3125 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual
6AP23025/SHIP_NY_32K_100_11.25-5.8125-4.3125 | 6AP23025 | SHIP_NY_32K_100_11.25-5.8125-4.3125 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AP23025/SHIP_NY_32K_100_11.25-5.8125-4.3125 | 6AP23025 | SHIP_NY_32K_100_11.25-5.8125-4.3125 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6AP23025/SHIP_NY_32K_100_11.25-5.8125-4.3125 | 6AP23025 | SHIP_NY_32K_100_11.25-5.8125-4.3125 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AP23025/SHIP_NY_32K_100_11.25-5.8125-4.3125 | 6AP23025 | SHIP_NY_32K_100_11.25-5.8125-4.3125 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 4.375 | in | FDA,4 | 4.125 | 4.5 | TM 141 / Visual
6AP23025/SHIP_NY_32K_100_11.25-5.8125-4.3125 | 6AP23025 | SHIP_NY_32K_100_11.25-5.8125-4.3125 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.25 | in | FDA,4 | 11.0625 | 11.4375 | TM 141 / Visual
6AP23025/SHIP_NY_32K_100_11.25-5.8125-4.3125 | 6AP23025 | SHIP_NY_32K_100_11.25-5.8125-4.3125 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AP23025/SHIP_NY_32K_100_11.25-5.8125-4.3125 | 6AP23025 | SHIP_NY_32K_100_11.25-5.8125-4.3125 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AP23025/SHIP_NY_32K_100_11.25-5.8125-4.3125 | 6AP23025 | SHIP_NY_32K_100_11.25-5.8125-4.3125 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 5.75 | in | FDA,4 | 5.625 | 6 | TM 141 / Visual
6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | 6AP24284 | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | 6AP24284 | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 40 | Lbs./In. | FDA,0 | 40 |  | TM 141 / Visual
6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | 6AP24284 | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | 6AP24284 | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 9.625 | in | FDA,4 | 9.4375 | 9.8125 | TM 141 / Visual
6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | 6AP24284 | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.375 | in | FDA,4 | 15.1875 | 15.5625 | TM 141 / Visual
6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | 6AP24284 | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | 6AP24284 | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | 6AP24284 | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 10.5625 | in | FDA,4 | 10.375 | 10.75 | TM 141 / Visual
6AP24372/SHIP_NY_44K_100_15.25-10.375-15.4375 | 6AP24372 | SHIP_NY_44K_100_15.25-10.375-15.4375 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AP24372/SHIP_NY_44K_100_15.25-10.375-15.4375 | 6AP24372 | SHIP_NY_44K_100_15.25-10.375-15.4375 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual
6AP24372/SHIP_NY_44K_100_15.25-10.375-15.4375 | 6AP24372 | SHIP_NY_44K_100_15.25-10.375-15.4375 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AP24372/SHIP_NY_44K_100_15.25-10.375-15.4375 | 6AP24372 | SHIP_NY_44K_100_15.25-10.375-15.4375 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 15.4375 | in | FDA,4 | 15.25 | 15.625 | TM 141 / Visual
6AP24372/SHIP_NY_44K_100_15.25-10.375-15.4375 | 6AP24372 | SHIP_NY_44K_100_15.25-10.375-15.4375 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.25 | in | FDA,4 | 15.0625 | 15.4375 | TM 141 / Visual
6AP24372/SHIP_NY_44K_100_15.25-10.375-15.4375 | 6AP24372 | SHIP_NY_44K_100_15.25-10.375-15.4375 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AP24372/SHIP_NY_44K_100_15.25-10.375-15.4375 | 6AP24372 | SHIP_NY_44K_100_15.25-10.375-15.4375 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AP24372/SHIP_NY_44K_100_15.25-10.375-15.4375 | 6AP24372 | SHIP_NY_44K_100_15.25-10.375-15.4375 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 10.375 | in | FDA,4 | 10.1875 | 10.5625 | TM 141 / Visual
6AP24684/SHIP_NY_40K_100_15.375-10.5625-9.625 | 6AP24684 | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual
6AP27284/SHIP_NY_44K_100_15.375-10.5625-9.625 | 6AP27284 | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AP27284/SHIP_NY_44K_100_15.375-10.5625-9.625 | 6AP27284 | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual
6AP27284/SHIP_NY_44K_100_15.375-10.5625-9.625 | 6AP27284 | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AP27284/SHIP_NY_44K_100_15.375-10.5625-9.625 | 6AP27284 | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 9.625 | in | FDA,4 | 9.4375 | 9.8125 | TM 141 / Visual
6AP27284/SHIP_NY_44K_100_15.375-10.5625-9.625 | 6AP27284 | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.25 | in | FDA,4 | 15.1875 | 15.5625 | TM 141 / Visual
6AP27284/SHIP_NY_44K_100_15.375-10.5625-9.625 | 6AP27284 | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AP27284/SHIP_NY_44K_100_15.375-10.5625-9.625 | 6AP27284 | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AP27284/SHIP_NY_44K_100_15.375-10.5625-9.625 | 6AP27284 | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 10.437 | in | FDA,4 | 10.375 | 10.75 | TM 141 / Visual
6AP30895/SHIP_32_1BI | 6AP30895 | SHIP_32_1BI | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AP30895/SHIP_32_1BI | 6AP30895 | SHIP_32_1BI | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6AP30895/SHIP_32_1BI | 6AP30895 | SHIP_32_1BI | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AP30895/SHIP_32_1BI | 6AP30895 | SHIP_32_1BI | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7.875 | in | FDA,4 | 7.6875 | 8.0625 | TM 141 / Visual
6AP30895/SHIP_32_1BI | 6AP30895 | SHIP_32_1BI | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 12.5 | in | FDA,4 | 12.3125 | 12.6875 | TM 141 / Visual
6AP30895/SHIP_32_1BI | 6AP30895 | SHIP_32_1BI | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AP30895/SHIP_32_1BI | 6AP30895 | SHIP_32_1BI | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AP30895/SHIP_32_1BI | 6AP30895 | SHIP_32_1BI | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.25 | in | FDA,4 | 7.0625 | 7.4375 | TM 141 / Visual
6AP32825/SHIP_32_1P | 6AP32825 | SHIP_32_1P | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AP32825/SHIP_32_1P | 6AP32825 | SHIP_32_1P | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6AP32825/SHIP_32_1P | 6AP32825 | SHIP_32_1P | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AP32825/SHIP_32_1P | 6AP32825 | SHIP_32_1P | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 4.25 | in | FDA,4 | 4.125 | 4.5 | TM 141 / Visual
6AP32825/SHIP_32_1P | 6AP32825 | SHIP_32_1P | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.1875 | in | FDA,4 | 11.0625 | 11.4375 | TM 141 / Visual
6AP32825/SHIP_32_1P | 6AP32825 | SHIP_32_1P | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AP32825/SHIP_32_1P | 6AP32825 | SHIP_32_1P | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AP32825/SHIP_32_1P | 6AP32825 | SHIP_32_1P | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6 | in | FDA,4 | 5.625 | 6 | TM 141 / Visual
6AP71310/SHIP_NY_44K_100_18.875-11.625-9 | 6AP71310 | SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual
6AP71310/SHIP_NY_44K_100_18.875-11.625-9 | 6AP71310 | SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AP71310/SHIP_NY_44K_100_18.875-11.625-9 | 6AP71310 | SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual
6AP71310/SHIP_NY_44K_100_18.875-11.625-9 | 6AP71310 | SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AP71310/SHIP_NY_44K_100_18.875-11.625-9 | 6AP71310 | SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 9 | in | FDA,4 | 8.8125 | 9.1875 | TM 141 / Visual
6AP71310/SHIP_NY_44K_100_18.875-11.625-9 | 6AP71310 | SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 18.75 | in | FDA,4 | 18.6875 | 19.0625 | TM 141 / Visual
6AP71310/SHIP_NY_44K_100_18.875-11.625-9 | 6AP71310 | SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AP71310/SHIP_NY_44K_100_18.875-11.625-9 | 6AP71310 | SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AP71310/SHIP_NY_44K_100_18.875-11.625-9 | 6AP71310 | SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 11.625 | in | FDA,4 | 11.4375 | 11.8125 | TM 141 / Visual
6AP80625/SHIP_NY_32K_100_11.25-5.8125-8.625 | 6AP80625 | SHIP_NY_32K_100_11.25-5.8125-8.625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual
6AP80625/SHIP_NY_32K_100_11.25-5.8125-8.625 | 6AP80625 | SHIP_NY_32K_100_11.25-5.8125-8.625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AP80625/SHIP_NY_32K_100_11.25-5.8125-8.625 | 6AP80625 | SHIP_NY_32K_100_11.25-5.8125-8.625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6AP80625/SHIP_NY_32K_100_11.25-5.8125-8.625 | 6AP80625 | SHIP_NY_32K_100_11.25-5.8125-8.625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AP80625/SHIP_NY_32K_100_11.25-5.8125-8.625 | 6AP80625 | SHIP_NY_32K_100_11.25-5.8125-8.625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 8.625 | in | FDA,4 | 8.4375 | 8.8125 | TM 141 / Visual
6AP80625/SHIP_NY_32K_100_11.25-5.8125-8.625 | 6AP80625 | SHIP_NY_32K_100_11.25-5.8125-8.625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.25 | in | FDA,4 | 11.0625 | 11.4375 | TM 141 / Visual
6AP80625/SHIP_NY_32K_100_11.25-5.8125-8.625 | 6AP80625 | SHIP_NY_32K_100_11.25-5.8125-8.625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AP80625/SHIP_NY_32K_100_11.25-5.8125-8.625 | 6AP80625 | SHIP_NY_32K_100_11.25-5.8125-8.625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AP80625/SHIP_NY_32K_100_11.25-5.8125-8.625 | 6AP80625 | SHIP_NY_32K_100_11.25-5.8125-8.625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 5.8125 | in | FDA,4 | 5.625 | 6 | TM 141 / Visual
6AS27350/SHIP_44_2O | 6AS27350 | SHIP_44_2O | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AS27350/SHIP_44_2O | 6AS27350 | SHIP_44_2O | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual
6AS27350/SHIP_44_2O | 6AS27350 | SHIP_44_2O | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AS27350/SHIP_44_2O | 6AS27350 | SHIP_44_2O | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 9 | in | FDA,4 | 8.8125 | 9.1875 | TM 141 / Visual
6AS27350/SHIP_44_2O | 6AS27350 | SHIP_44_2O | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.375 | in | FDA,4 | 15.1875 | 15.5625 | TM 141 / Visual
6AS27350/SHIP_44_2O | 6AS27350 | SHIP_44_2O | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AS27350/SHIP_44_2O | 6AS27350 | SHIP_44_2O | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AS27350/SHIP_44_2O | 6AS27350 | SHIP_44_2O | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 11.375 | in | FDA,4 | 11.1875 | 11.5625 | TM 141 / Visual
6AS40750/SHIP_32_1AP | 6AS40750 | SHIP_32_1AP | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AS40750/SHIP_32_1AP | 6AS40750 | SHIP_32_1AP | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6AS40750/SHIP_32_1AP | 6AS40750 | SHIP_32_1AP | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AS40750/SHIP_32_1AP | 6AS40750 | SHIP_32_1AP | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 13.4375 | in | FDA,4 | 13.25 | 13.625 | TM 141 / Visual
6AS40750/SHIP_32_1AP | 6AS40750 | SHIP_32_1AP | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.625 | in | FDA,4 | 15.4375 | 15.8125 | TM 141 / Visual
6AS40750/SHIP_32_1AP | 6AS40750 | SHIP_32_1AP | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AS40750/SHIP_32_1AP | 6AS40750 | SHIP_32_1AP | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AS40750/SHIP_32_1AP | 6AS40750 | SHIP_32_1AP | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.25 | in | FDA,4 | 7.0625 | 7.4375 | TM 141 / Visual
6AS41950/SHIP_40_1O | 6AS41950 | SHIP_40_1O | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AS41950/SHIP_40_1O | 6AS41950 | SHIP_40_1O | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 40 | Lbs./In. | FDA,0 | 40 |  | TM 141 / Visual
6AS41950/SHIP_40_1O | 6AS41950 | SHIP_40_1O | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AS41950/SHIP_40_1O | 6AS41950 | SHIP_40_1O | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 8.3125 | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual
6AS41950/SHIP_40_1O | 6AS41950 | SHIP_40_1O | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 21.125 | in | FDA,4 | 20.9375 | 21.3125 | TM 141 / Visual
6AS41950/SHIP_40_1O | 6AS41950 | SHIP_40_1O | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AS41950/SHIP_40_1O | 6AS41950 | SHIP_40_1O | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AS41950/SHIP_40_1O | 6AS41950 | SHIP_40_1O | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 13.5625 | in | FDA,4 | 13.375 | 13.75 | TM 141 / Visual
6AU12310/SHIP_32_1U | 6AU12310 | SHIP_32_1U | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AU12310/SHIP_32_1U | 6AU12310 | SHIP_32_1U | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6AU12310/SHIP_32_1U | 6AU12310 | SHIP_32_1U | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AU12310/SHIP_32_1U | 6AU12310 | SHIP_32_1U | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 8 | in | FDA,4 | 7.8125 | 8.1875 | TM 141 / Visual
6AU12310/SHIP_32_1U | 6AU12310 | SHIP_32_1U | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 12.125 | in | FDA,4 | 11.8125 | 12.1875 | TM 141 / Visual
6AU12310/SHIP_32_1U | 6AU12310 | SHIP_32_1U | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AU12310/SHIP_32_1U | 6AU12310 | SHIP_32_1U | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AU12310/SHIP_32_1U | 6AU12310 | SHIP_32_1U | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6 | in | FDA,4 | 5.9375 | 6.3125 | TM 141 / Visual
6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | 6AU130TY | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual
6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | 6AU130TY | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | 6AU130TY | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | 6AU130TY | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | 6AU130TY | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 4.5 | in | FDA,4 | 4.375 | 4.625 | TM 141 / Visual
6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | 6AU130TY | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 6.625 | in | FDA,4 | 6.375 | 6.625 | TM 141 / Visual
6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | 6AU130TY | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | White (Done) |  |  |  |  | TM 141 / Visual
6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | 6AU130TY | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | Fold Over Tray w/header (Done) |  |  |  |  | TM 141 / Visual
6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | 6AU130TY | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 9.625 | in | FDA,4 | 9.5 | 9.75 | TM 141 / Visual
6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | 6AU25206 | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual
6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | 6AU25206 | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | No Go - Does not meet Specifications (OOS) |  |  |  |  | TM 141 / Visual
6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | 6AU25206 | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | 6AU25206 | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | 6AU25206 | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 4.75 | in | FDA,4 | 4.375 | 4.75 | TM 141 / Visual
6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | 6AU25206 | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.375 | in | FDA,4 | 11.0625 | 11.4375 | TM 141 / Visual
6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | 6AU25206 | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | 6AU25206 | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | 6AU25206 | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 8.5 | in | FDA,4 | 8.25 | 8.625 | TM 141 / Visual
6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 6AU53601 | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 50 |  | FDA,0 |  |  | TM 141 / Visual
6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 6AU53601 | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 6AU53601 | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 6AU53601 | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 6AU53601 | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 8.5 | in | FDA,4 | 8.3125 | 8.6875 | TM 141 / Visual
6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 6AU53601 | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.875 | in | FDA,4 | 11.625 | 12 | TM 141 / Visual
6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 6AU53601 | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 6AU53601 | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 6AU53601 | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 9.5 | in | FDA,4 | 9.1875 | 9.5625 | TM 141 / Visual
6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 6AU56400 | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 50 |  | FDA,0 |  |  | TM 141 / Visual
6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 6AU56400 | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 6AU56400 | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 6AU56400 | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 6AU56400 | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7.8125 | in | FDA,4 | 7.5005 | 7.8755 | TM 141 / Visual
6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 6AU56400 | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.75 | in | FDA,4 | 11.5625 | 11.9375 | TM 141 / Visual
6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 6AU56400 | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 6AU56400 | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 6AU56400 | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 9.625 | in | FDA,4 | 9.4375 | 9.8125 | TM 141 / Visual
6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | 6AU57314 | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual
6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | 6AU57314 | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | 6AU57314 | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | 6AU57314 | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | 6AU57314 | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 8.0625 | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual
6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | 6AU57314 | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 7.5625 | in | FDA,4 | 7.5625 | 7.9375 | TM 141 / Visual
6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | 6AU57314 | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | 6AU57314 | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | 6AU57314 | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.375 | in | FDA,4 | 6.0625 | 6.4375 | TM 141 / Visual
6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 6AU71230 | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 50 |  | FDA,0 |  |  | TM 141 / Visual
6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 6AU71230 | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 6AU71230 | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 6AU71230 | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 6AU71230 | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7.75 | in | FDA,4 | 7.5625 | 7.9375 | TM 141 / Visual
6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 6AU71230 | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 9.5625 | in | FDA,4 | 9.3755 | 9.7505 | TM 141 / Visual
6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 6AU71230 | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 6AU71230 | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 6AU71230 | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 9.188 | in | FDA,4 | 9.0005 | 9.3755 | TM 141 / Visual
6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 6BA709SWN | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual
6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 6BA709SWN | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 6BA709SWN | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 40 | Lbs./In. | FDA,0 | 40 |  | TM 141 / Visual
6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 6BA709SWN | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 6BA709SWN | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6.125 | in | FDA,4 | 5.9375 | 6.3125 | TM 141 / Visual
6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 6BA709SWN | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 14.75 | in | FDA,4 | 14.5625 | 14.9375 | TM 141 / Visual
6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 6BA709SWN | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 6BA709SWN | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 6BA709SWN | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 9.375 | in | FDA,4 | 9.1875 | 9.5625 | TM 141 / Visual
6BB10800/SHIP_32_1S | 6BB10800 | SHIP_32_1S | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6BB10800/SHIP_32_1S | 6BB10800 | SHIP_32_1S | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6BB10800/SHIP_32_1S | 6BB10800 | SHIP_32_1S | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6BB10800/SHIP_32_1S | 6BB10800 | SHIP_32_1S | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 5.3125 | in | FDA,4 | 5.125 | 5.5 | TM 141 / Visual
6BB10800/SHIP_32_1S | 6BB10800 | SHIP_32_1S | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.875 | in | FDA,4 | 11.6875 | 12.0625 | TM 141 / Visual
6BB10800/SHIP_32_1S | 6BB10800 | SHIP_32_1S | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6BB10800/SHIP_32_1S | 6BB10800 | SHIP_32_1S | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6BB10800/SHIP_32_1S | 6BB10800 | SHIP_32_1S | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.25 | in | FDA,4 | 6.0625 | 6.4375 | TM 141 / Visual
6BD17385/SHIP_44_2AB | 6BD17385 | SHIP_44_2AB | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6BD17385/SHIP_44_2AB | 6BD17385 | SHIP_44_2AB | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual
6BD17385/SHIP_44_2AB | 6BD17385 | SHIP_44_2AB | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6BD17385/SHIP_44_2AB | 6BD17385 | SHIP_44_2AB | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6.5 | in | FDA,4 | 6.3125 | 6.6875 | TM 141 / Visual
6BD17385/SHIP_44_2AB | 6BD17385 | SHIP_44_2AB | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 20 | in | FDA,4 | 19.8125 | 20.1875 | TM 141 / Visual
6BD17385/SHIP_44_2AB | 6BD17385 | SHIP_44_2AB | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6BD17385/SHIP_44_2AB | 6BD17385 | SHIP_44_2AB | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6BD17385/SHIP_44_2AB | 6BD17385 | SHIP_44_2AB | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 12 | in | FDA,4 | 11.8125 | 12.1875 | TM 141 / Visual
6BP21195/SHIP_NY_44K_100_14.375-12.1875-9.25 | 6BP21195 | SHIP_NY_44K_100_14.375-12.1875-9.25 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6BP21195/SHIP_NY_44K_100_14.375-12.1875-9.25 | 6BP21195 | SHIP_NY_44K_100_14.375-12.1875-9.25 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual
6BP21195/SHIP_NY_44K_100_14.375-12.1875-9.25 | 6BP21195 | SHIP_NY_44K_100_14.375-12.1875-9.25 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6BP21195/SHIP_NY_44K_100_14.375-12.1875-9.25 | 6BP21195 | SHIP_NY_44K_100_14.375-12.1875-9.25 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 9.125 | in | FDA,4 | 9.0625 | 9.4375 | TM 141 / Visual
6BP21195/SHIP_NY_44K_100_14.375-12.1875-9.25 | 6BP21195 | SHIP_NY_44K_100_14.375-12.1875-9.25 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 14.375 | in | FDA,4 | 14.1875 | 14.5625 | TM 141 / Visual
6BP21195/SHIP_NY_44K_100_14.375-12.1875-9.25 | 6BP21195 | SHIP_NY_44K_100_14.375-12.1875-9.25 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6BP21195/SHIP_NY_44K_100_14.375-12.1875-9.25 | 6BP21195 | SHIP_NY_44K_100_14.375-12.1875-9.25 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6BP21195/SHIP_NY_44K_100_14.375-12.1875-9.25 | 6BP21195 | SHIP_NY_44K_100_14.375-12.1875-9.25 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 12.25 | in | FDA,4 | 12 | 12.375 | TM 141 / Visual
6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 6BP25372 | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 6BP25372 | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | TM 141 / Visual
6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 6BP25372 | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual
6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 6BP25372 | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 6BP25372 | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 15.25 | in | FDA,4 | 15 | 15.375 | TM 141 / Visual
6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 6BP25372 | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.125 | in | FDA,4 | 15 | 15.375 | TM 141 / Visual
6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 6BP25372 | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 6BP25372 | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 6BP25372 | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 10.25 | in | FDA,4 | 10 | 10.375 | TM 141 / Visual
6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | 6BP30895 | SHIP_NY_32K_100_12.5-7.25-7.875 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual
6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | 6BP30895 | SHIP_NY_32K_100_12.5-7.25-7.875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | 6BP30895 | SHIP_NY_32K_100_12.5-7.25-7.875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | 6BP30895 | SHIP_NY_32K_100_12.5-7.25-7.875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | 6BP30895 | SHIP_NY_32K_100_12.5-7.25-7.875 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7.6875 | in | FDA,4 | 7.6875 | 8.0625 | TM 141 / Visual
6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | 6BP30895 | SHIP_NY_32K_100_12.5-7.25-7.875 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 12.375 | in | FDA,4 | 12.3125 | 12.6875 | TM 141 / Visual
6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | 6BP30895 | SHIP_NY_32K_100_12.5-7.25-7.875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | 6BP30895 | SHIP_NY_32K_100_12.5-7.25-7.875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | 6BP30895 | SHIP_NY_32K_100_12.5-7.25-7.875 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.125 | in | FDA,4 | 7.0625 | 7.4375 | TM 141 / Visual
6BP54072/SHIP_44_2L | 6BP54072 | SHIP_44_2L | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6BP54072/SHIP_44_2L | 6BP54072 | SHIP_44_2L | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual
6BP54072/SHIP_44_2L | 6BP54072 | SHIP_44_2L | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6BP54072/SHIP_44_2L | 6BP54072 | SHIP_44_2L | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 15.4375 | in | FDA,4 | 15.25 | 15.625 | TM 141 / Visual
6BP54072/SHIP_44_2L | 6BP54072 | SHIP_44_2L | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.25 | in | FDA,4 | 15.0625 | 15.4375 | TM 141 / Visual
6BP54072/SHIP_44_2L | 6BP54072 | SHIP_44_2L | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6BP54072/SHIP_44_2L | 6BP54072 | SHIP_44_2L | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6BP54072/SHIP_44_2L | 6BP54072 | SHIP_44_2L | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 10.375 | in | FDA,4 | 10.1875 | 10.5625 | TM 141 / Visual
6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | 6BU13195 | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual
6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | 6BU13195 | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | 6BU13195 | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | 6BU13195 | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | 6BU13195 | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 12.25 | in | FDA,4 | 12.0625 | 12.4375 | TM 141 / Visual
6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | 6BU13195 | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 13.5 | in | FDA,4 | 13.4375 | 13.8125 | TM 141 / Visual
6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | 6BU13195 | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | 6BU13195 | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | 6BU13195 | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 4.375 | in | FDA,4 | 4.3125 | 4.6875 | TM 141 / Visual
6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | 6BU25206 | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 10 |  | FDA,0 |  |  | TM 141 / Visual
6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | 6BU25206 | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | 6BU25206 | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | 6BU25206 | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | 6BU25206 | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 4.5 | in | FDA,4 | 4.4375 | 4.8125 | TM 141 / Visual
6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | 6BU25206 | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.3125 | in | FDA,4 | 11.125 | 11.5 | TM 141 / Visual
6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | 6BU25206 | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | 6BU25206 | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | 6BU25206 | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 8.25 | in | FDA,4 | 8.25 | 8.625 | TM 141 / Visual
6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | 6CA972CK | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 32 |  | FDA,0 |  |  | TM 141 / Visual
6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | 6CA972CK | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | 6CA972CK | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual
6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | 6CA972CK | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | 6CA972CK | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 10.25 | in | FDA,4 | 10.0625 | 10.4375 | TM 141 / Visual
6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | 6CA972CK | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 18 | in | FDA,4 | 17.8125 | 18.1875 | TM 141 / Visual
6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | 6CA972CK | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | 6CA972CK | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | 6CA972CK | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 8.25 | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual
6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | 6CJ21396 | SHIP_NY_44K_100_16.625-9.375-7 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | 6CJ21396 | SHIP_NY_44K_100_16.625-9.375-7 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual
6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | 6CJ21396 | SHIP_NY_44K_100_16.625-9.375-7 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | 6CJ21396 | SHIP_NY_44K_100_16.625-9.375-7 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7.1875 | in | FDA,4 | 6.8125 | 7.1875 | TM 141 / Visual
6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | 6CJ21396 | SHIP_NY_44K_100_16.625-9.375-7 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 16.75 | in | FDA,4 | 16.4375 | 16.8125 | TM 141 / Visual
6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | 6CJ21396 | SHIP_NY_44K_100_16.625-9.375-7 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | 6CJ21396 | SHIP_NY_44K_100_16.625-9.375-7 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | 6CJ21396 | SHIP_NY_44K_100_16.625-9.375-7 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 9.5 | in | FDA,4 | 9.1875 | 9.5625 | TM 141 / Visual
6CP02724/SHIP_NY_32K_100_11-5.8125-4.0625 | 6CP02724 | SHIP_NY_32K_100_11-5.8125-4.0625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6CP02724/SHIP_NY_32K_100_11-5.8125-4.0625 | 6CP02724 | SHIP_NY_32K_100_11-5.8125-4.0625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6CP02724/SHIP_NY_32K_100_11-5.8125-4.0625 | 6CP02724 | SHIP_NY_32K_100_11-5.8125-4.0625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6CP02724/SHIP_NY_32K_100_11-5.8125-4.0625 | 6CP02724 | SHIP_NY_32K_100_11-5.8125-4.0625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 4.25 | in | FDA,4 | 3.875 | 4.25 | TM 141 / Visual
6CP02724/SHIP_NY_32K_100_11-5.8125-4.0625 | 6CP02724 | SHIP_NY_32K_100_11-5.8125-4.0625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11 | in | FDA,4 | 10.8125 | 11.1875 | TM 141 / Visual
6CP02724/SHIP_NY_32K_100_11-5.8125-4.0625 | 6CP02724 | SHIP_NY_32K_100_11-5.8125-4.0625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6CP02724/SHIP_NY_32K_100_11-5.8125-4.0625 | 6CP02724 | SHIP_NY_32K_100_11-5.8125-4.0625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6CP02724/SHIP_NY_32K_100_11-5.8125-4.0625 | 6CP02724 | SHIP_NY_32K_100_11-5.8125-4.0625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6 | in | FDA,4 | 5.625 | 6 | TM 141 / Visual
6CP22884/SHIP_NY_44K_100_15.375-10.5625-9.625 | 6CP22884 | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Shipper Inspection
6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 6DC22370 | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual
6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 6DC22370 | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 6DC22370 | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 6DC22370 | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 6DC22370 | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 10.375 | in | FDA,4 | 10.1875 | 10.5625 | TM 141 / Visual
6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 6DC22370 | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11 | in | FDA,4 | 10.9375 | 11.3125 | TM 141 / Visual
6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 6DC22370 | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 6DC22370 | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 6DC22370 | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.25 | in | FDA,4 | 6.1875 | 6.5625 | TM 141 / Visual
6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | 6DD33333 | SHIP_NY_44K_100_21-12.5-10.5 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | 6DD33333 | SHIP_NY_44K_100_21-12.5-10.5 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual
6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | 6DD33333 | SHIP_NY_44K_100_21-12.5-10.5 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | 6DD33333 | SHIP_NY_44K_100_21-12.5-10.5 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 10.5 | in | FDA,4 | 10.3125 | 10.6875 | TM 141 / Visual
6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | 6DD33333 | SHIP_NY_44K_100_21-12.5-10.5 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 21 | in | FDA,4 | 20.8125 | 21.1875 | TM 141 / Visual
6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | 6DD33333 | SHIP_NY_44K_100_21-12.5-10.5 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | 6DD33333 | SHIP_NY_44K_100_21-12.5-10.5 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | 6DD33333 | SHIP_NY_44K_100_21-12.5-10.5 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 12.625 | in | FDA,4 | 12.3125 | 12.6875 | TM 141 / Visual
6DP63884/SHIP_NY_44K_100_15.375-10.5625-9.625 | 6DP63884 | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual
6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | 6EP44584 | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 50 |  | FDA,0 |  |  | TM 141 / Visual
6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | 6EP44584 | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | 6EP44584 | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | 6EP44584 | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | 6EP44584 | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 9.75 | in | FDA,4 | 9.4375 | 9.8125 | TM 141 / Visual
6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | 6EP44584 | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.25 | in | FDA,4 | 15 | 15.375 | TM 141 / Visual
6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | 6EP44584 | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | 6EP44584 | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | 6EP44584 | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 10.25 | in | FDA,4 | 10 | 10.375 | TM 141 / Visual
6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 6FB71200 | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 50 |  | FDA,0 |  |  | TM 141 / Visual
6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 6FB71200 | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 6FB71200 | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 6FB71200 | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 6FB71200 | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 5.125 | in | FDA,4 | 4.8125 | 5.1875 | TM 141 / Visual
6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 6FB71200 | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 14.75 | in | FDA,4 | 14.4375 | 14.8125 | TM 141 / Visual
6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 6FB71200 | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 6FB71200 | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC w/ shyed flaps (Done) |  |  |  |  | TM 141 / Visual
6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 6FB71200 | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.344 | in | FDA,4 | 6.1875 | 6.5625 | TM 141 / Visual
6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | 6FJ14108 | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual
6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | 6FJ14108 | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | 6FJ14108 | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual
6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | 6FJ14108 | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | 6FJ14108 | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7 | in | FDA,4 | 6.8125 | 7.1875 | TM 141 / Visual
6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | 6FJ14108 | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 16.625 | in | FDA,4 | 16.4375 | 16.8125 | TM 141 / Visual
6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | 6FJ14108 | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | 6FJ14108 | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | 6FJ14108 | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 8.5 | in | FDA,4 | 8.4375 | 8.8125 | TM 141 / Visual
6FJ14143/SHIP_NY_44K_100_16.5-10.875-5.5 | 6FJ14143 | SHIP_NY_44K_100_16.5-10.875-5.5 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6FJ14143/SHIP_NY_44K_100_16.5-10.875-5.5 | 6FJ14143 | SHIP_NY_44K_100_16.5-10.875-5.5 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual
6FJ14143/SHIP_NY_44K_100_16.5-10.875-5.5 | 6FJ14143 | SHIP_NY_44K_100_16.5-10.875-5.5 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6FJ14143/SHIP_NY_44K_100_16.5-10.875-5.5 | 6FJ14143 | SHIP_NY_44K_100_16.5-10.875-5.5 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 5.5 | in | FDA,4 | 5.3125 | 5.6875 | TM 141 / Visual
6FJ14143/SHIP_NY_44K_100_16.5-10.875-5.5 | 6FJ14143 | SHIP_NY_44K_100_16.5-10.875-5.5 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 16.5 | in | FDA,4 | 16.3125 | 16.6875 | TM 141 / Visual
6FJ14143/SHIP_NY_44K_100_16.5-10.875-5.5 | 6FJ14143 | SHIP_NY_44K_100_16.5-10.875-5.5 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6FJ14143/SHIP_NY_44K_100_16.5-10.875-5.5 | 6FJ14143 | SHIP_NY_44K_100_16.5-10.875-5.5 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6FJ14143/SHIP_NY_44K_100_16.5-10.875-5.5 | 6FJ14143 | SHIP_NY_44K_100_16.5-10.875-5.5 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 10.875 | in | FDA,4 | 10.6875 | 11.0625 | TM 141 / Visual
6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | 6GB10300 | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual
6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | 6GB10300 | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | 6GB10300 | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | 6GB10300 | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | 6GB10300 | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 5 | in | FDA,4 | 4.8125 | 5.1875 | TM 141 / Visual
6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | 6GB10300 | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 14.75 | in | FDA,4 | 14.4375 | 14.8125 | TM 141 / Visual
6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | 6GB10300 | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | 6GB10300 | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC w/ shyed flaps (Done) |  |  |  |  | TM 141 / Visual
6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | 6GB10300 | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.25 | in | FDA,4 | 6.1875 | 6.5625 | TM 141 / Visual
6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 6GD35185 | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 50 |  | FDA,0 |  |  | TM 141 / Visual
6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 6GD35185 | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 6GD35185 | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 6GD35185 | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 6GD35185 | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 12.4375 | in | FDA,4 | 12.25 | 12.625 | TM 141 / Visual
6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 6GD35185 | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 13.25 | in | FDA,4 | 13.0625 | 13.4375 | TM 141 / Visual
6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 6GD35185 | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 6GD35185 | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 6GD35185 | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 12 | in | FDA,4 | 11.8125 | 12.1875 | TM 141 / Visual
6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | 6GS48450 | SHIP_NY_29K_100_17.125-6.6875-11.75 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | 6GS48450 | SHIP_NY_29K_100_17.125-6.6875-11.75 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 29 | Lbs./In. | FDA,0 | 29 |  | TM 141 / Visual
6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | 6GS48450 | SHIP_NY_29K_100_17.125-6.6875-11.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | 6GS48450 | SHIP_NY_29K_100_17.125-6.6875-11.75 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 11.875 | in | FDA,4 | 11.5625 | 11.9375 | TM 141 / Visual
6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | 6GS48450 | SHIP_NY_29K_100_17.125-6.6875-11.75 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 17.125 | in | FDA,4 | 16.9375 | 17.3125 | TM 141 / Visual
6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | 6GS48450 | SHIP_NY_29K_100_17.125-6.6875-11.75 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | 6GS48450 | SHIP_NY_29K_100_17.125-6.6875-11.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | 6GS48450 | SHIP_NY_29K_100_17.125-6.6875-11.75 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.75 | in | FDA,4 | 6.5 | 6.875 | TM 141 / Visual
6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | 6HB60307 | SHIP_NY_40K_100_16.5-12.3125-6.1875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | 6HB60307 | SHIP_NY_40K_100_16.5-12.3125-6.1875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 40 | Lbs./In. | FDA,0 | 40 |  | TM 141 / Visual
6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | 6HB60307 | SHIP_NY_40K_100_16.5-12.3125-6.1875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | 6HB60307 | SHIP_NY_40K_100_16.5-12.3125-6.1875 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6.312 | in | FDA,4 | 6 | 6.375 | TM 141 / Visual
6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | 6HB60307 | SHIP_NY_40K_100_16.5-12.3125-6.1875 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 16.5 | in | FDA,4 | 16.3125 | 16.6875 | TM 141 / Visual
6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | 6HB60307 | SHIP_NY_40K_100_16.5-12.3125-6.1875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | 6HB60307 | SHIP_NY_40K_100_16.5-12.3125-6.1875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | 6HB60307 | SHIP_NY_40K_100_16.5-12.3125-6.1875 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 12.25 | in | FDA,4 | 12.125 | 12.5 | TM 141 / Visual
6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 6HC69900 | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 50 |  | FDA,0 |  |  | TM 141 / Visual
6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 6HC69900 | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 6HC69900 | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 6HC69900 | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 6HC69900 | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 5.25 | in | FDA,4 | 5 | 5.375 | TM 141 / Visual
6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 6HC69900 | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.75 | in | FDA,4 | 15.5625 | 15.9375 | TM 141 / Visual
6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 6HC69900 | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 6HC69900 | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 6HC69900 | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 8.625 | in | FDA,4 | 8.4375 | 8.8125 | TM 141 / Visual
6JA580FW/Shipper_NY_Skip lot testing | 6JA580FW | Shipper_NY_Skip lot testing | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Shipper Inspection
6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | 6JA580FW | SHIP_NY_44K_100_16.750-8.125-10.750 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | 6JA580FW | SHIP_NY_44K_100_16.750-8.125-10.750 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual
6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | 6JA580FW | SHIP_NY_44K_100_16.750-8.125-10.750 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | 6JA580FW | SHIP_NY_44K_100_16.750-8.125-10.750 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 10.875 | in | FDA,4 | 10.5625 | 10.9375 | TM 141 / Visual
6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | 6JA580FW | SHIP_NY_44K_100_16.750-8.125-10.750 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 16.9375 | in | FDA,4 | 16.5625 | 16.9375 | TM 141 / Visual
6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | 6JA580FW | SHIP_NY_44K_100_16.750-8.125-10.750 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | 6JA580FW | SHIP_NY_44K_100_16.750-8.125-10.750 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | 6JA580FW | SHIP_NY_44K_100_16.750-8.125-10.750 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 8.1875 | in | FDA,4 | 7.9375 | 8.3125 | TM 141 / Visual
6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | 6LB40600 | SHIP_NY_29K_100_12.375-5.5-4.625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | 6LB40600 | SHIP_NY_29K_100_12.375-5.5-4.625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 29 | Lbs./In. | FDA,0 | 29 |  | TM 141 / Visual
6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | 6LB40600 | SHIP_NY_29K_100_12.375-5.5-4.625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | 6LB40600 | SHIP_NY_29K_100_12.375-5.5-4.625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 4.75 | in | FDA,4 | 4.4375 | 4.8125 | TM 141 / Visual
6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | 6LB40600 | SHIP_NY_29K_100_12.375-5.5-4.625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 12.437 | in | FDA,4 | 12.1875 | 12.5625 | TM 141 / Visual
6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | 6LB40600 | SHIP_NY_29K_100_12.375-5.5-4.625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | 6LB40600 | SHIP_NY_29K_100_12.375-5.5-4.625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | 6LB40600 | SHIP_NY_29K_100_12.375-5.5-4.625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 5.5 | in | FDA,4 | 5.3125 | 5.6875 | TM 141 / Visual
6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | 6OA404S48 | SHIP_NY_44K_100_15.625-12.5-6.1875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | 6OA404S48 | SHIP_NY_44K_100_15.625-12.5-6.1875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual
6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | 6OA404S48 | SHIP_NY_44K_100_15.625-12.5-6.1875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | 6OA404S48 | SHIP_NY_44K_100_15.625-12.5-6.1875 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6 | in | FDA,4 | 6 | 6.375 | TM 141 / Visual
6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | 6OA404S48 | SHIP_NY_44K_100_15.625-12.5-6.1875 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.625 | in | FDA,4 | 15.4375 | 15.8125 | TM 141 / Visual
6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | 6OA404S48 | SHIP_NY_44K_100_15.625-12.5-6.1875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | 6OA404S48 | SHIP_NY_44K_100_15.625-12.5-6.1875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | 6OA404S48 | SHIP_NY_44K_100_15.625-12.5-6.1875 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 12.5625 | in | FDA,4 | 12.3125 | 12.6875 | TM 141 / Visual
6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 6OA540S48 | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual
6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 6OA540S48 | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 6OA540S48 | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 6OA540S48 | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 6OA540S48 | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 12.625 | in | FDA,4 | 12.4375 | 12.8125 | TM 141 / Visual
6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 6OA540S48 | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.875 | in | FDA,4 | 11.6875 | 12.0625 | TM 141 / Visual
6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 6OA540S48 | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 6OA540S48 | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 6OA540S48 | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.75 | in | FDA,4 | 7.5625 | 7.9375 | TM 141 / Visual
6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 6OA632S60 | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual
6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 6OA632S60 | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 6OA632S60 | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 6OA632S60 | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 6OA632S60 | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 14.25 | in | FDA,4 | 13.9375 | 14.3125 | TM 141 / Visual
6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 6OA632S60 | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 19.5 | in | FDA,4 | 19.1875 | 19.5625 | TM 141 / Visual
6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 6OA632S60 | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 6OA632S60 | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | HSC (Done) |  |  |  |  | TM 141 / Visual
6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 6OA632S60 | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 8.25 | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual
6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 6OB123ST | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual
6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 6OB123ST | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 6OB123ST | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 6OB123ST | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 6OB123ST | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 5.875 | in | FDA,4 | 5.5625 | 5.9375 | TM 141 / Visual
6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 6OB123ST | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 12.5 | in | FDA,4 | 12.1875 | 12.5625 | TM 141 / Visual
6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 6OB123ST | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 6OB123ST | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 6OB123ST | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 10.75 | in | FDA,4 | 10.5625 | 10.9375 | TM 141 / Visual
6OB31400/SHIP_32_1AL | 6OB31400 | SHIP_32_1AL | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6OB31400/SHIP_32_1AL | 6OB31400 | SHIP_32_1AL | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6OB31400/SHIP_32_1AL | 6OB31400 | SHIP_32_1AL | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OB31400/SHIP_32_1AL | 6OB31400 | SHIP_32_1AL | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 5 | in | FDA,4 | 4.8125 | 5.1875 | TM 141 / Visual
6OB31400/SHIP_32_1AL | 6OB31400 | SHIP_32_1AL | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 14.625 | in | FDA,4 | 14.4375 | 14.8125 | TM 141 / Visual
6OB31400/SHIP_32_1AL | 6OB31400 | SHIP_32_1AL | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6OB31400/SHIP_32_1AL | 6OB31400 | SHIP_32_1AL | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6OB31400/SHIP_32_1AL | 6OB31400 | SHIP_32_1AL | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.375 | in | FDA,4 | 6.1875 | 6.5625 | TM 141 / Visual
6OB389PD/Ship_NA_3A | 6OB389PD | Ship_NA_3A | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OB389PD/Ship_NA_3A | 6OB389PD | Ship_NA_3A | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 14.5 | in | FDA,4 | 14.3125 | 14.6875 | TM 141 / Visual
6OB389PD/Ship_NA_3A | 6OB389PD | Ship_NA_3A | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | Layer Pad (Done) |  |  |  |  | TM 141 / Visual
6OB389PD/Ship_NA_3A | 6OB389PD | Ship_NA_3A | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.437 | in | FDA,4 | 6.0625 | 6.4375 | TM 141 / Visual
6OC22324/SHIP_32_1N | 6OC22324 | SHIP_32_1N | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6OC22324/SHIP_32_1N | 6OC22324 | SHIP_32_1N | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6OC22324/SHIP_32_1N | 6OC22324 | SHIP_32_1N | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OC22324/SHIP_32_1N | 6OC22324 | SHIP_32_1N | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 10.5 | in | FDA,4 | 10.1875 | 10.5625 | TM 141 / Visual
6OC22324/SHIP_32_1N | 6OC22324 | SHIP_32_1N | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.125 | in | FDA,4 | 10.9375 | 11.3125 | TM 141 / Visual
6OC22324/SHIP_32_1N | 6OC22324 | SHIP_32_1N | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6OC22324/SHIP_32_1N | 6OC22324 | SHIP_32_1N | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6OC22324/SHIP_32_1N | 6OC22324 | SHIP_32_1N | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.375 | in | FDA,4 | 6.1875 | 6.5625 | TM 141 / Visual
6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | 6OD01524 | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual
6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | 6OD01524 | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | 6OD01524 | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | 6OD01524 | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | 6OD01524 | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 4.75 | in | FDA,4 | 4.5 | 4.875 | TM 141 / Visual
6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | 6OD01524 | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 8.25 | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual
6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | 6OD01524 | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | 6OD01524 | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | 6OD01524 | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.25 | in | FDA,4 | 6.0625 | 6.4375 | TM 141 / Visual
6OD13500A/SHIP_32_1AD | 6OD13500A | SHIP_32_1AD | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6OD13500A/SHIP_32_1AD | 6OD13500A | SHIP_32_1AD | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6OD13500A/SHIP_32_1AD | 6OD13500A | SHIP_32_1AD | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OD13500A/SHIP_32_1AD | 6OD13500A | SHIP_32_1AD | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6.25 | in | FDA,4 | 5.9375 | 6.3125 | TM 141 / Visual
6OD13500A/SHIP_32_1AD | 6OD13500A | SHIP_32_1AD | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 13.25 | in | FDA,4 | 13.0625 | 13.4375 | TM 141 / Visual
6OD13500A/SHIP_32_1AD | 6OD13500A | SHIP_32_1AD | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6OD13500A/SHIP_32_1AD | 6OD13500A | SHIP_32_1AD | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6OD13500A/SHIP_32_1AD | 6OD13500A | SHIP_32_1AD | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 12 | in | FDA,4 | 11.8125 | 12.1875 | TM 141 / Visual
6OD239DR/SHIP_44_1A | 6OD239DR | SHIP_44_1A | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | Divider |  |  |  |  | Shipper Style
6OD239DR/SHIP_44_1A | 6OD239DR | SHIP_44_1A | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6OD239DR/SHIP_44_1A | 6OD239DR | SHIP_44_1A | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual
6OD239DR/SHIP_44_1A | 6OD239DR | SHIP_44_1A | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OD239DR/SHIP_44_1A | 6OD239DR | SHIP_44_1A | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 7.75 | in | FDA,4 | 7.4375 | 7.8125 | TM 141 / Visual
6OD239DR/SHIP_44_1A | 6OD239DR | SHIP_44_1A | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6OD239DR/SHIP_44_1A | 6OD239DR | SHIP_44_1A | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 15.375 | in | FDA,4 | 15.1875 | 15.5625 | TM 141 / Visual
6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | 6OD36055 | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual
6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | 6OD36055 | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | 6OD36055 | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | 6OD36055 | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | 6OD36055 | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 11.375 | in | FDA,4 | 11.0625 | 11.4375 | TM 141 / Visual
6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | 6OD36055 | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 10.5 | in | FDA,4 | 10.4375 | 10.8125 | TM 141 / Visual
6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | 6OD36055 | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Does not meet specifications (OOS) |  |  |  |  | TM 141 / Visual
6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | 6OD36055 | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC w/ shyed flaps (Done) |  |  |  |  | TM 141 / Visual
6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | 6OD36055 | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.375 | in | FDA,4 | 6.1875 | 6.5625 | TM 141 / Visual
6OH14324/SHIP_32_1Y | 6OH14324 | SHIP_32_1Y | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6OH14324/SHIP_32_1Y | 6OH14324 | SHIP_32_1Y | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6OH14324/SHIP_32_1Y | 6OH14324 | SHIP_32_1Y | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OH14324/SHIP_32_1Y | 6OH14324 | SHIP_32_1Y | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6.25 | in | FDA,4 | 6.0625 | 6.4375 | TM 141 / Visual
6OH14324/SHIP_32_1Y | 6OH14324 | SHIP_32_1Y | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 12.4375 | in | FDA,4 | 12.25 | 12.625 | TM 141 / Visual
6OH14324/SHIP_32_1Y | 6OH14324 | SHIP_32_1Y | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6OH14324/SHIP_32_1Y | 6OH14324 | SHIP_32_1Y | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6OH14324/SHIP_32_1Y | 6OH14324 | SHIP_32_1Y | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 9 | in | FDA,4 | 8.8125 | 9.1875 | TM 141 / Visual
6OH21112/SHIP_32_1F | 6OH21112 | SHIP_32_1F | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6OH21112/SHIP_32_1F | 6OH21112 | SHIP_32_1F | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6OH21112/SHIP_32_1F | 6OH21112 | SHIP_32_1F | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OH21112/SHIP_32_1F | 6OH21112 | SHIP_32_1F | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6.75 | in | FDA,4 | 6.5625 | 6.9375 | TM 141 / Visual
6OH21112/SHIP_32_1F | 6OH21112 | SHIP_32_1F | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 9.75 | in | FDA,4 | 9.5625 | 9.9375 | TM 141 / Visual
6OH21112/SHIP_32_1F | 6OH21112 | SHIP_32_1F | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6OH21112/SHIP_32_1F | 6OH21112 | SHIP_32_1F | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6OH21112/SHIP_32_1F | 6OH21112 | SHIP_32_1F | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.5 | in | FDA,4 | 7.4375 | 7.8125 | TM 141 / Visual
6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | 6OH45016 | SHIP_NY_32K_100_8.25-6.875-7.375 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | 6OH45016 | SHIP_NY_32K_100_8.25-6.875-7.375 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | 6OH45016 | SHIP_NY_32K_100_8.25-6.875-7.375 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | 6OH45016 | SHIP_NY_32K_100_8.25-6.875-7.375 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7.25 | in | FDA,4 | 7.1875 | 7.5625 | TM 141 / Visual
6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | 6OH45016 | SHIP_NY_32K_100_8.25-6.875-7.375 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 8.1562 | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual
6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | 6OH45016 | SHIP_NY_32K_100_8.25-6.875-7.375 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | 6OH45016 | SHIP_NY_32K_100_8.25-6.875-7.375 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | 6OH45016 | SHIP_NY_32K_100_8.25-6.875-7.375 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7 | in | FDA,4 | 6.6875 | 7.0625 | TM 141 / Visual
6OK75595/SHIP_32_1R | 6OK75595 | SHIP_32_1R | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6OK75595/SHIP_32_1R | 6OK75595 | SHIP_32_1R | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6OK75595/SHIP_32_1R | 6OK75595 | SHIP_32_1R | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OK75595/SHIP_32_1R | 6OK75595 | SHIP_32_1R | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7.75 | in | FDA,4 | 7.5 | 7.875 | TM 141 / Visual
6OK75595/SHIP_32_1R | 6OK75595 | SHIP_32_1R | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.75 | in | FDA,4 | 11.5625 | 11.9375 | TM 141 / Visual
6OK75595/SHIP_32_1R | 6OK75595 | SHIP_32_1R | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6OK75595/SHIP_32_1R | 6OK75595 | SHIP_32_1R | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6OK75595/SHIP_32_1R | 6OK75595 | SHIP_32_1R | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 9.5 | in | FDA,4 | 9.4375 | 9.8125 | TM 141 / Visual
6OM960CB/SHIP_NY_NA_200_2.5-2.5-33.75 | 6OM960CB | SHIP_NY_NA_200_2.5-2.5-33.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OM960CB/SHIP_NY_NA_200_2.5-2.5-33.75 | 6OM960CB | SHIP_NY_NA_200_2.5-2.5-33.75 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 33.625 | in | FDA,4 | 33.5625 | 33.9375 | TM 141 / Visual
6OM960CB/SHIP_NY_NA_200_2.5-2.5-33.75 | 6OM960CB | SHIP_NY_NA_200_2.5-2.5-33.75 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 2.5 | in | FDA,4 | 2.3125 | 2.6875 | TM 141 / Visual
6OM960CB/SHIP_NY_NA_200_2.5-2.5-33.75 | 6OM960CB | SHIP_NY_NA_200_2.5-2.5-33.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | ITW STACK MASTER 2600 (Done) |  |  |  |  | TM 141 / Visual
6OM960CB/SHIP_NY_NA_200_2.5-2.5-33.75 | 6OM960CB | SHIP_NY_NA_200_2.5-2.5-33.75 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 2.5 | in | FDA,4 | 2.3125 | 2.6875 | TM 141 / Visual
6OP027LP/SHIP_32_13A | 6OP027LP | SHIP_32_13A | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | Layer Pad |  |  |  |  | Shipper Style
6OP027LP/SHIP_32_13A | 6OP027LP | SHIP_32_13A | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6OP027LP/SHIP_32_13A | 6OP027LP | SHIP_32_13A | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6OP027LP/SHIP_32_13A | 6OP027LP | SHIP_32_13A | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OP027LP/SHIP_32_13A | 6OP027LP | SHIP_32_13A | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 10.75 | in | FDA,4 | 10.5625 | 10.9375 | TM 141 / Visual
6OP027LP/SHIP_32_13A | 6OP027LP | SHIP_32_13A | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6OP027LP/SHIP_32_13A | 6OP027LP | SHIP_32_13A | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 5.5625 | in | FDA,4 | 5.375 | 5.75 | TM 141 / Visual
6OP081LP/SHIP_NY_32NA_250_17.250-11.938 | 6OP081LP | SHIP_NY_32NA_250_17.250-11.938 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual
6OP081LP/SHIP_NY_32NA_250_17.250-11.938 | 6OP081LP | SHIP_NY_32NA_250_17.250-11.938 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6OP081LP/SHIP_NY_32NA_250_17.250-11.938 | 6OP081LP | SHIP_NY_32NA_250_17.250-11.938 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6OP081LP/SHIP_NY_32NA_250_17.250-11.938 | 6OP081LP | SHIP_NY_32NA_250_17.250-11.938 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OP081LP/SHIP_NY_32NA_250_17.250-11.938 | 6OP081LP | SHIP_NY_32NA_250_17.250-11.938 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 17.125 | in | FDA,3 | 17.125 | 17.375 | TM 141 / Visual
6OP081LP/SHIP_NY_32NA_250_17.250-11.938 | 6OP081LP | SHIP_NY_32NA_250_17.250-11.938 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | Inserts (Done) |  |  |  |  | TM 141 / Visual
6OP081LP/SHIP_NY_32NA_250_17.250-11.938 | 6OP081LP | SHIP_NY_32NA_250_17.250-11.938 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 11.875 | in | FDA,3 | 11.813 | 12.063 | TM 141 / Visual
6OP34524/SHIP_44_2G | 6OP34524 | SHIP_44_2G | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6OP34524/SHIP_44_2G | 6OP34524 | SHIP_44_2G | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual
6OP34524/SHIP_44_2G | 6OP34524 | SHIP_44_2G | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OP34524/SHIP_44_2G | 6OP34524 | SHIP_44_2G | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 8.25 | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual
6OP34524/SHIP_44_2G | 6OP34524 | SHIP_44_2G | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 13.5 | in | FDA,4 | 13.3125 | 13.6875 | TM 141 / Visual
6OP34524/SHIP_44_2G | 6OP34524 | SHIP_44_2G | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6OP34524/SHIP_44_2G | 6OP34524 | SHIP_44_2G | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6OP34524/SHIP_44_2G | 6OP34524 | SHIP_44_2G | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 11.75 | in | FDA,4 | 11.5625 | 11.9375 | TM 141 / Visual
6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 6OP75510 | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 32 |  | FDA,0 |  |  | TM 141 / Visual
6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 6OP75510 | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 6OP75510 | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual
6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 6OP75510 | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 6OP75510 | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 10.5625 | in | FDA,4 | 10.375 | 10.75 | TM 141 / Visual
6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 6OP75510 | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.375 | in | FDA,4 | 15.1875 | 15.5625 | TM 141 / Visual
6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 6OP75510 | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 6OP75510 | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 6OP75510 | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 11.5625 | in | FDA,4 | 11.375 | 11.75 | TM 141 / Visual
6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | 6OQ22272 | SHIP_NY_44K_100_15.188-10.188-15.188 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | 6OQ22272 | SHIP_NY_44K_100_15.188-10.188-15.188 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual
6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | 6OQ22272 | SHIP_NY_44K_100_15.188-10.188-15.188 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | 6OQ22272 | SHIP_NY_44K_100_15.188-10.188-15.188 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 15.188 | in | FDA,4 | 15.0005 | 15.3755 | TM 141 / Visual
6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | 6OQ22272 | SHIP_NY_44K_100_15.188-10.188-15.188 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.188 | in | FDA,4 | 15.0005 | 15.3755 | TM 141 / Visual
6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | 6OQ22272 | SHIP_NY_44K_100_15.188-10.188-15.188 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | 6OQ22272 | SHIP_NY_44K_100_15.188-10.188-15.188 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | 6OQ22272 | SHIP_NY_44K_100_15.188-10.188-15.188 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 10.188 | in | FDA,4 | 10.0005 | 10.3755 | TM 141 / Visual
6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 6OS12350 | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual
6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 6OS12350 | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 6OS12350 | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | TM 141 / Visual
6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 6OS12350 | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 6OS12350 | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 6OS12350 | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 5.938 | in | FDA,4 | 5.7505 | 6.1255 | TM 141 / Visual
6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 6OS12350 | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.75 | in | FDA,4 | 11.4375 | 11.8125 | TM 141 / Visual
6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 6OS12350 | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 6OS12350 | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 6OS12350 | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.375 | in | FDA,4 | 6.3125 | 6.6875 | TM 141 / Visual
6OS42850/SHIP_44_2N | 6OS42850 | SHIP_44_2N | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6OS42850/SHIP_44_2N | 6OS42850 | SHIP_44_2N | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual
6OS42850/SHIP_44_2N | 6OS42850 | SHIP_44_2N | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OS42850/SHIP_44_2N | 6OS42850 | SHIP_44_2N | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 5.625 | in | FDA,4 | 5.3125 | 5.6875 | TM 141 / Visual
6OS42850/SHIP_44_2N | 6OS42850 | SHIP_44_2N | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.375 | in | FDA,4 | 15.1875 | 15.5625 | TM 141 / Visual
6OS42850/SHIP_44_2N | 6OS42850 | SHIP_44_2N | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6OS42850/SHIP_44_2N | 6OS42850 | SHIP_44_2N | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6OS42850/SHIP_44_2N | 6OS42850 | SHIP_44_2N | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 11.375 | in | FDA,4 | 11.1875 | 11.5625 | TM 141 / Visual
6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 6OS45625 | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual
6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 6OS45625 | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 6OS45625 | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | TM 141 / Visual
6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 6OS45625 | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 6OS45625 | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 6OS45625 | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6 | in | FDA,4 | 5.7505 | 6.1255 | TM 141 / Visual
6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 6OS45625 | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 6.75 | in | FDA,4 | 6.6875 | 7.0625 | TM 141 / Visual
6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 6OS45625 | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 6OS45625 | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 6OS45625 | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.5 | in | FDA,4 | 6.3125 | 6.6875 | TM 141 / Visual
6OS46250/SHIP_29_1J | 6OS46250 | SHIP_29_1J | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6OS46250/SHIP_29_1J | 6OS46250 | SHIP_29_1J | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 29 | Lbs./In. | FDA,0 | 29 |  | TM 141 / Visual
6OS46250/SHIP_29_1J | 6OS46250 | SHIP_29_1J | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OS46250/SHIP_29_1J | 6OS46250 | SHIP_29_1J | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 11.75 | in | FDA,4 | 11.5625 | 11.9375 | TM 141 / Visual
6OS46250/SHIP_29_1J | 6OS46250 | SHIP_29_1J | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 17.125 | in | FDA,4 | 16.9375 | 17.3125 | TM 141 / Visual
6OS46250/SHIP_29_1J | 6OS46250 | SHIP_29_1J | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6OS46250/SHIP_29_1J | 6OS46250 | SHIP_29_1J | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6OS46250/SHIP_29_1J | 6OS46250 | SHIP_29_1J | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.6875 | in | FDA,4 | 6.5 | 6.875 | TM 141 / Visual
6OU12066/SHIP_40_1M | 6OU12066 | SHIP_40_1M | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6OU12066/SHIP_40_1M | 6OU12066 | SHIP_40_1M | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 40 | Lbs./In. | FDA,0 | 40 |  | TM 141 / Visual
6OU12066/SHIP_40_1M | 6OU12066 | SHIP_40_1M | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OU12066/SHIP_40_1M | 6OU12066 | SHIP_40_1M | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 9 | in | FDA,4 | 8.8125 | 9.1875 | TM 141 / Visual
6OU12066/SHIP_40_1M | 6OU12066 | SHIP_40_1M | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 17.75 | in | FDA,4 | 17.625 | 18 | TM 141 / Visual
6OU12066/SHIP_40_1M | 6OU12066 | SHIP_40_1M | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6OU12066/SHIP_40_1M | 6OU12066 | SHIP_40_1M | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6OU12066/SHIP_40_1M | 6OU12066 | SHIP_40_1M | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 15.125 | in | FDA,4 | 14.9375 | 15.3125 | TM 141 / Visual
6OU21470/SHIP_40_1J | 6OU21470 | SHIP_40_1J | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6OU21470/SHIP_40_1J | 6OU21470 | SHIP_40_1J | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 40 | Lbs./In. | FDA,0 | 40 |  | TM 141 / Visual
6OU21470/SHIP_40_1J | 6OU21470 | SHIP_40_1J | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OU21470/SHIP_40_1J | 6OU21470 | SHIP_40_1J | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 9.4375 | in | FDA,4 | 9.0625 | 9.4375 | TM 141 / Visual
6OU21470/SHIP_40_1J | 6OU21470 | SHIP_40_1J | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.5 | in | FDA,4 | 15.1875 | 15.5625 | TM 141 / Visual
6OU21470/SHIP_40_1J | 6OU21470 | SHIP_40_1J | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6OU21470/SHIP_40_1J | 6OU21470 | SHIP_40_1J | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6OU21470/SHIP_40_1J | 6OU21470 | SHIP_40_1J | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 11.25 | in | FDA,4 | 11.1875 | 11.5625 | TM 141 / Visual
6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 6OU31450 | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual
6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 6OU31450 | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 6OU31450 | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual
6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 6OU31450 | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 6OU31450 | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6.125 | in | FDA,4 | 5.9375 | 6.3125 | TM 141 / Visual
6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 6OU31450 | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 18.5 | in | FDA,4 | 18.3125 | 18.6875 | TM 141 / Visual
6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 6OU31450 | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 6OU31450 | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 6OU31450 | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 13.5 | in | FDA,4 | 13.3125 | 13.6875 | TM 141 / Visual
6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 6OU35199 | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual
6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 6OU35199 | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 6OU35199 | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 6OU35199 | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 6OU35199 | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 11.5 | in | FDA,4 | 11.25 | 11.5 | TM 141 / Visual
6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 6OU35199 | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 19.75 | in | FDA,4 | 19.625 | 19.875 | TM 141 / Visual
6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 6OU35199 | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 6OU35199 | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 6OU35199 | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 12 | in | FDA,4 | 11.875 | 12.125 | TM 141 / Visual
6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | 6OU351LP | SHIP_NY_32K_NA_19.500X11.750 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | Pad |  |  |  |  | Shipper Style
6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | 6OU351LP | SHIP_NY_32K_NA_19.500X11.750 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual
6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | 6OU351LP | SHIP_NY_32K_NA_19.500X11.750 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | 6OU351LP | SHIP_NY_32K_NA_19.500X11.750 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | 6OU351LP | SHIP_NY_32K_NA_19.500X11.750 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | 6OU351LP | SHIP_NY_32K_NA_19.500X11.750 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 19.5 | in | FDA,4 | 19.375 | 19.625 | TM 141 / Visual
6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | 6OU351LP | SHIP_NY_32K_NA_19.500X11.750 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | 6OU351LP | SHIP_NY_32K_NA_19.500X11.750 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 11.75 | in | FDA,4 | 11.625 | 11.875 | TM 141 / Visual
6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 6OU351TY | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | D/C Tray |  |  |  |  | Shipper Style
6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 6OU351TY | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual
6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 6OU351TY | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 6OU351TY | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 6OU351TY | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 6OU351TY | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 3.5 | in | FDA,4 | 3.4375 | 3.5625 | TM 141 / Visual
6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 6OU351TY | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 19.062 | in | FDA,4 | 19 | 19.125 | TM 141 / Visual
6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 6OU351TY | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 6OU351TY | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 2.812 | in | FDA,4 | 2.75 | 2.875 | TM 141 / Visual
6OU38510/SHIP_32_1BG | 6OU38510 | SHIP_32_1BG | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6OU38510/SHIP_32_1BG | 6OU38510 | SHIP_32_1BG | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6OU38510/SHIP_32_1BG | 6OU38510 | SHIP_32_1BG | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OU38510/SHIP_32_1BG | 6OU38510 | SHIP_32_1BG | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6 | in | FDA,4 | 5.8125 | 6.1875 | TM 141 / Visual
6OU38510/SHIP_32_1BG | 6OU38510 | SHIP_32_1BG | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 8 | in | FDA,4 | 7.8125 | 8.1875 | TM 141 / Visual
6OU38510/SHIP_32_1BG | 6OU38510 | SHIP_32_1BG | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6OU38510/SHIP_32_1BG | 6OU38510 | SHIP_32_1BG | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6OU38510/SHIP_32_1BG | 6OU38510 | SHIP_32_1BG | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.125 | in | FDA,4 | 5.9375 | 6.3125 | TM 141 / Visual
6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 6OU55510 | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual
6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 6OU55510 | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 6OU55510 | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 6OU55510 | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 6OU55510 | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 8 | in | FDA,4 | 7.8125 | 8.1875 | TM 141 / Visual
6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 6OU55510 | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.75 | in | FDA,4 | 11.5625 | 11.9375 | TM 141 / Visual
6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 6OU55510 | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 6OU55510 | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 6OU55510 | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.375 | in | FDA,4 | 6.0625 | 6.4375 | TM 141 / Visual
6OU62620/SHIP_NY_32K_100_15-9.5-8 | 6OU62620 | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual
6OU62620/SHIP_NY_32K_100_15-9.5-8 | 6OU62620 | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6OU62620/SHIP_NY_32K_100_15-9.5-8 | 6OU62620 | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6OU62620/SHIP_NY_32K_100_15-9.5-8 | 6OU62620 | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OU62620/SHIP_NY_32K_100_15-9.5-8 | 6OU62620 | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 8 | in | FDA,4 | 7.8125 | 8.1875 | TM 141 / Visual
6OU62620/SHIP_NY_32K_100_15-9.5-8 | 6OU62620 | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.125 | in | FDA,4 | 14.8125 | 15.1875 | TM 141 / Visual
6OU62620/SHIP_NY_32K_100_15-9.5-8 | 6OU62620 | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6OU62620/SHIP_NY_32K_100_15-9.5-8 | 6OU62620 | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual
6OU62620/SHIP_NY_32K_100_15-9.5-8 | 6OU62620 | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 9.375 | in | FDA,4 | 9.3125 | 9.6875 | TM 141 / Visual
6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 6OX46500 | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual
6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 6OX46500 | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 6OX46500 | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 6OX46500 | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 6OX46500 | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7.5 | in | FDA,4 | 7.3125 | 7.6875 | TM 141 / Visual
6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 6OX46500 | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 23.75 | in | FDA,4 | 23.5005 | 23.8755 | TM 141 / Visual
6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 6OX46500 | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual
6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 6OX46500 | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | Shipper (Done) |  |  |  |  | TM 141 / Visual
6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 6OX46500 | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 15.6 | in | FDA,4 | 15.5005 | 15.8755 | TM 141 / Visual
6RQ55172/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 6RQ55172 | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Shipper Inspection
6ZS505PD/SHIP_32_14A | 6ZS505PD | SHIP_32_14A | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual
6ZS505PD/SHIP_32_14A | 6ZS505PD | SHIP_32_14A | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual
6ZS505PD/SHIP_32_14A | 6ZS505PD | SHIP_32_14A | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual
6ZS505PD/SHIP_32_14A | 6ZS505PD | SHIP_32_14A | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.4375 | in | FDA,4 | 11.1875 | 11.5625 | TM 141 / Visual
6ZS505PD/SHIP_32_14A | 6ZS505PD | SHIP_32_14A | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | Cut Pad - Sheet (Done) |  |  |  |  | TM 141 / Visual
6ZS505PD/SHIP_32_14A | 6ZS505PD | SHIP_32_14A | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 5.75 | in | FDA,4 | 5.5625 | 5.9375 | TM 141 / Visual
7A4AA580/Chem_NY_Disodium Cocoamphodipropinate | 7A4AA580 | Chem_NY_Disodium Cocoamphodipropinate | IN-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA
7A4AA580/Chem_NY_Disodium Cocoamphodipropinate | 7A4AA580 | Chem_NY_Disodium Cocoamphodipropinate | IN-Analytical | 447 | pH as is | pH | pH | 9.92 |  | FDA,1 | 9 | 10 | TM 200
7A4AA580/Chem_NY_Disodium Cocoamphodipropinate | 7A4AA580 | Chem_NY_Disodium Cocoamphodipropinate | IN-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7A4AA580/Chem_NY_Disodium Cocoamphodipropinate | 7A4AA580 | Chem_NY_Disodium Cocoamphodipropinate | IN-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear to Hazy liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4AA580/Chem_NY_Disodium Cocoamphodipropinate | 7A4AA580 | Chem_NY_Disodium Cocoamphodipropinate | IN-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to Yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4AA580/Chem_NY_Disodium Cocoamphodipropinate | 7A4AA580 | Chem_NY_Disodium Cocoamphodipropinate | IN-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4AA580/Chem_NY_Disodium Cocoamphodipropinate | 7A4AA580 | Chem_NY_Disodium Cocoamphodipropinate | IN-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4AL930/Chem_08A | 7A4AL930 | Chem_08A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7A4AL930/Chem_08A | 7A4AL930 | Chem_08A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Pellets (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4AL930/Chem_08A | 7A4AL930 | Chem_08A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White/Grey (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4AL930/Chem_08A | 7A4AL930 | Chem_08A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4AL930/Chem_08A | 7A4AL930 | Chem_08A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4AL930/Chem_08A | 7A4AL930 | Chem_08A | NY-Analytical | 390 | Identification | Identification for Potassium (USP <191>) | Identification for Potassium | Pass (Done) |  |  |  |  | TM400/408b/421/422/USP
7A4AS162/Chem_09A | 7A4AS162 | Chem_09A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C
7A4AS162/Chem_09A | 7A4AS162 | Chem_09A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.2846 |  | FDA,3 | 1.249 |  | TM 202
7A4AS162/Chem_09A | 7A4AS162 | Chem_09A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7A4AS162/Chem_09A | 7A4AS162 | Chem_09A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, syrupy liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4AS162/Chem_09A | 7A4AS162 | Chem_09A | NY-Analytical | 792 | Chemical Inspection | Color (USP) | Color | Meet USP requirements (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4AS162/Chem_09A | 7A4AS162 | Chem_09A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4AS162/Chem_09A | 7A4AS162 | Chem_09A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4AS162/Chem_09A | 7A4AS162 | Chem_09A | NY-Analytical | 816 | Related Compounds | Individual Impurity | Individual Impurity | 0.01 | % | FDA,3 |  | 0.1 | USP <621>
7A4AS162/Chem_09A | 7A4AS162 | Chem_09A | NY-Analytical | 816 | Related Compounds | Total Impurities | Total Impurities | 0.02 | % | FDA,3 |  | 1 | USP <621>
7A4BL625/Chem_11A | 7A4BL625 | Chem_11A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification
7A4BL625/Chem_11A | 7A4BL625 | Chem_11A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7A4BL625/Chem_11A | 7A4BL625 | Chem_11A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4BL625/Chem_11A | 7A4BL625 | Chem_11A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Light Yellow to Yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4BL625/Chem_11A | 7A4BL625 | Chem_11A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4BL625/Chem_11A | 7A4BL625 | Chem_11A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Lard-Like (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4CD419/Chem_13A | 7A4CD419 | Chem_13A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification
7A4CD419/Chem_13A | 7A4CD419 | Chem_13A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7A4CD419/Chem_13A | 7A4CD419 | Chem_13A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Crystalline Powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4CD419/Chem_13A | 7A4CD419 | Chem_13A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White to off White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4CD419/Chem_13A | 7A4CD419 | Chem_13A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4CD419/Chem_13A | 7A4CD419 | Chem_13A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Faint phenolic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4FD110/Chem_Sodium Chloride USP/NF | 7A4FD110 | Chem_Sodium Chloride USP/NF | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7A4FD110/Chem_Sodium Chloride USP/NF | 7A4FD110 | Chem_Sodium Chloride USP/NF | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Crystals/crystalline powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4FD110/Chem_Sodium Chloride USP/NF | 7A4FD110 | Chem_Sodium Chloride USP/NF | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to white (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4FD110/Chem_Sodium Chloride USP/NF | 7A4FD110 | Chem_Sodium Chloride USP/NF | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4FD110/Chem_Sodium Chloride USP/NF | 7A4FD110 | Chem_Sodium Chloride USP/NF | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4FD110/Chem_Sodium Chloride USP/NF | 7A4FD110 | Chem_Sodium Chloride USP/NF | NY-Analytical | 390 | Identification | Identification for Chloride (USP <191>) | Identification for Chloride | Meets USP (Done) |  |  |  |  | TM400/408b/421/422/USP
7A4FD110/Chem_Sodium Chloride USP/NF | 7A4FD110 | Chem_Sodium Chloride USP/NF | NY-Analytical | 390 | Identification | Identification for Sodium (USP <191>) | Identification for Sodium | Meets USP (Done) |  |  |  |  | TM400/408b/421/422/USP
7A4FD110/Chem_Sodium Chloride USP/NF | 7A4FD110 | Chem_Sodium Chloride USP/NF | NY-Analytical | 802 | Loss on Drying | Loss on Drying | Loss on Drying |  | % | FDA,1 |  | 0.5 | USP
7A4FD110/Chem_Sodium Chloride USP/NF | 7A4FD110 | Chem_Sodium Chloride USP/NF | NY-Analytical | 802 | Loss on Drying | Loss on Drying Final | Loss on Drying Final |  | % | FDA,1 |  | 0.5 | USP
7A4FD110/Chem_Sodium Chloride USP/NF | 7A4FD110 | Chem_Sodium Chloride USP/NF | NY-Analytical | 796 | Assay | Assay on Dried Basis | Assay on Dried Basis | 99.7 | % | FDA,1 | 99 | 100.5 | USP <541>
7A4FD110/Chem_Sodium Chloride USP/NF | 7A4FD110 | Chem_Sodium Chloride USP/NF | NY-Analytical | 796 | Assay | Assay on Dried Basis Final | Assay on Dried Basis Final | 99.7 | % | FDA,1 | 99 | 100.5 | USP <541>
7A4OA347/Chem_NY_Sodium-L(+) Lactate 60% | 7A4OA347 | Chem_NY_Sodium-L(+) Lactate 60% | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH
7A4OA347/Chem_NY_Sodium-L(+) Lactate 60% | 7A4OA347 | Chem_NY_Sodium-L(+) Lactate 60% | NY-Analytical | 398 | pH | pH (1 in 20 solution) | pH (1 in 20 solution) | 6.86 |  | FDA,1 | 6.5 | 7.5 | TM 200
7A4OA347/Chem_NY_Sodium-L(+) Lactate 60% | 7A4OA347 | Chem_NY_Sodium-L(+) Lactate 60% | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Comparison of scans of sample vs standard (Done) |  |  |  |  | TM 411
7A4OA347/Chem_NY_Sodium-L(+) Lactate 60% | 7A4OA347 | Chem_NY_Sodium-L(+) Lactate 60% | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Transparent liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OA347/Chem_NY_Sodium-L(+) Lactate 60% | 7A4OA347 | Chem_NY_Sodium-L(+) Lactate 60% | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OA347/Chem_NY_Sodium-L(+) Lactate 60% | 7A4OA347 | Chem_NY_Sodium-L(+) Lactate 60% | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OA347/Chem_NY_Sodium-L(+) Lactate 60% | 7A4OA347 | Chem_NY_Sodium-L(+) Lactate 60% | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless or with a slight Characteristic Odor (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OA709/Chem_NY_Chamomile Extract # 11-19-2 | 7A4OA709 | Chem_NY_Chamomile Extract # 11-19-2 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH
7A4OA709/Chem_NY_Chamomile Extract # 11-19-2 | 7A4OA709 | Chem_NY_Chamomile Extract # 11-19-2 | NY-Analytical | 398 | pH | pH | pH | 4.92 |  | FDA,1 | 4 | 6.5 | TM 200
7A4OA709/Chem_NY_Chamomile Extract # 11-19-2 | 7A4OA709 | Chem_NY_Chamomile Extract # 11-19-2 | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.0482 |  | FDA,3 | 1.043 | 1.053 | TM 202
7A4OA709/Chem_NY_Chamomile Extract # 11-19-2 | 7A4OA709 | Chem_NY_Chamomile Extract # 11-19-2 | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.412 |  | FDA,3 | 1.409 | 1.419 | TM 204
7A4OA709/Chem_NY_Chamomile Extract # 11-19-2 | 7A4OA709 | Chem_NY_Chamomile Extract # 11-19-2 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None- free of foreign particles (Done) |  |  |  |  | TM 208
7A4OA709/Chem_NY_Chamomile Extract # 11-19-2 | 7A4OA709 | Chem_NY_Chamomile Extract # 11-19-2 | NY-Analytical | 792 | Chemical Inspection | Description (TM 208) | Description | Clear Amber to brown liquid having a characteristic odor (Done) |  |  |  |  | TM 208
7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | 7A4OA911 | Chem_NY_Fragrance Hugs & Kisses Sheer1 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Refractive Index at 25°C
7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | 7A4OA911 | Chem_NY_Fragrance Hugs & Kisses Sheer1 | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.4537 |  | FDA,4 | 1.4536 | 1.4636 | TM 204
7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | 7A4OA911 | Chem_NY_Fragrance Hugs & Kisses Sheer1 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Liquid (Done) |  |  |  |  | TM 208
7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | 7A4OA911 | Chem_NY_Fragrance Hugs & Kisses Sheer1 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to Slightly Yellow (Done) |  |  |  |  | TM 208
7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | 7A4OA911 | Chem_NY_Fragrance Hugs & Kisses Sheer1 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | 7A4OA911 | Chem_NY_Fragrance Hugs & Kisses Sheer1 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208
7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | 7A4OA911 | Chem_NY_Fragrance Hugs & Kisses Sheer1 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411
7A4OB223/Chem_189A | 7A4OB223 | Chem_189A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification
7A4OB223/Chem_189A | 7A4OB223 | Chem_189A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7A4OB223/Chem_189A | 7A4OB223 | Chem_189A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Translucent crystals or crystalline powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OB223/Chem_189A | 7A4OB223 | Chem_189A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to white (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OB223/Chem_189A | 7A4OB223 | Chem_189A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OB223/Chem_189A | 7A4OB223 | Chem_189A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis
7A4OB314/Chem_125A | 7A4OB314 | Chem_125A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Certificate of Analysis
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH
7A4OB314/Chem_125A | 7A4OB314 | Chem_125A | NY-Analytical | 398 | pH | pH (10% Solution) | pH (10% Solution) |  |  | FDA,1 | 6 | 8 | TM 200
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.7837 |  | FDA,3 | 0.783 | 0.787 | TM 202
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.7839 |  | FDA,3 | 0.783 | 0.787 | TM 202
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C |  |  | FDA,3 | 1.376 | 1.378 | TM 204
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, mobile liquid (Done) |  |  |  |  | TM 208
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM 208
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | IPA - Characteristic (Done) |  |  |  |  | TM 208
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 797 | Nonvolatile Residue | End Weight (g) | End Weight |  | g | FDA,4 |  |  | TM 226
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 797 | Nonvolatile Residue | Nonvolatile Residue (calc) | Nonvolatile Residue |  | % | FDA,3 |  | 0.005 | TM 226
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 797 | Nonvolatile Residue | Nonvolatile Residue Final (calc) | Nonvolatile Residue Final |  | % | FDA,3 |  | 0.005 | TM 226
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 797 | Nonvolatile Residue | Start Weight (g) | Start Weight |  | g | FDA,4 |  |  | TM 226
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 797 | Nonvolatile Residue | Weight of Residue (g) | Weight of Residue |  | g | FDA,4 |  |  | TM 226
7A4OB314/Chem_125A | 7A4OB314 | Chem_125A | NY-Analytical | 796 | Assay | Assay | Assay |  | % | FDA,1 | 76 | 84 | TM 301/434
7A4OB314/Chem_125A | 7A4OB314 | Chem_125A | NY-Analytical | 796 | Assay | Assay Final | Assay Final |  | % | FDA,1 | 76 | 84 | TM 301/434
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 795 | Acidity | Acidity (calc) | Acidity |  | mL | FDA,2 |  | 0.7 | TM 375
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 795 | Acidity | Acidity Final (Calc) | Acidity Final (Calc) |  | mL | FDA,2 |  | 0.7 | TM 375
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 795 | Acidity | Titrant End Volume (mL) | Titrant End Volume |  | mL | FDA,2 |  |  | TM 375
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 795 | Acidity | Titrant Start Volume (mL) | Titrant Start Volume |  | mL | FDA,2 |  |  | TM 375
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 796 | Assay | Assay | Assay |  | % | FDA,1 | 99 |  | TM 376
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 796 | Assay | Assay | Assay | 99.96 | % | FDA,1 | 99 |  | TM 376
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 796 | Assay | Assay Final | Assay Final |  | % | FDA,1 | 99 |  | TM 376
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 796 | Assay | Assay Final | Assay Final |  | % | FDA,1 | 99 |  | TM 376
7A4OB314/Chem_125A | 7A4OB314 | Chem_125A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | [5 empty] | TM 411
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Match standard, per USP (Done) |  |  |  |  | TM 411
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 390 | Identification | IR Identification | IR Identification | [5 empty] | TM 411
7A4OB314/Chem_125A | 7A4OB314 | Chem_125A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear viscous liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | [5 empty] | TM208/220/230/USP/Kao LM
7A4OB314/Chem_125A | 7A4OB314 | Chem_125A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | [5 empty] | TM208/220/230/USP/Kao LM
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | [5 empty] | TM208/220/230/USP/Kao LM
7A4OB314/Chem_125A | 7A4OB314 | Chem_125A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | [5 empty] | TM208/220/230/USP/Kao LM
7A4OB314/Chem_125A | 7A4OB314 | Chem_125A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | [5 empty] | TM208/220/230/USP/Kao LM
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | 7A4OB314 | Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | [5 empty] | TM208/220/230/USP/Kao LM
7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB339 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.376 |  | FDA,3 | 1.376 | 1.378 | TM 204
7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB339 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 797 | Nonvolatile Residue | End Weight (g) | End Weight | 92.0885 | g | FDA,4 |  |  | TM 226
7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB339 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 797 | Nonvolatile Residue | Nonvolatile Residue (calc) | Nonvolatile Residue | 0 | % | FDA,3 |  | 0.005 | TM 226
7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB339 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 797 | Nonvolatile Residue | Nonvolatile Residue Final (calc) | Nonvolatile Residue Final | 0 | % | FDA,3 |  | 0.005 | TM 226
7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB339 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 797 | Nonvolatile Residue | Start Weight (g) | Start Weight | 92.0885 | g | FDA,4 |  |  | TM 226
7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB339 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 797 | Nonvolatile Residue | Weight of Residue (g) | Weight of Residue | 0 | g | FDA,4 |  |  | TM 226
7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB339 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 795 | Acidity | Acidity (calc) | Acidity | 0.4 | mL | FDA,2 |  | 0.7 | TM 375
7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB339 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 795 | Acidity | Acidity Final (Calc) | Acidity Final (Calc) | 0.425 | mL | FDA,2 |  | 0.7 | TM 375
7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB339 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 795 | Acidity | Titrant End Volume (mL) | Titrant End Volume | 0.5 | mL | FDA,2 |  |  | TM 375
7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB339 | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 795 | Acidity | Titrant Start Volume (mL) | Titrant Start Volume | 0 | mL | FDA,2 |  |  | TM 375
7A4OB397/Chem_04A | 7A4OB397 | Chem_04A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Assay
7A4OB397/Chem_04A | 7A4OB397 | Chem_04A | NY-Analytical | 796 | Assay | Assay | Assay | 55.85 | % | FDA,1 | 47.5 | 52.5 | TM 301/434
7A4OB397/Chem_04A | 7A4OB397 | Chem_04A | NY-Analytical | 796 | Assay | Assay Final | Assay Final | 55.77 | % | FDA,1 | 47.5 | 52.5 | TM 301/434
7A4OB397/Chem_04A | 7A4OB397 | Chem_04A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear viscous liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OB397/Chem_04A | 7A4OB397 | Chem_04A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to light yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OB397/Chem_04A | 7A4OB397 | Chem_04A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OB397/Chem_04A | 7A4OB397 | Chem_04A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Mild almond (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OB397/Chem_04A | 7A4OB397 | Chem_04A | NY-Analytical | 390 | Identification | Identification (TM 400) | Identification | Formation of precipitate per test (Done) |  |  |  |  | TM400/408b/421/422/USP
7A4OB690/Chem_01A | 7A4OB690 | Chem_01A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification
7A4OB690/Chem_01A | 7A4OB690 | Chem_01A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7A4OB690/Chem_01A | 7A4OB690 | Chem_01A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, mobile liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OB690/Chem_01A | 7A4OB690 | Chem_01A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OB690/Chem_01A | 7A4OB690 | Chem_01A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OB690/Chem_01A | 7A4OB690 | Chem_01A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Solvent (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OF602/Chem_19A | 7A4OF602 | Chem_19A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.968 |  | FDA,3 | 0.964 | 0.968 | TM 202
7A4OF602/Chem_19A | 7A4OF602 | Chem_19A | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.458 |  | FDA,3 | 1.455 | 1.461 | TM 204
7A4OF602/Chem_19A | 7A4OF602 | Chem_19A | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411
7A4OF602/Chem_19A | 7A4OF602 | Chem_19A | NY-Analytical | 792 | Chemical Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OF602/Chem_19A | 7A4OF602 | Chem_19A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OG169/Chem_20A | 7A4OG169 | Chem_20A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification
7A4OG169/Chem_20A | 7A4OG169 | Chem_20A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7A4OG169/Chem_20A | 7A4OG169 | Chem_20A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OG169/Chem_20A | 7A4OG169 | Chem_20A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Amber (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OG169/Chem_20A | 7A4OG169 | Chem_20A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OH247/Chem_21A | 7A4OH247 | Chem_21A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7A4OH247/Chem_21A | 7A4OH247 | Chem_21A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | White granular material (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OH247/Chem_21A | 7A4OH247 | Chem_21A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Match reference (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OH247/Chem_21A | 7A4OH247 | Chem_21A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OH247/Chem_21A | 7A4OH247 | Chem_21A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match reference (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | Chem_NY_Salicylic Acid USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | Chem_NY_Salicylic Acid USP | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | White crystalline powder (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | Chem_NY_Salicylic Acid USP | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | Chem_NY_Salicylic Acid USP | NY-Analytical | 832 | Chloride | Chloride | Chloride | <0.014 | % | FDA,3 |  | 0.014 | USP
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | Chem_NY_Salicylic Acid USP | NY-Analytical | 833 | Heavy Metals | Heavy Metals | Heavy Metals | <20 | ug/g | FDA,0 |  | <20 | USP
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | Chem_NY_Salicylic Acid USP | NY-Analytical | 816 | Related Compounds | Impurity | Impurity | 0 | % | FDA,4 |  | 0.05 | USP
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | Chem_NY_Salicylic Acid USP | NY-Analytical | 802 | Loss on Drying | Loss on Drying | Loss on Drying | 0.03 | % | FDA,3 |  | 0.5 | USP
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | Chem_NY_Salicylic Acid USP | NY-Analytical | 816 | Related Compounds | Phenol | Phenol | 0 | % | FDA,4 |  | 0.01 | USP
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | Chem_NY_Salicylic Acid USP | NY-Analytical | 816 | Related Compounds | Related Compound A | Related Compound A | 0.043 | % | FDA,4 |  | 0.1 | USP
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | Chem_NY_Salicylic Acid USP | NY-Analytical | 816 | Related Compounds | Related Compound B | Related Compound B | 0.037 | % | FDA,4 |  | 0.05 | USP
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | Chem_NY_Salicylic Acid USP | NY-Analytical | 834 | Residue on Ignition | Residue on Ignition | Residue on Ignition | 0.02 | % | FDA,2 |  | 0.05 | USP
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | Chem_NY_Salicylic Acid USP | NY-Analytical | 835 | Sulfate Content | Sulfate Content | Sulfate Content | <0.020 | % | FDA,3 |  | 0.02 | USP
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | Chem_NY_Salicylic Acid USP | NY-Analytical | 816 | Related Compounds | Total Impurities | Total Impurities | 0.081 | % | FDA,4 |  | 0.2 | USP
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | Chem_NY_Salicylic Acid USP | NY-Analytical | 796 | Assay | Assay | Assay | 99.7 | % | FDA,1 | 98 | 102 | USP/NF
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | Chem_NY_Salicylic Acid USP | NY-Analytical | 455 | Identification A | Identification A | Identification A | Passes Test (Done) |  |  |  |  | USP/NF
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | Chem_NY_Salicylic Acid USP | NY-Analytical | 456 | Identification B | Identification B | Identification B | Passes Test (Done) |  |  |  |  | USP/NF
7A4OM002/Chem_NY_ Dissolvine GL-47-S | 7A4OM002 | Chem_NY_ Dissolvine GL-47-S | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH
7A4OM002/Chem_NY_ Dissolvine GL-47-S | 7A4OM002 | Chem_NY_ Dissolvine GL-47-S | NY-Analytical | 398 | pH | pH (1% w/v aqueous solution) | pH (1% w/v aqueous solution) | 11.14 |  | FDA,1 | 11 | 11.8 | TM 200
7A4OM002/Chem_NY_ Dissolvine GL-47-S | 7A4OM002 | Chem_NY_ Dissolvine GL-47-S | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM 208
7A4OM002/Chem_NY_ Dissolvine GL-47-S | 7A4OM002 | Chem_NY_ Dissolvine GL-47-S | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Characteristic (Done) |  |  |  |  | TM 208
7A4OM002/Chem_NY_ Dissolvine GL-47-S | 7A4OM002 | Chem_NY_ Dissolvine GL-47-S | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7A4OM002/Chem_NY_ Dissolvine GL-47-S | 7A4OM002 | Chem_NY_ Dissolvine GL-47-S | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) |  |  |  |  | TM 208
7A4OM002/Chem_NY_ Dissolvine GL-47-S | 7A4OM002 | Chem_NY_ Dissolvine GL-47-S | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411
7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP027 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP027 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | A clear colorless to pale straw colored liquid. (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP027 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 792 | Chemical Inspection | Odor (J&J TM-003188) | Odor | Passes Test (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP027 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP027 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013
7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP027 | MASTER - Microbiology - J&J Chemical Testing) | NY-Microbiology | 524 | J&J Fungal | Fungal | Fungal | 0 | CFU/g | FDA,0 | 0 | 200 | TM 869
7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP027 | MASTER - Microbiology - J&J Chemical Testing) | NY-Microbiology | 700 | J&J SDA Controls | Negative Control | Negative Control | [5 empty] | TM 869
7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP027 | MASTER - Microbiology - J&J Chemical Testing) | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | NegativeControl | No Growth TSB |  |  |  |  | TM 869
7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP027 | MASTER - Microbiology - J&J Chemical Testing) | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | TSA JJ | 0 | cfu/mL | FDA,0 | 0 | 500 | TM 869
7A4OP065/Chem_NY_ Fragrance Plaisir Gourmand 23 | 7A4OP065 | Chem_NY_ Fragrance Plaisir Gourmand 23 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7A4OP065/Chem_NY_ Fragrance Plaisir Gourmand 23 | 7A4OP065 | Chem_NY_ Fragrance Plaisir Gourmand 23 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM 208
7A4OP065/Chem_NY_ Fragrance Plaisir Gourmand 23 | 7A4OP065 | Chem_NY_ Fragrance Plaisir Gourmand 23 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Pale Yellow to Yellow (Done) |  |  |  |  | TM 208
7A4OP065/Chem_NY_ Fragrance Plaisir Gourmand 23 | 7A4OP065 | Chem_NY_ Fragrance Plaisir Gourmand 23 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7A4OP065/Chem_NY_ Fragrance Plaisir Gourmand 23 | 7A4OP065 | Chem_NY_ Fragrance Plaisir Gourmand 23 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208
7A4OP065/Chem_NY_ Fragrance Plaisir Gourmand 23 | 7A4OP065 | Chem_NY_ Fragrance Plaisir Gourmand 23 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411
7A4OP073/Chem_NY_Fragrance Bali Leia 5 | 7A4OP073 | Chem_NY_Fragrance Bali Leia 5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7A4OP073/Chem_NY_Fragrance Bali Leia 5 | 7A4OP073 | Chem_NY_Fragrance Bali Leia 5 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Liquid (Done) |  |  |  |  | TM 208
7A4OP073/Chem_NY_Fragrance Bali Leia 5 | 7A4OP073 | Chem_NY_Fragrance Bali Leia 5 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Dark yellow to lemon yellow (Done) |  |  |  |  | TM 208
7A4OP073/Chem_NY_Fragrance Bali Leia 5 | 7A4OP073 | Chem_NY_Fragrance Bali Leia 5 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7A4OP073/Chem_NY_Fragrance Bali Leia 5 | 7A4OP073 | Chem_NY_Fragrance Bali Leia 5 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208
7A4OP073/Chem_NY_Fragrance Bali Leia 5 | 7A4OP073 | Chem_NY_Fragrance Bali Leia 5 | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411
7A4OP081/Chem_NY_Fragrance SC 2705 LF | 7A4OP081 | Chem_NY_Fragrance SC 2705 LF | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7A4OP081/Chem_NY_Fragrance SC 2705 LF | 7A4OP081 | Chem_NY_Fragrance SC 2705 LF | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Liquid (Done) |  |  |  |  | TM 208
7A4OP081/Chem_NY_Fragrance SC 2705 LF | 7A4OP081 | Chem_NY_Fragrance SC 2705 LF | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to very light yellow (Done) |  |  |  |  | TM 208
7A4OP081/Chem_NY_Fragrance SC 2705 LF | 7A4OP081 | Chem_NY_Fragrance SC 2705 LF | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7A4OP081/Chem_NY_Fragrance SC 2705 LF | 7A4OP081 | Chem_NY_Fragrance SC 2705 LF | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208
7A4OP081/Chem_NY_Fragrance SC 2705 LF | 7A4OP081 | Chem_NY_Fragrance SC 2705 LF | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411
7A4OP093/Chem_NY_Essential Oil 27A Fragrance | 7A4OP093 | Chem_NY_Essential Oil 27A Fragrance | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7A4OP093/Chem_NY_Essential Oil 27A Fragrance | 7A4OP093 | Chem_NY_Essential Oil 27A Fragrance | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Liquid (Done) |  |  |  |  | TM 208
7A4OP093/Chem_NY_Essential Oil 27A Fragrance | 7A4OP093 | Chem_NY_Essential Oil 27A Fragrance | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Amber (Done) |  |  |  |  | TM 208
7A4OP093/Chem_NY_Essential Oil 27A Fragrance | 7A4OP093 | Chem_NY_Essential Oil 27A Fragrance | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7A4OP093/Chem_NY_Essential Oil 27A Fragrance | 7A4OP093 | Chem_NY_Essential Oil 27A Fragrance | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208
7A4OP093/Chem_NY_Essential Oil 27A Fragrance | 7A4OP093 | Chem_NY_Essential Oil 27A Fragrance | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411
7A4OP113/Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | 7A4OP113 | Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH
7A4OP113/Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | 7A4OP113 | Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | NY-Analytical | 398 | pH | pH | pH | 4.11 |  | FDA,1 | 2.5 | 6 | TM 200
7A4OP113/Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | 7A4OP113 | Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear to slightly cloudy liquid (Done) |  |  |  |  | TM 208
7A4OP113/Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | 7A4OP113 | Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM 208
7A4OP113/Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | 7A4OP113 | Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7A4OP113/Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | 7A4OP113 | Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) |  |  |  |  | TM 208
7A4OP197/Chem_NY_D-Panthenol 75W_01 | 7A4OP197 | Chem_NY_D-Panthenol 75W_01 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH
7A4OP197/Chem_NY_D-Panthenol 75W_01 | 7A4OP197 | Chem_NY_D-Panthenol 75W_01 | NY-Analytical | 398 | pH | pH | pH | 6.96 |  | FDA,1 | 6 | 7.5 | TM 200
7A4OP197/Chem_NY_D-Panthenol 75W_01 | 7A4OP197 | Chem_NY_D-Panthenol 75W_01 | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.466 |  | FDA,3 | 1.466 | 1.475 | TM 204
7A4OP197/Chem_NY_D-Panthenol 75W_01 | 7A4OP197 | Chem_NY_D-Panthenol 75W_01 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM 208
7A4OP197/Chem_NY_D-Panthenol 75W_01 | 7A4OP197 | Chem_NY_D-Panthenol 75W_01 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Slightly Yellowish or Colorless (Done) |  |  |  |  | TM 208
7A4OP197/Chem_NY_D-Panthenol 75W_01 | 7A4OP197 | Chem_NY_D-Panthenol 75W_01 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7A4OP197/Chem_NY_D-Panthenol 75W_01 | 7A4OP197 | Chem_NY_D-Panthenol 75W_01 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208
7A4OP197/Chem_NY_D-Panthenol 75W_01 | 7A4OP197 | Chem_NY_D-Panthenol 75W_01 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411
7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | 7A4OP205 | Chem_NY_Cucumis Sativus Fruit Extract | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | 7A4OP205 | Chem_NY_Cucumis Sativus Fruit Extract | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Transparent Liquid to slightly Turbid (Done) |  |  |  |  | TM 208
7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | 7A4OP205 | Chem_NY_Cucumis Sativus Fruit Extract | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to pale yellow (Done) |  |  |  |  | TM 208
7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | 7A4OP205 | Chem_NY_Cucumis Sativus Fruit Extract | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | 7A4OP205 | Chem_NY_Cucumis Sativus Fruit Extract | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208
7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | 7A4OP205 | Chem_NY_Cucumis Sativus Fruit Extract | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411
7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | 7A4OP205 | Chem_NY_Cucumis Sativus Fruit Extract | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram) | TAMC (Average per Gram) | 0 | CFU/g | FDA,1 |  | 100 | TM 610
7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | 7A4OP205 | Chem_NY_Cucumis Sativus Fruit Extract | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | 7A4OP205 | Chem_NY_Cucumis Sativus Fruit Extract | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram) | TMYC (Average per Gram) | 0 | CFU/g | FDA,1 |  | 100 | TM 610
7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | 7A4OP205 | Chem_NY_Cucumis Sativus Fruit Extract | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
7A4OP345/Chem_195A | 7A4OP345 | Chem_195A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH as is
7A4OP345/Chem_195A | 7A4OP345 | Chem_195A | NY-Analytical | 447 | pH as is | pH | pH | 3.92 |  | FDA,1 | 3.5 | 4.5 | TM 200
7A4OP345/Chem_195A | 7A4OP345 | Chem_195A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7A4OP345/Chem_195A | 7A4OP345 | Chem_195A | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram) | TAMC (Average per Gram) | 0 | CFU/g | FDA,1 |  |  | TM 610
7A4OP345/Chem_195A | 7A4OP345 | Chem_195A | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
7A4OP345/Chem_195A | 7A4OP345 | Chem_195A | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram) | TMYC (Average per Gram) | 0 | CFU/g | FDA,1 |  |  | TM 610
7A4OP345/Chem_195A | 7A4OP345 | Chem_195A | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
7A4OP345/Chem_195A | 7A4OP345 | Chem_195A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Hazy liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OP345/Chem_195A | 7A4OP345 | Chem_195A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Light green/amber (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OP345/Chem_195A | 7A4OP345 | Chem_195A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OP345/Chem_195A | 7A4OP345 | Chem_195A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OP448/Chem_208A | 7A4OP448 | Chem_208A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis
7A4OP448/Chem_197A | 7A4OP448 | Chem_197A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH
7A4OP448/Chem_197A | 7A4OP448 | Chem_197A | NY-Analytical | 398 | pH | pH (10% in water) | pH (10% in water) | 5.9 |  | FDA,1 | 9 | 11 | TM 200
7A4OP448/Chem_208A | 7A4OP448 | Chem_208A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.9801 |  | FDA,4 | 0.9694 | 0.9894 | TM 202
7A4OP448/Chem_208A | 7A4OP448 | Chem_208A | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.469 |  | FDA,4 | 1.4655 | 1.4855 | TM 204
7A4OP448/Chem_197A | 7A4OP448 | Chem_197A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7A4OP448/Chem_208A | 7A4OP448 | Chem_208A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7A4OP448/Chem_197A | 7A4OP448 | Chem_197A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OP448/Chem_208A | 7A4OP448 | Chem_208A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear to pale yellow liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OP448/Chem_197A | 7A4OP448 | Chem_197A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OP448/Chem_197A | 7A4OP448 | Chem_197A | NY-Analytical | 792 | Chemical Inspection | Color (TM 230) | Color |  |  | FDA,1 | 0 | 3 | TM208/220/230/USP/Kao LM
7A4OP448/Chem_197A | 7A4OP448 | Chem_197A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OP448/Chem_208A | 7A4OP448 | Chem_208A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OP448/Chem_197A | 7A4OP448 | Chem_197A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Fatty odor (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OP448/Chem_208A | 7A4OP448 | Chem_208A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Fresh Bouquet (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OP567/Chem_NY_Surfonic L12-6_Surfactant | 7A4OP567 | Chem_NY_Surfonic L12-6_Surfactant | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA
7A4OP567/Chem_NY_Surfonic L12-6_Surfactant | 7A4OP567 | Chem_NY_Surfonic L12-6_Surfactant | NY-Analytical | 398 | pH | pH (1%) in 10:6 IPA: H2O | pH (1%) in 10:6 IPA: H2O | 6.9 |  | FDA,1 | 6.5 | 8 | TM 200
7A4OP567/Chem_NY_Surfonic L12-6_Surfactant | 7A4OP567 | Chem_NY_Surfonic L12-6_Surfactant | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear to slightly turbid, free of suspended matter (Done) |  |  |  |  | TM 208
7A4OP567/Chem_NY_Surfonic L12-6_Surfactant | 7A4OP567 | Chem_NY_Surfonic L12-6_Surfactant | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to Slightly Yellow (Done) |  |  |  |  | TM 208
7A4OP567/Chem_NY_Surfonic L12-6_Surfactant | 7A4OP567 | Chem_NY_Surfonic L12-6_Surfactant | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7A4OP567/Chem_NY_Surfonic L12-6_Surfactant | 7A4OP567 | Chem_NY_Surfonic L12-6_Surfactant | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | To match standard (Done) |  |  |  |  | TM 208
7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | 7A4OP601 | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Iodine Value
7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | 7A4OP601 | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 828 | Iodine Value | Iodine Value | Iodine Value | 2.2 | g I/100g | FDA,2 |  | 5 | J&J TM-003029
7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | 7A4OP601 | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 828 | Iodine Value | Iodine Value Final | Iodine Value Final | 2.2 | g I/100g | FDA,2 |  | 5 | J&J TM-003029
7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | 7A4OP601 | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Light yellow pasty fluid (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | 7A4OP601 | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | 7A4OP601 | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013
7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | 7A4OP601 | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 398 | pH | pH (5% Solution in DI Water:Methanol 1:1) | pH (5% Solution in DI Water:Methanol 1:1) | 7 |  | FDA,1 | 6.5 | 7.5 | J&J TMWW78014
7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | 7A4OP601 | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 810 | Saponification Value | Saponification Value | Saponification Value | 54.2 | KOH/g | FDA,1 | 47 | 57 | J&J TMWW78016
7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | 7A4OP601 | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 810 | Saponification Value | Saponification Value Final | Saponification Value Final | 53.35 | KOH/g | FDA,1 | 47 | 57 | J&J TMWW78016
7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | 7A4OP601 | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 801 | % Water | Water Content | Water Content | 35.15 | % | FDA,2 | 33 | 37 | J&J TMWW78017
7A4OP806/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP806 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7A4OP806/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP806 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C | 1.1155 |  | FDA,4 | 1.11 | 1.12 | J&J TMWW78012
7A4OP806/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP806 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.391 |  | FDA,4 | 1.386 | 1.396 | J&J TMWW78015
7A4OP806/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP806 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 825 | Limit of Ethylene Glycol and Diethylene Glycol | Diethylene Glycol | Diethylene Glycol | 0 | % | FDA,2 |  | 0.1 | USP
7A4OP806/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP806 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 825 | Limit of Ethylene Glycol and Diethylene Glycol | Ethylene Glycol | Ethylene Glycol | 0 | % | FDA,2 |  | 0.1 | USP
7A4OQ028/Chem_NY_Emulgin VL-75 | 7A4OQ028 | Chem_NY_Emulgin VL-75 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH as is
7A4OQ028/Chem_NY_Emulgin VL-75 | 7A4OQ028 | Chem_NY_Emulgin VL-75 | NY-Analytical | 447 | pH as is | pH | pH | 5.76 |  | FDA,1 | 5 | 6 | TM 200
7A4OQ028/Chem_NY_Emulgin VL-75 | 7A4OQ028 | Chem_NY_Emulgin VL-75 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7A4OQ028/Chem_NY_Emulgin VL-75 | 7A4OQ028 | Chem_NY_Emulgin VL-75 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear to slightly Turbid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OQ028/Chem_NY_Emulgin VL-75 | 7A4OQ028 | Chem_NY_Emulgin VL-75 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Yellowish Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OQ028/Chem_NY_Emulgin VL-75 | 7A4OQ028 | Chem_NY_Emulgin VL-75 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OQ028/Chem_NY_Emulgin VL-75 | 7A4OQ028 | Chem_NY_Emulgin VL-75 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OQ084/Chem_23A | 7A4OQ084 | Chem_23A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification
7A4OQ084/Chem_23A | 7A4OQ084 | Chem_23A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7A4OQ084/Chem_23A | 7A4OQ084 | Chem_23A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OQ084/Chem_23A | 7A4OQ084 | Chem_23A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | None or Pale Straw (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OQ084/Chem_23A | 7A4OQ084 | Chem_23A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OQ084/Chem_23A | 7A4OQ084 | Chem_23A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Little or no distinguisable odor (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OQ132/Chem_25A | 7A4OQ132 | Chem_25A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification
7A4OQ132/Chem_25A | 7A4OQ132 | Chem_25A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7A4OQ132/Chem_25A | 7A4OQ132 | Chem_25A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Transparent liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OQ132/Chem_25A | 7A4OQ132 | Chem_25A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to pale yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OQ132/Chem_25A | 7A4OQ132 | Chem_25A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OQ132/Chem_25A | 7A4OQ132 | Chem_25A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Mild (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OQ203/MASTER - Microbiology - J&J Chemical Testing) | 7A4OQ203 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 802 | Loss on Drying | Loss on Drying | Loss on Drying | 0.75 | % | FDA,2 |  | 2 | J&J TM7221
7A4OQ203/MASTER - Microbiology - J&J Chemical Testing) | 7A4OQ203 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 802 | Loss on Drying | Loss on Drying Final | Loss on Drying Final | 0.75 | % | FDA,2 |  | 2 | J&J TM7221
7A4OQ203/MASTER - Microbiology - J&J Chemical Testing) | 7A4OQ203 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 747 | Viscosity | Viscosity (Spindle 7@20RPM@30 sec 0.5% solution) | Viscosity (Spindle 7@20RPM@30 sec 0.5% solution) | 60667 | cps | FDA,0 | 45000 | 65000 | J&J TM7423
7A4OQ229/Chem_Trisodium Citrate Dihydrate USP | 7A4OQ229 | Chem_Trisodium Citrate Dihydrate USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA
7A4OQ229/Chem_NY_Witch Hazel | 7A4OQ229 | Chem_NY_Witch Hazel | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C
7A4OQ229/Chem_NY_Witch Hazel | 7A4OQ229 | Chem_NY_Witch Hazel | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.981 |  | FDA,3 | 0.979 | 0.983 | TM 202
7A4OQ229/Chem_Trisodium Citrate Dihydrate USP | 7A4OQ229 | Chem_Trisodium Citrate Dihydrate USP | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Fine granular (Done) |  |  |  |  | TM 208
7A4OQ229/Chem_Trisodium Citrate Dihydrate USP | 7A4OQ229 | Chem_Trisodium Citrate Dihydrate USP | NY-Analytical | 792 | Chemical Inspection | Color | Color | White (Done) |  |  |  |  | TM 208
7A4OQ229/Chem_Trisodium Citrate Dihydrate USP | 7A4OQ229 | Chem_Trisodium Citrate Dihydrate USP | NY-Analytical | 792 | Chemical Inspection | Contamination | Contamination | None (Done) |  |  |  |  | TM 208
7A4OQ229/Chem_Trisodium Citrate Dihydrate USP | 7A4OQ229 | Chem_Trisodium Citrate Dihydrate USP | NY-Analytical | 792 | Chemical Inspection | Odor | Odor | Odorless (Done) |  |  |  |  | TM 208
7A4OQ229/Chem_Trisodium Citrate Dihydrate USP | 7A4OQ229 | Chem_Trisodium Citrate Dihydrate USP | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7A4OQ229/Chem_NY_Witch Hazel | 7A4OQ229 | Chem_NY_Witch Hazel | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OQ229/Chem_NY_Witch Hazel | 7A4OQ229 | Chem_NY_Witch Hazel | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OQ229/Chem_NY_Witch Hazel | 7A4OQ229 | Chem_NY_Witch Hazel | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OQ229/Chem_NY_Witch Hazel | 7A4OQ229 | Chem_NY_Witch Hazel | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic witch hazel, free from acetous odor (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OQ229/Chem_NY_Witch Hazel | 7A4OQ229 | Chem_NY_Witch Hazel | NY-Analytical | 797 | Nonvolatile Residue | End Weight (g) | End Weight | 240.3127 | g | FDA,0 |  |  | USP
7A4OQ229/Chem_NY_Witch Hazel | 7A4OQ229 | Chem_NY_Witch Hazel | NY-Analytical | 797 | Nonvolatile Residue | Start Weight (g) | Start Weight | 240.3107 | g | FDA,0 |  |  | USP
7A4OQ229/Chem_NY_Witch Hazel | 7A4OQ229 | Chem_NY_Witch Hazel | NY-Analytical | 797 | Nonvolatile Residue | Total non-volatile reside (calc) | Total non-volatile reside | 2 | mg/100 mL | FDA,0 |  | 25 | USP
7A4OQ229/Chem_NY_Witch Hazel | 7A4OQ229 | Chem_NY_Witch Hazel | NY-Analytical | 447 | pH as is | pH | pH | 3.64 |  | FDA,1 | 3 | 5 | USP <791>
7A4OQ260/Chem_29A | 7A4OQ260 | Chem_29A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C
7A4OQ260/Chem_29A | 7A4OQ260 | Chem_29A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.9659 |  | FDA,3 | 0.954 | 0.974 | TM 202
7A4OQ260/Chem_29A | 7A4OQ260 | Chem_29A | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.48 |  | FDA,3 | 1.464 | 1.484 | TM 204
7A4OQ260/Chem_29A | 7A4OQ260 | Chem_29A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7A4OQ260/Chem_29A | 7A4OQ260 | Chem_29A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OQ260/Chem_29A | 7A4OQ260 | Chem_29A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to pale yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OQ260/Chem_29A | 7A4OQ260 | Chem_29A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OQ260/Chem_29A | 7A4OQ260 | Chem_29A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Powder Rose (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OQ518/Chem_31A | 7A4OQ518 | Chem_31A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH
7A4OQ518/Chem_31A | 7A4OQ518 | Chem_31A | NY-Analytical | 398 | pH | pH (10% Solution) | pH (10% Solution) | 7.57 |  | FDA,1 | 6 | 9 | TM 200
7A4OQ518/Chem_31A | 7A4OQ518 | Chem_31A | NY-Analytical | 796 | Assay | Assay | Assay | 51.97 | % | FDA,1 | 50 | 54 | TM 301/434
7A4OQ518/Chem_31A | 7A4OQ518 | Chem_31A | NY-Analytical | 796 | Assay | Assay Final | Assay Final | 51.98 | % | FDA,1 | 50 | 54 | TM 301/434
7A4OQ518/Chem_31A | 7A4OQ518 | Chem_31A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear viscous liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OQ518/Chem_31A | 7A4OQ518 | Chem_31A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to light straw (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OQ518/Chem_31A | 7A4OQ518 | Chem_31A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OQ518/Chem_31A | 7A4OQ518 | Chem_31A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Pleasant almond (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OQ890/Chem_34A | 7A4OQ890 | Chem_34A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7A4OQ890/Chem_34A | 7A4OQ890 | Chem_34A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Granular Powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OQ890/Chem_34A | 7A4OQ890 | Chem_34A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OQ890/Chem_34A | 7A4OQ890 | Chem_34A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OQ890/Chem_34A | 7A4OQ890 | Chem_34A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 792 | Chemical Inspection | Description | Description | Transparent, colorless, mobile, volatile liquid (Done) |  |  |  |  | ORGANOLEPTIC
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 795 | Acidity | Acidity (calc) | Acidity | 0.4 | mL | FDA,2 |  | 0.7 | USP
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 795 | Acidity | Acidity Final (Calc) | Acidity Final (Calc) | 0.45 | mL | FDA,2 |  | 0.7 | USP
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 796 | Assay | Assay | Assay | 100 | % | FDA,1 | 99 |  | USP
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 796 | Assay | Assay Final | Assay Final | 99.975 | % | FDA,1 | 99 |  | USP
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 797 | Nonvolatile Residue | End Weight (g) | End Weight | 90.0835 | g | FDA,4 |  |  | USP
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 797 | Nonvolatile Residue | Nonvolatile Residue (calc) | Nonvolatile Residue | 0 | % | FDA,3 |  | 0.005 | USP
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 797 | Nonvolatile Residue | Nonvolatile Residue Final (calc) | Nonvolatile Residue Final | 0.0008 | % | FDA,3 |  | 0.005 | USP
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.376 |  | FDA,3 | 1.376 | 1.378 | USP
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 797 | Nonvolatile Residue | Start Weight (g) | Start Weight | 90.7647 | g | FDA,4 |  |  | USP
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 795 | Acidity | Titrant End Volume (mL) | Titrant End Volume | 0.9 | mL | FDA,2 |  |  | USP
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 795 | Acidity | Titrant Start Volume (mL) | Titrant Start Volume | 0.5 | mL | FDA,2 |  |  | USP
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 797 | Nonvolatile Residue | Weight of Residue (g) | Weight of Residue | 0.0004 | g | FDA,4 |  |  | USP
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR conform to standard (Done) |  |  |  |  | USP <197F>
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.7832 |  | FDA,3 | 0.783 | 0.787 | USP/NF <841>
7A4OS411/Chem_36A | 7A4OS411 | Chem_36A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7A4OS411/Chem_36A | 7A4OS411 | Chem_36A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Free Flowing Powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OS411/Chem_36A | 7A4OS411 | Chem_36A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Brownish-Red (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OS411/Chem_36A | 7A4OS411 | Chem_36A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OS411/Chem_36A | 7A4OS411 | Chem_36A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7A4OS411/Chem_36A | 7A4OS411 | Chem_36A | NY-Analytical | 796 | Assay | Assay as is (Without Drying) | Assay as is (Without Drying) | 10.17 | % | FDA,1 |  |  | USP
7A4OS411/Chem_36A | 7A4OS411 | Chem_36A | NY-Analytical | 796 | Assay | Assay as is Final (Without Drying) | Assay as is Final (Without Drying) | 10.15 | % | FDA,1 |  |  | USP
7A4OS411/Chem_36A | 7A4OS411 | Chem_36A | NY-Analytical | 796 | Assay | Assay on Dried Basis | Assay on Dried Basis | 10.59 | % | FDA,1 | 9 | 12 | USP
7A4OS411/Chem_36A | 7A4OS411 | Chem_36A | NY-Analytical | 796 | Assay | Assay on Dried Basis Final | Assay on Dried Basis Final | 10.54 | % | FDA,1 | 9 | 12 | USP
7A4OS411/Chem_36A | 7A4OS411 | Chem_36A | NY-Analytical | 802 | Loss on Drying | Loss on Drying | Loss on Drying | 5.55 | % | FDA,1 |  | 8 | USP
7A4OS411/Chem_36A | 7A4OS411 | Chem_36A | NY-Analytical | 802 | Loss on Drying | Loss on Drying Final | Loss on Drying Final | 5.55 | % | FDA,1 |  | 8 | USP
7A4OU351/Chem_NY_Dehydrated Alcohol USP | 7A4OU351 | Chem_NY_Dehydrated Alcohol USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA
7A4OU351/Chem_NY_Dehydrated Alcohol USP | 7A4OU351 | Chem_NY_Dehydrated Alcohol USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis (ELLI) | Certificate of Analysis | COA Meets Specification (Done) |  |  |  |  | COA
7A4OU351/Chem_NY_Dehydrated Alcohol USP | 7A4OU351 | Chem_NY_Dehydrated Alcohol USP | NY-Analytical | 822 | Specific Gravity @ 15.56°C/ 60° F | Specific Gravity @ 15.56°C | Specific Gravity at 15.56°C | 0.794 |  | FDA,4 |  | 0.7962 | TM 202
7A4OU351/Chem_NY_Dehydrated Alcohol USP | 7A4OU351 | Chem_NY_Dehydrated Alcohol USP | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM 208
7A4OU351/Chem_NY_Dehydrated Alcohol USP | 7A4OU351 | Chem_NY_Dehydrated Alcohol USP | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Clear to Colorless and Match standard (Done) |  |  |  |  | TM 208
7A4OU351/Chem_NY_Dehydrated Alcohol USP | 7A4OU351 | Chem_NY_Dehydrated Alcohol USP | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7A4OU351/Chem_NY_Dehydrated Alcohol USP | 7A4OU351 | Chem_NY_Dehydrated Alcohol USP | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Alcohol (Done) |  |  |  |  | TM 208
7A4OU351/Chem_NY_Dehydrated Alcohol USP | 7A4OU351 | Chem_NY_Dehydrated Alcohol USP | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Compare to reference Standard (Done) |  |  |  |  | TM 411
7A4OU555/Chem_NY_Phenylephrine Reduced | 7A4OU555 | Chem_NY_Phenylephrine Reduced | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7A4OU555/Chem_NY_Phenylephrine Reduced | 7A4OU555 | Chem_NY_Phenylephrine Reduced | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Crystalline Powder or Crystals (Done) |  |  |  |  | TM 208
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Crystalline Powder or Crystals (Done) |  |  |  |  | TM 208
7A4OU555/Chem_NY_Phenylephrine Reduced | 7A4OU555 | Chem_NY_Phenylephrine Reduced | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White to almost white (Done) |  |  |  |  | TM 208
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White to almost white (Done) |  |  |  |  | TM 208
7A4OU555/Chem_NY_Phenylephrine Reduced | 7A4OU555 | Chem_NY_Phenylephrine Reduced | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7A4OU555/Chem_NY_Phenylephrine Reduced | 7A4OU555 | Chem_NY_Phenylephrine Reduced | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7A4OU555/Chem_NY_Phenylephrine Reduced | 7A4OU555 | Chem_NY_Phenylephrine Reduced | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 812 | Organic Impurites | Any individual unspecified impuritv | Any individual unspecified impuritv |  | % | FDA,2 |  | 0.1 | USP
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 796 | Assay | Assay | Assay | 99.4 | % | FDA,1 | 98 | 102 | USP
7A4OU555/Chem_NY_Phenylephrine Reduced | 7A4OU555 | Chem_NY_Phenylephrine Reduced | NY-Analytical | 796 | Assay | Assay | Assay | 99.9 | % | FDA,1 | 98 | 102 | USP
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 802 | Loss on Drying | Loss on Drying | Loss on Drying |  | % | FDA,1 |  | 1 | USP
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 802 | Loss on Drying | Loss on Drying Final | Loss on Drying Final |  | % | FDA,1 |  | 1 | USP
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 812 | Organic Impurites | Norphenylephrine | Norphenylephrine |  | % | FDA,2 |  | 0.1 | USP
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 1043 | Optical Rotation | Optical Rotation | Optical Rotation |  | ° | FDA,0 | -47 | -43 | USP
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 812 | Organic Impurites | Phenylephrine Related Compound C | Phenylephrine Related Compound C |  | % | FDA,2 |  | 0.1 | USP
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 812 | Organic Impurites | Phenylephrine Related Compound D | Phenylephrine Related Compound D |  | % | FDA,2 |  | 0.1 | USP
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 812 | Organic Impurites | Phenylephrine Related Compound E | Phenylephrine Related Compound E |  | % | FDA,2 |  | 0.1 | USP
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 834 | Residue on Ignition | Residue on Ignition | Residue on Ignition |  | % | FDA,1 |  | 0.2 | USP
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 835 | Sulfate Content | Sulfate | Sulfate |  | % | FDA,2 |  | 0.2 | USP
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Chem_NY_Phenylephrine HCL | NY-Analytical | 812 | Organic Impurites | Total Impurities | Total Impurities |  | % | FDA,2 |  | 0.2 | USP
7A4OX110/Chem_39A | 7A4OX110 | Chem_39A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification
7A4OX110/Chem_39A | 7A4OX110 | Chem_39A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | [5 empty] | TM 411
7A4OX110/Chem_39A | 7A4OX110 | Chem_39A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | [5 empty] | TM208/220/230/USP/Kao LM
7A4OX110/Chem_39A | 7A4OX110 | Chem_39A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | [5 empty] | TM208/220/230/USP/Kao LM
7A4OX110/Chem_39A | 7A4OX110 | Chem_39A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | [5 empty] | TM208/220/230/USP/Kao LM
7A4OX110/Chem_39A | 7A4OX110 | Chem_39A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | [5 empty] | TM208/220/230/USP/Kao LM
7A4OX231/Chem_NY_D&C Yellow 10 | 7A4OX231 | Chem_NY_D&C Yellow 10 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Appearance
7A4OX231/Chem_NY_D&C Yellow 10 | 7A4OX231 | Chem_NY_D&C Yellow 10 | NY-Analytical | 397 | Appearance | Physical Properties | Physical Properties | Yellow Powder (Done) |  |  |  |  | TM 208
7A4OX231/Chem_NY_D&C Yellow 10 | 7A4OX231 | Chem_NY_D&C Yellow 10 | NY-Analytical | 390 | Identification | Spectrophotometric | Spectrophotometric | 410.88 | λ | FDA,0 | 408 | 416 | TM 3014
7A4OX231/Chem_NY_D&C Yellow 10 | 7A4OX231 | Chem_NY_D&C Yellow 10 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram) | TAMC (Average per Gram) | 0 | CFU/g | FDA,1 |  |  | TM 610
7A4OX231/Chem_NY_D&C Yellow 10 | 7A4OX231 | Chem_NY_D&C Yellow 10 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
7A4OX231/Chem_NY_D&C Yellow 10 | 7A4OX231 | Chem_NY_D&C Yellow 10 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram) | TMYC (Average per Gram) | 0 | CFU/g | FDA,1 |  |  | TM 610
7A4OX231/Chem_NY_D&C Yellow 10 | 7A4OX231 | Chem_NY_D&C Yellow 10 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
7ACHE61P/Chem_NY_Sodium Benzoate | 7ACHE61P | Chem_NY_Sodium Benzoate | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7ACHE61P/Chem_NY_Sodium Benzoate | 7ACHE61P | Chem_NY_Sodium Benzoate | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Extruded Pellets (Done) |  |  |  |  | TM 208
7ACHE61P/Chem_NY_Sodium Benzoate | 7ACHE61P | Chem_NY_Sodium Benzoate | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) |  |  |  |  | TM 208
7ACHE61P/Chem_NY_Sodium Benzoate | 7ACHE61P | Chem_NY_Sodium Benzoate | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7ACHE61P/Chem_NY_Sodium Benzoate | 7ACHE61P | Chem_NY_Sodium Benzoate | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless (Done) |  |  |  |  | TM 208
7ACHE61P/Chem_NY_Sodium Benzoate | 7ACHE61P | Chem_NY_Sodium Benzoate | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7B4AA580/Chem_NY_Acetic Acid - USP | 7B4AA580 | Chem_NY_Acetic Acid - USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7B4AA580/Chem_NY_Acetic Acid - USP | 7B4AA580 | Chem_NY_Acetic Acid - USP | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4AA580/Chem_NY_Acetic Acid - USP | 7B4AA580 | Chem_NY_Acetic Acid - USP | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4AA580/Chem_NY_Acetic Acid - USP | 7B4AA580 | Chem_NY_Acetic Acid - USP | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4AA580/Chem_NY_Acetic Acid - USP | 7B4AA580 | Chem_NY_Acetic Acid - USP | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Pungent (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4AL930/Chem_42A | 7B4AL930 | Chem_42A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH as is
7B4AL930/Chem_42A | 7B4AL930 | Chem_42A | NY-Analytical | 447 | pH as is | pH | pH |  |  | FDA,1 | 1.1 | 1.5 | TM 200
7B4AL930/Chem_42A | 7B4AL930 | Chem_42A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.2444 |  | FDA,2 | 1.21 | 1.27 | TM 202
7B4AL930/Chem_42A | 7B4AL930 | Chem_42A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7B4AL930/Chem_42A | 7B4AL930 | Chem_42A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4AL930/Chem_42A | 7B4AL930 | Chem_42A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Light Yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4AL930/Chem_42A | 7B4AL930 | Chem_42A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4AL930/Chem_42A | 7B4AL930 | Chem_42A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4FD110/Chem_44A | 7B4FD110 | Chem_44A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification
7B4FD110/Chem_44A | 7B4FD110 | Chem_44A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7B4FD110/Chem_44A | 7B4FD110 | Chem_44A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | White granular powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4FD110/Chem_44A | 7B4FD110 | Chem_44A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4FD110/Chem_44A | 7B4FD110 | Chem_44A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4GB598/Chem_45A | 7B4GB598 | Chem_45A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7B4GB598/Chem_45A | 7B4GB598 | Chem_45A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Grain Powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4GB598/Chem_45A | 7B4GB598 | Chem_45A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Off-White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4GB598/Chem_45A | 7B4GB598 | Chem_45A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4GB598/Chem_45A | 7B4GB598 | Chem_45A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OA242/Chem_48A | 7B4OA242 | Chem_48A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification
7B4OA242/Chem_48A | 7B4OA242 | Chem_48A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7B4OA242/Chem_48A | 7B4OA242 | Chem_48A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Crystals/crystalline powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OA242/Chem_48A | 7B4OA242 | Chem_48A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to white (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OA242/Chem_48A | 7B4OA242 | Chem_48A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OA242/Chem_48A | 7B4OA242 | Chem_48A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | 7B4OA347 | Chem_NY_Calming Waters INCI w/symfresh | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C
7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | 7B4OA347 | Chem_NY_Calming Waters INCI w/symfresh | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.964 |  | FDA,3 | 0.959 | 0.969 | TM 202
7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | 7B4OA347 | Chem_NY_Calming Waters INCI w/symfresh | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.472 |  | FDA,3 | 1.465 | 1.475 | TM 204
7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | 7B4OA347 | Chem_NY_Calming Waters INCI w/symfresh | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM 208
7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | 7B4OA347 | Chem_NY_Calming Waters INCI w/symfresh | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to pale yellow (Done) |  |  |  |  | TM 208
7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | 7B4OA347 | Chem_NY_Calming Waters INCI w/symfresh | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | 7B4OA347 | Chem_NY_Calming Waters INCI w/symfresh | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208
7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | 7B4OA347 | Chem_NY_Calming Waters INCI w/symfresh | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411
7B4OA709/Chem_NY_Bell fragrance | 7B4OA709 | Chem_NY_Bell fragrance | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C
7B4OA709/Chem_NY_Bell fragrance | 7B4OA709 | Chem_NY_Bell fragrance | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.0639 |  | FDA,3 | 1.06 | 1.07 | TM 202
7B4OA709/Chem_NY_Bell fragrance | 7B4OA709 | Chem_NY_Bell fragrance | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.468 |  | FDA,3 | 1.467 | 1.477 | TM 204
7B4OA709/Chem_NY_Bell fragrance | 7B4OA709 | Chem_NY_Bell fragrance | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM 208
7B4OA709/Chem_NY_Bell fragrance | 7B4OA709 | Chem_NY_Bell fragrance | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Pale to light yellow (Done) |  |  |  |  | TM 208
7B4OA709/Chem_NY_Bell fragrance | 7B4OA709 | Chem_NY_Bell fragrance | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7B4OA709/Chem_NY_Bell fragrance | 7B4OA709 | Chem_NY_Bell fragrance | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Conforms to Standard (Done) |  |  |  |  | TM 208
7B4OB690/Chem_50A | 7B4OB690 | Chem_50A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7B4OB690/Chem_50A | 7B4OB690 | Chem_50A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, mobile liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OB690/Chem_50A | 7B4OB690 | Chem_50A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OB690/Chem_50A | 7B4OB690 | Chem_50A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OB690/Chem_50A | 7B4OB690 | Chem_50A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Solvent (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | CHE ANTIFOAM AF EMULSION | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | CHE ANTIFOAM AF EMULSION | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | CHE ANTIFOAM AF EMULSION | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per 10mL) | TAMC (Average per 10mL) | 0 | cfu/10mL | FDA,1 |  | 0 | TM 610
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | CHE ANTIFOAM AF EMULSION | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per mL - Undiluted) | TAMC (Average per mL - Undiluted) | 0 | cfu/mL |  |  | 0 | TM 610
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | CHE ANTIFOAM AF EMULSION | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  | 0 | TM 610
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | CHE ANTIFOAM AF EMULSION | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per 10mL) | TMYC (Average per 10mL) | 0 | cfu/10mL | FDA,1 |  | 0 | TM 610
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | CHE ANTIFOAM AF EMULSION | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per mL - Undiluted) | TMYC (Average per mL - Undiluted) | 0 | cfu/mL |  |  | 0 | TM 610
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | CHE ANTIFOAM AF EMULSION | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 | cfu | FDA,0 |  | 0 | TM 610
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | CHE ANTIFOAM AF EMULSION | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Homogeneous Liquid, Free of Dark Particles (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | CHE ANTIFOAM AF EMULSION | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White to off White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | CHE ANTIFOAM AF EMULSION | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | CHE ANTIFOAM AF EMULSION | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | CHE ANTIFOAM AF EMULSION | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | [5 empty] | WI QAS 076
7B4OD232/Chem_52A | 7B4OD232 | Chem_52A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C
7B4OD232/Chem_52A | 7B4OD232 | Chem_52A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.0361 |  | FDA,3 | 1.035 | 1.037 | TM 202
7B4OD232/Chem_52A | 7B4OD232 | Chem_52A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7B4OD232/Chem_52A | 7B4OD232 | Chem_52A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear viscous liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OD232/Chem_52A | 7B4OD232 | Chem_52A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OD232/Chem_52A | 7B4OD232 | Chem_52A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OD232/Chem_52A | 7B4OD232 | Chem_52A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | 7B4OM799 | Chem_NY_Lipotec_Cucumber Extract in Glycerin | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C
7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | 7B4OM799 | Chem_NY_Lipotec_Cucumber Extract in Glycerin | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.2627 |  | FDA,3 | 1.25 | 1.29 | TM 202
7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | 7B4OM799 | Chem_NY_Lipotec_Cucumber Extract in Glycerin | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.473 |  | FDA,3 | 1.45 | 1.49 | TM 204
7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | 7B4OM799 | Chem_NY_Lipotec_Cucumber Extract in Glycerin | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear viscous liquid (Done) |  |  |  |  | TM 208
7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | 7B4OM799 | Chem_NY_Lipotec_Cucumber Extract in Glycerin | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Light to Medium Yellow (Done) |  |  |  |  | TM 208
7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | 7B4OM799 | Chem_NY_Lipotec_Cucumber Extract in Glycerin | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | 7B4OM799 | Chem_NY_Lipotec_Cucumber Extract in Glycerin | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) |  |  |  |  | TM 208
7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | 7B4OM799 | Chem_NY_Lipotec_Cucumber Extract in Glycerin | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Chem_NY_Abil B88183 PH | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Chem_NY_Abil B88183 PH | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear to slightly turbid.Colorless to pale yellow,slightly viscous liquid (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Chem_NY_Abil B88183 PH | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Chem_NY_Abil B88183 PH | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Chem_NY_Abil B88183 PH | NY-Microbiology | 509 | J&J_Pathogens | Biggy or SDA | Biggy or SDA | No Growth_SDA |  |  |  |  | TM 869
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Chem_NY_Abil B88183 PH | NY-Microbiology | 509 | J&J_Pathogens | Cetrimide | Cetrimide | No Growth |  |  |  |  | TM 869
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Chem_NY_Abil B88183 PH | NY-Microbiology | 524 | J&J Fungal | Fungal | Fungal | 0 | CFU/g | FDA,0 | 0 | 200 | TM 869
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Chem_NY_Abil B88183 PH | NY-Microbiology | 509 | J&J_Pathogens | MacConkey | MacConkey | No Growth |  |  |  |  | TM 869
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Chem_NY_Abil B88183 PH | NY-Microbiology | 509 | J&J_Pathogens | MSA or VJ | MSA or VJ | No Growth_MSA |  |  |  |  | TM 869
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Chem_NY_Abil B88183 PH | NY-Microbiology | 700 | J&J SDA Controls | Negative Control | Negative Control | No Growth SDA |  |  |  |  | TM 869
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Chem_NY_Abil B88183 PH | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | NegativeControl | No Growth TSB |  |  |  |  | TM 869
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Chem_NY_Abil B88183 PH | NY-Microbiology | 509 | J&J_Pathogens | TSA | TSA | No Growth |  |  |  |  | TM 869
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Chem_NY_Abil B88183 PH | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | TSA JJ | 0 | cfu/mL | FDA,0 | 0 | 200 | TM 869
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Chem_NY_Abil B88183 PH | NY-Microbiology | 517 | J&J_Enrichment | TSB | TSB | Not Turbid |  |  |  |  | TM 869
7B4OP065/Chem_NY_Chamomile Flower Extract | 7B4OP065 | Chem_NY_Chamomile Flower Extract | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7B4OP065/Chem_NY_Chamomile Flower Extract | 7B4OP065 | Chem_NY_Chamomile Flower Extract | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Liquid (Done) |  |  |  |  | TM 208
7B4OP065/Chem_NY_Chamomile Flower Extract | 7B4OP065 | Chem_NY_Chamomile Flower Extract | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Greenish Yellow to Yellow Orange (Done) |  |  |  |  | TM 208
7B4OP065/Chem_NY_Chamomile Flower Extract | 7B4OP065 | Chem_NY_Chamomile Flower Extract | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7B4OP065/Chem_NY_Chamomile Flower Extract | 7B4OP065 | Chem_NY_Chamomile Flower Extract | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208
7B4OP065/Chem_NY_Chamomile Flower Extract | 7B4OP065 | Chem_NY_Chamomile Flower Extract | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411
7B4OP065/Chem_NY_Chamomile Flower Extract | 7B4OP065 | Chem_NY_Chamomile Flower Extract | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram) | TAMC (Average per Gram) | 0 | CFU/g | FDA,1 |  | 100 | TM 610
7B4OP065/Chem_NY_Chamomile Flower Extract | 7B4OP065 | Chem_NY_Chamomile Flower Extract | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
7B4OP065/Chem_NY_Chamomile Flower Extract | 7B4OP065 | Chem_NY_Chamomile Flower Extract | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram) | TMYC (Average per Gram) |  | CFU/g | FDA,1 |  | 100 | TM 610
7B4OP065/Chem_NY_Chamomile Flower Extract | 7B4OP065 | Chem_NY_Chamomile Flower Extract | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  | cfu | FDA,0 |  |  | TM 610
7B4OP065/Chem_NY_Chamomile Flower Extract | 7B4OP065 | Chem_NY_Chamomile Flower Extract | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Non-Subculturable Growth (Done) |  |  |  |  | WI QAS 076
7B4OP093/Chem_NY_PEG-4 | 7B4OP093 | Chem_NY_PEG-4 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH
7B4OP093/Chem_NY_PEG-4 | 7B4OP093 | Chem_NY_PEG-4 | NY-Analytical | 398 | pH | pH (5% aqueous @25 degree C) | pH (5% aqueous @25 degree C) | 6.01 |  | FDA,1 | 4.5 | 7.5 | TM 200
7B4OP093/Chem_NY_PEG-4 | 7B4OP093 | Chem_NY_PEG-4 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Not More Than Slightly Hazy (Done) |  |  |  |  | TM 208
7B4OP093/Chem_NY_PEG-4 | 7B4OP093 | Chem_NY_PEG-4 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM 208
7B4OP093/Chem_NY_PEG-4 | 7B4OP093 | Chem_NY_PEG-4 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7B4OP093/Chem_NY_PEG-4 | 7B4OP093 | Chem_NY_PEG-4 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208
7B4OP093/Chem_NY_PEG-4 | 7B4OP093 | Chem_NY_PEG-4 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7B4OP113/Chem_NY_Mentha Piperita Extract_01 | 7B4OP113 | Chem_NY_Mentha Piperita Extract_01 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH
7B4OP113/Chem_NY_Mentha Piperita Extract_01 | 7B4OP113 | Chem_NY_Mentha Piperita Extract_01 | NY-Analytical | 398 | pH | pH | pH | 4.24 |  | FDA,1 | 4 | 5 | TM 200
7B4OP113/Chem_NY_Mentha Piperita Extract_01 | 7B4OP113 | Chem_NY_Mentha Piperita Extract_01 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear aqueous solution (Done) |  |  |  |  | TM 208
7B4OP113/Chem_NY_Mentha Piperita Extract_01 | 7B4OP113 | Chem_NY_Mentha Piperita Extract_01 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Amber (Done) |  |  |  |  | TM 208
7B4OP113/Chem_NY_Mentha Piperita Extract_01 | 7B4OP113 | Chem_NY_Mentha Piperita Extract_01 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7B4OP113/Chem_NY_Mentha Piperita Extract_01 | 7B4OP113 | Chem_NY_Mentha Piperita Extract_01 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) |  |  |  |  | TM 208
7B4OP203/Chem_NY_Diazolidinyl Urea – RI-1157 | 7B4OP203 | Chem_NY_Diazolidinyl Urea – RI-1157 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH
7B4OP203/Chem_NY_Diazolidinyl Urea – RI-1157 | 7B4OP203 | Chem_NY_Diazolidinyl Urea – RI-1157 | NY-Analytical | 398 | pH | pH (1% w/v aqueous solution) | pH (1% w/v aqueous solution) | 7.02 |  | FDA,1 | 6.5 | 8 | TM 200
7B4OP203/Chem_NY_Diazolidinyl Urea – RI-1157 | 7B4OP203 | Chem_NY_Diazolidinyl Urea – RI-1157 | NY-Analytical | 891 | Solutions and Solubility | Solubility (1% w/v aqueous solution) | Solubility (1% w/v aqueous solution) | Clear and Colorless (Done) |  |  |  |  | TM 222
7B4OP203/Chem_NY_Diazolidinyl Urea – RI-1157 | 7B4OP203 | Chem_NY_Diazolidinyl Urea – RI-1157 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7B4OP203/Chem_NY_Diazolidinyl Urea – RI-1157 | 7B4OP203 | Chem_NY_Diazolidinyl Urea – RI-1157 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Free flowing powder, free of foreign material (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OP203/Chem_NY_Diazolidinyl Urea – RI-1157 | 7B4OP203 | Chem_NY_Diazolidinyl Urea – RI-1157 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OP203/Chem_NY_Diazolidinyl Urea – RI-1157 | 7B4OP203 | Chem_NY_Diazolidinyl Urea – RI-1157 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Similar in character and intensity to standard, slight pungent odor (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OP209/MASTER - Microbiology - J&J Chemical Testing) | 7B4OP209 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 808 | Absorption | Absorption (1%/1cm) | Absorption (1%/1cm) | 404.16 |  | FDA,1 | 401.8 |  | J&J TM-002747
7B4OP209/MASTER - Microbiology - J&J Chemical Testing) | 7B4OP209 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 398 | pH | pH (1% Solution) | pH (1% Solution) | 6.2 |  | FDA,1 | 4.5 | 7 | J&J TMWW78014
7B4OP230/Chem_NY_Cosvat (Elestab CPN Ultra Pure) | 7B4OP230 | Chem_NY_Cosvat (Elestab CPN Ultra Pure) | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7B4OP230/Chem_NY_Cosvat (Elestab CPN Ultra Pure) | 7B4OP230 | Chem_NY_Cosvat (Elestab CPN Ultra Pure) | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | White to off - white powder or crystalline aggregates. (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7B4OP230/Chem_NY_Cosvat (Elestab CPN Ultra Pure) | 7B4OP230 | Chem_NY_Cosvat (Elestab CPN Ultra Pure) | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013
7B4OP230/Chem_NY_Cosvat (Elestab CPN Ultra Pure) | 7B4OP230 | Chem_NY_Cosvat (Elestab CPN Ultra Pure) | NY-Analytical | 801 | % Water | Water Content | Water Content | 0.2 | % | FDA,1 |  | 1 | J&J TMWW78017
7B4OP230/Chem_NY_Cosvat (Elestab CPN Ultra Pure) | 7B4OP230 | Chem_NY_Cosvat (Elestab CPN Ultra Pure) | NY-Analytical | 796 | Assay | Assay | Assay | 100.7 | % | FDA,1 | 99 |  | TM-002291
7B4OP230/Chem_NY_Cosvat (Elestab CPN Ultra Pure) | 7B4OP230 | Chem_NY_Cosvat (Elestab CPN Ultra Pure) | NY-Analytical | 796 | Assay | Assay Final | Assay Final | 100.7 | % | FDA,1 | 99 |  | TM-002291
7B4OP284/Chem_NY_Sodium Citrate USP | 7B4OP284 | Chem_NY_Sodium Citrate USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Odor
7B4OP284/Chem_NY_Sodium Citrate USP | 7B4OP284 | Chem_NY_Sodium Citrate USP | NY-Analytical | 396 | Odor | Odor | Odor | Passes Test (Done) |  |  |  |  | J&J TM-003188
7B4OP284/Chem_NY_Sodium Citrate USP | 7B4OP284 | Chem_NY_Sodium Citrate USP | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Colorless crystals, or as a white crystalline powder (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7B4OP284/Chem_NY_Sodium Citrate USP | 7B4OP284 | Chem_NY_Sodium Citrate USP | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7B4OP284/Chem_NY_Sodium Citrate USP | 7B4OP284 | Chem_NY_Sodium Citrate USP | NY-Analytical | 813 | Identification - USP | Identification A - Sodium | Identification A - Sodium | Passes Test (Done) |  |  |  |  | USP
7B4OP284/Chem_NY_Sodium Citrate USP | 7B4OP284 | Chem_NY_Sodium Citrate USP | NY-Analytical | 813 | Identification - USP | Identification B - Citrate | Identification B - Citrate | Passes Test (Done) |  |  |  |  | USP
7B4OP284/Chem_NY_Sodium Citrate USP | 7B4OP284 | Chem_NY_Sodium Citrate USP | NY-Analytical | 813 | Identification - USP | Identification C | Identification C | Passes Test (Done) |  |  |  |  | USP
7B4OP661/Chem_NY_Pemulen TR-2 | 7B4OP661 | Chem_NY_Pemulen TR-2 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA
7B4OP661/Chem_NY_Pemulen TR-2 | 7B4OP661 | Chem_NY_Pemulen TR-2 | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7B4OP661/Chem_NY_Pemulen TR-2 | 7B4OP661 | Chem_NY_Pemulen TR-2 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OP661/Chem_NY_Pemulen TR-2 | 7B4OP661 | Chem_NY_Pemulen TR-2 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OP661/Chem_NY_Pemulen TR-2 | 7B4OP661 | Chem_NY_Pemulen TR-2 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OP661/Chem_NY_Pemulen TR-2 | 7B4OP661 | Chem_NY_Pemulen TR-2 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Slightly acetic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OQ028/Chem_NY_Sorbitan Monolaurate | 7B4OQ028 | Chem_NY_Sorbitan Monolaurate | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification
7B4OQ028/Chem_NY_Sorbitan Monolaurate | 7B4OQ028 | Chem_NY_Sorbitan Monolaurate | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411
7B4OQ028/Chem_NY_Sorbitan Monolaurate | 7B4OQ028 | Chem_NY_Sorbitan Monolaurate | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OQ028/Chem_NY_Sorbitan Monolaurate | 7B4OQ028 | Chem_NY_Sorbitan Monolaurate | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OQ028/Chem_NY_Sorbitan Monolaurate | 7B4OQ028 | Chem_NY_Sorbitan Monolaurate | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OQ028/Chem_NY_Sorbitan Monolaurate | 7B4OQ028 | Chem_NY_Sorbitan Monolaurate | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OQ084/Chem_60A | 7B4OQ084 | Chem_60A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7B4OQ084/Chem_60A | 7B4OQ084 | Chem_60A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Slightly Opaque Viscous Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OQ084/Chem_60A | 7B4OQ084 | Chem_60A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Slightly White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OQ084/Chem_60A | 7B4OQ084 | Chem_60A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OQ084/Chem_60A | 7B4OQ084 | Chem_60A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Mild (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OQ209/Chem_61A | 7B4OQ209 | Chem_61A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification
7B4OQ209/Chem_61A | 7B4OQ209 | Chem_61A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | [5 empty] | TM 411
7B4OQ209/Chem_61A | 7B4OQ209 | Chem_61A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Free flowing granules or powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OQ209/Chem_61A | 7B4OQ209 | Chem_61A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OQ209/Chem_61A | 7B4OQ209 | Chem_61A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OQ209/Chem_61A | 7B4OQ209 | Chem_61A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OQ214/Chem_62A | 7B4OQ214 | Chem_62A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification
7B4OQ214/Chem_62A | 7B4OQ214 | Chem_62A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7B4OQ214/Chem_62A | 7B4OQ214 | Chem_62A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Free flowing crystalline powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OQ214/Chem_62A | 7B4OQ214 | Chem_62A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OQ214/Chem_62A | 7B4OQ214 | Chem_62A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OQ214/Chem_62A | 7B4OQ214 | Chem_62A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OQ595/Chem_63A | 7B4OQ595 | Chem_63A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification
7B4OQ595/Chem_63A | 7B4OQ595 | Chem_63A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7B4OQ595/Chem_63A | 7B4OQ595 | Chem_63A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Crystalline Powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OQ595/Chem_63A | 7B4OQ595 | Chem_63A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White to almost white (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OQ595/Chem_63A | 7B4OQ595 | Chem_63A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OQ595/Chem_63A | 7B4OQ595 | Chem_63A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Faint (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Chem_NY_Disodium EDTA-USP | NY-Analytical | 841 | Residual Solvents | Residual Solvents | Residual Solvents | Supplier Attestation Letter |  |  |  |  | Residual Solvents
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Chem_NY_Disodium EDTA-USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Chem_NY_Disodium EDTA-USP | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | White crystalline powder (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Chem_NY_Disodium EDTA-USP | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Chem_NY_Disodium EDTA-USP | NY-Analytical | 398 | pH | pH (1 in 20 solution) | pH (1 in 20 solution) | 4.9 |  | FDA,1 | 4 | 6 | J&J TMWW78014
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Chem_NY_Disodium EDTA-USP | NY-Analytical | 796 | Assay | Assay | Assay | 99.8 | % | FDA,1 | 99 | 101 | USP
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Chem_NY_Disodium EDTA-USP | NY-Analytical | 796 | Assay | Assay Final | Assay Final | 99.8 | % | FDA,1 | 99 | 101 | USP
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Chem_NY_Disodium EDTA-USP | NY-Analytical | 843 | Calcium | Calcium | Calcium | Passes Test (Done) |  |  |  |  | USP
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Chem_NY_Disodium EDTA-USP | NY-Analytical | 833 | Heavy Metals | Heavy Metals | Heavy Metals | <0.005 | % | FDA,3 |  | 0.005 | USP
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Chem_NY_Disodium EDTA-USP | NY-Analytical | 813 | Identification - USP | Identification A | Identification A | Passes Test (Done) |  |  |  |  | USP
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Chem_NY_Disodium EDTA-USP | NY-Analytical | 813 | Identification - USP | Identification B | Identification B | Passes Test (Done) |  |  |  |  | USP
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Chem_NY_Disodium EDTA-USP | NY-Analytical | 813 | Identification - USP | Identification C | Identification C | Passes Test (Done) |  |  |  |  | USP
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Chem_NY_Disodium EDTA-USP | NY-Analytical | 802 | Loss on Drying | Loss on Drying | Loss on Drying | 9.6 | % | FDA,1 | 8.7 | 11.4 | USP
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Chem_NY_Disodium EDTA-USP | NY-Analytical | 802 | Loss on Drying | Loss on Drying Final | Loss on Drying Final | 9.6 | % | FDA,1 | 8.7 | 11.4 | USP
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Chem_NY_Disodium EDTA-USP | NY-Analytical | 842 | Nitrilotriacetic Acid | Nitrilotriacetic Acid | Nitrilotriacetic Acid | 0 | % | FDA,1 |  | 0.1 | USP
7B4OX232/Chem_40A | 7B4OX232 | Chem_40A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Apparent proof
7B4OX232/Chem_40A | 7B4OX232 | Chem_40A | NY-Analytical | 936 | Apparent proof | Apparent proof | Apparent proof | 189.7 |  | FDA,1 | 189.7 | 190.2 | STM 101
7B4OX232/Chem_40A | 7B4OX232 | Chem_40A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.8091 |  | FDA,4 | 0.8086 | 0.8098 | TM 202
7B4OX232/Chem_40A | 7B4OX232 | Chem_40A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Meets USP requirement (Done) |  |  |  |  | TM 411
7B4OX232/Chem_40A | 7B4OX232 | Chem_40A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, mobile liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OX232/Chem_40A | 7B4OX232 | Chem_40A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OX232/Chem_40A | 7B4OX232 | Chem_40A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7B4OX232/Chem_40A | 7B4OX232 | Chem_40A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Alcohol (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | 7C4AA580 | Chem_NY_Fragrance-Grapefruit Blossom &Orange | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C
7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | 7C4AA580 | Chem_NY_Fragrance-Grapefruit Blossom &Orange | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.0662 |  | FDA,4 | 1.0561 | 1.0761 | TM 202
7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | 7C4AA580 | Chem_NY_Fragrance-Grapefruit Blossom &Orange | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | 7C4AA580 | Chem_NY_Fragrance-Grapefruit Blossom &Orange | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | 7C4AA580 | Chem_NY_Fragrance-Grapefruit Blossom &Orange | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Pale Yellow – Light Yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | 7C4AA580 | Chem_NY_Fragrance-Grapefruit Blossom &Orange | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | 7C4AA580 | Chem_NY_Fragrance-Grapefruit Blossom &Orange | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7C4FD110/Chem_69A | 7C4FD110 | Chem_69A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7C4FD110/Chem_69A | 7C4FD110 | Chem_69A | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | White or nearly white, hard, brittle material (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7C4FD110/Chem_69A | 7C4FD110 | Chem_69A | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7C4FD110/Chem_69A | 7C4FD110 | Chem_69A | NY-Analytical | 390 | Identification | Identification B | Identification B | Pass (Done) |  |  |  |  | USP
7C4FD110/Chem_69A | 7C4FD110 | Chem_69A | NY-Analytical | 390 | Identification | Identification for Sodium (USP <191>) | Identification for Sodium | Positive response for Sodium (Done) |  |  |  |  | USP
7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 7C4OA709 | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 961 | Peroxide Value | Peroxide Value | Peroxide Value |  |  | FDA,1 |  | 1 | AOCS Cd 8-53
7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 7C4OA709 | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 944 | color (Gardner) | color (Gardner) | color (Gardner) |  |  | FDA,0 |  | 3 | AOCS Td 1a-64)
7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 7C4OA709 | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA
7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 7C4OA709 | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 826 | Acid Value | Acid Value | Acid Value | 1.7 | mg KOH/g | FDA,2 |  | 2 | IUPAC 2.201(M)
7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 7C4OA709 | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 826 | Acid Value | Acid Value Final | Acid Value Final | 1.7 | mg KOH/g | FDA,2 |  | 2 | IUPAC 2.201(M)
7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 7C4OA709 | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 810 | Saponification Value | Saponification Value | Saponification Value | 22.1 | mg/g | FDA,1 | 17 | 25 | IUPAC 2.202
7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 7C4OA709 | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 810 | Saponification Value | Saponification Value Final | Saponification Value Final | 22.1 | mg/g | FDA,1 | 17 | 25 | IUPAC 2.202
7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 7C4OA709 | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 1000 | Cloudpoint | Cloudpoint | Cloudpoint | 67.4 | C | FDA,1 | 66 | 71 | TC276/WG2N30 930809C
7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 7C4OA709 | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 959 | Water Content | Water Content | Water Content |  | % | FDA,2 |  | 1.5 | Vatten LB
7C4OD232/Chem_NY_Methylparaben NF/USP | 7C4OD232 | Chem_NY_Methylparaben NF/USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis
7C4OD232/Chem_NY_Propylparaben USP | 7C4OD232 | Chem_NY_Propylparaben USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification
7C4OD232/Chem_NY_Methylparaben NF/USP | 7C4OD232 | Chem_NY_Methylparaben NF/USP | NY-Analytical | 390 | Identification | IR Identification | IR Identification | [5 empty] | TM 411
7C4OD232/Chem_NY_Methylparaben NF/USP | 7C4OD232 | Chem_NY_Methylparaben NF/USP | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Colorless crystals or white crystalline powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7C4OD232/Chem_NY_Propylparaben USP | 7C4OD232 | Chem_NY_Propylparaben USP | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Small, colorless crystals or white crystalline powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7C4OD232/Chem_NY_Methylparaben NF/USP | 7C4OD232 | Chem_NY_Methylparaben NF/USP | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | [5 empty] | TM208/220/230/USP/Kao LM
7C4OD232/Chem_NY_Propylparaben USP | 7C4OD232 | Chem_NY_Propylparaben USP | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7C4OD232/Chem_NY_Methylparaben NF/USP | 7C4OD232 | Chem_NY_Methylparaben NF/USP | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | [5 empty] | TM208/220/230/USP/Kao LM
7C4OD232/Chem_NY_Propylparaben USP | 7C4OD232 | Chem_NY_Propylparaben USP | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7C4OD232/Chem_NY_Propylparaben USP | 7C4OD232 | Chem_NY_Propylparaben USP | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | USP <197M>
7C4OD232/Chem_NY_Propylparaben USP | 7C4OD232 | Chem_NY_Propylparaben USP | NY-Analytical | 796 | Assay | Assay | Assay | 99.4 | % | FDA,1 | 98 | 102 | USP/NF
7C4OD232/Chem_NY_Methylparaben NF/USP | 7C4OD232 | Chem_NY_Methylparaben NF/USP | NY-Analytical | 796 | Assay | Assay | Assay |  | % | FDA,1 | 98 | 102 | USP/NF
7C4OD232/Chem_NY_Propylparaben USP | 7C4OD232 | Chem_NY_Propylparaben USP | NY-Analytical | 796 | Assay | Assay Final | Assay Final | 99.4 | % | FDA,1 | 98 | 102 | USP/NF
7C4OD232/Chem_NY_Methylparaben NF/USP | 7C4OD232 | Chem_NY_Methylparaben NF/USP | NY-Analytical | 796 | Assay | Assay Final | Assay Final |  | % | FDA,1 | 98 | 102 | USP/NF
7C4OG169/Chem_75A | 7C4OG169 | Chem_75A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7C4OG169/Chem_75A | 7C4OG169 | Chem_75A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Crystals (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7C4OG169/Chem_75A | 7C4OG169 | Chem_75A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7C4OG169/Chem_75A | 7C4OG169 | Chem_75A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7C4OG169/Chem_75A | 7C4OG169 | Chem_75A | NY-Analytical | 390 | Identification | Identification for Ammonium (USP <191>) | Identification for Ammonium | Pass (Done) |  |  |  |  | TM400/408b/421/422/USP
7C4OG169/Chem_75A | 7C4OG169 | Chem_75A | NY-Analytical | 390 | Identification | Identification for Chloride (USP <191>) | Identification for Chloride | Pass (Done) |  |  |  |  | TM400/408b/421/422/USP
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Chem_NY_Camellia Oleifera Leaf Extract | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Chem_NY_Camellia Oleifera Leaf Extract | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Light yellow to amber liquid (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Chem_NY_Camellia Oleifera Leaf Extract | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Chem_NY_Camellia Oleifera Leaf Extract | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Chem_NY_Camellia Oleifera Leaf Extract | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 509 | J&J_Pathogens | Biggy or SDA | Biggy or SDA | No Growth_SDA |  |  |  |  | TM 869
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 509 | J&J_Pathogens | Cetrimide | Cetrimide | No Growth |  |  |  |  | TM 869
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 524 | J&J Fungal | Fungal | Fungal | 0 | CFU/g | FDA,0 | 0 | 200 | TM 869
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 509 | J&J_Pathogens | MacConkey | MacConkey | No Growth |  |  |  |  | TM 869
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 509 | J&J_Pathogens | MSA or VJ | MSA or VJ | No Growth_MSA |  |  |  |  | TM 869
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 723 | J&J TSB Controls - TSA | Negative Control | Negative Control | No Growth TSB |  |  |  |  | TM 869
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | NegativeControl | No Growth TSB |  |  |  |  | TM 869
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 509 | J&J_Pathogens | TSA | TSA | No Growth |  |  |  |  | TM 869
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | TSA JJ | 0 | cfu/mL | FDA,0 | 0 | 500 | TM 869
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 517 | J&J_Enrichment | TSB | TSB | Not Turbid |  |  |  |  | TM 869
7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | 7C4OP065 | Chem_NY_ Sodium Lauroamphoacetate 30% | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH
7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | 7C4OP065 | Chem_NY_ Sodium Lauroamphoacetate 30% | NY-Analytical | 398 | pH | pH (1 in 20 solution) | pH (1 in 20 solution) | 9.5 |  | FDA,1 | 9 | 9.5 | TM 200
7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | 7C4OP065 | Chem_NY_ Sodium Lauroamphoacetate 30% | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Clean Liquid (Done) |  |  |  |  | TM 208
7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | 7C4OP065 | Chem_NY_ Sodium Lauroamphoacetate 30% | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to light yellow (Done) |  |  |  |  | TM 208
7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | 7C4OP065 | Chem_NY_ Sodium Lauroamphoacetate 30% | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | 7C4OP065 | Chem_NY_ Sodium Lauroamphoacetate 30% | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208
7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | 7C4OP065 | Chem_NY_ Sodium Lauroamphoacetate 30% | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411
7C4OP230/MASTER - Microbiology - J&J Chemical Testing) | 7C4OP230 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 826 | Acid Value | Acid Value | Acid Value | 0.01 |  | FDA,2 | 0 | 0.2 | J&J TMWW78011
7C4OP230/MASTER - Microbiology - J&J Chemical Testing) | 7C4OP230 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 831 | Density @ 20° C | Density | Density | 0.8916 |  | FDA,4 | 0.89 | 0.893 | J&J TMWW78012
7C4OP284/Chem_NY_Grape Fruit Extract | 7C4OP284 | Chem_NY_Grape Fruit Extract | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification
7C4OP284/Chem_NY_Grape Fruit Extract | 7C4OP284 | Chem_NY_Grape Fruit Extract | NY-Analytical | 390 | Identification | Identification (FTIR) | Identification (FTIR) | Matches that of the Reference Standard (Done) |  |  |  |  | J&J TMWW78013
7C4OP284/Chem_NY_Grape Fruit Extract | 7C4OP284 | Chem_NY_Grape Fruit Extract | NY-Analytical | 792 | Chemical Inspection | Appearance (J&J TMWW78010) | Appearance | Clear to Pale Yellow Liquid (Done) |  |  |  |  | JJTM7252
7C4OP284/Chem_NY_Grape Fruit Extract | 7C4OP284 | Chem_NY_Grape Fruit Extract | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | JJTM7252
7C4OP601/MASTER - Microbiology - J&J Chemical Testing) | 7C4OP601 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 748 | Visual Inspection | Appearance of Packaging | Appearance of Packaging | Meets specifications (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7C4OQ028/Chem_NY_Isohexadecane | 7C4OQ028 | Chem_NY_Isohexadecane | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Refractive Index at 20°C
7C4OQ028/Chem_NY_Isohexadecane | 7C4OQ028 | Chem_NY_Isohexadecane | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.439 |  | FDA,1 | 1.4 | 1.5 | TM 204
7C4OQ028/Chem_NY_Isohexadecane | 7C4OQ028 | Chem_NY_Isohexadecane | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411
7C4OQ028/Chem_NY_Isohexadecane | 7C4OQ028 | Chem_NY_Isohexadecane | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7C4OQ028/Chem_NY_Isohexadecane | 7C4OQ028 | Chem_NY_Isohexadecane | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7C4OQ028/Chem_NY_Isohexadecane | 7C4OQ028 | Chem_NY_Isohexadecane | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7C4OQ028/Chem_NY_Isohexadecane | 7C4OQ028 | Chem_NY_Isohexadecane | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7C4OQ595/Chem_81A | 7C4OQ595 | Chem_81A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7C4OQ595/Chem_81A | 7C4OQ595 | Chem_81A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Free flowing granules or powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7C4OQ595/Chem_81A | 7C4OQ595 | Chem_81A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7C4OQ595/Chem_81A | 7C4OQ595 | Chem_81A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7C4OQ595/Chem_81A | 7C4OQ595 | Chem_81A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7C4OQ595/Chem_81A | 7C4OQ595 | Chem_81A | NY-Analytical | 390 | Identification | Identification B | Identification B | Pass (Done) |  |  |  |  | TM400/408b/421/422/USP
7C4OQ595/Chem_81A | 7C4OQ595 | Chem_81A | NY-Analytical | 390 | Identification | Identification for Potassium (USP <191>) | Identification for Potassium | Pass (Done) |  |  |  |  | TM400/408b/421/422/USP
7C4OX110/Chem_84A | 7C4OX110 | Chem_84A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C
7C4OX110/Chem_84A | 7C4OX110 | Chem_84A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.8529 |  | FDA,3 | 0.845 | 0.858 | TM 202
7C4OX110/Chem_84A | 7C4OX110 | Chem_84A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7C4OX110/Chem_84A | 7C4OX110 | Chem_84A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7C4OX110/Chem_84A | 7C4OX110 | Chem_84A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7C4OX110/Chem_84A | 7C4OX110 | Chem_84A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7D4AA580/Chem_NY_Monoethanolamine | 7D4AA580 | Chem_NY_Monoethanolamine | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification
7D4AA580/Chem_NY_Monoethanolamine | 7D4AA580 | Chem_NY_Monoethanolamine | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411
7D4AA580/Chem_NY_Monoethanolamine | 7D4AA580 | Chem_NY_Monoethanolamine | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, Colorless, Viscous Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7D4AA580/Chem_NY_Monoethanolamine | 7D4AA580 | Chem_NY_Monoethanolamine | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7D4AA580/Chem_NY_Monoethanolamine | 7D4AA580 | Chem_NY_Monoethanolamine | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7D4AA580/Chem_NY_Monoethanolamine | 7D4AA580 | Chem_NY_Monoethanolamine | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Ammoniacal (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7D4OG169/Chem_88A | 7D4OG169 | Chem_88A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C
7D4OG169/Chem_88A | 7D4OG169 | Chem_88A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.068 |  | FDA,4 | 1.0594 | 1.0794 | TM 202
7D4OG169/Chem_88A | 7D4OG169 | Chem_88A | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.4755 |  | FDA,4 | 1.4695 | 1.4895 | TM 204
7D4OG169/Chem_88A | 7D4OG169 | Chem_88A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7D4OG169/Chem_88A | 7D4OG169 | Chem_88A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7D4OG169/Chem_88A | 7D4OG169 | Chem_88A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Pale straw (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7D4OG169/Chem_88A | 7D4OG169 | Chem_88A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7D4OP065/Chem_NY_ Glydant Plus Liquid | 7D4OP065 | Chem_NY_ Glydant Plus Liquid | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH
7D4OP065/Chem_NY_ Glydant Plus Liquid | 7D4OP065 | Chem_NY_ Glydant Plus Liquid | NY-Analytical | 398 | pH | pH | pH | 7.66 |  | FDA,1 | 6 | 8 | TM 200
7D4OP065/Chem_NY_ Glydant Plus Liquid | 7D4OP065 | Chem_NY_ Glydant Plus Liquid | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM 208
7D4OP065/Chem_NY_ Glydant Plus Liquid | 7D4OP065 | Chem_NY_ Glydant Plus Liquid | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Water-White to light yellow (Done) |  |  |  |  | TM 208
7D4OP065/Chem_NY_ Glydant Plus Liquid | 7D4OP065 | Chem_NY_ Glydant Plus Liquid | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7D4OP065/Chem_NY_ Glydant Plus Liquid | 7D4OP065 | Chem_NY_ Glydant Plus Liquid | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) |  |  |  |  | TM 208
7D4OP065/Chem_NY_ Glydant Plus Liquid | 7D4OP065 | Chem_NY_ Glydant Plus Liquid | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411
7D4OP230/Chem_NY_Zemea_01 | 7D4OP230 | Chem_NY_Zemea_01 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Appearance
7D4OP230/Chem_NY_Zemea_01 | 7D4OP230 | Chem_NY_Zemea_01 | NY-Analytical | 397 | Appearance | Physical Properties | Physical Properties | A Clear, Colorless Liquid (Done) |  |  |  |  | TM 208
7D4OP230/Chem_NY_Zemea_01 | 7D4OP230 | Chem_NY_Zemea_01 | NY-Analytical | 390 | Identification | Identification A | Identification A | Passes test. (Done) |  |  |  |  | USP/NF
7D4OP230/Chem_NY_Zemea_01 | 7D4OP230 | Chem_NY_Zemea_01 | NY-Analytical | 390 | Identification | Identification B | Identification B | Passes test. (Done) |  |  |  |  | USP/NF
7D4OP601/Chem_91A | 7D4OP601 | Chem_91A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7D4OP601/Chem_91A | 7D4OP601 | Chem_91A | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear, yellow liquid. (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7D4OP601/Chem_91A | 7D4OP601 | Chem_91A | NY-Analytical | 792 | Chemical Inspection | Odor (J&J TM-003188) | Odor | Meets Odor Specification or of Characteristic Odor. [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7D4OP601/Chem_91A | 7D4OP601 | Chem_91A | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7D4OP601/Chem_91A | 7D4OP601 | Chem_91A | NY-Analytical | 831 | Density @ 20° C | Density | Density | 1.1193 |  | FDA,3 | 1.111 | 1.131 | J&J TMWW78012
7D4OP601/Chem_91A | 7D4OP601 | Chem_91A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013
7D4OP601/Chem_91A | 7D4OP601 | Chem_91A | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.53 |  | FDA,3 | 1.526 | 1.538 | J&J TMWW78015
7D4OP601/Chem_91A | 7D4OP601 | Chem_91A | NY-Analytical | 844 | % Benzoic Acid Phenoxyethanol & Dehydroacetic Acid | Benzoic Acid | Benzoic Acid | 12.13 | % | FDA,1 | 11 | 13 | TM 882
7D4OP601/Chem_91A | 7D4OP601 | Chem_91A | NY-Analytical | 844 | % Benzoic Acid Phenoxyethanol & Dehydroacetic Acid | Dehydroacetic Acid | Dehydroacetic Acid | 6.91 | % | FDA,1 | 6.5 | 7.5 | TM 882
7D4OP601/Chem_91A | 7D4OP601 | Chem_91A | NY-Analytical | 844 | % Benzoic Acid Phenoxyethanol & Dehydroacetic Acid | Phenoxyethanol | Phenoxyethanol | 76.51 | % | FDA,1 | 71 | 77 | TM 882
7D4OQ028/Chem_NY_Dicaprylyl Ether | 7D4OQ028 | Chem_NY_Dicaprylyl Ether | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Refractive Index at 20°C
7D4OQ028/Chem_NY_Dicaprylyl Ether | 7D4OQ028 | Chem_NY_Dicaprylyl Ether | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.431 |  | FDA,3 | 1.431 | 1.435 | TM 204
7D4OQ028/Chem_NY_Dicaprylyl Ether | 7D4OQ028 | Chem_NY_Dicaprylyl Ether | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411
7D4OQ028/Chem_NY_Dicaprylyl Ether | 7D4OQ028 | Chem_NY_Dicaprylyl Ether | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7D4OQ028/Chem_NY_Dicaprylyl Ether | 7D4OQ028 | Chem_NY_Dicaprylyl Ether | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7D4OQ028/Chem_NY_Dicaprylyl Ether | 7D4OQ028 | Chem_NY_Dicaprylyl Ether | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7D4OQ028/Chem_NY_Dicaprylyl Ether | 7D4OQ028 | Chem_NY_Dicaprylyl Ether | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7D4OQ209/Chem_92A | 7D4OQ209 | Chem_92A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C
7D4OQ209/Chem_92A | 7D4OQ209 | Chem_92A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.0512 |  | FDA,4 | 1.0399 | 1.0599 | TM 202
7D4OQ209/Chem_92A | 7D4OQ209 | Chem_92A | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.511 |  | FDA,4 | 1.507 | 1.5272 | TM 204
7D4OQ209/Chem_92A | 7D4OQ209 | Chem_92A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7D4OQ209/Chem_92A | 7D4OQ209 | Chem_92A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Pale Yellow Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7D4OQ209/Chem_92A | 7D4OQ209 | Chem_92A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7D4OQ209/Chem_92A | 7D4OQ209 | Chem_92A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Baby Powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7D4OQ214/Chem_93A | 7D4OQ214 | Chem_93A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White to Off-White |  |  |  |  | Color (TM 208)
7D4OQ214/Chem_93A | 7D4OQ214 | Chem_93A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH
7D4OQ214/Chem_93A | 7D4OQ214 | Chem_93A | NY-Analytical | 398 | pH | pH (5% Solution) | pH (5% Solution) | 4.72 |  | FDA,1 | 4 | 6 | TM 200
7D4OQ214/Chem_93A | 7D4OQ214 | Chem_93A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7D4OQ214/Chem_93A | 7D4OQ214 | Chem_93A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7D4OQ214/Chem_93A | 7D4OQ214 | Chem_93A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7D4OQ214/Chem_93A | 7D4OQ214 | Chem_93A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7D4OQ967/Chem_94A | 7D4OQ967 | Chem_94A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification
7D4OQ967/Chem_94A | 7D4OQ967 | Chem_94A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7D4OQ967/Chem_94A | 7D4OQ967 | Chem_94A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear viscous liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7D4OQ967/Chem_94A | 7D4OQ967 | Chem_94A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Yellow to amber (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7D4OQ967/Chem_94A | 7D4OQ967 | Chem_94A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7E4OD232/Chem_97A | 7E4OD232 | Chem_97A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification
7E4OD232/Chem_97A | 7E4OD232 | Chem_97A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7E4OD232/Chem_97A | 7E4OD232 | Chem_97A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Granular Solid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7E4OD232/Chem_97A | 7E4OD232 | Chem_97A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7E4OD232/Chem_97A | 7E4OD232 | Chem_97A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7E4OD232/Chem_97A | 7E4OD232 | Chem_97A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7E4OP065/Chem_NY_Benzyl Alcohol | 7E4OP065 | Chem_NY_Benzyl Alcohol | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7E4OP065/Chem_NY_Benzyl Alcohol | 7E4OP065 | Chem_NY_Benzyl Alcohol | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Colorless Liquid (Done) |  |  |  |  | TM 208
7E4OP065/Chem_NY_Benzyl Alcohol | 7E4OP065 | Chem_NY_Benzyl Alcohol | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM 208
7E4OP065/Chem_NY_Benzyl Alcohol | 7E4OP065 | Chem_NY_Benzyl Alcohol | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7E4OP065/Chem_NY_Benzyl Alcohol | 7E4OP065 | Chem_NY_Benzyl Alcohol | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208
7E4OP065/Chem_NY_Benzyl Alcohol | 7E4OP065 | Chem_NY_Benzyl Alcohol | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411
7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | 7E4OP203 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH
7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | 7E4OP203 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Analytical | 398 | pH | pH (10% w/v aqueous solution) | pH (10% w/v aqueous solution) | 6.93 |  | FDA,1 | 6.5 | 7.5 | TM 200
7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | 7E4OP203 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | 7E4OP203 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Microbiology | 867 | Celsis_Rapid | Bacteria Count (Obeserved) | Bacteria Count (Obeserved) | 0 | cfu | FDA,0 |  | 100 | TM 603
7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | 7E4OP203 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Microbiology | 867 | Celsis_Rapid | Bacteria Count (Reported) | Bacteria Count (Reported) | 0 | CFU/g | FDA,0 |  | 100 | TM 603
7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | 7E4OP203 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Microbiology | 867 | Celsis_Rapid | Celsis_ Result | Celsis Result | Negative ATP (Done) |  |  |  |  | TM 603
7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | 7E4OP203 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Microbiology | 867 | Celsis_Rapid | Mold Count (Observed) | Mold Count (Observed) | 0 | cfu | FDA,0 |  |  | TM 603
7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | 7E4OP203 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Microbiology | 867 | Celsis_Rapid | Mold Count (Reported) | Mold Count (Reported) | 0 | CFU/g | FDA,0 |  |  | TM 603
7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | 7E4OP203 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear viscous liquid, free of foreign matter. (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | 7E4OP203 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Analytical | 792 | Chemical Inspection | Color (TM 230) | Color | 2 |  | FDA,0 |  | 6 | TM208/220/230/USP/Kao LM
7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | 7E4OP203 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Similar in character and intensity to standard. Herbal Odor (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Chem_NY_Lexfeel Natural | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Chem_NY_Lexfeel Natural | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear colorless to pale yellow liquid. (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Chem_NY_Lexfeel Natural | NY-Analytical | 792 | Chemical Inspection | Colour (J&J TM7377) | Colour | 2 |  | FDA,0 |  | 50 | J&J TMWW78010/J&J TM7252
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Chem_NY_Lexfeel Natural | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Chem_NY_Lexfeel Natural | NY-Analytical | 826 | Acid Value | Acid Value | Acid Value | 0.03 | mg KOH/g | FDA,1 |  | 0.1 | J&J TMWW78011
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Chem_NY_Lexfeel Natural | NY-Analytical | 826 | Acid Value | Acid Value Final | Acid Value Final | 0.03 | mg KOH/g | FDA,1 |  | 0.1 | J&J TMWW78011
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Chem_NY_Lexfeel Natural | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Chem_NY_Lexfeel Natural | NY-Analytical | 801 | % Water | Moisture Content | Moisture Content | 0.04 | % | FDA,1 |  | 0.1 | J&J TMWW78017
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Chem_NY_Lexfeel Natural | NY-Microbiology | 524 | J&J Fungal | Fungal | Fungal | 0 | CFU/g | FDA,0 | 0 | 200 | TM 869
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Chem_NY_Lexfeel Natural | NY-Microbiology | 701 | J&J TSA Controls | Negative Control | Negative Control | No Growth TSA |  |  |  |  | TM 869
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Chem_NY_Lexfeel Natural | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | NegativeControl | No Growth TSB |  |  |  |  | TM 869
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Chem_NY_Lexfeel Natural | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | TSA JJ | 0 | cfu/mL | FDA,0 | 0 | 2000 | TM 869
7E4OP601/Chem_102A | 7E4OP601 | Chem_102A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Assay
7E4OP601/Chem_102A | 7E4OP601 | Chem_102A | NY-Analytical | 796 | Assay | Assay | Assay | 10.16 | % | FDA,2 | 9 | 11 | J&J TM-002528
7E4OP601/Chem_102A | 7E4OP601 | Chem_102A | NY-Analytical | 796 | Assay | Assay Final | Assay Final | 10.16 | % | FDA,1 | 9 | 11 | J&J TM-002528
7E4OP601/Chem_102A | 7E4OP601 | Chem_102A | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear, yellow liquid. (Done) |  |  |  |  | J&J TMWW78010
7E4OP601/Chem_102A | 7E4OP601 | Chem_102A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013
7E4OP601/Chem_102A | 7E4OP601 | Chem_102A | NY-Analytical | 801 | % Water | Water Content | Water Content | 0.4 | % | FDA,2 |  | 1 | J&J TMWW78017
7E4OP806/Chem_58B | 7E4OP806 | Chem_58B | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7E4OP806/Chem_58B | 7E4OP806 | Chem_58B | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear colorless liquid free from suspended materials (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7E4OP806/Chem_58B | 7E4OP806 | Chem_58B | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7E4OP806/Chem_58B | 7E4OP806 | Chem_58B | NY-Analytical | 831 | Density @ 20° C | Density | Density | 0.9221 |  | FDA,3 | 0.92 | 0.923 | J&J TMWW78012
7E4OP806/Chem_58B | 7E4OP806 | Chem_58B | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013
7E4OP806/Chem_58B | 7E4OP806 | Chem_58B | NY-Analytical | 801 | % Water | Moisture Content (10 grams of sample) | Moisture Content (10 grams of sample) | 0.06 | % | FDA,3 |  | 0.1 | J&J TMWW78017
7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | 7E4OQ028 | Chem_NY_Ethylhexyl Hydroxystearate | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C
7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | 7E4OQ028 | Chem_NY_Ethylhexyl Hydroxystearate | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.8964 |  | FDA,4 | 0.892 | 0.898 | TM 202
7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | 7E4OQ028 | Chem_NY_Ethylhexyl Hydroxystearate | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.4555 |  | FDA,4 | 1.4555 | 1.458 | TM 204
7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | 7E4OQ028 | Chem_NY_Ethylhexyl Hydroxystearate | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411
7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | 7E4OQ028 | Chem_NY_Ethylhexyl Hydroxystearate | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear to slightly hazy liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | 7E4OQ028 | Chem_NY_Ethylhexyl Hydroxystearate | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | 7E4OQ028 | Chem_NY_Ethylhexyl Hydroxystearate | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | 7E4OQ028 | Chem_NY_Ethylhexyl Hydroxystearate | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Slightly Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7F4BP869/Chem_NY_Citric Acid Anhydrous - USP | 7F4BP869 | Chem_NY_Citric Acid Anhydrous - USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA
7F4BP869/Chem_NY_Citric Acid Anhydrous - USP | 7F4BP869 | Chem_NY_Citric Acid Anhydrous - USP | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Crystals/crystalline powder (Done) |  |  |  |  | TM 208
7F4BP869/Chem_NY_Citric Acid Anhydrous - USP | 7F4BP869 | Chem_NY_Citric Acid Anhydrous - USP | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to white (Done) |  |  |  |  | TM 208
7F4BP869/Chem_NY_Citric Acid Anhydrous - USP | 7F4BP869 | Chem_NY_Citric Acid Anhydrous - USP | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7F4BP869/Chem_NY_Citric Acid Anhydrous - USP | 7F4BP869 | Chem_NY_Citric Acid Anhydrous - USP | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless (Done) |  |  |  |  | TM 208
7F4BP869/Chem_NY_Citric Acid Anhydrous - USP | 7F4BP869 | Chem_NY_Citric Acid Anhydrous - USP | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | 7F4OA160 | Chem_NY_PEG-40 Hydrogenated Castor Oil | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH
7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | 7F4OA160 | Chem_NY_PEG-40 Hydrogenated Castor Oil | NY-Analytical | 398 | pH | pH (10% in water) | pH (10% in water) | 6.99 |  | FDA,1 | 6 | 7 | TM 200
7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | 7F4OA160 | Chem_NY_PEG-40 Hydrogenated Castor Oil | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | 7F4OA160 | Chem_NY_PEG-40 Hydrogenated Castor Oil | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear viscous liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | 7F4OA160 | Chem_NY_PEG-40 Hydrogenated Castor Oil | NY-Analytical | 792 | Chemical Inspection | Color / Appearance (TM 208) | Color / Appearance | White to yellowish soft or flowing paste. Close Match to Standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | 7F4OA160 | Chem_NY_PEG-40 Hydrogenated Castor Oil | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | 7F4OA160 | Chem_NY_PEG-40 Hydrogenated Castor Oil | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Close match to standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7F4OD110/Chem_106A | 7F4OD110 | Chem_106A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH
7F4OD110/Chem_106A | 7F4OD110 | Chem_106A | NY-Analytical | 398 | pH | pH (1% Solution) | pH (1% Solution) | 10.6 |  | FDA,1 | 10.5 | 11.5 | TM 200
7F4OD110/Chem_106A | 7F4OD110 | Chem_106A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7F4OD110/Chem_106A | 7F4OD110 | Chem_106A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Crystalline Powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7F4OD110/Chem_106A | 7F4OD110 | Chem_106A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7F4OD110/Chem_106A | 7F4OD110 | Chem_106A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7F4OD110/Chem_106A | 7F4OD110 | Chem_106A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | Chem_NY_Arlasilk Phospholipid PTC | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | color (Gardner)
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | Chem_NY_Arlasilk Phospholipid PTC | NY-Analytical | 944 | color (Gardner) | color (Gardner) | color (Gardner) | 3 |  | FDA,0 |  | 4 | J&J TM701
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | Chem_NY_Arlasilk Phospholipid PTC | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear amber liquid. (Done) |  |  |  |  | J&J TMWW78010
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | Chem_NY_Arlasilk Phospholipid PTC | NY-Analytical | 792 | Chemical Inspection | Odor (J&J TM-003188) | Odor | Passes Test (Done) |  |  |  |  | J&J TMWW78010
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | Chem_NY_Arlasilk Phospholipid PTC | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | Chem_NY_Arlasilk Phospholipid PTC | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 509 | J&J_Pathogens | Biggy or SDA | Biggy or SDA | No Growth_SDA |  |  |  |  | TM 869
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 509 | J&J_Pathogens | Cetrimide | Cetrimide | No Growth |  |  |  |  | TM 869
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 524 | J&J Fungal | Fungal | Fungal | 0 | CFU/g | FDA,0 | 0 | 200 | TM 869
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 509 | J&J_Pathogens | MacConkey | MacConkey | No Growth |  |  |  |  | TM 869
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 509 | J&J_Pathogens | MSA or VJ | MSA or VJ | No Growth_MSA |  |  |  |  | TM 869
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 700 | J&J SDA Controls | Negative Control | Negative Control | No Growth SDA |  |  |  |  | TM 869
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | NegativeControl | No Growth TSB |  |  |  |  | TM 869
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 509 | J&J_Pathogens | TSA | TSA | No Growth |  |  |  |  | TM 869
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | TSA JJ | 0 | cfu/mL | FDA,0 | 0 | 2000 | TM 869
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 517 | J&J_Enrichment | TSB | TSB | Turbid |  |  |  |  | TM 869
7F4OP601/Chem_109A | 7F4OP601 | Chem_109A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Odor
7F4OP601/MASTER - Microbiology - J&J Chemical Testing) | 7F4OP601 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 396 | Odor | Odor | Odor | Passes Test (Done) |  |  |  |  | J&J TM-003188
7F4OP601/Chem_109A | 7F4OP601 | Chem_109A | NY-Analytical | 396 | Odor | Odor | Odor | Passes Test (Done) |  |  |  |  | J&J TM-003188
7F4OP601/MASTER - Microbiology - J&J Chemical Testing) | 7F4OP601 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear colorless liquid with oily consistency (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7F4OP601/Chem_109A | 7F4OP601 | Chem_109A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear colorless liquid with oily consistency (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7F4OP601/Chem_109A | 7F4OP601 | Chem_109A | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7F4OP601/Chem_109A | 7F4OP601 | Chem_109A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013
7F4OP806/Chem_NY_Phenoxyethanol_7F4OP806 | 7F4OP806 | Chem_NY_Phenoxyethanol_7F4OP806 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7F4OP806/Chem_NY_Phenoxyethanol_7F4OP806 | 7F4OP806 | Chem_NY_Phenoxyethanol_7F4OP806 | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear, thin colorless liquid (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7F4OP806/Chem_NY_Phenoxyethanol_7F4OP806 | 7F4OP806 | Chem_NY_Phenoxyethanol_7F4OP806 | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7F4OP806/Chem_NY_Phenoxyethanol_7F4OP806 | 7F4OP806 | Chem_NY_Phenoxyethanol_7F4OP806 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013
7F4OP806/Chem_NY_Phenoxyethanol_7F4OP806 | 7F4OP806 | Chem_NY_Phenoxyethanol_7F4OP806 | NY-Analytical | 905 | Phenol | Phenol | Phenol | 21.5 | ppm | FDA,0 |  | 25 | USP/NF
7F4OP806/Chem_NY_Phenoxyethanol_7F4OP806 | 7F4OP806 | Chem_NY_Phenoxyethanol_7F4OP806 | NY-Analytical | 986 | Phenoxyethanol Content | Phenoxyethanol Assay | Phenoxyethanol Assay | 100.9 | % | FDA,1 | 98 | 102 | USP/NF
7F4OP806/Chem_NY_Phenoxyethanol_7F4OP806 | 7F4OP806 | Chem_NY_Phenoxyethanol_7F4OP806 | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C | 1.1095 |  | FDA,3 | 1.105 | 1.11 | USP/NF
7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | 7F4OQ028 | Chem_NY_Octyldodecyl Stearoyl Stearate | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C
7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | 7F4OQ028 | Chem_NY_Octyldodecyl Stearoyl Stearate | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.8747 |  | FDA,4 | 0.865 | 0.885 | TM 202
7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | 7F4OQ028 | Chem_NY_Octyldodecyl Stearoyl Stearate | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.4585 |  | FDA,4 | 1.455 | 1.46 | TM 204
7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | 7F4OQ028 | Chem_NY_Octyldodecyl Stearoyl Stearate | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411
7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | 7F4OQ028 | Chem_NY_Octyldodecyl Stearoyl Stearate | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | 7F4OQ028 | Chem_NY_Octyldodecyl Stearoyl Stearate | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Light Yellow Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | 7F4OQ028 | Chem_NY_Octyldodecyl Stearoyl Stearate | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | 7F4OQ028 | Chem_NY_Octyldodecyl Stearoyl Stearate | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Mild Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | 7G4OP027 | Chem_NY_SD Alcohol 40-B 200 Proof_ADM | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | 7G4OP027 | Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Identity of Denatonium Benzoate
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | 7G4OP027 | Chem_NY_SD Alcohol 40-B 200 Proof_ADM | NY-Analytical | 824 | Identity of Denatonium Benzoate | Identification | Identification | Passes test with 5cm Quartz (Done) |  |  |  |  | J&J TM4028
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | 7G4OP027 | Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | NY-Analytical | 824 | Identity of Denatonium Benzoate | Identification | Identification | [5 empty] | J&J TM4028
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | 7G4OP027 | Chem_NY_SD Alcohol 40-B 200 Proof_ADM | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear colorless liquid. (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | 7G4OP027 | Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | [5 empty] | J&J TMWW78010/J&J TM7252
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | 7G4OP027 | Chem_NY_SD Alcohol 40-B 200 Proof_ADM | NY-Analytical | 792 | Chemical Inspection | Odor (J&J TM-003188) | Odor | Passes Test (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | 7G4OP027 | Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | NY-Analytical | 792 | Chemical Inspection | Odor (J&J TM-003188) | Odor | [5 empty] | J&J TMWW78010/J&J TM7252
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | 7G4OP027 | Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | [5 empty] | J&J TMWW78010/J&J TM7252
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | 7G4OP027 | Chem_NY_SD Alcohol 40-B 200 Proof_ADM | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | 7G4OP027 | Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C |  |  | FDA,4 | 0.7905 | 0.792 | J&J TMWW78012
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | 7G4OP027 | Chem_NY_SD Alcohol 40-B 200 Proof_ADM | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C |  |  | FDA,4 | 0.7905 | 0.792 | J&J TMWW78012
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | 7G4OP027 | Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | NY-Analytical | 822 | Specific Gravity @ 15.56°C/ 60° F | Specific Gravity | Specific Gravity at 15.56°C | 0.7942 |  | FDA,4 |  | 0.7952 | J&J TMWW78012
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | 7G4OP027 | Chem_NY_SD Alcohol 40-B 200 Proof_ADM | NY-Analytical | 822 | Specific Gravity @ 15.56°C/ 60° F | Specific Gravity | Specific Gravity at 15.56°C | 0.7941 |  | FDA,4 |  | 0.7952 | J&J TMWW78012
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | 7G4OP027 | Chem_NY_SD Alcohol 40-B 200 Proof_ADM | NY-Analytical | 390 | Identification | IR Identification | IR Identification | [5 empty] | J&J TMWW78013
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | 7G4OP027 | Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | NY-Analytical | 390 | Identification | IR Identification | IR Identification | [5 empty] | J&J TMWW78013
7G4OP308/Chem_111A | 7G4OP308 | Chem_111A | NY-Analytical | 792 | Chemical Inspection | Odor | Odor | Matches |  |  |  |  | TM 208
7G4OP308/Chem_111A | 7G4OP308 | Chem_111A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7G4OP308/Chem_111A | 7G4OP308 | Chem_111A | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Colorless to Light Yellow Liquid, Free of Particulate Matter (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7G4OP308/Chem_111A | 7G4OP308 | Chem_111A | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7G4OP308/Chem_111A | 7G4OP308 | Chem_111A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.9464 |  | FDA,3 | 0.942 | 0.952 | J&J TMWW78012
7G4OP308/Chem_111A | 7G4OP308 | Chem_111A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013
7G4OP308/Chem_111A | 7G4OP308 | Chem_111A | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.4632 |  | FDA,3 | 1.463 | 1.473 | J&J TMWW78015
7G4OP601/MASTER - Microbiology - J&J Chemical Testing) | 7G4OP601 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 828 | Iodine Value | Iodine Value | Iodine Value | 0.6 | g I2/100g | FDA,1 |  | 2 | J&J TM-003029
7G4OP601/MASTER - Microbiology - J&J Chemical Testing) | 7G4OP601 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 828 | Iodine Value | Iodine Value Final | Iodine Value Final | 0.6 | g I2/100g | FDA,1 |  | 2 | J&J TM-003029
7G4OP601/MASTER - Microbiology - J&J Chemical Testing) | 7G4OP601 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 792 | Chemical Inspection | Color (J&J TM-000839) | Color | 1 |  | FDA,1 |  | 1 | J&J TMWW78010/J&J TM7252
7G4OP601/MASTER - Microbiology - J&J Chemical Testing) | 7G4OP601 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 826 | Acid Value | Acid Value Final | Acid Value Final | 0.1 | mg KOH/g | FDA,1 |  | 2 | J&J TMWW78011
7G4OP601/MASTER - Microbiology - J&J Chemical Testing) | 7G4OP601 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.451 |  | FDA,4 | 1.451 | 1.453 | J&J TMWW78015
7G4OP601/MASTER - Microbiology - J&J Chemical Testing) | 7G4OP601 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 810 | Saponification Value | Saponification Value | Saponification Value | 108.1 | mg KOH/g | FDA,1 | 100 | 115 | J&J TMWW78016
7G4OP601/MASTER - Microbiology - J&J Chemical Testing) | 7G4OP601 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 810 | Saponification Value | Saponification Value Final | Saponification Value Final | 108.1 | mg KOH/g | FDA,1 | 100 | 115 | J&J TMWW78016
7G4OQ028/Chem_Aminomethylpropanol | 7G4OQ028 | Chem_Aminomethylpropanol | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification
7G4OQ028/Chem_Aminomethylpropanol | 7G4OQ028 | Chem_Aminomethylpropanol | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7G4OQ028/Chem_Aminomethylpropanol | 7G4OQ028 | Chem_Aminomethylpropanol | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7G4OQ028/Chem_Aminomethylpropanol | 7G4OQ028 | Chem_Aminomethylpropanol | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7G4OQ028/Chem_Aminomethylpropanol | 7G4OQ028 | Chem_Aminomethylpropanol | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7G4OQ028/Chem_Aminomethylpropanol | 7G4OQ028 | Chem_Aminomethylpropanol | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Chem_NY_Feverfew CMP-02A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Chem_NY_Feverfew CMP-02A | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Red Brown Liquid (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Chem_NY_Feverfew CMP-02A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Chem_NY_Feverfew CMP-02A | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Chem_NY_Feverfew CMP-02A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 509 | J&J_Pathogens | Biggy or SDA | Biggy or SDA | No Growth_SDA |  |  |  |  | TM 869
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 509 | J&J_Pathogens | Cetrimide | Cetrimide | No Growth |  |  |  |  | TM 869
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 524 | J&J Fungal | Fungal | Fungal | 0 | CFU/g | FDA,0 | 0 | 200 | TM 869
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 509 | J&J_Pathogens | MacConkey | MacConkey | No Growth |  |  |  |  | TM 869
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 509 | J&J_Pathogens | MSA or VJ | MSA or VJ | No Growth_MSA |  |  |  |  | TM 869
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 700 | J&J SDA Controls | Negative Control | Negative Control | No Growth SDA |  |  |  |  | TM 869
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | NegativeControl | No Growth TSB |  |  |  |  | TM 869
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 509 | J&J_Pathogens | TSA | TSA | No Growth |  |  |  |  | TM 869
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | TSA JJ | 0 | cfu/mL | FDA,0 | 0 | 200 | TM 869
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 517 | J&J_Enrichment | TSB | TSB | Not Turbid |  |  |  |  | TM 869
7H4OP601/Chem_NY_DUB PTO | 7H4OP601 | Chem_NY_DUB PTO | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7H4OP601/Chem_NY_DUB PTO | 7H4OP601 | Chem_NY_DUB PTO | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Colorless to pale yellow liquid (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7H4OP601/Chem_NY_DUB PTO | 7H4OP601 | Chem_NY_DUB PTO | NY-Analytical | 792 | Chemical Inspection | Colour (J&J TM7377) | Colour | 3 |  | FDA,0 |  | 60 | J&J TMWW78010/J&J TM7252
7H4OP601/Chem_NY_DUB PTO | 7H4OP601 | Chem_NY_DUB PTO | NY-Analytical | 792 | Chemical Inspection | Odor (J&J TM-003188) | Odor | Passes Test (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7H4OP601/Chem_NY_DUB PTO | 7H4OP601 | Chem_NY_DUB PTO | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7H4OP601/Chem_NY_DUB PTO | 7H4OP601 | Chem_NY_DUB PTO | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013
7H4OP601/Chem_NY_DUB PTO | 7H4OP601 | Chem_NY_DUB PTO | NY-Microbiology | 524 | J&J Fungal | Fungal | Fungal |  | CFU/g | FDA,0 | 0 | 200 | TM 869
7H4OP601/J&J Chemical CHE DUB PTO | 7H4OP601 | J&J Chemical CHE DUB PTO | NY-Microbiology | 524 | J&J Fungal | Fungal | Fungal |  | CFU/g | FDA,0 | 0 | 200 | TM 869
7H4OP601/Chem_NY_DUB PTO | 7H4OP601 | Chem_NY_DUB PTO | NY-Microbiology | 701 | J&J TSA Controls | Negative Control | Negative Control | [5 empty] | TM 869
7H4OP601/J&J Chemical CHE DUB PTO | 7H4OP601 | J&J Chemical CHE DUB PTO | NY-Microbiology | 701 | J&J TSA Controls | Negative Control | Negative Control | [5 empty] | TM 869
7H4OP601/Chem_NY_DUB PTO | 7H4OP601 | Chem_NY_DUB PTO | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | NegativeControl | No Growth TSB |  |  |  |  | TM 869
7H4OP601/J&J Chemical CHE DUB PTO | 7H4OP601 | J&J Chemical CHE DUB PTO | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | NegativeControl | [5 empty] | TM 869
7H4OP601/Chem_NY_DUB PTO | 7H4OP601 | Chem_NY_DUB PTO | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | TSA JJ | 0 | cfu/mL | FDA,0 | 0 | 2000 | TM 869
7H4OP601/J&J Chemical CHE DUB PTO | 7H4OP601 | J&J Chemical CHE DUB PTO | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | TSA JJ |  | cfu/mL | FDA,0 | 0 | 2000 | TM 869
7H4OP806/Chem_NY_Kelp & Sea Salt Fragrance | 7H4OP806 | Chem_NY_Kelp & Sea Salt Fragrance | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | GC Profile
7H4OP806/Chem_NY_Kelp & Sea Salt Fragrance | 7H4OP806 | Chem_NY_Kelp & Sea Salt Fragrance | NY-Analytical | 839 | GC Profile | GC Profile | GC Profile | To match standard. Not required when performing IR test. (Done) |  |  |  |  | J&J TM-003030
7H4OP806/Chem_NY_Kelp & Sea Salt Fragrance | 7H4OP806 | Chem_NY_Kelp & Sea Salt Fragrance | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear yellow to dark yellow liquid (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7H4OP806/Chem_NY_Kelp & Sea Salt Fragrance | 7H4OP806 | Chem_NY_Kelp & Sea Salt Fragrance | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7H4OP806/Chem_NY_Kelp & Sea Salt Fragrance | 7H4OP806 | Chem_NY_Kelp & Sea Salt Fragrance | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C | 0.9621 |  | FDA,3 | 0.957 | 0.967 | J&J TMWW78012
7H4OP806/Chem_NY_Kelp & Sea Salt Fragrance | 7H4OP806 | Chem_NY_Kelp & Sea Salt Fragrance | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013
7H4OP806/Chem_NY_Kelp & Sea Salt Fragrance | 7H4OP806 | Chem_NY_Kelp & Sea Salt Fragrance | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.458 |  | FDA,4 | 1.456 | 1.466 | J&J TMWW78015
7J4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7J4OP027 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 792 | Chemical Inspection | Solubility | Solubility | Passes Test Soluble in any proportion of water (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7J4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7J4OP027 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 447 | pH as is | pH | pH | 5.59 |  | FDA,1 | 4 | 6.5 | TM WW78014
7J4OP485/Chem_NY_Butylene Glycol_7J4OP485 | 7J4OP485 | Chem_NY_Butylene Glycol_7J4OP485 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7J4OP485/Chem_NY_Butylene Glycol_7J4OP485 | 7J4OP485 | Chem_NY_Butylene Glycol_7J4OP485 | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Colorless viscous liquid (Done) |  |  |  |  | J&J TM-003188
7J4OP485/Chem_NY_Butylene Glycol_7J4OP485 | 7J4OP485 | Chem_NY_Butylene Glycol_7J4OP485 | NY-Analytical | 792 | Chemical Inspection | Colour (J&J TMP-552) | Colour | 1 |  | FDA,0 |  | 10 | J&J TM-003188
7J4OP485/Chem_NY_Butylene Glycol_7J4OP485 | 7J4OP485 | Chem_NY_Butylene Glycol_7J4OP485 | NY-Analytical | 792 | Chemical Inspection | Odor (J&J TM-003188) | Odor | Passes Test (Done) |  |  |  |  | J&J TM-003188
7J4OP485/Chem_NY_Butylene Glycol_7J4OP485 | 7J4OP485 | Chem_NY_Butylene Glycol_7J4OP485 | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TM-003188
7J4OP485/Chem_NY_Butylene Glycol_7J4OP485 | 7J4OP485 | Chem_NY_Butylene Glycol_7J4OP485 | NY-Analytical | 795 | Acidity | Acidity as Acetic Acid(10ml burette 0.05ml subdiv) | Acidity as Acetic Acid(10ml burette 0.05ml subdiv) | 0.002 | % | FDA,3 |  | 0.005 | J&J TM-003589
7J4OP485/Chem_NY_Butylene Glycol_7J4OP485 | 7J4OP485 | Chem_NY_Butylene Glycol_7J4OP485 | NY-Analytical | 796 | Assay | Assay | Assay | 99.82 | % | FDA,2 | 99.5 |  | J&J TM4075
7J4OP485/Chem_NY_Butylene Glycol_7J4OP485 | 7J4OP485 | Chem_NY_Butylene Glycol_7J4OP485 | NY-Analytical | 796 | Assay | Assay Final | Assay Final | 99.82 | % | FDA,2 | 99.5 |  | J&J TM4075
7J4OP485/Chem_NY_Butylene Glycol_7J4OP485 | 7J4OP485 | Chem_NY_Butylene Glycol_7J4OP485 | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C | 1.006 |  | FDA,3 | 1.004 | 1.006 | J&J TMWW78012
7J4OP485/Chem_NY_Butylene Glycol_7J4OP485 | 7J4OP485 | Chem_NY_Butylene Glycol_7J4OP485 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013
7J4OP485/Chem_NY_Butylene Glycol_7J4OP485 | 7J4OP485 | Chem_NY_Butylene Glycol_7J4OP485 | NY-Analytical | 801 | % Water | Water Content | Water Content | 0.02 | % | FDA,1 |  | 0.5 | J&J TMWW78017
7K4OP027/Chem_116A | 7K4OP027 | Chem_116A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7K4OP027/Chem_116A | 7K4OP027 | Chem_116A | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Pale yellow to yellow liquid (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7K4OP027/Chem_116A | 7K4OP027 | Chem_116A | NY-Analytical | 792 | Chemical Inspection | Odor (CoA) | Odor (CoA) | Certified on CoA as Meets Odor Specification or of Characteristic Odor. (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7K4OP027/Chem_116A | 7K4OP027 | Chem_116A | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7K4OP027/Chem_116A | 7K4OP027 | Chem_116A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Colorless to white needle crystal or granule (Done) |  |  |  |  | JJTMWW78010 / JJTM7252
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | JJTMWW78010 / JJTM7252
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 809 | Melting Range | Melting Range - End | Melting Range - End | 42.5 | C | FDA,0 | 41 | 44 | USP
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 809 | Melting Range | Melting Range - Start | Melting Range - Start | 41.2 | C | FDA,0 | 41 | 44 | USP
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 797 | Nonvolatile Residue | Nonvolatile Residue | Nonvolatile Residue | 0.02 | % w/v | FDA,2 |  | 0.05 | USP
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 863 | Specific Rotation | Specific Rotation | Specific Rotation | -49.9 | ° | FDA,0 | -51 | -45 | USP
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 796 | Assay | Assay | Assay | 100.4 | % | FDA,1 | 98 | 102 | USP/NF
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 796 | Assay | Assay Final | Assay Final | 100.4 | % | FDA,1 | 98 | 102 | USP/NF
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 455 | Identification A | Identification A | Identification A | Passes Test (Done) |  |  |  |  | USP/NF
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 456 | Identification B | Identification B | Identification B | Passes Test (Done) |  |  |  |  | USP/NF
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 964 | Individual Impurity | Iso-Menthone | Iso-Menthone | 0.3 | % | FDA,1 |  | 0.3 | USP/NF
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 964 | Individual Impurity | Iso-Pulegol | Iso-Pulegol | 0.14 | % | FDA,1 |  | 0.3 | USP/NF
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 964 | Individual Impurity | Menthone | Menthone | 0.04 | % | FDA,1 |  | 0.3 | USP/NF
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 964 | Individual Impurity | Neo-Menthol | Neo-Menthol | 0.18 | % | FDA,1 |  | 0.3 | USP/NF
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 964 | Individual Impurity | Pulegone | Pulegone | 0.08 | % | FDA,1 |  | 0.3 | USP/NF
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 965 | Total Impurities | Total Impurities | Total Impurities | 0.53 | % | FDA,1 |  | 2 | USP/NF
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Chem_NY_Menthol USP | NY-Analytical | 964 | Individual Impurity | Unknown impurities | Unknown impurities | 0.05 | % | FDA,1 |  | 0.1 | USP/NF
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 801 | % Water | Water | Water | 0.09 | % | FDA,1 |  | 5 | APAM 6.018
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, syrupy liquid, free of foreign matter. (Done) |  |  |  |  | See Analyte for Method #
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) |  |  |  |  | See Analyte for Method #
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Similar in character and intensity to standard, practically odorless (Done) |  |  |  |  | See Analyte for Method #
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.273 |  | FDA,3 | 1.249 |  | TM 202
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Chem_NY_Glycerin - RI-0420 | NY-Microbiology | 867 | Celsis_Rapid | Bacteria Count (Obeserved) | Bacteria Count (Obeserved) | 0 | cfu | FDA,0 |  | 100 | TM 603
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Chem_NY_Glycerin - RI-0420 | NY-Microbiology | 867 | Celsis_Rapid | Bacteria Count (Reported) | Bacteria Count (Reported) | 0 | CFU/g | FDA,0 |  | 100 | TM 603
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Chem_NY_Glycerin - RI-0420 | NY-Microbiology | 867 | Celsis_Rapid | Celsis_ Result | Celsis Result | Negative ATP (Done) |  |  |  |  | TM 603
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Chem_NY_Glycerin - RI-0420 | NY-Microbiology | 867 | Celsis_Rapid | Mold Count (Observed) | Mold Count (Observed) | 0 | cfu | FDA,0 |  |  | TM 603
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Chem_NY_Glycerin - RI-0420 | NY-Microbiology | 867 | Celsis_Rapid | Mold Count (Reported) | Mold Count (Reported) | 0 | CFU/g | FDA,0 |  |  | TM 603
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 825 | Limit of Ethylene Glycol and Diethylene Glycol | Diethylene Glycol | Diethylene Glycol | 0 | % | FDA,1 |  | 0.1 | USP
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 825 | Limit of Ethylene Glycol and Diethylene Glycol | Ethylene Glycol | Ethylene Glycol | 0 | % | FDA,1 |  | 0.1 | USP
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 825 | Limit of Ethylene Glycol and Diethylene Glycol | Individual Impurity | Individual Impurity | 0.01 | % | FDA,1 |  | 0.1 | USP
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 825 | Limit of Ethylene Glycol and Diethylene Glycol | Total Impurities | Total Impurities | 0.02 | % | FDA,0 |  | 1 | USP
7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | 7L4OP485 | Chem_NY_Polysorbate 20_7L4OP485 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Odor
7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | 7L4OP485 | Chem_NY_Polysorbate 20_7L4OP485 | NY-Analytical | 396 | Odor | Odor | Odor | Passes Test (Done) |  |  |  |  | J&J TM-003188
7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | 7L4OP485 | Chem_NY_Polysorbate 20_7L4OP485 | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | A pale yellow to an amber colored liquid, free of foreign matter. (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | 7L4OP485 | Chem_NY_Polysorbate 20_7L4OP485 | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | 7L4OP485 | Chem_NY_Polysorbate 20_7L4OP485 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013
7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | 7L4OP485 | Chem_NY_Polysorbate 20_7L4OP485 | NY-Microbiology | 524 | J&J Fungal | Fungal | Fungal | 0 | CFU/g | FDA,0 | 0 | 200 | TM 869
7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | 7L4OP485 | Chem_NY_Polysorbate 20_7L4OP485 | NY-Microbiology | 723 | J&J TSB Controls - TSA | Negative Control | Negative Control | No Growth TSB |  |  |  |  | TM 869
7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | 7L4OP485 | Chem_NY_Polysorbate 20_7L4OP485 | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | NegativeControl | No Growth TSB |  |  |  |  | TM 869
7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | 7L4OP485 | Chem_NY_Polysorbate 20_7L4OP485 | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | TSA JJ | 0 | CFU/g | FDA,0 | 0 | 2000 | TM 869
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear colorless liquid. (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 796 | Assay | Assay | Assay | 48.2 | % | FDA,1 | 47.5 | 52.5 | USP/NF
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 796 | Assay | Assay Final | Assay Final | 48.2 | % | FDA,1 | 47.5 | 52.5 | USP/NF
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 846 | Ratio of Alkyl Compounds | C12 + C14 Content of Total | C12 + C14 Content of Total | 95.7 | % | FDA,1 | 70 |  | USP/NF
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 846 | Ratio of Alkyl Compounds | C12 Content of Total | C12 Content of Total | 56.6 | % | FDA,1 | 40 |  | USP/NF
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 846 | Ratio of Alkyl Compounds | C14 Content of Total | C14 Content of Total | 39.1 | % | FDA,1 | 20 |  | USP/NF
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 390 | Identification | Identification A | Identification A | Passes test. (Done) |  |  |  |  | USP/NF
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 390 | Identification | Identification B | Identification B | Passes test. (Done) |  |  |  |  | USP/NF
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 390 | Identification | Identification C | Identification C | Passes test. (Done) |  |  |  |  | USP/NF
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 390 | Identification | Identification D | Identification D | Passes test. (Done) |  |  |  |  | USP/NF
7O4AA580/Chem_NY_Ethylene Glycol Monohexyl Ether | 7O4AA580 | Chem_NY_Ethylene Glycol Monohexyl Ether | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification
7O4AA580/Chem_NY_Ethylene Glycol Monohexyl Ether | 7O4AA580 | Chem_NY_Ethylene Glycol Monohexyl Ether | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411
7O4AA580/Chem_NY_Ethylene Glycol Monohexyl Ether | 7O4AA580 | Chem_NY_Ethylene Glycol Monohexyl Ether | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, Colorless, Viscous Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4AA580/Chem_NY_Ethylene Glycol Monohexyl Ether | 7O4AA580 | Chem_NY_Ethylene Glycol Monohexyl Ether | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4AA580/Chem_NY_Ethylene Glycol Monohexyl Ether | 7O4AA580 | Chem_NY_Ethylene Glycol Monohexyl Ether | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4AA580/Chem_NY_Ethylene Glycol Monohexyl Ether | 7O4AA580 | Chem_NY_Ethylene Glycol Monohexyl Ether | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Pungent (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4AQ084/Chem_162A | 7O4AQ084 | Chem_162A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7O4AQ084/Chem_162A | 7O4AQ084 | Chem_162A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4AQ084/Chem_162A | 7O4AQ084 | Chem_162A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Clear/Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4AQ084/Chem_162A | 7O4AQ084 | Chem_162A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4AQ084/Chem_162A | 7O4AQ084 | Chem_162A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Light Amine (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4BJ131/Chem_NY_Acticide CBM-2 | 7O4BJ131 | Chem_NY_Acticide CBM-2 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification
7O4BJ131/Chem_NY_Acticide CBM-2 | 7O4BJ131 | Chem_NY_Acticide CBM-2 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7O4BJ131/Chem_NY_Acticide CBM-2 | 7O4BJ131 | Chem_NY_Acticide CBM-2 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Off-white - greenish dispersion (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4BJ131/Chem_NY_Acticide CBM-2 | 7O4BJ131 | Chem_NY_Acticide CBM-2 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4BJ131/Chem_NY_Acticide CBM-2 | 7O4BJ131 | Chem_NY_Acticide CBM-2 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4BJ131/Chem_NY_Acticide CBM-2 | 7O4BJ131 | Chem_NY_Acticide CBM-2 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4BM301/Chem_126A | 7O4BM301 | Chem_126A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification
7O4BM301/Chem_126A | 7O4BM301 | Chem_126A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7O4BM301/Chem_126A | 7O4BM301 | Chem_126A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4BM301/Chem_126A | 7O4BM301 | Chem_126A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to Amber (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4BM301/Chem_126A | 7O4BM301 | Chem_126A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4BM301/Chem_126A | 7O4BM301 | Chem_126A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4DH247/Chem_201A | 7O4DH247 | Chem_201A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C
7O4DH247/Chem_201A | 7O4DH247 | Chem_201A | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C | 1.2883 |  | FDA,3 | 1.27 | 1.31 | TM 202
7O4DH247/Chem_201A | 7O4DH247 | Chem_201A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear solution, free from solids (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4DH247/Chem_201A | 7O4DH247 | Chem_201A | NY-Analytical | 792 | Chemical Inspection | Color (TM 230) | Color | 2 |  | FDA,1 |  | 2 | TM208/220/230/USP/Kao LM
7O4DH247/Chem_201A | 7O4DH247 | Chem_201A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4DH247/Chem_201A | 7O4DH247 | Chem_201A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Bleach/chlorine (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4DX220/Chem_NY_Neutralizing Agent | 7O4DX220 | Chem_NY_Neutralizing Agent | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C
7O4DX220/Chem_NY_Neutralizing Agent | 7O4DX220 | Chem_NY_Neutralizing Agent | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.0999 |  | FDA,4 | 1.088 | 1.118 | TM 202
7O4DX220/Chem_NY_Neutralizing Agent | 7O4DX220 | Chem_NY_Neutralizing Agent | NY-Analytical | 451 | Refractive Index | Refractive Index | Refractive Index at 25°C | 1.4685 |  | FDA,4 | 1.46 | 1.48 | TM 204
7O4DX220/Chem_NY_Neutralizing Agent | 7O4DX220 | Chem_NY_Neutralizing Agent | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7O4DX220/Chem_NY_Neutralizing Agent | 7O4DX220 | Chem_NY_Neutralizing Agent | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4DX220/Chem_NY_Neutralizing Agent | 7O4DX220 | Chem_NY_Neutralizing Agent | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Light Yellow Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4DX220/Chem_NY_Neutralizing Agent | 7O4DX220 | Chem_NY_Neutralizing Agent | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4DX220/Chem_NY_Neutralizing Agent | 7O4DX220 | Chem_NY_Neutralizing Agent | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic – comparable to standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OA234/Chem_131A | 7O4OA234 | Chem_131A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH as is
7O4OA234/Chem_131A | 7O4OA234 | Chem_131A | NY-Analytical | 447 | pH as is | pH | pH | 3.22 |  | FDA,2 | 3 | 3.3 | TM 200
7O4OA234/Chem_131A | 7O4OA234 | Chem_131A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7O4OA234/Chem_131A | 7O4OA234 | Chem_131A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Opaque, Thin Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OA234/Chem_131A | 7O4OA234 | Chem_131A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OA234/Chem_131A | 7O4OA234 | Chem_131A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OA234/Chem_131A | 7O4OA234 | Chem_131A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OA347/Chem_NY_Lauryl Glucoside | 7O4OA347 | Chem_NY_Lauryl Glucoside | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7O4OA347/Chem_NY_Lauryl Glucoside | 7O4OA347 | Chem_NY_Lauryl Glucoside | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Match standard (Done) |  |  |  |  | TM 208
7O4OA347/Chem_NY_Lauryl Glucoside | 7O4OA347 | Chem_NY_Lauryl Glucoside | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Match standard (Done) |  |  |  |  | TM 208
7O4OA347/Chem_NY_Lauryl Glucoside | 7O4OA347 | Chem_NY_Lauryl Glucoside | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7O4OA347/Chem_NY_Lauryl Glucoside | 7O4OA347 | Chem_NY_Lauryl Glucoside | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208
7O4OA347/Chem_NY_Lauryl Glucoside | 7O4OA347 | Chem_NY_Lauryl Glucoside | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Comparison of scans of sample vs standard (Done) |  |  |  |  | TM 411
7O4OA401/Chem_132A | 7O4OA401 | Chem_132A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C
7O4OA401/Chem_132A | 7O4OA401 | Chem_132A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.0225 |  | FDA,4 | 1.0115 | 1.0315 | TM 202
7O4OA401/Chem_132A | 7O4OA401 | Chem_132A | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.485 |  | FDA,3 | 1.4794 | 1.4994 | TM 204
7O4OA401/Chem_132A | 7O4OA401 | Chem_132A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7O4OA401/Chem_132A | 7O4OA401 | Chem_132A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OA401/Chem_132A | 7O4OA401 | Chem_132A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to light yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OA401/Chem_132A | 7O4OA401 | Chem_132A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OA401/Chem_132A | 7O4OA401 | Chem_132A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Fresh Floral (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OA490/Chem_133A | 7O4OA490 | Chem_133A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C
7O4OA490/Chem_133A | 7O4OA490 | Chem_133A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.9182 |  | FDA,3 | 0.909 | 0.929 | TM 202
7O4OA490/Chem_133A | 7O4OA490 | Chem_133A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7O4OA490/Chem_133A | 7O4OA490 | Chem_133A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OA490/Chem_133A | 7O4OA490 | Chem_133A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OA490/Chem_133A | 7O4OA490 | Chem_133A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OA490/Chem_133A | 7O4OA490 | Chem_133A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Ether (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OA580/Chem_NY_Glycacil SG | 7O4OA580 | Chem_NY_Glycacil SG | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification
7O4OA580/Chem_NY_Glycacil SG | 7O4OA580 | Chem_NY_Glycacil SG | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7O4OA580/Chem_NY_Glycacil SG | 7O4OA580 | Chem_NY_Glycacil SG | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OA580/Chem_NY_Glycacil SG | 7O4OA580 | Chem_NY_Glycacil SG | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to Amber (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OA580/Chem_NY_Glycacil SG | 7O4OA580 | Chem_NY_Glycacil SG | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OA580/Chem_NY_Glycacil SG | 7O4OA580 | Chem_NY_Glycacil SG | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic Match Standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | 7O4OA634 | Chem_NY_Water Soluble Flowering Herbs Fragrance | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C
7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | 7O4OA634 | Chem_NY_Water Soluble Flowering Herbs Fragrance | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.0629 |  | FDA,4 | 1.0531 | 1.0831 | TM 202
7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | 7O4OA634 | Chem_NY_Water Soluble Flowering Herbs Fragrance | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.47 |  | FDA,4 | 1.4655 | 1.4855 | TM 204
7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | 7O4OA634 | Chem_NY_Water Soluble Flowering Herbs Fragrance | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | 7O4OA634 | Chem_NY_Water Soluble Flowering Herbs Fragrance | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | 7O4OA634 | Chem_NY_Water Soluble Flowering Herbs Fragrance | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Pale Yellow to Straw (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | 7O4OA634 | Chem_NY_Water Soluble Flowering Herbs Fragrance | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | 7O4OA634 | Chem_NY_Water Soluble Flowering Herbs Fragrance | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Floral (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OA709/Chem_NY_Cucumber Extract # 11-19-3 | 7O4OA709 | Chem_NY_Cucumber Extract # 11-19-3 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Specific Gravity at 25°C
7O4OA709/Chem_NY_Cucumber Extract # 11-19-3 | 7O4OA709 | Chem_NY_Cucumber Extract # 11-19-3 | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C |  |  | FDA,3 | 1.034 | 1.044 | TM 202
7O4OA709/Chem_NY_Cucumber Extract # 11-19-3 | 7O4OA709 | Chem_NY_Cucumber Extract # 11-19-3 | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.425 |  | FDA,3 | 1.42 | 1.43 | TM 204
7O4OA709/Chem_NY_Cucumber Extract # 11-19-3 | 7O4OA709 | Chem_NY_Cucumber Extract # 11-19-3 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7O4OA709/Chem_NY_Cucumber Extract # 11-19-3 | 7O4OA709 | Chem_NY_Cucumber Extract # 11-19-3 | NY-Analytical | 792 | Chemical Inspection | Description (TM 208) | Description | Clear colorless liquid having a characteristic odor (Done) |  |  |  |  | TM 208
7O4OA911/Chem_NY_Zinc Sulfate Heptahydrate | 7O4OA911 | Chem_NY_Zinc Sulfate Heptahydrate | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7O4OA911/Chem_NY_Zinc Sulfate Heptahydrate | 7O4OA911 | Chem_NY_Zinc Sulfate Heptahydrate | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Crystals/crystalline powder (Done) |  |  |  |  | TM 208
7O4OA911/Chem_NY_Zinc Sulfate Heptahydrate | 7O4OA911 | Chem_NY_Zinc Sulfate Heptahydrate | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to white (Done) |  |  |  |  | TM 208
7O4OA911/Chem_NY_Zinc Sulfate Heptahydrate | 7O4OA911 | Chem_NY_Zinc Sulfate Heptahydrate | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7O4OA911/Chem_NY_Zinc Sulfate Heptahydrate | 7O4OA911 | Chem_NY_Zinc Sulfate Heptahydrate | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208
7O4OA911/Chem_NY_Zinc Sulfate Heptahydrate | 7O4OA911 | Chem_NY_Zinc Sulfate Heptahydrate | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411
7O4OD351/Chem_NY_Benzalkonium Chloride 50% | 7O4OD351 | Chem_NY_Benzalkonium Chloride 50% | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [5 empty] | Appearance
7O4OD351/Chem_NY_Benzalkonium Chloride 50% | 7O4OD351 | Chem_NY_Benzalkonium Chloride 50% | NY-Analytical | 397 | Appearance | Appearance | Appearance | Clear liquid, free from foreign material (FFFM) (Done) |  |  |  |  | TM 208
7O4OD351/Chem_NY_Benzalkonium Chloride 50% | 7O4OD351 | Chem_NY_Benzalkonium Chloride 50% | NY-Analytical | 938 | Ethanol | Ethanol | Ethanol |  | % v/v | FDA,2 | 5.9 | 6.5 | TM 323
7O4OD351/Chem_NY_Benzalkonium Chloride 50% | 7O4OD351 | Chem_NY_Benzalkonium Chloride 50% | NY-Analytical | 938 | Ethanol | Ethanol Final | Ethanol Final |  | % v/v | FDA,2 | 5.9 | 6.5 | TM 323
7O4OD351/Chem_NY_Benzalkonium Chloride 50% | 7O4OD351 | Chem_NY_Benzalkonium Chloride 50% | NY-Analytical | 455 | Identification A | Identification A - Chloride | Identification A - Chloride | [5 empty] | TM 4001
7O4OD351/Chem_NY_Benzalkonium Chloride 50% | 7O4OD351 | Chem_NY_Benzalkonium Chloride 50% | NY-Analytical | 456 | Identification B | Identification B - Benzalkonium | Identification B - Benzalkonium | [5 empty] | TM 4001
7O4OD351/Chem_NY_Benzalkonium Chloride 50% | 7O4OD351 | Chem_NY_Benzalkonium Chloride 50% | NY-Analytical | 1064 | Total % ADBAC | % BZK and Dual-Quat | % BZK and Dual-Quat |  | % | FDA,1 | 47.5 | 52.5 | TM 434
7O4OD363/Chem_137A | 7O4OD363 | Chem_137A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH as is
7O4OD363/Chem_137A | 7O4OD363 | Chem_137A | NY-Analytical | 447 | pH as is | pH | pH | 5.06 |  | FDA,1 | 4.5 | 5.5 | TM 200
7O4OD363/Chem_137A | 7O4OD363 | Chem_137A | NY-Analytical | 806 | Total Solids | Total Solids (2 g for 3 hr @ 105°C) | Total Solids (2 g for 3 hr @ 105°C) | 43.5 | % | FDA,1 | 41 | 45 | TM 206
7O4OD363/Chem_137A | 7O4OD363 | Chem_137A | NY-Analytical | 806 | Total Solids | Total Solids Final (2 g for 3 hr @ 105°C) | Total Solids Final (2 g for 3 hr @ 105°C) | 43.45 | % | FDA,1 | 41 | 45 | TM 206
7O4OD363/Chem_137A | 7O4OD363 | Chem_137A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7O4OD363/Chem_137A | 7O4OD363 | Chem_137A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OD363/Chem_137A | 7O4OD363 | Chem_137A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to pale yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OD419/Chem_138A | 7O4OD419 | Chem_138A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA does not meet specifications (OOS) |  |  |  |  | Identification
7O4OD419/Chem_138A | 7O4OD419 | Chem_138A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7O4OD419/Chem_138A | 7O4OD419 | Chem_138A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Semi Solid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OD419/Chem_138A | 7O4OD419 | Chem_138A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White to light yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OD419/Chem_138A | 7O4OD419 | Chem_138A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OD419/Chem_138A | 7O4OD419 | Chem_138A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Fatty (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OD748/Chem_139A | 7O4OD748 | Chem_139A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification
7O4OD748/Chem_139A | 7O4OD748 | Chem_139A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7O4OD748/Chem_139A | 7O4OD748 | Chem_139A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | White solid crystals (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OD748/Chem_139A | 7O4OD748 | Chem_139A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White Crystals (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OD748/Chem_139A | 7O4OD748 | Chem_139A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OD748/Chem_139A | 7O4OD748 | Chem_139A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | To match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OF575/Che_NY_Plantaren 2000N - UP | 7O4OF575 | Che_NY_Plantaren 2000N - UP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification
7O4OF575/Che_NY_Plantaren 2000N - UP | 7O4OF575 | Che_NY_Plantaren 2000N - UP | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7O4OF575/Che_NY_Plantaren 2000N - UP | 7O4OF575 | Che_NY_Plantaren 2000N - UP | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Hazy, Pale Yellow Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OF575/Che_NY_Plantaren 2000N - UP | 7O4OF575 | Che_NY_Plantaren 2000N - UP | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OF602/Chem_141A | 7O4OF602 | Chem_141A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification
7O4OF602/Chem_141A | 7O4OF602 | Chem_141A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7O4OF602/Chem_141A | 7O4OF602 | Chem_141A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OG169/Chem_142A | 7O4OG169 | Chem_142A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C
7O4OG169/Chem_142A | 7O4OG169 | Chem_142A | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C | 1.08 |  | FDA,3 | 1.08 | 1.09 | TM 202
7O4OG169/Chem_142A | 7O4OG169 | Chem_142A | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.435 |  | FDA,3 | 1.435 | 1.437 | TM 204
7O4OG169/Chem_142A | 7O4OG169 | Chem_142A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7O4OG169/Chem_142A | 7O4OG169 | Chem_142A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OG169/Chem_142A | 7O4OG169 | Chem_142A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OG169/Chem_142A | 7O4OG169 | Chem_142A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OH247/Chem_143A | 7O4OH247 | Chem_143A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH
7O4OH247/Chem_143A | 7O4OH247 | Chem_143A | NY-Analytical | 398 | pH | pH (1% w/v aqueous solution) | pH (1% w/v aqueous solution) | 2 |  | FDA,1 |  | 2 | TM 200
7O4OH247/Chem_143A | 7O4OH247 | Chem_143A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7O4OH247/Chem_143A | 7O4OH247 | Chem_143A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear viscous liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OH247/Chem_143A | 7O4OH247 | Chem_143A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to dark amber (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OH247/Chem_143A | 7O4OH247 | Chem_143A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OH460/MASTER - Microbiology - J&J Chemical Testing) | 7O4OH460 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 801 | % Water | Moisture Content | Moisture Content | 0.6 | % | FDA,1 |  | 1.5 | J&J TMWW78017
7O4OJ123/Chem_NY_NET-KOB | 7O4OJ123 | Chem_NY_NET-KOB | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification
7O4OJ123/Chem_NY_NET-KOB | 7O4OJ123 | Chem_NY_NET-KOB | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7O4OJ123/Chem_NY_NET-KOB | 7O4OJ123 | Chem_NY_NET-KOB | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Viscous Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OJ123/Chem_NY_NET-KOB | 7O4OJ123 | Chem_NY_NET-KOB | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to pale yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OJ123/Chem_NY_NET-KOB | 7O4OJ123 | Chem_NY_NET-KOB | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OJ123/Chem_NY_NET-KOB | 7O4OJ123 | Chem_NY_NET-KOB | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OM165/Chem_146A | 7O4OM165 | Chem_146A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification
7O4OM165/Chem_146A | 7O4OM165 | Chem_146A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7O4OM165/Chem_146A | 7O4OM165 | Chem_146A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, slightly oily viscous liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OM165/Chem_146A | 7O4OM165 | Chem_146A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Light Yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OM165/Chem_146A | 7O4OM165 | Chem_146A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OM165/Chem_146A | 7O4OM165 | Chem_146A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | To match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OM666/Chem_NY_Sodium Benzoate EDF NF/FCC | 7O4OM666 | Chem_NY_Sodium Benzoate EDF NF/FCC | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7O4OM666/Chem_NY_Sodium Benzoate EDF NF/FCC | 7O4OM666 | Chem_NY_Sodium Benzoate EDF NF/FCC | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Extruded Pellets (Done) |  |  |  |  | TM 208
7O4OM666/Chem_NY_Sodium Benzoate EDF NF/FCC | 7O4OM666 | Chem_NY_Sodium Benzoate EDF NF/FCC | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) |  |  |  |  | TM 208
7O4OM666/Chem_NY_Sodium Benzoate EDF NF/FCC | 7O4OM666 | Chem_NY_Sodium Benzoate EDF NF/FCC | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7O4OM666/Chem_NY_Sodium Benzoate EDF NF/FCC | 7O4OM666 | Chem_NY_Sodium Benzoate EDF NF/FCC | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless (Done) |  |  |  |  | TM 208
7O4OM666/Chem_NY_Sodium Benzoate EDF NF/FCC | 7O4OM666 | Chem_NY_Sodium Benzoate EDF NF/FCC | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7O4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7O4OP027 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 827 | Hydroxyl Value | Hydroxyl Value | Hydroxyl Value | 38.5 | mg KOH/g | FDA,1 | 35 | 50 | J&J TMWW7671
7O4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7O4OP027 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 827 | Hydroxyl Value | Hydroxyl Value Final | Hydroxyl Value Final | 38.5 | mg KOH/g | FDA,1 | 35 | 50 | J&J TMWW7671
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Chem_NY_Actipone White Tea | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Density @ 20° C
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Chem_NY_Actipone White Tea | NY-Analytical | 831 | Density @ 20° C | Density | Density | 1.1173 |  | FDA,3 | 1.112 | 1.122 | TM 202
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Chem_NY_Actipone White Tea | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.392 |  | FDA,3 | 1.388 | 1.398 | TM 204
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Chem_NY_Actipone White Tea | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear to slightly cloudy liquid (Done) |  |  |  |  | TM 208
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Chem_NY_Actipone White Tea | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Yellow to Dark Brown (Done) |  |  |  |  | TM 208
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Chem_NY_Actipone White Tea | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Chem_NY_Actipone White Tea | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Chem_NY_Actipone White Tea | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Chem_NY_Actipone White Tea | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram) | TAMC (Average per Gram) | 0 | CFU/g | FDA,1 |  | 100 | TM 610
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Chem_NY_Actipone White Tea | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Chem_NY_Actipone White Tea | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram) | TMYC (Average per Gram) |  | CFU/g | FDA,1 |  | 100 | TM 610
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Chem_NY_Actipone White Tea | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  | cfu | FDA,0 |  |  | TM 610
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Chem_NY_Actipone White Tea | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Negative Rods (OOS) |  |  |  |  | WI QAS 076
7O4OP075/Chem_NY_Magique Lift 02LF | 7O4OP075 | Chem_NY_Magique Lift 02LF | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA does not meet specifications (OOS) |  |  |  |  | Specific Gravity at 25°C
7O4OP075/Chem_NY_Magique Lift 02LF | 7O4OP075 | Chem_NY_Magique Lift 02LF | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.9516 |  | FDA,3 | 0.938 | 0.948 | TM 202
7O4OP075/Chem_NY_Magique Lift 02LF | 7O4OP075 | Chem_NY_Magique Lift 02LF | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.48 |  | FDA,3 | 1.474 | 1.484 | TM 204
7O4OP075/Chem_NY_Magique Lift 02LF | 7O4OP075 | Chem_NY_Magique Lift 02LF | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7O4OP075/Chem_NY_Magique Lift 02LF | 7O4OP075 | Chem_NY_Magique Lift 02LF | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OP075/Chem_NY_Magique Lift 02LF | 7O4OP075 | Chem_NY_Magique Lift 02LF | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to very light yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OP075/Chem_NY_Magique Lift 02LF | 7O4OP075 | Chem_NY_Magique Lift 02LF | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OP075/Chem_NY_Magique Lift 02LF | 7O4OP075 | Chem_NY_Magique Lift 02LF | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OP093/Chem_NY_PEG-11 | 7O4OP093 | Chem_NY_PEG-11 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA
7O4OP093/Chem_NY_PEG-11 | 7O4OP093 | Chem_NY_PEG-11 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, Light liquid (Done) |  |  |  |  | TM 208
7O4OP093/Chem_NY_PEG-11 | 7O4OP093 | Chem_NY_PEG-11 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Match standard (Done) |  |  |  |  | TM 208
7O4OP093/Chem_NY_PEG-11 | 7O4OP093 | Chem_NY_PEG-11 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7O4OP093/Chem_NY_PEG-11 | 7O4OP093 | Chem_NY_PEG-11 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208
7O4OP093/Chem_NY_PEG-11 | 7O4OP093 | Chem_NY_PEG-11 | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411
7O4OP113/Chem_NY_Robertet The Matcha I | 7O4OP113 | Chem_NY_Robertet The Matcha I | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C
7O4OP113/Chem_NY_Robertet The Matcha I | 7O4OP113 | Chem_NY_Robertet The Matcha I | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C | 0.9652 |  | FDA,3 | 0.96 | 0.97 | TM 202
7O4OP113/Chem_NY_Robertet The Matcha I | 7O4OP113 | Chem_NY_Robertet The Matcha I | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.47 |  | FDA,3 | 1.469 | 1.479 | TM 204
7O4OP113/Chem_NY_Robertet The Matcha I | 7O4OP113 | Chem_NY_Robertet The Matcha I | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Oily liquid (Done) |  |  |  |  | TM 208
7O4OP113/Chem_NY_Robertet The Matcha I | 7O4OP113 | Chem_NY_Robertet The Matcha I | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Pale Yellow – Light Yellow (Done) |  |  |  |  | TM 208
7O4OP113/Chem_NY_Robertet The Matcha I | 7O4OP113 | Chem_NY_Robertet The Matcha I | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7O4OP113/Chem_NY_Robertet The Matcha I | 7O4OP113 | Chem_NY_Robertet The Matcha I | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208
7O4OP113/Chem_NY_Robertet The Matcha I | 7O4OP113 | Chem_NY_Robertet The Matcha I | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411
7O4OP197/Chem_NY_Emulgade Smart/MB_01 | 7O4OP197 | Chem_NY_Emulgade Smart/MB_01 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH
7O4OP197/Chem_NY_Emulgade Smart/MB_01 | 7O4OP197 | Chem_NY_Emulgade Smart/MB_01 | NY-Analytical | 398 | pH | pH | pH | 3 |  | FDA,1 | 3 | 3.5 | TM 200
7O4OP197/Chem_NY_Emulgade Smart/MB_01 | 7O4OP197 | Chem_NY_Emulgade Smart/MB_01 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Match standard (Done) |  |  |  |  | TM 208
7O4OP197/Chem_NY_Emulgade Smart/MB_01 | 7O4OP197 | Chem_NY_Emulgade Smart/MB_01 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White Liquid (Done) |  |  |  |  | TM 208
7O4OP197/Chem_NY_Emulgade Smart/MB_01 | 7O4OP197 | Chem_NY_Emulgade Smart/MB_01 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7O4OP197/Chem_NY_Emulgade Smart/MB_01 | 7O4OP197 | Chem_NY_Emulgade Smart/MB_01 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208
7O4OP197/Chem_NY_Emulgade Smart/MB_01 | 7O4OP197 | Chem_NY_Emulgade Smart/MB_01 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | An opalescent, pale yellow to pale brown viscous liquid. (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Analytical | 792 | Chemical Inspection | Odor (J&J TM-003188) | Odor | Passes Test (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C | 1.1158 |  | FDA,3 | 1.1 | 1.14 | J&J TMWW78012
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Microbiology | 509 | J&J_Pathogens | Biggy or SDA | Biggy or SDA | No Growth_SDA |  |  |  |  | TM 869
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Microbiology | 509 | J&J_Pathogens | Cetrimide | Cetrimide | No Growth |  |  |  |  | TM 869
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Microbiology | 524 | J&J Fungal | Fungal | Fungal | 0 | CFU/g | FDA,0 | 0 | 200 | TM 869
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Microbiology | 509 | J&J_Pathogens | MacConkey | MacConkey | No Growth |  |  |  |  | TM 869
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Microbiology | 509 | J&J_Pathogens | MSA or VJ | MSA or VJ | No Growth_MSA |  |  |  |  | TM 869
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Microbiology | 700 | J&J SDA Controls | Negative Control | Negative Control | No Growth SDA |  |  |  |  | TM 869
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | NegativeControl | No Growth TSB |  |  |  |  | TM 869
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Microbiology | 509 | J&J_Pathogens | TSA | TSA | No Growth |  |  |  |  | TM 869
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | TSA JJ | 0 | cfu/mL | FDA,0 | 0 | 2000 | TM 869
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Microbiology | 517 | J&J_Enrichment | TSB | TSB | Not Turbid |  |  |  |  | TM 869
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Analytical | 447 | pH as is | pH | pH | 4.87 |  | FDA,1 | 4.5 | 5.5 | TM WW78014
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Analytical | 917 | Total Sugar | Total Sugar | Total Sugar | 8.2 | % | FDA,1 | 6.5 | 9 | TM-002568
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Chem_NY_Hydromanil | NY-Analytical | 916 | Reducing Sugars | Reducing Sugars | Reducing Sugars | 3.3 | % | FDA,1 | 2.5 | 3.6 | TM-002569
7O4OP253/Chem_NY_BTC 1010 | 7O4OP253 | Chem_NY_BTC 1010 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH
7O4OP253/Chem_NY_BTC 1010 | 7O4OP253 | Chem_NY_BTC 1010 | NY-Analytical | 398 | pH | pH (10% in water) | pH (10% in water) | 7.19 |  | FDA,1 | 6 | 9 | TM 200
7O4OP253/Chem_NY_BTC 1010 | 7O4OP253 | Chem_NY_BTC 1010 | NY-Analytical | 796 | Assay | Assay | Assay | 50.8 | % | FDA,0 | 50 | 55 | TM 301/434
7O4OP253/Chem_NY_BTC 1010 | 7O4OP253 | Chem_NY_BTC 1010 | NY-Analytical | 796 | Assay | Assay Final | Assay Final | 50.85 | % | FDA,0 | 50 | 55 | TM 301/434
7O4OP284/Chem_155A | 7O4OP284 | Chem_155A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | GC Profile
7O4OP284/Chem_155A | 7O4OP284 | Chem_155A | NY-Analytical | 839 | GC Profile | GC Profile | GC Profile | To match standard. Not required when performing IR test. (Done) |  |  |  |  | J&J TM-001862
7O4OP284/Chem_155A | 7O4OP284 | Chem_155A | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear, yellow liquid. (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7O4OP284/Chem_155A | 7O4OP284 | Chem_155A | NY-Analytical | 792 | Chemical Inspection | Odor (J&J TM-003188) | Odor | Supplier certifies on C of A that odor matches standard (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7O4OP284/Chem_155A | 7O4OP284 | Chem_155A | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7O4OP284/Chem_155A | 7O4OP284 | Chem_155A | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C | 0.9722 |  | FDA,3 | 0.963 | 0.973 | J&J TMWW78012
7O4OP284/Chem_155A | 7O4OP284 | Chem_155A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013
7O4OP284/Chem_155A | 7O4OP284 | Chem_155A | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.442 |  | FDA,4 | 1.442 | 1.452 | J&J TMWW78015
7O4OP345/Chem_157A | 7O4OP345 | Chem_157A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification
7O4OP345/Chem_157A | 7O4OP345 | Chem_157A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7O4OP345/Chem_157A | 7O4OP345 | Chem_157A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OP345/Chem_157A | 7O4OP345 | Chem_157A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OP345/Chem_157A | 7O4OP345 | Chem_157A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OP345/Chem_157A | 7O4OP345 | Chem_157A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Phenoxyethanol-like (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OP567/Chem_NY_BTC 1210-80% | 7O4OP567 | Chem_NY_BTC 1210-80% | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA
7O4OP567/Chem_NY_BTC 1210-80% | 7O4OP567 | Chem_NY_BTC 1210-80% | NY-Analytical | 398 | pH | pH (10% in water) | pH (10% in water) | 7.79 |  | FDA,1 | 6 | 8.5 | TM 200
7O4OP567/Chem_NY_BTC 1210-80% | 7O4OP567 | Chem_NY_BTC 1210-80% | NY-Analytical | 397 | Appearance | Appearance | Appearance | Colorless to slight yellow (Done) |  |  |  |  | TM 208
7O4OP567/Chem_NY_BTC 1210-80% | 7O4OP567 | Chem_NY_BTC 1210-80% | NY-Analytical | 1015 | Active (BTC 1210-80%) | Active (BTC 1210-80%) | Active (BTC 1210-80%) | 80.9 | % | FDA,1 | 80 | 82 | TM 3005
7O4OP601/Chem_107A | 7O4OP601 | Chem_107A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | GC Profile
7O4OP601/Chem_107A | 7O4OP601 | Chem_107A | NY-Analytical | 839 | GC Profile | GC Profile | GC Profile | To match standard. Not required when performing IR test. (Done) |  |  |  |  | J&J TM-003030
7O4OP601/Chem_107A | 7O4OP601 | Chem_107A | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Colorless to pale yellow clear liquid. (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7O4OP601/Chem_107A | 7O4OP601 | Chem_107A | NY-Analytical | 792 | Chemical Inspection | Odor (J&J TM-003188) | Odor | Meets Odor Specification or of Characteristic Odor. [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7O4OP601/Chem_107A | 7O4OP601 | Chem_107A | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7O4OP601/Chem_107A | 7O4OP601 | Chem_107A | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C | 0.9921 |  | FDA,3 | 0.986 | 0.996 | J&J TMWW78012
7O4OP601/Chem_107A | 7O4OP601 | Chem_107A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013
7O4OP601/Chem_107A | 7O4OP601 | Chem_107A | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.452 |  | FDA,4 | 1.452 | 1.462 | J&J TMWW78015
7O4OP638/Chem_192A | 7O4OP638 | Chem_192A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7O4OP638/Chem_192A | 7O4OP638 | Chem_192A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, colorless to light yellow liquid / White solid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OP705/Chem_NY_Cetaphil Concentrate | 7O4OP705 | Chem_NY_Cetaphil Concentrate | NY-Analytical | 1051 | Temptale Reading | Temptale Reading | Temptale Reading |  | ºC | FDA,0 | 5 | 40 | Temptale Reading
7O4OP705/Chem_NY_Cetaphil Concentrate | 7O4OP705 | Chem_NY_Cetaphil Concentrate | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH as is
7O4OP705/Chem_NY_Cetaphil Concentrate | 7O4OP705 | Chem_NY_Cetaphil Concentrate | NY-Analytical | 447 | pH as is | pH | pH | 7 |  | FDA,1 | 5.5 | 7 | TM 200
7O4OP705/Chem_NY_Cetaphil Concentrate | 7O4OP705 | Chem_NY_Cetaphil Concentrate | NY-Analytical | 747 | Viscosity | Viscosity (Spindle 3 @ 60RPM in 300 ml beaker) | Viscosity (Spindle 3 @ 60RPM in 300 ml beaker) | 915 | cps | FDA,0 | 400 | 1600 | TM 203
7O4OP705/Chem_NY_Cetaphil Concentrate | 7O4OP705 | Chem_NY_Cetaphil Concentrate | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7O4OP705/Chem_NY_Cetaphil Concentrate | 7O4OP705 | Chem_NY_Cetaphil Concentrate | NY-Analytical | 792 | Chemical Inspection | Description (TM 208) | Description | Viscous, slightly translucent, offwhite liquid: cetyl alcohol odor (Done) |  |  |  |  | TM 208
7O4OP806/Chem_NY_Carbopol 2020 | 7O4OP806 | Chem_NY_Carbopol 2020 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Loss on Drying
7O4OP806/Chem_NY_Carbopol 2020 | 7O4OP806 | Chem_NY_Carbopol 2020 | NY-Analytical | 802 | Loss on Drying | Loss on Drying (80C Vacuum 1 hour) | Loss on Drying (80C Vacuum 1 hour) | 0.49 | % | FDA,2 |  | 2 | J&J TM7221
7O4OP806/Chem_NY_Carbopol 2020 | 7O4OP806 | Chem_NY_Carbopol 2020 | NY-Analytical | 802 | Loss on Drying | Loss on Drying (80C Vacuum 1 hour) Final | Loss on Drying (80C Vacuum 1 hour) Final | 0.49 | % | FDA,2 |  | 2 | J&J TM7221
7O4OP806/Chem_NY_Carbopol 2020 | 7O4OP806 | Chem_NY_Carbopol 2020 | NY-Analytical | 747 | Viscosity | Viscosity | Viscosity | 48667 | cps | FDA,0 | 47000 | 77000 | J&J TM7423
7O4OP806/Chem_NY_Carbopol 2020 | 7O4OP806 | Chem_NY_Carbopol 2020 | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Fluffy white powder free of lumps (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7O4OP806/Chem_NY_Carbopol 2020 | 7O4OP806 | Chem_NY_Carbopol 2020 | NY-Analytical | 748 | Visual Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7O4OP806/Chem_NY_Carbopol 2020 | 7O4OP806 | Chem_NY_Carbopol 2020 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013
7O4OQ028/Chem_NY_Hexylene Glycol | 7O4OQ028 | Chem_NY_Hexylene Glycol | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis
7O4OQ028/Chem_NY_Euxyl K703 | 7O4OQ028 | Chem_NY_Euxyl K703 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Density @ 20° C
7O4OQ028/Chem_NY_Euxyl K703 | 7O4OQ028 | Chem_NY_Euxyl K703 | NY-Analytical | 831 | Density @ 20° C | Density | Density | 1.1282 |  | FDA,3 | 1.124 | 1.128 | TM 202
7O4OQ028/Chem_NY_Hexylene Glycol | 7O4OQ028 | Chem_NY_Hexylene Glycol | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C |  |  | FDA,4 | 0.92 | 0.923 | TM 202
7O4OQ028/Chem_NY_Euxyl K703 | 7O4OQ028 | Chem_NY_Euxyl K703 | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.539 |  | FDA,3 | 1.539 | 1.543 | TM 204
7O4OQ028/Chem_NY_Euxyl K703 | 7O4OQ028 | Chem_NY_Euxyl K703 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7O4OQ028/Chem_NY_Hexylene Glycol | 7O4OQ028 | Chem_NY_Hexylene Glycol | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411
7O4OQ028/Chem_NY_Hexylene Glycol | 7O4OQ028 | Chem_NY_Hexylene Glycol | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OQ028/Chem_NY_Euxyl K703 | 7O4OQ028 | Chem_NY_Euxyl K703 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OQ028/Chem_NY_Hexylene Glycol | 7O4OQ028 | Chem_NY_Hexylene Glycol | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OQ028/Chem_NY_Euxyl K703 | 7O4OQ028 | Chem_NY_Euxyl K703 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OQ028/Chem_NY_Hexylene Glycol | 7O4OQ028 | Chem_NY_Hexylene Glycol | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OQ028/Chem_NY_Euxyl K703 | 7O4OQ028 | Chem_NY_Euxyl K703 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OQ028/Chem_NY_Hexylene Glycol | 7O4OQ028 | Chem_NY_Hexylene Glycol | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OQ028/Chem_NY_Euxyl K703 | 7O4OQ028 | Chem_NY_Euxyl K703 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Phenoxyethanol-like (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OQ243/Chem_165A | 7O4OQ243 | Chem_165A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification
7O4OQ243/Chem_165A | 7O4OQ243 | Chem_165A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7O4OQ243/Chem_165A | 7O4OQ243 | Chem_165A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Fine, free-flowing powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OQ243/Chem_165A | 7O4OQ243 | Chem_165A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OQ243/Chem_165A | 7O4OQ243 | Chem_165A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OQ243/Chem_165A | 7O4OQ243 | Chem_165A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristically mild (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OQ342/Chem_166A | 7O4OQ342 | Chem_166A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C
7O4OQ342/Chem_166A | 7O4OQ342 | Chem_166A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.9558 |  | FDA,4 | 0.9455 | 0.9655 | TM 202
7O4OQ342/Chem_166A | 7O4OQ342 | Chem_166A | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.458 |  | FDA,4 | 1.4505 | 1.4705 | TM 204
7O4OQ342/Chem_166A | 7O4OQ342 | Chem_166A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, mobile liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OQ342/Chem_166A | 7O4OQ342 | Chem_166A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Very pale yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OQ342/Chem_166A | 7O4OQ342 | Chem_166A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OQ342/Chem_166A | 7O4OQ342 | Chem_166A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Citrus Fresh (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OQ437/Chem_168A | 7O4OQ437 | Chem_168A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25°C
7O4OQ437/Chem_168A | 7O4OQ437 | Chem_168A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.944 |  | FDA,4 | 0.938 | 0.9465 | TM 202
7O4OQ437/Chem_168A | 7O4OQ437 | Chem_168A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7O4OQ437/Chem_168A | 7O4OQ437 | Chem_168A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Mobile Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OQ437/Chem_168A | 7O4OQ437 | Chem_168A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OQ437/Chem_168A | 7O4OQ437 | Chem_168A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OQ437/Chem_168A | 7O4OQ437 | Chem_168A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Amine (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OQ890/Chem_170A | 7O4OQ890 | Chem_170A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification
7O4OQ890/Chem_170A | 7O4OQ890 | Chem_170A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7O4OQ890/Chem_170A | 7O4OQ890 | Chem_170A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OQ890/Chem_170A | 7O4OQ890 | Chem_170A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OQ890/Chem_170A | 7O4OQ890 | Chem_170A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OQ890/Chem_170A | 7O4OQ890 | Chem_170A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Ether-like (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 792 | Chemical Inspection | Description | Description | An almost colorless or paleyellowish liquid (Done) |  |  |  |  | TM 208
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Any individual unspecified impuritv | Any individual unspecified impuritv | 0.077 | % | FDA,1 |  | 0.1 | USP
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Ch!orhexidine oxazinone analog | Ch!orhexidine oxazinone analog | 0.042 | % | FDA,1 |  | 0.2 | USP
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Chlorhexidine | Chlorhexidine | 93.338 | % | FDA,1 |  |  | USP
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Chlorhexidine urea | Chlorhexidine urea | 0.085 | % | FDA,1 |  | 0.2 | USP
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Chlorhexidine amine | Chlorhexidine amine | 0.047 | % | FDA,1 |  | 0.3 | USP
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Chlorhexidine dimer | Chlorhexidine dimer | 0.09 | % | FDA,1 |  | 0.5 | USP
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Chlorhexidine glucityl triazine | Chlorhexidine glucityl triazine | 0.083 | % | FDA,1 |  | 0.4 | USP
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Chlorhexidine guanidine | Chlorhexidine guanidine | 0.083 | % | FDA,1 |  | 0.3 | USP
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Chlorhexidine nitrile | Chlorhexidine nitrile | 0.114 | % | FDA,1 |  | 0.4 | USP
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 390 | Identification | Identification A (USP <197k>) | Identification A | Conforms to USP (Done) |  |  |  |  | USP
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 390 | Identification | Identification B (USP <201>) | Identification B | Conforms to USP (Done) |  |  |  |  | USP
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | o-Chlorhexidine and specified unidentified impurit | o-Chlorhexidine and specified unidentified impurit | 0.164 | % | FDA,1 |  | 0.4 | USP
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Oxochlorhexidine | Oxochlorhexidine | 0.288 | % | FDA,1 |  | 0.4 | USP
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 812 | Organic Impurites | p-Chloroaniline | p-Chloroaniline | 0 | ug/mL | FDA,0 |  | 100 | USP
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 812 | Organic Impurites | p-Chloroaniline Final | p-Chloroaniline Final | 0 | ug/mL | FDA,0 |  | 100 | USP
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | p-Chlorophenvl urea | p-Chlorophenvl urea | 0 | % | FDA,1 |  | 0.2 | USP
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Any individual unspecified impuritv | RRT of Any individual unspecified impuritv | 1.13 |  | FDA,2 |  |  | USP
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Ch!orhexidine oxazinone analog | RRT of Ch!orhexidine oxazinone analog | 0.16 |  | FDA,2 |  |  | USP
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Chlorhexidine | RRT of Chlorhexidine | 1.1 |  | FDA,2 |  |  | USP
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Chlorhexidine urea | RRT of Chlorhexidine urea | 0.32 |  | FDA,2 |  |  | USP
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Chlorhexidine amine | RRT of Chlorhexidine amine | 0.21 |  | FDA,2 |  |  | USP
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Chlorhexidine dimer | RRT of Chlorhexidine dimer | 0.87 |  | FDA,2 |  |  | USP
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Chlorhexidine glucityl triazine | RRT of Chlorhexidine glucityl triazine | 0.95 |  | FDA,2 |  |  | USP
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Chlorhexidine guanidine | RRT of Chlorhexidine guanidine | 0.31 |  | FDA,2 |  |  | USP
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Chlorhexidine nitrile | RRT of Chlorhexidine nitrile | 0.52 |  | FDA,2 |  |  | USP
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of o-Chlorhexidine and specified unidentified | RRT of o-Chlorhexidine and specified unidentified | 0.9 |  | FDA,2 |  |  | USP
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Oxochlorhexidine | RRT of Oxochlorhexidine | 1.43 |  | FDA,2 |  |  | USP
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of p-Chlorophenvl urea | RRT of p-Chlorophenvl urea | 0 |  | FDA,2 |  |  | USP
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Specified unidentified impuritv 1 | RRT of Specified unidentified impuritv 1 | 0.17 |  | FDA,2 |  |  | USP
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Specified unidentified impuritv 1 | Specified unidentified impuritv 1 | 0.011 | % | FDA,1 |  | 0.2 | USP
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Total Impurities | Total Impurities | 0.983 | % | FDA,1 |  | 3 | USP
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 398 | pH | pH (Diluted 1 in 20) | pH (Diluted 1 in 20) | 5.55 |  | FDA,1 | 5.5 | 7 | USP <791>
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 796 | Assay | Assay | Assay | 20.1 | % | FDA,1 | 19 | 21 | USP-HPLC
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 796 | Assay | Assay Final | Assay Final | 20.1 | % | FDA,1 | 19 | 21 | USP-HPLC
7O4OS407/Chem_173A | 7O4OS407 | Chem_173A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.0648 |  | FDA,2 | 1.06 | 1.07 | USP/NF <841>
7O4OU181/Chem_NY_Isopropanol 70% | 7O4OU181 | Chem_NY_Isopropanol 70% | IN-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA
7O4OU181/Chem_NY_Isopropanol 70% | 7O4OU181 | Chem_NY_Isopropanol 70% | IN-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear (Done) |  |  |  |  | TM 208
7O4OU181/Chem_NY_Isopropanol 70% | 7O4OU181 | Chem_NY_Isopropanol 70% | IN-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) |  |  |  |  | TM 208
7O4OU181/Chem_NY_Isopropanol 70% | 7O4OU181 | Chem_NY_Isopropanol 70% | IN-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7O4OU181/Chem_NY_Isopropanol 70% | 7O4OU181 | Chem_NY_Isopropanol 70% | IN-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208
7O4OU181/Chem_NY_Isopropanol 70% | 7O4OU181 | Chem_NY_Isopropanol 70% | IN-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411
7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 7O4OU225 | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH
7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 7O4OU225 | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 398 | pH | pH (10% Solution) | pH (10% Solution) | 11.8 |  | FDA,1 | 10.5 | 13 | TM 200
7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 7O4OU225 | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 807 | % NaOCl | % NaOCl by weight (calc) | % NaOCl by weight | 9.9316 | % | FDA,2 | 11.5 | 14 | TM 366
7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 7O4OU225 | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 807 | % NaOCl | % NaOCl by weight Final (calc) | % NaOCl by weight Final | 10.0533 | % | FDA,2 | 11.5 | 14 | TM 366
7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 7O4OU225 | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 807 | % NaOCl | Sample Weight (g) | Sample Weight | 1.0023 | g | FDA,4 |  |  | TM 366
7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 7O4OU225 | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 807 | % NaOCl | Titrant End Volume (mL) | Titrant End Volume | 27.5 | mL | FDA,1 |  |  | TM 366
7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 7O4OU225 | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 807 | % NaOCl | Titrant Start Volume (mL) | Titrant Start Volume | 0 | mL | FDA,1 |  |  | TM 366
7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 7O4OU225 | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 807 | % NaOCl | Total Volume of Titrant (mL) | Total Volume of Titrant | 27.5 | mL | FDA,1 |  |  | TM 366
7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 7O4OU225 | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear liquid with a chlorine odor (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 7O4OU225 | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | straw-yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM
7O4OU351/Chem_NY_Glycopyrronium Tosylate Monohydrate | 7O4OU351 | Chem_NY_Glycopyrronium Tosylate Monohydrate | NY-Analytical | 1051 | Temptale Reading | Shipment Temperatutre | Shipment Temperatutre | Send Temperature data to Dermira (Done) |  |  |  |  | Shipment Temperatutre
7O4OU351/Chem_NY_Glycopyrronium Tosylate Monohydrate | 7O4OU351 | Chem_NY_Glycopyrronium Tosylate Monohydrate | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA
7O4OU351/Chem_NY_Glycopyrronium Tosylate Monohydrate | 7O4OU351 | Chem_NY_Glycopyrronium Tosylate Monohydrate | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis (ELLI) | Certificate of Analysis | COA Meets Specification (Done) |  |  |  |  | COA
7O4OU555/Chem_NY_Lidocaine Reduced testing | 7O4OU555 | Chem_NY_Lidocaine Reduced testing | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Chem_NY_Lidocaine Base USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Chem_NY_Lidocaine Base USP | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Match standard (Done) |  |  |  |  | TM 208
7O4OU555/Chem_NY_Lidocaine Reduced testing | 7O4OU555 | Chem_NY_Lidocaine Reduced testing | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Match standard (Done) |  |  |  |  | TM 208
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Chem_NY_Lidocaine Base USP | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Match standard (Done) |  |  |  |  | TM 208
7O4OU555/Chem_NY_Lidocaine Reduced testing | 7O4OU555 | Chem_NY_Lidocaine Reduced testing | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Match standard (Done) |  |  |  |  | TM 208
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Chem_NY_Lidocaine Base USP | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7O4OU555/Chem_NY_Lidocaine Reduced testing | 7O4OU555 | Chem_NY_Lidocaine Reduced testing | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) |  |  |  |  | TM 208
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Chem_NY_Lidocaine Base USP | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208
7O4OU555/Chem_NY_Lidocaine Reduced testing | 7O4OU555 | Chem_NY_Lidocaine Reduced testing | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) |  |  |  |  | TM 208
7O4OU555/Chem_NY_Lidocaine Reduced testing | 7O4OU555 | Chem_NY_Lidocaine Reduced testing | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Chem_NY_Lidocaine Base USP | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411
7O4OU555/Chem_NY_Lidocaine Reduced testing | 7O4OU555 | Chem_NY_Lidocaine Reduced testing | NY-Analytical | 796 | Assay | Assay | Assay | 100.8 | % | FDA,1 | 97.5 | 102.5 | USP
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Chem_NY_Lidocaine Base USP | NY-Analytical | 796 | Assay | Assay | Assay | 100.4 | % | FDA,1 | 97.5 | 102.5 | USP
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Chem_NY_Lidocaine Base USP | NY-Analytical | 832 | Chloride | Chloride | Chloride |  | % | FDA,4 |  | 0.0035 | USP
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Chem_NY_Lidocaine Base USP | NY-Analytical | 833 | Heavy Metals | Heavy Metals | Heavy Metals |  | ug/g | FDA,3 |  | 0.002 | USP
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Chem_NY_Lidocaine Base USP | NY-Analytical | 809 | Melting Range | Melting Range - End | Melting Range - End |  | C | FDA,1 | 66 | 69 | USP
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Chem_NY_Lidocaine Base USP | NY-Analytical | 809 | Melting Range | Melting Range - Start | Melting Range - Start |  | C | FDA,1 | 66 | 69 | USP
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Chem_NY_Lidocaine Base USP | NY-Analytical | 834 | Residue on Ignition | Residue on Ignition | Residue on Ignition |  | % | FDA,1 |  | 0.1 | USP
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Chem_NY_Lidocaine Base USP | NY-Analytical | 835 | Sulfate Content | Sulfate | Sulfate |  | % | FDA,2 |  | 0.1 | USP
7O4OX231/Chem_NY_FD&C Red 40 | 7O4OX231 | Chem_NY_FD&C Red 40 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Appearance
7O4OX231/Chem_NY_FD&C Red 40 | 7O4OX231 | Chem_NY_FD&C Red 40 | NY-Analytical | 397 | Appearance | Physical Properties | Physical Properties | Brown Powder with Purple/red tint Powder (Done) |  |  |  |  | TM 208
7O4OX231/Chem_NY_FD&C Red 40 | 7O4OX231 | Chem_NY_FD&C Red 40 | NY-Analytical | 1061 | Total Color | Spectrophotometric | Spectrophotometric | 89.4 | % | FDA,0 | 85 |  | TM 3014
7R4OP485/Chem_NY_Phenoxyethanol_7R4OP485 | 7R4OP485 | Chem_NY_Phenoxyethanol_7R4OP485 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection
7R4OP485/Chem_NY_Phenoxyethanol_7R4OP485 | 7R4OP485 | Chem_NY_Phenoxyethanol_7R4OP485 | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear, thin colorless liquid (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7R4OP485/Chem_NY_Phenoxyethanol_7R4OP485 | 7R4OP485 | Chem_NY_Phenoxyethanol_7R4OP485 | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252
7R4OP485/Chem_NY_Phenoxyethanol_7R4OP485 | 7R4OP485 | Chem_NY_Phenoxyethanol_7R4OP485 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013
7R4OP485/Chem_NY_Phenoxyethanol_7R4OP485 | 7R4OP485 | Chem_NY_Phenoxyethanol_7R4OP485 | NY-Analytical | 892 | Limit of Phenol | Phenol | Phenol | 18 | ppm | FDA,0 |  | 25 | USP/NF
7R4OP485/Chem_NY_Phenoxyethanol_7R4OP485 | 7R4OP485 | Chem_NY_Phenoxyethanol_7R4OP485 | NY-Analytical | 849 | Purity | Purity | Purity | 99.4 | % | FDA,1 | 98 | 102 | USP/NF
7R4OP485/Chem_NY_Phenoxyethanol_7R4OP485 | 7R4OP485 | Chem_NY_Phenoxyethanol_7R4OP485 | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C | 1.1077 |  | FDA,3 | 1.105 | 1.11 | USP/NF
8AA654SI/Label_A10 | 8AA654SI | Label_A10 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection
8AA654SI/Label_A10 | 8AA654SI | Label_A10 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8AA654SI/Label_NY_Skip lot testing | 8AA654SI | Label_NY_Skip lot testing | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | [5 empty] | TM 143 / Visual
8AA654SI/Label_A10 | 8AA654SI | Label_A10 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8AA654SI/Label_NY_Skip lot testing | 8AA654SI | Label_NY_Skip lot testing | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8AA654SI/Label_A10 | 8AA654SI | Label_A10 | NY-Analytical | 770 | Label Inspection | Length | Length | 1.375 | in | FDA,4 | 1.3125 | 1.4375 | TM 143 / Visual
8AA654SI/Label_NY_Skip lot testing | 8AA654SI | Label_NY_Skip lot testing | NY-Analytical | 770 | Label Inspection | Length | Length | 1.375 | in | FDA,4 | 1.3125 | 1.4375 | TM 143 / Visual
8AA654SI/Label_A10 | 8AA654SI | Label_A10 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8AA654SI/Label_NY_Skip lot testing | 8AA654SI | Label_NY_Skip lot testing | NY-Analytical | 770 | Label Inspection | Odor | Odor | [5 empty] | TM 143 / Visual
8AA654SI/Label_A10 | 8AA654SI | Label_A10 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8AA654SI/Label_NY_Skip lot testing | 8AA654SI | Label_NY_Skip lot testing | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8AA654SI/Label_A10 | 8AA654SI | Label_A10 | NY-Analytical | 770 | Label Inspection | Width | Width | 2.5625 | in | FDA,3 | 2.5 | 2.625 | TM 143 / Visual
8AA654SI/Label_NY_Skip lot testing | 8AA654SI | Label_NY_Skip lot testing | NY-Analytical | 770 | Label Inspection | Width | Width | 2.562 | in | FDA,3 | 2.5 | 2.625 | TM 143 / Visual
8AA70900/Label_NY_NP9990_2.5-4.625 | 8AA70900 | Label_NY_NP9990_2.5-4.625 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection
8AA70900/Label_NY_NP9990_2.5-4.625 | 8AA70900 | Label_NY_NP9990_2.5-4.625 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8AA70900/Label_NY_NP9990_2.5-4.625 | 8AA70900 | Label_NY_NP9990_2.5-4.625 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8AA70900/Label_NY_NP9990_2.5-4.625 | 8AA70900 | Label_NY_NP9990_2.5-4.625 | NY-Analytical | 770 | Label Inspection | Length | Length | 2.5 | in | FDA,4 | 2.4375 | 2.5625 | TM 143 / Visual
8AA70900/Label_NY_NP9990_2.5-4.625 | 8AA70900 | Label_NY_NP9990_2.5-4.625 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8AA70900/Label_NY_NP9990_2.5-4.625 | 8AA70900 | Label_NY_NP9990_2.5-4.625 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8AA70900/Label_NY_NP9990_2.5-4.625 | 8AA70900 | Label_NY_NP9990_2.5-4.625 | NY-Analytical | 770 | Label Inspection | Width | Width | 4.625 | in | FDA,4 | 4.5625 | 4.6875 | TM 143 / Visual
8AH12880/Label_NY_NP7030_9-2.625 | 8AH12880 | Label_NY_NP7030_9-2.625 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection
8AH12880/Label_NY_NP7030_9-2.625 | 8AH12880 | Label_NY_NP7030_9-2.625 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8AH12880/Label_NY_NP7030_9-2.625 | 8AH12880 | Label_NY_NP7030_9-2.625 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8AH12880/Label_NY_NP7030_9-2.625 | 8AH12880 | Label_NY_NP7030_9-2.625 | NY-Analytical | 770 | Label Inspection | Length | Length | 9 | in | FDA,4 | 8.9375 | 9.0625 | TM 143 / Visual
8AH12880/Label_NY_NP7030_9-2.625 | 8AH12880 | Label_NY_NP7030_9-2.625 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8AH12880/Label_NY_NP7030_9-2.625 | 8AH12880 | Label_NY_NP7030_9-2.625 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8AH12880/Label_NY_NP7030_9-2.625 | 8AH12880 | Label_NY_NP7030_9-2.625 | NY-Analytical | 770 | Label Inspection | Width | Width | 2.625 | in | FDA,4 | 2.5625 | 2.6875 | TM 143 / Visual
8AP02700RF/Label_NY_NP9815_73-85 | 8AP02700RF | Label_NY_NP9815_73-85 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection
8AP02700RF/Label_NY_NP9815_73-85 | 8AP02700RF | Label_NY_NP9815_73-85 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8AP02700RF/Label_NY_NP9815_73-85 | 8AP02700RF | Label_NY_NP9815_73-85 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8AP02700RF/Label_NY_NP9815_73-85 | 8AP02700RF | Label_NY_NP9815_73-85 | NY-Analytical | 770 | Label Inspection | Length | Length | 73 | mm | FDA,4 | 71.4125 | 74.5875 | TM 143 / Visual
8AP02700RF/Label_NY_NP9815_73-85 | 8AP02700RF | Label_NY_NP9815_73-85 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8AP02700RF/Label_NY_NP9815_73-85 | 8AP02700RF | Label_NY_NP9815_73-85 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8AP02700RF/Label_NY_NP9815_73-85 | 8AP02700RF | Label_NY_NP9815_73-85 | NY-Analytical | 770 | Label Inspection | Width | Width | 85 | mm | FDA,4 | 83.4125 | 86.5875 | TM 143 / Visual
8AP034RSL/Label_NY_NP9800_3.375-2.3437 | 8AP034RSL | Label_NY_NP9800_3.375-2.3437 | IN-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8AP034RSL/Label_NY_NP9800_3.375-2.3437 | 8AP034RSL | Label_NY_NP9800_3.375-2.3437 | IN-Analytical | 770 | Label Inspection | Certificate of Conformance | Certificate of Conformance | CofC meets specifications (Done) |  |  |  |  | TM 143 / Visual
8AP034RSL/Label_NY_NP9800_3.375-2.3437 | 8AP034RSL | Label_NY_NP9800_3.375-2.3437 | IN-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8AP034RSL/Label_NY_NP9800_3.375-2.3437 | 8AP034RSL | Label_NY_NP9800_3.375-2.3437 | IN-Analytical | 770 | Label Inspection | Length | Length | 3.375 | in | FDA,4 | 3.3125 | 3.4375 | TM 143 / Visual
8AP034RSL/Label_NY_NP9800_3.375-2.3437 | 8AP034RSL | Label_NY_NP9800_3.375-2.3437 | IN-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8AP034RSL/Label_NY_NP9800_3.375-2.3437 | 8AP034RSL | Label_NY_NP9800_3.375-2.3437 | IN-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8AP034RSL/Label_NY_NP9800_3.375-2.3437 | 8AP034RSL | Label_NY_NP9800_3.375-2.3437 | IN-Analytical | 770 | Label Inspection | Width | Width | 2.3437 | in | FDA,4 | 2.2812 | 2.4062 | TM 143 / Visual
8AP13800/Label_NY_NP7200_6.25-4 | 8AP13800 | Label_NY_NP7200_6.25-4 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection
8AP13800/Label_NY_NP7200_6.25-4 | 8AP13800 | Label_NY_NP7200_6.25-4 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8AP13800/Label_NY_NP7200_6.25-4 | 8AP13800 | Label_NY_NP7200_6.25-4 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8AP13800/Label_NY_NP7200_6.25-4 | 8AP13800 | Label_NY_NP7200_6.25-4 | NY-Analytical | 770 | Label Inspection | Length | Length | 6.1875 | in | FDA,4 | 6.1875 | 6.3125 | TM 143 / Visual
8AP13800/Label_NY_NP7200_6.25-4 | 8AP13800 | Label_NY_NP7200_6.25-4 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8AP13800/Label_NY_NP7200_6.25-4 | 8AP13800 | Label_NY_NP7200_6.25-4 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8AP13800/Label_NY_NP7200_6.25-4 | 8AP13800 | Label_NY_NP7200_6.25-4 | NY-Analytical | 770 | Label Inspection | Width | Width | 3.9375 | in | FDA,4 | 3.9375 | 4.0625 | TM 143 / Visual
8AP21100A/Label_NY_NP9890_85-73 | 8AP21100A | Label_NY_NP9890_85-73 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection
8AP21100A/Label_NY_NP9890_85-73 | 8AP21100A | Label_NY_NP9890_85-73 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8AP21100A/Label_NY_NP9890_85-73 | 8AP21100A | Label_NY_NP9890_85-73 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8AP21100A/Label_NY_NP9890_85-73 | 8AP21100A | Label_NY_NP9890_85-73 | NY-Analytical | 770 | Label Inspection | Length | Length | 85 | mm | FDA,4 | 83.4125 | 86.5875 | TM 143 / Visual
8AP21100A/Label_NY_NP9890_85-73 | 8AP21100A | Label_NY_NP9890_85-73 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8AP21100A/Label_NY_NP9890_85-73 | 8AP21100A | Label_NY_NP9890_85-73 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8AP21100A/Label_NY_NP9890_85-73 | 8AP21100A | Label_NY_NP9890_85-73 | NY-Analytical | 770 | Label Inspection | Width | Width | 73 | mm | FDA,4 | 71.4125 | 74.5875 | TM 143 / Visual
8AP22800/Label_NY_NP7200_5.8125-6.5 | 8AP22800 | Label_NY_NP7200_5.8125-6.5 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | [5 empty] | Label Inspection
8AP22800/Label_NY_NP7200_5.8125-6.5 | 8AP22800 | Label_NY_NP7200_5.8125-6.5 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8AP22800/Label_NY_NP7200_5.8125-6.5 | 8AP22800 | Label_NY_NP7200_5.8125-6.5 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8AP22800/Label_NY_NP7200_5.8125-6.5 | 8AP22800 | Label_NY_NP7200_5.8125-6.5 | NY-Analytical | 770 | Label Inspection | Length | Length | 5.75 | in | FDA,4 | 5.75 | 5.875 | TM 143 / Visual
8AP22800/Label_NY_NP7200_5.8125-6.5 | 8AP22800 | Label_NY_NP7200_5.8125-6.5 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8AP22800/Label_NY_NP7200_5.8125-6.5 | 8AP22800 | Label_NY_NP7200_5.8125-6.5 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8AP22800/Label_NY_NP7200_5.8125-6.5 | 8AP22800 | Label_NY_NP7200_5.8125-6.5 | NY-Analytical | 770 | Label Inspection | Width | Width | 6.5 | in | FDA,4 | 6.4375 | 6.5625 | TM 143 / Visual
8AP32800/Label_NY_NP98372_73-85 | 8AP32800 | Label_NY_NP98372_73-85 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection
8AP32800/Label_NY_NP98372_73-85 | 8AP32800 | Label_NY_NP98372_73-85 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8AP32800/Label_NY_NP98372_73-85 | 8AP32800 | Label_NY_NP98372_73-85 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8AP32800/Label_NY_NP98372_73-85 | 8AP32800 | Label_NY_NP98372_73-85 | NY-Analytical | 770 | Label Inspection | Length | Length | 73 | mm | FDA,4 | 71.4125 | 74.5875 | TM 143 / Visual
8AP32800/Label_NY_NP98372_73-85 | 8AP32800 | Label_NY_NP98372_73-85 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8AP32800/Label_NY_NP98372_73-85 | 8AP32800 | Label_NY_NP98372_73-85 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8AP32800/Label_NY_NP98372_73-85 | 8AP32800 | Label_NY_NP98372_73-85 | NY-Analytical | 770 | Label Inspection | Width | Width | 85 | mm | FDA,4 | 83.4125 | 86.5875 | TM 143 / Visual
8AP60100RF/Label_NY_NP9890_73-85 | 8AP60100RF | Label_NY_NP9890_73-85 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection
8AP60100RF/Label_NY_NP9890_73-85 | 8AP60100RF | Label_NY_NP9890_73-85 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8AP60100RF/Label_NY_NP9890_73-85 | 8AP60100RF | Label_NY_NP9890_73-85 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8AP60100RF/Label_NY_NP9890_73-85 | 8AP60100RF | Label_NY_NP9890_73-85 | NY-Analytical | 770 | Label Inspection | Length | Length | 72 | mm | FDA,4 | 71.4125 | 74.5875 | TM 143 / Visual
8AP60100RF/Label_NY_NP9890_73-85 | 8AP60100RF | Label_NY_NP9890_73-85 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8AP60100RF/Label_NY_NP9890_73-85 | 8AP60100RF | Label_NY_NP9890_73-85 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8AP60100RF/Label_NY_NP9890_73-85 | 8AP60100RF | Label_NY_NP9890_73-85 | NY-Analytical | 770 | Label Inspection | Width | Width | 85 | mm | FDA,4 | 83.4125 | 86.5875 | TM 143 / Visual
8AU120CK/Label_NY_NP7030_6-4 | 8AU120CK | Label_NY_NP7030_6-4 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection
8AU120CK/Label_NY_NP7030_6-4 | 8AU120CK | Label_NY_NP7030_6-4 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8AU120CK/Label_NY_NP7030_6-4 | 8AU120CK | Label_NY_NP7030_6-4 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8AU120CK/Label_NY_NP7030_6-4 | 8AU120CK | Label_NY_NP7030_6-4 | NY-Analytical | 770 | Label Inspection | Length | Length | 6 | in | FDA,4 | 5.9375 | 6.0625 | TM 143 / Visual
8AU120CK/Label_NY_NP7030_6-4 | 8AU120CK | Label_NY_NP7030_6-4 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8AU120CK/Label_NY_NP7030_6-4 | 8AU120CK | Label_NY_NP7030_6-4 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8AU120CK/Label_NY_NP7030_6-4 | 8AU120CK | Label_NY_NP7030_6-4 | NY-Analytical | 770 | Label Inspection | Width | Width | 4 | in | FDA,4 | 3.9375 | 4.0625 | TM 143 / Visual
8BB196CK/Label_NY_NP7030_6-5 | 8BB196CK | Label_NY_NP7030_6-5 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection
8BB196CK/Label_NY_NP7030_6-5 | 8BB196CK | Label_NY_NP7030_6-5 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8BB196CK/Label_NY_NP7030_6-5 | 8BB196CK | Label_NY_NP7030_6-5 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8BB196CK/Label_NY_NP7030_6-5 | 8BB196CK | Label_NY_NP7030_6-5 | NY-Analytical | 770 | Label Inspection | Length | Length | 5.99 | in | FDA,4 | 5.9375 | 6.0625 | TM 143 / Visual
8BB196CK/Label_NY_NP7030_6-5 | 8BB196CK | Label_NY_NP7030_6-5 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8BB196CK/Label_NY_NP7030_6-5 | 8BB196CK | Label_NY_NP7030_6-5 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8BB196CK/Label_NY_NP7030_6-5 | 8BB196CK | Label_NY_NP7030_6-5 | NY-Analytical | 770 | Label Inspection | Width | Width | 5.01 | in | FDA,4 | 4.9375 | 5.0625 | TM 143 / Visual
8BB512CL/Label_NY_NP7030_6-4 | 8BB512CL | Label_NY_NP7030_6-4 | NY-Analytical | 770 | Label Inspection | AQL Sample Size | AQL Sample Size | 50 |  |  |  |  | TM 143 / Visual
8BB512CL/Label_NY_NP7030_6-4 | 8BB512CL | Label_NY_NP7030_6-4 | NY-Analytical | 770 | Label Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Sample(s) all meet specifications (Done) | in |  |  |  | TM 143 / Visual
8BP12801/Label_NY_NP7200_4.5-4 | 8BP12801 | Label_NY_NP7200_4.5-4 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection
8BP12801/Label_NY_NP7200_4.5-4 | 8BP12801 | Label_NY_NP7200_4.5-4 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8BP12801/Label_NY_NP7200_4.5-4 | 8BP12801 | Label_NY_NP7200_4.5-4 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8BP12801/Label_NY_NP7200_4.5-4 | 8BP12801 | Label_NY_NP7200_4.5-4 | NY-Analytical | 770 | Label Inspection | Length | Length | 4.48 | in | FDA,4 | 4.4375 | 4.5625 | TM 143 / Visual
8BP12801/Label_NY_NP7200_4.5-4 | 8BP12801 | Label_NY_NP7200_4.5-4 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8BP12801/Label_NY_NP7200_4.5-4 | 8BP12801 | Label_NY_NP7200_4.5-4 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8BP12801/Label_NY_NP7200_4.5-4 | 8BP12801 | Label_NY_NP7200_4.5-4 | NY-Analytical | 770 | Label Inspection | Width | Width | 3.98 | in | FDA,4 | 3.9375 | 4.0625 | TM 143 / Visual
8BP24500/Label_NY_NP7200_4-4.5 | 8BP24500 | Label_NY_NP7200_4-4.5 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection
8BP24500/Label_NY_NP7200_4-4.5 | 8BP24500 | Label_NY_NP7200_4-4.5 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8BP24500/Label_NY_NP7200_4-4.5 | 8BP24500 | Label_NY_NP7200_4-4.5 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8BP24500/Label_NY_NP7200_4-4.5 | 8BP24500 | Label_NY_NP7200_4-4.5 | NY-Analytical | 770 | Label Inspection | Length | Length | 4 | in | FDA,4 | 3.9375 | 4.0625 | TM 143 / Visual
8BP24500/Label_NY_NP7200_4-4.5 | 8BP24500 | Label_NY_NP7200_4-4.5 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8BP24500/Label_NY_NP7200_4-4.5 | 8BP24500 | Label_NY_NP7200_4-4.5 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8BP24500/Label_NY_NP7200_4-4.5 | 8BP24500 | Label_NY_NP7200_4-4.5 | NY-Analytical | 770 | Label Inspection | Width | Width | 4.5 | in | FDA,4 | 4.4375 | 4.5625 | TM 143 / Visual
8BQ76500/Label_NY_NP7200_6-3.75 | 8BQ76500 | Label_NY_NP7200_6-3.75 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection
8BQ76500/Label_NY_NP7200_6-3.75 | 8BQ76500 | Label_NY_NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8BQ76500/Label_NY_NP7200_6-3.75 | 8BQ76500 | Label_NY_NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8BQ76500/Label_NY_NP7200_6-3.75 | 8BQ76500 | Label_NY_NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Length | Length | 6 | in | FDA,4 | 5.9375 | 6.0625 | TM 143 / Visual
8BQ76500/Label_NY_NP7200_6-3.75 | 8BQ76500 | Label_NY_NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8BQ76500/Label_NY_NP7200_6-3.75 | 8BQ76500 | Label_NY_NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8BQ76500/Label_NY_NP7200_6-3.75 | 8BQ76500 | Label_NY_NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Width | Width | 3.6875 | in | FDA,4 | 3.6875 | 3.8125 | TM 143 / Visual
8BQ76501/Label_NY_NP7200_3.75-6 | 8BQ76501 | Label_NY_NP7200_3.75-6 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection
8BQ76501/Label_NY_NP7200_3.75-6 | 8BQ76501 | Label_NY_NP7200_3.75-6 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8BQ76501/Label_NY_NP7200_3.75-6 | 8BQ76501 | Label_NY_NP7200_3.75-6 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8BQ76501/Label_NY_NP7200_3.75-6 | 8BQ76501 | Label_NY_NP7200_3.75-6 | NY-Analytical | 770 | Label Inspection | Length | Length | 3.6875 | in | FDA,4 | 3.6875 | 3.8125 | TM 143 / Visual
8BQ76501/Label_NY_NP7200_3.75-6 | 8BQ76501 | Label_NY_NP7200_3.75-6 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8BQ76501/Label_NY_NP7200_3.75-6 | 8BQ76501 | Label_NY_NP7200_3.75-6 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8BQ76501/Label_NY_NP7200_3.75-6 | 8BQ76501 | Label_NY_NP7200_3.75-6 | NY-Analytical | 770 | Label Inspection | Width | Width | 6 | in | FDA,4 | 5.9375 | 6.0625 | TM 143 / Visual
8CH211CK/Label_A12 | 8CH211CK | Label_A12 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection
8CH211CK/Label_A12 | 8CH211CK | Label_A12 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8CH211CK/Label_A12 | 8CH211CK | Label_A12 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8CH211CK/Label_A12 | 8CH211CK | Label_A12 | NY-Analytical | 770 | Label Inspection | Length | Length | 9 | in | FDA,4 | 8.9375 | 9.0625 | TM 143 / Visual
8CH211CK/Label_A12 | 8CH211CK | Label_A12 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8CH211CK/Label_A12 | 8CH211CK | Label_A12 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8CH211CK/Label_A12 | 8CH211CK | Label_A12 | NY-Analytical | 770 | Label Inspection | Width | Width | 2.625 | in | FDA,4 | 2.5625 | 2.6875 | TM 143 / Visual
8CP13400/Label_NY_NP7100_4.5-4 | 8CP13400 | Label_NY_NP7100_4.5-4 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection
8CP13400/Label_NY_NP7100_4.5-4 | 8CP13400 | Label_NY_NP7100_4.5-4 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8CP13400/Label_NY_NP7100_4.5-4 | 8CP13400 | Label_NY_NP7100_4.5-4 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8CP13400/Label_NY_NP7100_4.5-4 | 8CP13400 | Label_NY_NP7100_4.5-4 | NY-Analytical | 770 | Label Inspection | Length | Length | 4.5 | in | FDA,4 | 4.4375 | 4.5625 | TM 143 / Visual
8CP13400/Label_NY_NP7100_4.5-4 | 8CP13400 | Label_NY_NP7100_4.5-4 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8CP13400/Label_NY_NP7100_4.5-4 | 8CP13400 | Label_NY_NP7100_4.5-4 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8CP13400/Label_NY_NP7100_4.5-4 | 8CP13400 | Label_NY_NP7100_4.5-4 | NY-Analytical | 770 | Label Inspection | Width | Width | 4 | in | FDA,4 | 3.9375 | 4.0625 | TM 143 / Visual
8CP15900/Label_NY_NP7100/NP7200_6-3.75 | 8CP15900 | Label_NY_NP7100/NP7200_6-3.75 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection
8CP15900/Label_NY_NP7100/NP7200_6-3.75 | 8CP15900 | Label_NY_NP7100/NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8CP15900/Label_NY_NP7100/NP7200_6-3.75 | 8CP15900 | Label_NY_NP7100/NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8CP15900/Label_NY_NP7100/NP7200_6-3.75 | 8CP15900 | Label_NY_NP7100/NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Length | Length | 5.99 | in | FDA,4 | 5.9375 | 6.0625 | TM 143 / Visual
8CP15900/Label_NY_NP7100/NP7200_6-3.75 | 8CP15900 | Label_NY_NP7100/NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8CP15900/Label_NY_NP7100/NP7200_6-3.75 | 8CP15900 | Label_NY_NP7100/NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8CP15900/Label_NY_NP7100/NP7200_6-3.75 | 8CP15900 | Label_NY_NP7100/NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Width | Width | 3.705 | in | FDA,4 | 3.6875 | 3.8125 | TM 143 / Visual
8CP63801/Label_A20 | 8CP63801 | Label_A20 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection
8CP63801/Label_A20 | 8CP63801 | Label_A20 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | [5 empty] | TM 143 / Visual
8CP63801/Label_A20 | 8CP63801 | Label_A20 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | [5 empty] | TM 143 / Visual
8CP63801/Label_A20 | 8CP63801 | Label_A20 | NY-Analytical | 770 | Label Inspection | Length | Length |  | in | FDA,4 | 5.9375 | 6.0625 | TM 143 / Visual
8CP63801/Label_A20 | 8CP63801 | Label_A20 | NY-Analytical | 770 | Label Inspection | Odor | Odor | [5 empty] | TM 143 / Visual
8CP63801/Label_A20 | 8CP63801 | Label_A20 | NY-Analytical | 770 | Label Inspection | Print | Print | [5 empty] | TM 143 / Visual
8CP63801/Label_A20 | 8CP63801 | Label_A20 | NY-Analytical | 770 | Label Inspection | Width | Width |  | in | FDA,4 | 3.6875 | 3.8125 | TM 143 / Visual
8DP92000/Label_NY_NP7200_6.5-5.8125 | 8DP92000 | Label_NY_NP7200_6.5-5.8125 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | [5 empty] | Label Inspection
8DP92000/Label_NY_NP7200_6.5-5.8125 | 8DP92000 | Label_NY_NP7200_6.5-5.8125 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8DP92000/Label_NY_NP7200_6.5-5.8125 | 8DP92000 | Label_NY_NP7200_6.5-5.8125 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8DP92000/Label_NY_NP7200_6.5-5.8125 | 8DP92000 | Label_NY_NP7200_6.5-5.8125 | NY-Analytical | 770 | Label Inspection | Length | Length | 6.5 | in | FDA,4 | 6.4375 | 6.5625 | TM 143 / Visual
8DP92000/Label_NY_NP7200_6.5-5.8125 | 8DP92000 | Label_NY_NP7200_6.5-5.8125 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8DP92000/Label_NY_NP7200_6.5-5.8125 | 8DP92000 | Label_NY_NP7200_6.5-5.8125 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8DP92000/Label_NY_NP7200_6.5-5.8125 | 8DP92000 | Label_NY_NP7200_6.5-5.8125 | NY-Analytical | 770 | Label Inspection | Width | Width | 5.8125 | in | FDA,4 | 5.75 | 5.875 | TM 143 / Visual
8DQ85400/Label_A41 | 8DQ85400 | Label_A41 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection
8DQ85400/Label_A41 | 8DQ85400 | Label_A41 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8DQ85400/Label_A41 | 8DQ85400 | Label_A41 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8DQ85400/Label_A41 | 8DQ85400 | Label_A41 | NY-Analytical | 770 | Label Inspection | Length | Length | 5.5625 | in | FDA,4 | 5.5 | 5.625 | TM 143 / Visual
8DQ85400/Label_A41 | 8DQ85400 | Label_A41 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8DQ85400/Label_A41 | 8DQ85400 | Label_A41 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8DQ85400/Label_A41 | 8DQ85400 | Label_A41 | NY-Analytical | 770 | Label Inspection | Width | Width | 6.25 | in | FDA,4 | 6.1875 | 6.3125 | TM 143 / Visual
8DQ85900/Label_NY_NP7200_5.5625-6.25 | 8DQ85900 | Label_NY_NP7200_5.5625-6.25 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection
8DQ85900/Label_NY_NP7200_5.5625-6.25 | 8DQ85900 | Label_NY_NP7200_5.5625-6.25 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8DQ85900/Label_NY_NP7200_5.5625-6.25 | 8DQ85900 | Label_NY_NP7200_5.5625-6.25 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8DQ85900/Label_NY_NP7200_5.5625-6.25 | 8DQ85900 | Label_NY_NP7200_5.5625-6.25 | NY-Analytical | 770 | Label Inspection | Length | Length | 5.5625 | in | FDA,4 | 5.5 | 5.625 | TM 143 / Visual
8DQ85900/Label_NY_NP7200_5.5625-6.25 | 8DQ85900 | Label_NY_NP7200_5.5625-6.25 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8DQ85900/Label_NY_NP7200_5.5625-6.25 | 8DQ85900 | Label_NY_NP7200_5.5625-6.25 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8DQ85900/Label_NY_NP7200_5.5625-6.25 | 8DQ85900 | Label_NY_NP7200_5.5625-6.25 | NY-Analytical | 770 | Label Inspection | Width | Width | 6.25 | in | FDA,4 | 6.1875 | 6.3125 | TM 143 / Visual
8DQ914SL/Label_NY_NP7030_9-7 | 8DQ914SL | Label_NY_NP7030_9-7 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | [5 empty] | Label Inspection
8DQ914SL/Label_NY_NP7030_9-7 | 8DQ914SL | Label_NY_NP7030_9-7 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8DQ914SL/Label_NY_NP7030_9-7 | 8DQ914SL | Label_NY_NP7030_9-7 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8DQ914SL/Label_NY_NP7030_9-7 | 8DQ914SL | Label_NY_NP7030_9-7 | NY-Analytical | 770 | Label Inspection | Length | Length | 9 | in | FDA,4 | 8.9375 | 9.0625 | TM 143 / Visual
8DQ914SL/Label_NY_NP7030_9-7 | 8DQ914SL | Label_NY_NP7030_9-7 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8DQ914SL/Label_NY_NP7030_9-7 | 8DQ914SL | Label_NY_NP7030_9-7 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8DQ914SL/Label_NY_NP7030_9-7 | 8DQ914SL | Label_NY_NP7030_9-7 | NY-Analytical | 770 | Label Inspection | Width | Width | 7 | in | FDA,4 | 6.9375 | 7.0625 | TM 143 / Visual
8GQ772SK/Label_B03 | 8GQ772SK | Label_B03 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection
8GQ772SK/Label_B03 | 8GQ772SK | Label_B03 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8GQ772SK/Label_B03 | 8GQ772SK | Label_B03 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8GQ772SK/Label_B03 | 8GQ772SK | Label_B03 | NY-Analytical | 770 | Label Inspection | Diameter | Diameter | 1.875 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual
8GQ772SK/Label_B03 | 8GQ772SK | Label_B03 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8GQ772SK/Label_B03 | 8GQ772SK | Label_B03 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8OH217UP/Label_A25 | 8OH217UP | Label_A25 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection
8OH217UP/Label_A25 | 8OH217UP | Label_A25 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8OH217UP/Label_A25 | 8OH217UP | Label_A25 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8OH217UP/Label_A25 | 8OH217UP | Label_A25 | NY-Analytical | 770 | Label Inspection | Length | Length | 10 | in | FDA,4 | 9.9375 | 10.0625 | TM 143 / Visual
8OH217UP/Label_A25 | 8OH217UP | Label_A25 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8OH217UP/Label_A25 | 8OH217UP | Label_A25 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8OH217UP/Label_A25 | 8OH217UP | Label_A25 | NY-Analytical | 770 | Label Inspection | Width | Width | 4.0625 | in | FDA,4 | 3.9375 | 4.0625 | TM 143 / Visual
8OJ22750/Label_A11 | 8OJ22750 | Label_A11 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection
8OJ22750/Label_A11 | 8OJ22750 | Label_A11 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8OJ22750/Label_A11 | 8OJ22750 | Label_A11 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8OJ22750/Label_A11 | 8OJ22750 | Label_A11 | NY-Analytical | 770 | Label Inspection | Length | Length | 5.35 | in | FDA,4 | 5.3125 | 5.4375 | TM 143 / Visual
8OJ22750/Label_A11 | 8OJ22750 | Label_A11 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8OJ22750/Label_A11 | 8OJ22750 | Label_A11 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8OJ22750/Label_A11 | 8OJ22750 | Label_A11 | NY-Analytical | 770 | Label Inspection | Width | Width | 2.61 | in | FDA,4 | 2.5625 | 2.6875 | TM 143 / Visual
8OP03700/Label_NY_NP9840_87-75 | 8OP03700 | Label_NY_NP9840_87-75 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection
8OP03700/Label_NY_NP9840_87-75 | 8OP03700 | Label_NY_NP9840_87-75 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8OP03700/Label_NY_NP9840_87-75 | 8OP03700 | Label_NY_NP9840_87-75 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8OP03700/Label_NY_NP9840_87-75 | 8OP03700 | Label_NY_NP9840_87-75 | NY-Analytical | 770 | Label Inspection | Length | Length | 86.97 | mm | FDA,4 | 85.4125 | 88.5875 | TM 143 / Visual
8OP03700/Label_NY_NP9840_87-75 | 8OP03700 | Label_NY_NP9840_87-75 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8OP03700/Label_NY_NP9840_87-75 | 8OP03700 | Label_NY_NP9840_87-75 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8OP03700/Label_NY_NP9840_87-75 | 8OP03700 | Label_NY_NP9840_87-75 | NY-Analytical | 770 | Label Inspection | Width | Width | 75.06 | mm | FDA,4 | 73.4125 | 76.5875 | TM 143 / Visual
8OP065RSL/Label_NY_NP9995_2.625-3.375 | 8OP065RSL | Label_NY_NP9995_2.625-3.375 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection
8OP065RSL/Label_NY_NP9995_2.625-3.375 | 8OP065RSL | Label_NY_NP9995_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8OP065RSL/Label_NY_NP9995_2.625-3.375 | 8OP065RSL | Label_NY_NP9995_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8OP065RSL/Label_NY_NP9995_2.625-3.375 | 8OP065RSL | Label_NY_NP9995_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Length | Length | 2.625 | in | FDA,4 | 2.5625 | 2.6875 | TM 143 / Visual
8OP065RSL/Label_NY_NP9995_2.625-3.375 | 8OP065RSL | Label_NY_NP9995_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8OP065RSL/Label_NY_NP9995_2.625-3.375 | 8OP065RSL | Label_NY_NP9995_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8OP065RSL/Label_NY_NP9995_2.625-3.375 | 8OP065RSL | Label_NY_NP9995_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Width | Width | 3.375 | in | FDA,4 | 3.3125 | 3.4375 | TM 143 / Visual
8OP08700/Label_NY_NP9800_3.375-2.625 | 8OP08700 | Label_NY_NP9800_3.375-2.625 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection
8OP08700/Label_NY_NP9800_3.375-2.625 | 8OP08700 | Label_NY_NP9800_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8OP08700/Label_NY_NP9800_3.375-2.625 | 8OP08700 | Label_NY_NP9800_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8OP08700/Label_NY_NP9800_3.375-2.625 | 8OP08700 | Label_NY_NP9800_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Length | Length | 3.375 | in | FDA,4 | 3.3125 | 3.4375 | TM 143 / Visual
8OP08700/Label_NY_NP9800_3.375-2.625 | 8OP08700 | Label_NY_NP9800_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8OP08700/Label_NY_NP9800_3.375-2.625 | 8OP08700 | Label_NY_NP9800_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8OP08700/Label_NY_NP9800_3.375-2.625 | 8OP08700 | Label_NY_NP9800_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Width | Width | 2.5625 | in | FDA,4 | 2.5625 | 2.6875 | TM 143 / Visual
8OP09100/Label_NY_NP9991_3.375-2.625 | 8OP09100 | Label_NY_NP9991_3.375-2.625 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Label Inspection
8OP09100/Label_NY_NP9991_3.375-2.625 | 8OP09100 | Label_NY_NP9991_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8OP09100/Label_NY_NP9991_3.375-2.625 | 8OP09100 | Label_NY_NP9991_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8OP09100/Label_NY_NP9991_3.375-2.625 | 8OP09100 | Label_NY_NP9991_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Length | Length | 3.375 | in | FDA,4 | 3.3125 | 3.4375 | TM 143 / Visual
8OP09100/Label_NY_NP9991_3.375-2.625 | 8OP09100 | Label_NY_NP9991_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8OP09100/Label_NY_NP9991_3.375-2.625 | 8OP09100 | Label_NY_NP9991_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8OP09100/Label_NY_NP9991_3.375-2.625 | 8OP09100 | Label_NY_NP9991_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Width | Width | 2.625 | in | FDA,4 | 2.5625 | 2.6875 | TM 143 / Visual
8OP09300/Label_NY_NP7110_2.55-3 | 8OP09300 | Label_NY_NP7110_2.55-3 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8OP09300/Label_NY_NP7110_2.55-3 | 8OP09300 | Label_NY_NP7110_2.55-3 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8OP09300/Label_NY_NP7110_2.55-3 | 8OP09300 | Label_NY_NP7110_2.55-3 | NY-Analytical | 770 | Label Inspection | Length | Length | 2.5 | in | FDA,4 | 2.4875 | 2.6125 | TM 143 / Visual
8OP09300/Label_NY_NP7110_2.55-3 | 8OP09300 | Label_NY_NP7110_2.55-3 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8OP09300/Label_NY_NP7110_2.55-3 | 8OP09300 | Label_NY_NP7110_2.55-3 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8OP09300/Label_NY_NP7110_2.55-3 | 8OP09300 | Label_NY_NP7110_2.55-3 | NY-Analytical | 770 | Label Inspection | Width | Width | 3 | in | FDA,4 | 2.9375 | 3.0625 | TM 143 / Visual
8OP09300/Label_NY_NP7110_2.55-3 | 8OP09300 | Label_NY_NP7110_2.55-3 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Visual
8OP09301/Label_NY_NP7110_2.55-5 | 8OP09301 | Label_NY_NP7110_2.55-5 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8OP09301/Label_NY_NP7110_2.55-5 | 8OP09301 | Label_NY_NP7110_2.55-5 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8OP09301/Label_NY_NP7110_2.55-5 | 8OP09301 | Label_NY_NP7110_2.55-5 | NY-Analytical | 770 | Label Inspection | Length | Length | 2.5313 | in | FDA,4 | 2.4875 | 2.6125 | TM 143 / Visual
8OP09301/Label_NY_NP7110_2.55-5 | 8OP09301 | Label_NY_NP7110_2.55-5 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8OP09301/Label_NY_NP7110_2.55-5 | 8OP09301 | Label_NY_NP7110_2.55-5 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8OP09301/Label_NY_NP7110_2.55-5 | 8OP09301 | Label_NY_NP7110_2.55-5 | NY-Analytical | 770 | Label Inspection | Width | Width | 4.9687 | in | FDA,4 | 4.9375 | 5.0625 | TM 143 / Visual
8OP09301/Label_NY_NP7110_2.55-5 | 8OP09301 | Label_NY_NP7110_2.55-5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Visual
8OP227RSL/Label_NY_NP9997_2.625-3.375 | 8OP227RSL | Label_NY_NP9997_2.625-3.375 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | [5 empty] | Label Inspection
8OP227RSL/Label_NY_NP9997_2.625-3.375 | 8OP227RSL | Label_NY_NP9997_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8OP227RSL/Label_NY_NP9997_2.625-3.375 | 8OP227RSL | Label_NY_NP9997_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8OP227RSL/Label_NY_NP9997_2.625-3.375 | 8OP227RSL | Label_NY_NP9997_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Length | Length | 2.63 | in | FDA,4 | 2.5625 | 2.6875 | TM 143 / Visual
8OP227RSL/Label_NY_NP9997_2.625-3.375 | 8OP227RSL | Label_NY_NP9997_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8OP227RSL/Label_NY_NP9997_2.625-3.375 | 8OP227RSL | Label_NY_NP9997_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8OP227RSL/Label_NY_NP9997_2.625-3.375 | 8OP227RSL | Label_NY_NP9997_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Width | Width | 3.37 | in | FDA,4 | 3.3125 | 3.4375 | TM 143 / Visual
8OP22900RF/Label_NY_NP98371_78-61 | 8OP22900RF | Label_NY_NP98371_78-61 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection
8OP22900RF/Label_NY_NP98371_78-61 | 8OP22900RF | Label_NY_NP98371_78-61 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8OP22900RF/Label_NY_NP98371_78-61 | 8OP22900RF | Label_NY_NP98371_78-61 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8OP22900RF/Label_NY_NP98371_78-61 | 8OP22900RF | Label_NY_NP98371_78-61 | NY-Analytical | 770 | Label Inspection | Length | Length | 78.75 | mm | FDA,4 | 76.4125 | 79.5875 | TM 143 / Visual
8OP22900RF/Label_NY_NP98371_78-61 | 8OP22900RF | Label_NY_NP98371_78-61 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8OP22900RF/Label_NY_NP98371_78-61 | 8OP22900RF | Label_NY_NP98371_78-61 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8OP22900RF/Label_NY_NP98371_78-61 | 8OP22900RF | Label_NY_NP98371_78-61 | NY-Analytical | 770 | Label Inspection | Width | Width | 61.36 | mm | FDA,4 | 59.4125 | 62.5875 | TM 143 / Visual
8OP260RSL/Label_NY_NP7700_3.375-2.3437 | 8OP260RSL | Label_NY_NP7700_3.375-2.3437 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection
8OP260RSL/Label_NY_NP7700_3.375-2.3437 | 8OP260RSL | Label_NY_NP7700_3.375-2.3437 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8OP260RSL/Label_NY_NP7700_3.375-2.3437 | 8OP260RSL | Label_NY_NP7700_3.375-2.3437 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8OP260RSL/Label_NY_NP7700_3.375-2.3437 | 8OP260RSL | Label_NY_NP7700_3.375-2.3437 | NY-Analytical | 770 | Label Inspection | Length | Length | 3.3125 | in | FDA,4 | 3.3125 | 3.4375 | TM 143 / Visual
8OP260RSL/Label_NY_NP7700_3.375-2.3437 | 8OP260RSL | Label_NY_NP7700_3.375-2.3437 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8OP260RSL/Label_NY_NP7700_3.375-2.3437 | 8OP260RSL | Label_NY_NP7700_3.375-2.3437 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8OP260RSL/Label_NY_NP7700_3.375-2.3437 | 8OP260RSL | Label_NY_NP7700_3.375-2.3437 | NY-Analytical | 770 | Label Inspection | Width | Width | 2.2812 | in | FDA,4 | 2.2812 | 2.4062 | TM 143 / Visual
8OP263CK/Label_NY_NP7030_8.5-6.5 | 8OP263CK | Label_NY_NP7030_8.5-6.5 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection
8OP263CK/Label_NY_NP7030_8.5-6.5 | 8OP263CK | Label_NY_NP7030_8.5-6.5 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8OP263CK/Label_NY_NP7030_8.5-6.5 | 8OP263CK | Label_NY_NP7030_8.5-6.5 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8OP263CK/Label_NY_NP7030_8.5-6.5 | 8OP263CK | Label_NY_NP7030_8.5-6.5 | NY-Analytical | 770 | Label Inspection | Length | Length | 8.5 | in | FDA,4 | 8.4375 | 8.5625 | TM 143 / Visual
8OP263CK/Label_NY_NP7030_8.5-6.5 | 8OP263CK | Label_NY_NP7030_8.5-6.5 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8OP263CK/Label_NY_NP7030_8.5-6.5 | 8OP263CK | Label_NY_NP7030_8.5-6.5 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8OP263CK/Label_NY_NP7030_8.5-6.5 | 8OP263CK | Label_NY_NP7030_8.5-6.5 | NY-Analytical | 770 | Label Inspection | Width | Width | 6.5 | in | FDA,4 | 6.4375 | 6.5625 | TM 143 / Visual
8OP308SIRF/Label_NY_NP9300_2-1.375 | 8OP308SIRF | Label_NY_NP9300_2-1.375 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection
8OP308SIRF/Label_NY_NP9300_2-1.375 | 8OP308SIRF | Label_NY_NP9300_2-1.375 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8OP308SIRF/Label_NY_NP9300_2-1.375 | 8OP308SIRF | Label_NY_NP9300_2-1.375 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8OP308SIRF/Label_NY_NP9300_2-1.375 | 8OP308SIRF | Label_NY_NP9300_2-1.375 | NY-Analytical | 770 | Label Inspection | Length | Length | 2 | in | FDA,4 | 1.9375 | 2.0625 | TM 143 / Visual
8OP308SIRF/Label_NY_NP9300_2-1.375 | 8OP308SIRF | Label_NY_NP9300_2-1.375 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8OP308SIRF/Label_NY_NP9300_2-1.375 | 8OP308SIRF | Label_NY_NP9300_2-1.375 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8OP308SIRF/Label_NY_NP9300_2-1.375 | 8OP308SIRF | Label_NY_NP9300_2-1.375 | NY-Analytical | 770 | Label Inspection | Width | Width | 1.375 | in | FDA,4 | 1.3125 | 1.4375 | TM 143 / Visual
8OP585SIRF/Label_NY_NP9300_1.375-2 | 8OP585SIRF | Label_NY_NP9300_1.375-2 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection
8OP585SIRF/Label_NY_NP9300_1.375-2 | 8OP585SIRF | Label_NY_NP9300_1.375-2 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8OP585SIRF/Label_NY_NP9300_1.375-2 | 8OP585SIRF | Label_NY_NP9300_1.375-2 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8OP585SIRF/Label_NY_NP9300_1.375-2 | 8OP585SIRF | Label_NY_NP9300_1.375-2 | NY-Analytical | 770 | Label Inspection | Length | Length | 1.37 | in | FDA,4 | 1.3125 | 1.4375 | TM 143 / Visual
8OP585SIRF/Label_NY_NP9300_1.375-2 | 8OP585SIRF | Label_NY_NP9300_1.375-2 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8OP585SIRF/Label_NY_NP9300_1.375-2 | 8OP585SIRF | Label_NY_NP9300_1.375-2 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8OP585SIRF/Label_NY_NP9300_1.375-2 | 8OP585SIRF | Label_NY_NP9300_1.375-2 | NY-Analytical | 770 | Label Inspection | Width | Width | 1.99 | in | FDA,4 | 1.9375 | 2.0625 | TM 143 / Visual
8OP63000RF/Label_NY_NP9837_85-73 | 8OP63000RF | Label_NY_NP9837_85-73 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection
8OP63000RF/Label_NY_NP9837_85-73 | 8OP63000RF | Label_NY_NP9837_85-73 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8OP63000RF/Label_NY_NP9837_85-73 | 8OP63000RF | Label_NY_NP9837_85-73 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8OP63000RF/Label_NY_NP9837_85-73 | 8OP63000RF | Label_NY_NP9837_85-73 | NY-Analytical | 770 | Label Inspection | Length | Length | 84 | mm | FDA,4 | 83.4125 | 86.5875 | TM 143 / Visual
8OP63000RF/Label_NY_NP9837_85-73 | 8OP63000RF | Label_NY_NP9837_85-73 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8OP63000RF/Label_NY_NP9837_85-73 | 8OP63000RF | Label_NY_NP9837_85-73 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8OP63000RF/Label_NY_NP9837_85-73 | 8OP63000RF | Label_NY_NP9837_85-73 | NY-Analytical | 770 | Label Inspection | Width | Width | 73 | mm | FDA,4 | 71.4125 | 74.5875 | TM 143 / Visual
8OX351CK/Label_NY_50#W_5_4 | 8OX351CK | Label_NY_50#W_5_4 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection
8OX351CK/Label_NY_50#W_5_4 | 8OX351CK | Label_NY_50#W_5_4 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual
8OX351CK/Label_NY_50#W_5_4 | 8OX351CK | Label_NY_50#W_5_4 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual
8OX351CK/Label_NY_50#W_5_4 | 8OX351CK | Label_NY_50#W_5_4 | NY-Analytical | 770 | Label Inspection | Length | Length | 5.01 | in | FDA,4 | 4.9375 | 5.0625 | TM 143 / Visual
8OX351CK/Label_NY_50#W_5_4 | 8OX351CK | Label_NY_50#W_5_4 | NY-Analytical | 770 | Label Inspection | NP Structure | NP Structure | 50# White Paper Stock (Done) | in |  |  |  | TM 143 / Visual
8OX351CK/Label_NY_50#W_5_4 | 8OX351CK | Label_NY_50#W_5_4 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual
8OX351CK/Label_NY_50#W_5_4 | 8OX351CK | Label_NY_50#W_5_4 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual
8OX351CK/Label_NY_50#W_5_4 | 8OX351CK | Label_NY_50#W_5_4 | NY-Analytical | 770 | Label Inspection | Width | Width | 3.99 | in | FDA,4 | 3.9375 | 4.0625 | TM 143 / Visual
9OX465GL/NY_Glue_01 | 9OX465GL | NY_Glue_01 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Visual
9OX465GL/NY_Glue_01 | 9OX465GL | NY_Glue_01 | NY-Analytical | 395 | Color | Color | Color | Very Pale Amber (Done) |  |  |  |  | Visual
/
/

## Plans
plan
2AA180LI/Lid_14A
2AA284R/Web_NP8504_R_11.0
2AA284SLI/Lid_01B
2AA3341R/Web_NP8535_13
2AA347/Web_NP8550_10.25
2AA351SLI/Lid_01U
2AA402RNY/Web_NP8504_R_11.0
2AA402SLI/Lid_21A
2AA412/Tub_NY_01
2AA437/Web_NP8530_15.5
2AA437SLI/Lid_01A
2AA461R-1/Web_NP8535_16.5
2AA466R/Web_NP8530_11.0
2AA466SLI/Lid_01D
2AA507/Web_NP8530_12.5
2AA520/Web_NP8530_11.0
2AA520SLI/Lid_01M
2AA571SLI/Lid_01S
2AA572R/Web_NP8530_12.375
2AA572SLI/Lid_01T
2AA580SLI/Lid_01A
2AA592/Web_NP8504_11.0
2AA592SLI/Lid_NY_NP7600_22R1
2AA630/Web_NP8535_15.5
2AA630B/Web_NP8530_14.375
2AA630SLI/Lid_01A
2AA632/Web_NP8504_11.75
2AA634SLI/Lid_16B
2AA657SLI/Lid_01H
2AA680R/Web_NP8530_11.0
2AA682R/Web_NP8530_11.0
2AA682SLI/Lid_01B
2AA694R/Web_NP8530_10.75
2AA709/Web_NP8504_11.0
2AA714SLI/Lid_01V
2AA759/Web_NP8530_11.0
2AA759SLI/Lid_01F
2AA8031NU/Web_NP8550_14.25
2AA922SLI/Lid_01V
2AA941/Web_NP8530_15.75
2AA941SLI/Lid_NY_NP7600_22C
2AA972/Web_NP8540_13.5
2AB114/Web_NP3860_5.0
2AB196/Web_NP5114_41.0
2AB293/Web_NP5410_8.0
2AB512/Web_NP3870_4.5
2AB810/Web_NP5115_4.0
2AC801/Web_NP3870_6.0
2AD015/Web_NP5115_7.5
2AD173/Web_NP5115_4.5
2AD246/Web_NP5115U_4.5
2AD314/Web_NP5115_7.5
2AD399/Web_NP5115_7.5
2AD436/Web_NP5115_4.5
2AD504/Web_NP5115U_4.5
2AD507/Web_NP5116_4.5
2AD576/Web_NP5115_4.5
2AD644/Web_NP5115_7.5
2AD725/Web_NP5115_4.5
2AD766/Web_NP5115_4.5
2AG306/Web_NP9503_26
2AG451/Web_NP9503_26.0
2AH115/Web_NP5115_7.5
2AH126/Web_NP5115_7.5
2AH128/Web_NP5115_7.5
2AH129/Web_NP5115_7.5
2AH130/Web_NP5115_7.5
2AH131/Web_NP5115_7.5
2AH132/Web_NP5115_7.5
2AH140/Web_NP5115_7.5
2AH211/Web_NP5115_6.0
2AH217/Web_NP5115_7.5
2AH220/Web_NP5115_7.5
2AH225/Web_NP5115_7.5
2AH241/Web_NP5115_7.5
2AH592/Web_NP3895_6.0
2AJ123/Web_NP8550_11.0
2AJ123SLI/Lid_NY_NP7600_22U
2AJ137/Web_NP8504_10.25
2AJ143/Web_NP8530_16.25
2AJ143SLI/Lid_NY_NP7600_01C
2AJ165/Web_NP8535_14
2AJ170/Web_NP8530_11.0
2AJ170SLI/Lid_NY_NP7600_22FA
2AJ190/Web_NP8530_11.0
2AJ190SLI/Lid_NY_NP7600_22V
2AJ200/Web_NP8530_11.0
2AJ200SLI/Lid_NY_NP7600_22W
2AJ213/Web_NP8535_18.0
2AJ213SLI/Lid_NY_NP7600_22A
2AJ230/Web_NP8530_11.0
2AJ230SLI/Lid_NY_NP7600_22FB
2AJ324/Web_NP8530_11.0
2AJ324SLI/Lid_NY_NP7600_22X
2AJ660/Web_NP8535_13.25
2AM225SLI/Lid_01B
2AM233SLI/Lid_01B
2AM306S3B/Web_NP9225_12.125
2AM3371R/Web_NP8535_15
2AM420SF3B/Web_NP9225_13.75
2AM667R/Web_NP8530_15.0
2AM8076XF6B/Web_NP9225_18.75
2AM945SLI/Lid_14OA
2AM985NU/Web_NP8595_16.5
2AMC1412RJ/Web_NP8535_13.875
2AP027/Web_NP8512_310
2AP034/Web_NP8530_8.25
2AP075/Web_NP8542_12.205
2AP087/Web_NP8550_12.205
2AP091/Web_NP8550_12.205
2AP097/Web_NP8508_12.0
2AP104/Web_NP8508_12.0
2AP113/Web_NP8542_11.0
2AP136/Web_NP8508_6.88
2AP140/Web_NP8512_12.205
2AP147/Web_NP8512_12.205
2AP149/Web_NP8512_12.625
2AP156/Web_NP8508_6.25
2AP158/Web_NP8512_12.205
2AP159LI/Lid_03K
2AP161/Web_NP8512_12.205
2AP162SLI/Lid_19AD
2AP164/Web_NP8508_8.875
2AP169/Web_NP8512_12.205
2AP169SLI/Lid_01P
2AP172/Web_NP8512_13.25
2AP176/Web_NP8512_12.205
2AP181/Web_NP8512_12.205
2AP203/Web_NP8535_12.75
2AP210/Web_NP8542_11.0
2AP211/Web_NP8512_12.205
2AP225/Web_NP8512_14.5
2AP225SLI/Lid_19AC
2AP227/Web_NP8542_11.0
2AP229/Web_NP8512_9.5
2AP237LB/Canister_04A
2AP260/Web_NP8535_12.75
2AP279/Web_NP8512_12.205
2AP284/Web_NP8512_12.205
2AP308/Web_NP8512_12.205
2AP308SLI/Lid_19AA
2AP328/Web_NP8512_12.205
2AP337/Web_NP8512_310
2AP345/Web_NP8508_12.0
2AP352/Reduced Testing_NY_Can 2AP352
2AP374/Web_NP8535_10.0
2AP382/Web_NP8530_8.25
2AP423/Web_NP8508_8.5
2AP438/Web_NP8508_6.88
2AP485/Web_NP8512_12.205
2AP487/Web_NP8508_9.5
2AP550LB/Canister_04B
2AP563/Web_NP8530_12.705
2AP585/Web_NP8512_12.205
2AP585SLI/Lid_NY_NP7550_01
2AP601/Web_NP8512_12.205
2AP705/Web_NP8530_8.25
2AP713/Web_NP8530_8.25
2AP721/Web_NP8512_12.205
2AP755/Web_NP8530_8.25
2AP765TPA/Web_NP8550_8.75
2AP794/Web_NP8508_6.25
2AP806/Web_NP8430_310
2AP870/Web_NP8508_10.25
2AQ006BN/Web_NP8535_14.375
2AQ038BN/Web_NP8535_14.375
2AQ053R/Web_NP8530_13.125
2AQ053TP/Web_NP8508_10.25
2AQ337LI/Lid_04A_560/_673
2AQ387/Web_NP8530_12.125
2AQ406BN/Web_NP8530_14.375
2AQ412R/Web_NP8535_12.625
2AQ467SLI/Lid_01E
2AQ471LI/Lid_04B_560/_673
2AQ5058R/Web_NP8530_12.375
2AQ654/Canister_12A
2AQ810SLI/Lid_01Q
2AQ828LI/Lid_09A
2AS273/Web_NP3860_15.5
2AS407/Web_NP3885_13.0
2AS419/Web_NP3885_16.0
2AS420/Web_NP3870_11.5
2AS423/Web_NP3870_11.5
2AU191/Web_NP5115_7.5
2AU212/Web_NP5115_7.5
2AU214/Web_NP5116_9.0
2AU252/Web_NP3459_8.0
2AU265/Web_NP3895_8.5
2AU275/Web_NP3895_8.5
2AU351/Web_NP5211_7.0
2AU372W/Web_NP5115_7.5
2AU385/Web_NP5115_7.5
2AU444/Web_NP5115_7.5
2AU555/Web_NP3460_7.5
2AX123CAP/Lid_NY_white Polypropylene Cap with Foam Liner_02
2BA284R/Web_NP8504_11.0
2BA357R/Web_NP8600_12.0
2BA402R/Web_NP8504_R_11.0
2BA426SLI/Lid_01A
2BA436R/Web_NP8535_12.375
2BA436SLI/Lid_01B
2BA436XF2B/Web_NP9225_8.5
2BA446TP/Web_NP8508_9.375
2BA461R-1/Web_NP8535_16.5
2BA466R/Web_NP8530_11.0
2BA466SLI/Lid_01G
2BA490SLI/Lid_14E
2BA580SLI/Lid_01B
2BA585SLI/Lid_19AB
2BA634R/Web_NP8530_13.625
2BA680R/Web_NP8530_11.0
2BA682R/Web_NP8530_11.0
2BA682SLI/Lid_01B
2BA735/Web_NP8810_12.125
2BA766SLI/Lid_14M
2BA819R/Web_NP8535_12.25
2BA849SLI/Lid_01B
2BA868/Canister_10A
2BA911/Web_NP8810_12.125
2BA922/Web_NP8810_12.125
2BA941/Web_NP8530_15.75
2BA941SLI/Lid_NY_NP7600_22M
2BA972/Web_NP8540_13.5
2BB196/Web_NP5114_41.0
2BB293/Web_NP5410_8.0
2BB320/Web_NP5115_4.0
2BB3978/Web_NP5410_8.0
2BB512/Web_NP3870_4.5
2BB6038/Web_NP5410_8.0
2BB8568/Web_NP5410_8.0
2BC425/Web_NP5410_6.0
2BC801/Web_NP3870_6.0
2BD246/Web_NP5115_4.5
2BD249/Web_NP5115_4.5
2BD365/Web_NP5115_4.5
2BD399/Web_NP5115_7.5
2BD454/Web_NP5115_7.5
2BD724/Web_NP5115_4.5
2BG002/Web_NP9503_26.0
2BG524/Web_NP9503_26.0
2BH115/Web_NP5115_7.5
2BH126/Web_NP5115_7.5
2BH128/Web_NP5115_7.5
2BH140/Web_NP5115_7.5
2BH201/Web_NP5115_7.5
2BH211/Web_NP5116_6.0
2BH217/Web_NP5115_7.5
2BH220/Web_NP5115_7.5
2BH225/Web_NP5115_7.5
2BH592/Web_NP3895_6.0
2BJ123/Web_NP8550_11
2BJ143/Web_NP8530_16.25
2BJ213/Web_NP8530_18
2BJ227/Web_NP8530_15.75
2BM233SLI/Lid_01J
2BM420R/Web_NP8535_15.0
2BM466R/Web_NP8595_16.25
2BM924SLI/Lid_01V
2BP043/Analytical -BUCKET_ Table Turners printed Bucket
2BP156/Web_NP8508_6.88
2BP308SLI/Lid_Aveeno Ultra-Calming MUR
2BP378/Web_NP8508_10.25
2BP423/Web_NP8508_8.5
2BP601/Web_NP8512_12.205
2BP617LB/Canister_04A
2BP721/Web_NP8512_12.205
2BP870/Web_NP8508_10.25
2BQ041R/Web_NP8530_10.375
2BQ084LI/Lid_11A
2BQ387/Web_NP8530_12.125
2BQ471LI/Lid_04F_560
2BQ551LI/Lid_13A
2BS420/Web_NP3870_11.5
2BS423/Web_NP3870_11.5
2BS428/Web_NP3860_8.0
2BU212/Web_NP5115_7.5
2BU214/Web_NP5116_9.0
2BU275/Web_NP3895_8.5
2BU351/Web_NP5211_7.0
2BU385/Web_NP5115_7.5
2BU451/Web_NP5116_9.0
2C3870/Web_NP3870_12
2CA284R/Web_NP8504_11.0
2CA437/Web_NP8530_14.375
2CA580/Web_NP8540_17.0
2CA580SLI/Lid_01R
2CA630/Web_NP8530_14.375
2CA682R/Web_NP8530_11.0
2CA682SLI/Lid_NY_NP7600_22D
2CA735/Web_NP8810_12.125
2CA911/Web_NP8810_12.125
2CA922/Web_NP8810_12.125
2CA941SLI/Lid_NY_NP7600_22M1
2CB100/Web_NP3870_4.5
2CB196/Web_NP5114_41.0
2CB3798/Web_NP5410_8.0
2CB512/Web_NP3870_4.5
2CB6038/Web_NP5410_8.0
2CC334/Web_NP5410_6.0
2CD246/Web_NP5115_4.5
2CD333/Web_NP5115_4.5
2CD454/Web_NP5115_4.5
2CD724/Web_NP5115_4.5
2CH126/Web_NP5115_7.5
2CH128/Web_NP5115_7.5
2CH211/Web_NP5116_6.0
2CH217/Web_NP5115_7.5
2CH225/Web_NP5115_7.5
2CH247/Web_NP3895_6.0
2CH540/Web_NP3895_6.0
2CK457/Web_NP5130_7.0
2CM225SLI/Lid_01K
2CM321R/Web_NP8530_16.375
2CM321SLI/Lid_14C
2CP156/Web_NP8508_6.88
2CP172/Web_NP8512_13
2CQ024R/Web_NP8530_12.25
2CQ181TP/Web_NP8535_10.25
2CQ349R/Web_NP8530_12.25
2CQ467R/Web_NP8550_11.75
2CQ496/Web_NP8508_11.125
2CU132/Web_NP3895_8.5
2CU214/Web_NP5116_9.0
2CU351/Web_NP5211_7.0
2D5115U/Web_NP5115U_9.0
2D5130U/Web_NP5130U_9
2DA437/Web_NP8530_14.375
2DA580/Web_NP8540_17.0
2DA630/Web_NP8530_14.375
2DB007/Web_NP5410_8.0
2DB010/Web_NP5115_4.5
2DB108/Web_NP5114_5.0
2DB114/Web_NP3860_5.0
2DD198/Web_NP5115_4.5
2DD249/Web_NP5116_4.5
2DD306/Web_NP5115_4.5
2DD310/Web_NP5410_4.5
2DD346/Web_NP5410_4.5
2DD419/Web_NP5410_4.5
2DD436/Web_NP5115_4.5
2DD438/Web_NP5115_4.5
2DH070/Web_NP5115_6.0
2DH128/Web_NP5115_7.5
2DH143/Web_NP5115_7.5
2DH211/Web_NP5116_6.0
2DH217/Web_NP5115_7.5
2DJ141R/Web_NP8535_16.0
2DMC288SLI/Lid_14D
2DP710LB/Canister_04A
2DS324/Web_NP3860_12.5
2DS428/Web_NP3860_8.0
2DU351/Web_NP5211_7.0
2DU385/Web_NP5115_7.5
2EB103/Web_NP5115_4.5
2EB339/Web_NP5410_8.0
2EB470/Web_NP5115_4.5
2EB6038/Web_NP5410_8.0
2ED110/Web_NP5410_4.5
2ED363/Web_NP5115_4.5
2ED419/Web_NP5410_4.5
2ED436/Web_NP5115_4.5
2ED454/Web_NP5115_4.5
2ED655/Web_NP5410_4.5
2ED716/Web_NP5115_4.5
2EH070/Web_NP5115_6.0
2EH143/Web_NP5115_7.5
2EJ141R/Web_NP8535_16
2EQ345/Web_NP8535_13.5
2EQ432RA/Web_NP8550_12.375
2EQ700/Web_NP8530_13.5
2EU351/Web_NP5211_7.0
2FB103/Web_NP5115_4.5
2FB108/Web_NP5114_5.0
2FB114/Web_NP3860_5.0
2FB470/Web_NP5115_4.5
2FB6038/Web_NP5410_8.0
2FB688/Web_NP5115_4.0
2FC124/Web_NP3870_6.0
2FC699/Web_NP5410_6.0
2FD748/Web_NP5115_4.5
2FH040/Web_NP5115_6.0
2FH070/Web_NP5115_6.0
2FJ141R/Web_NP8535_16
2FQ345/Web_NP8530_12.125
2FQ700/Web_NP8530_12.125
2FS324/Web_NP3860_12.5
2FS407/Web_NP3860_13.0
2FS419/Web_NP3885_16.0
2FS428_/Web_NP3860_18.0
2FS484/Web_NP5130_11.5
2FS849/Web_NP5130_11.5
2GA580/Web_NP8540_17.0
2GB108/Web_NP5114_5.0
2GB114/Web_NP3860_5.0
2GB349/Web_NP5115U_4.5
2GB470/Web_NP5115_4.5
2GB803/Web_NP5115U_4.0
2GC124/Web_NP3870_6.0
2GC223/Web_NP3870_6.0
2GC699/Web_NP5410_6.0
2GD748/Web_NP5115_4.5
2GH040/Web_NP5115_6.0
2GH356/Web_NP5115_6.0
2GQ345/Web_NP8530_12.125
2GQ700/Web_NP8530_12.125
2GS162/Web_NP5130_11.5
2GS419/Web_NP3885_16.0
2GS428/Web_NP3860_8.0
2GS484/Web_NP5130_11.5
2GU120/Web_NP5115_8.5
2HA580/Web_NP8540_17.0
2HB055/Web_NP5410_8.0
2HB349/Web_NP5115U_4.5
2HB598/Web_NP5115_4.5
2HB687/Web_NP5115_4.5
2HC124/Web_NP3870_6.0
2HC223/Web_NP3870_6.0
2HC699/Web_NP5410_6.0
2HD748/Web_NP5115_4.5
2HH040/Web_NP5115_6.0
2HQ345/Web_NP8530_12.125
2HQ700/Web_NP8530_12.125
2HS162/Web_NP5130_11.5
2HS273/Web_NP3860_15.5
2HS424/Web_NP5130_11.5
2HS484/Web_NP5130_11.5
2HS821/Web_NP3870_11.5
2HU120/Web_NP5115_8.5
2JB108/Web_NP5114_5.0
2JB114/Web_NP3860_5.0
2JB164/Web_NP5410_4.0
2JB406/Web_NP3870_4.5
2JB687/Web_NP5115_4.5
2JC124/Web_NP3870_6.0
2JC223/Web_NP3870_6.0
2JC699/Web_NP5410_6.0
2JH040/Web_NP5115_6.0
2JQ345/Web_NP8530_12.125
2JS162/Web_NP5130_11.5
2JS324/Web_NP3860_12.5
2JS324/Web_NP3860_12.5
2JS407/Web_NP3885_13.0
2JS413/Web_NP3870_11.5
2JS419/Web_NP3885_16.0
2JS424/Web_NP5130_11.5
2JS428/Web_NP3860_8.0
2JS821/Web_NP3870_11.5
2KB108/Web_NP5114_41.0
2KB111/Web_NP5115_4.5
2KB114/Web_NP3860_41
2KB406/Web_NP3870_10.0
2KB598/Web_NP5115_4.5
2KB712/Web_NP5115_4.5
2KC296/Web_NP5410_6.0
2KH040/Web_NP5115_6.0
2KS273/Web_NP3860_15.5
2KS413/Web_NP3870_11.5
2KS419/Web_NP3885_16
2KS424/Web_NP5130_11.5
2KS428/Web_NP3860_8.0
2KS480/Web_NP3870_11.5
2KS821/Web_NP3870_11.5
2KU872/Web_NP5115_8.5
2L5115U/Web_NP5115U_12.0
2LB108/Web_NP5114_41.0
2LB111/Web_NP5115_4.5
2LB114/Web_NP3860_5.0
2LB164/Web_NP5410_4.0
2LB406/Web_NP3870_10.0
2LB598/Web_NP5115_4.5
2LB712/Web_NP5115_4.5
2LD351/Web_NP5410_4.5
2LS324/Web_NP3860_12.5
2LS407/Web_NP3885_13.0
2LS413/Web_NP3870_11.5
2LS419/Web_NP3885_16.0
2LS428/Web_NP3860_8.0
2LS821/Web_NP3870_11.5
2LU872/Web_NP5115_8.5
2MB108/Web_NP5114_41.0
2MB114/Web_NP3860_5.0
2MB114X/Web_NP3860_5.0
2MB164/Web_NP5410_8.0
2MB712/Web_NP5115_4.5
2MD351/Web_NP5116_4.5
2MS273/Web_NP3860_15.5
2MS324/Web_NP3860_12.5
2MS407/Web_NP3885_13.0
2MS411/Web_NP3870_11.5
2MS413/Web_NP3870_11.5
2MS419/Web_NP3885_16
2MS428/Web_NP3860_8.0
2MS480/Web_NP3870_11.5
2MU872/Web_NP5115_8.5
2NB108/Web_NP5114_41.0
2NB114X/Web_NP3860_5.0
2ND351/Web_NP5116_4.5
2NQ594/Web_NP8508_6.88
2NS273/Web_NP3860_15.5
2NS324/Web_NP3860_12.5
2NS407/Web_NP3885_13.0
2NS411/Web_NP3870_11.5
2NS419/Web_NP3885_16
2NS428/Web_NP3860_8.0
2NT001/Web_NP3870_6.0
2NT002/Web_NP3870_8.5
2NU872/Web_NP5115_8.5
2OA105SF/Not tested
2OA119R/Web_NP8530_12.375
2OA119S3B/Web_NP9225_11.0
2OA119SLI/Lid_14B
2OA226BG/Web_NP9225_7.75
2OA226R/Web_NP8535_12.625
2OA284SLI/Lid_22A
2OA302/Web_NP8530_15.75
2OA302SLI/Lid_01B
2OA347/Web_NP8550_10.25
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle
2OA366RNY/Web_NP8530_11.0
2OA402SLI/Lid_01A
2OA405/Web_NP8530_15.75
2OA405SLI/Lid_01B
2OA436FR/Web_NP8535_12.375
2OA437B/Web_NP8530_14.375
2OA445SLI/Lid_16A
2OA500SLI/Lid_01B
2OA520/Web_NP8530_11.0
2OA520SLI/Lid_01B
2OA527R/Web_NP8535_12.5
2OA565R/Web_NP8535_13
2OA578SLI/Lid_14L
2OA592/Web_NP8504_11.0
2OA592SLI/Lid_NY_NP7600_22S
2OA621SLI/Lid_01A
2OA630B/Web_NP8530_14.375
2OA632/Web_NP8504_11.75
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes
2OA641SLI/Lid_14H
2OA657/Web_NP8530_11.0
2OA657SLI/Lid_01B
2OA662R/Web_NP8535_11.5
2OA664/Web_NP8530_15.75
2OA664SLI/Lid_NY_NP7600_22M2
2OA670/Web_NP8535_14.5
2OA670SLI/Lid_20A
2OA671SLI/Lid_14K
2OA682SLI/Lid_01C
2OA709/Web_NP8504_11.0
2OA711/Web_NP8530_11.0
2OA712/Web_NP8530_11.0
2OA712SLI/Lid_01L
2OA714SLI/Lid_01V
2OA721SLI/Lid_01B
2OA743R/Web_NP8530_10.5
2OA759/Web_NP8530_11.0
2OA8026NU/Web_NP8530_11.5
2OA8026T/Tub with Hinged Lid
2OA8101S3B/Web_NP9225_10
2OA8121R/Web_NP8530_11.5
2OA8121SLI/Lid_NY_Kidgets Flushable Toddler
2OA863SLI/Lid_NY_NP7600_22L
2OA899R/Web_NP8530_15.0
2OA9009SLI/Lid_NY_Summer's Eve Aloe Love
2OA9010SLI/Lid_NY_Summer's Eve Island Splash
2OA9011SLI/Lid_NY_Summer's Eve Sheer Floral
2OA9012SLI/Lid_NY_Summer's Eve Simply Sensitive
2OA9013SLI/Lid_NY_Summer's Eve Aloe Love (Canada)
2OA9016SLI/Lid_NY_Summer's Eve Simply Sensitive (Canada)
2OA9021R/Web_NP8550_10
2OA9021SLI/Lid_NY_CVS Health Flushable Cleansing Cloth
2OA911/Web_NP8810_12.125
2OA922/Web_NP8810_12.125
2OA922SLI/Lid_01U
2OA940/Web_NP8530_11.687
2OA941/Web_NP8530_15.75
2OA941SLI/Lid_NY_NP7600_22B
2OA961CS/Clamshell_NY_218.4-119.5-28.8-55
2OA972/Web_NP8540_13.25
2OB123/Web_NP5114_41.0
2OB196/Web_NP3860_41.0
2OB293/Web_NP5410_8.0
2OB347/Web_NP5410_8.0
2OB412/Web_NP5410_8.0
2OB530/Web_NP5115_4.5
2OB714/Web_NP5115_4.5
2OB963/Web_NP5115_4.5
2OC421/Web_NP3870_6.0
2OC455/Web_NP5115_6.0
2OC610/Web_NP5115_6.0
2OC801/Web_NP3870_6.0
2OD015/Web_NP5115_7.5
2OD024/Web_NP5115_7.5
2OD029/Web_NP5115_7.5
2OD177/Web_NP5115_4.5
2OD239/Web_NP5115_7.5
2OD240/Web_NP5115_4.5
2OD242/Web_NP5410_4.5
2OD247/Web_NP5115_7.5
2OD333/Web_NP5115_4.5
2OD384/Web_NP5115_4.5
2OD399/Web_NP5115_7.5
2OD454/Web_NP5115_4.5
2OD470/Web_NP5115_4.5
2OD507/Web_NP5115_4.5
2OD535/Web_NP5115U_4.5
2OD612/Web_NP5115_7.5
2OD621/Web_NP5115_4.5
2OD644/Web_NP5115_7.5
2OD646/Web_NP5115_4.5
2OD703/Web_NP5115_4.5
2OD712/Web_NP5115_4.5
2OD718/Web_NP5115_4.5
2OD725/Web_NP5115_4.5
2OD752/Web_NP5115U_4.5
2OD766/Web_NP5115_4.5
2OD785/Web_NP5115_4.5
2OG201R/Web_NP5115_14.0
2OG306/Web_NP9503_26.0
2OH061/Web_NP3895_6.0
2OH126/Web_NP5115_7.5
2OH128/Web_NP5115_7.5
2OH129/Web_NP5115_7.5
2OH130/Web_NP5115_7.5
2OH131/Web_NP5115_7.5
2OH132/Web_NP5115_7.5
2OH140/Web_NP5115_7.5
2OH146/Web_NP5115_7.5
2OH158/Web_NP5116_6.0
2OH191/Web_NP5115_7.5
2OH201/Web_NP5115_7.5
2OH215/Web_NP5115_7.5
2OH220/Web_NP5115_7.5
2OH225/Web_NP5116_7.5
2OH269/Web_NP5115_7.5
2OH301/Web_NP5115_7.5
2OH441/Web_NP5115_7.5
2OH450/Web_NP5115_7.5
2OH460/Web_NP3450_7.5
2OH510/Web_NP5115_7.5
2OH565/Web_NP5115U_7.5
2OH581/Web_NP3895_6.0
2OH592/Web_NP5115_6.0
2OH617/Web_NP5115_7.5
2OH725/Web_NP5116_6.0
2OJ112/Web_NP8535_16
2OJ112SLI/Lid_01VA
2OJ123/Web_NP8550_11.0
2OJ123SLI/Lid_NY_NP7600_22E
2OJ124/Web_NP8530_11.0
2OJ124SLI/Lid_NY_Cetaphil Gentle Skin cleansing wipes solo
2OJ128/Web_NP8530_11.0
2OJ128SLI/Lid_NY_NP7600_22T
2OJ130/Web_NP8530_11.0
2OJ130SLI/Lid_NY_NP7600_22Y
2OJ131/Web_NP8530_14
2OJ132/Web_NP8530_14
2OJ133/Web_NP8530_8.5
2OJ137/Web_NP8530_10.25
2OJ137SLI/Lid_NY_NP7600_22Q
2OJ141LI/Lid_02A
2OJ143/Web_NP8535_16.25
2OJ143SLI/Lid_01B
2OJ151/Web_NP8530_11.0
2OJ151SLI/Lid_NY_NP7600_22R
2OJ165/Web_NP8535_14
2OJ170/Web_NP8530_11.0
2OJ170SLI/Lid_NY_NP7600_22F
2OJ171/Web_NP8615_9.625
2OJ171SLI/Lid_NY_NP7600_22G
2OJ190/Web_NP8530_11.0
2OJ190SLI/Lid_NY_NP7600_22J
2OJ200/Web_NP8530_11.0
2OJ200SLI/Lid_NY_NP7600_22N
2OJ213SLI/Lid_01A
2OJ227/Web_NP8530_15.75
2OJ230/Web_NP8530_11.0
2OJ230SLI/Lid_NY_NP7600_22P
2OJ291SLI/Lid_01A
2OJ315/Web_NP8535_13.25
2OJ315SLI/Lid_20A
2OJ324/Web_NP8530_11.0
2OJ324SLI/Lid_NY_NP7600_22K
2OJ660/Web_NP8535_13.25
2OJ660SLI/Lid_01B
2OJ701/Web_NP8540_10.125
2OM157R/Web_NP8530_15.5
2OM159SLI/Solo Lid w/Signature care_01
2OM225SLI/Lid_01A
2OM233SLI/Lid_01A
2OM304/Web_NP8540_16.0
2OM304SLI/Lid_No-Rinse sanitizing Solo
2OM347SLI/Solo Lid w/CVS_01
2OM418SLI/Lid_14B
2OM423SLI/Lid_14F
2OM439SLI/Lid_14G
2OM456SLI/Lid_14N
2OM615SLI/Lid_14A
2OM701/Web_NP8540_16.0
2OM701SLI/Lid_No-Rinse sanitizing Solo_01
2OM8002SL/Lid_14OB
2OM8014LID/Lid_NY_Tub Cover Assembly (Pearl)
2OM8014T/Tub_NY_02
2OM8023LID/Lid_NY_Tub Cover Assembly (Blue)
2OM8052S3B/Web_NP9225_13.5
2OM8095NUB/Web_NP8550_14
2OM814SLI/Lid_14J
2OM9000SL/Lid_NY_Oval Solo Lid
2OM9007S3B/Web_NP9250_12.75
2OM9017SL/Lid_NY_Blue Solo Lid w/kirkland Debossing
2OM924R/Web_NP8760_17.125
2OM924SLI/Lid_01B
2OM945SLI/Lid_14O
2OMC288SLI/Lid_14A/D
2OMC4527NU/Web_NP8530_15.75
2OMC799R/Web_NP8535_13.875
2OMC8317/Web_NP8420_15.2
2ON135/Web_NP5115_7.0
2OP017LI/Lid_18B
2OP023LI/Lid_08A
2OP027/Web_NP8512_310
2OP028/Web_NP8530_8.25
2OP034/Web_NP8530_8.25
2OP036/Reduced Testing_NY_Can 2OP036
2OP036C/Reduced Testing_NY_Can 2OP036C
2OP036LI/Lid_04J_560
2OP037/Web_NP8511_245
2OP043/Analytical -BUCKET_ Jet 38 3P Gladator bucket w/handle
2OP043LI/Lid_NY_DET 200-15P Gladiator lid
2OP051/Web_NP8541_11
2OP062/Web_NP8508_8.5
2OP063/Web_NP8511_320
2OP065/Web_NP8536_11.875
2OP067/Web_NP8536_11.875
2OP072/Web_NP8511_245
2OP073/Web_NP8536_11.875
2OP075/Web_NP8541_12.205
2OP079/Web_NP8541_12.205
2OP081/Web_NP8536_11.875
2OP087/Web_NP8550_12.205
2OP091/Web_NP8550_12.205
2OP092/Web_NP8536_10
2OP093/Analytical -Jar _10 oz Blue HDPE Grade Jar
2OP093LI/Lid_NY_Threaded Blue Cap
2OP097/Web_NP8508_12.0
2OP104/Web_NP8508_12.0
2OP109/Web_NP8530_8.25
2OP113/Web_NP8541_11
2OP116LI/Lid_18B
2OP125LI/Lid_06A
2OP127LI/Lid_06B
2OP131/Canister_01A
2OP134/Reduced Testing_NY_Can 2OP134
2OP134LI/Lid_03L_560
2OP135LI/Lid_03A
2OP138/Reduced Testing_NY_Can 2OP138
2OP138C/Reduced Testing_NY_Can 2OP138C
2OP140/Web_NP8512_310
2OP144/Web_NP8512_12.205
2OP144SLI/Lid_01G
2OP147/Web_NP8512_310
2OP147SLI/Lid_01F
2OP149/Web_NP8512_12.625
2OP150/Web_NP8512_310
2OP150LI/Lid_04E_560
2OP150SLI/Lid_01F
2OP155/Web_NP8512_13.25
2OP155SLI/Lid_01G
2OP155TY/Analytical -Tray_NY_10.38-5-1.5-14.17
2OP158/Web_NP8512_310
2OP159LI/Lid_03B
2OP160/Web_NP8512_310
2OP161/Web_NP8512_310
2OP162/Web_NP8512_310
2OP162SLI/Lid_01O
2OP163LI/Lid_18C
2OP169/Web_NP8512_12.205
2OP169SLI/Lid_01N
2OP172/Web_NP8512_13.25
2OP176/Web_NP8512_12.205
2OP177/Web_NP8512_12.205
2OP178/Web_NP8512_12.205
2OP181/Web_NP8512_12.205
2OP183/Web_NP8541_11
2OP189/Web_NP8536_9
2OP190/Web_NP8512_12.205
2OP195/Web_NP8511_320
2OP197/Web_NP8511_320
2OP203/Web_NP8535_12.0
2OP204/Web_NP8530_11.0
2OP204SLI/Lid_NY_NP7600_22H
2OP205/Web_NP8536_11.875
2OP210/Web_NP8541_11
2OP211/Web_NP8512_12.205
2OP213/Web_NP8541_11
2OP214/Web_NP8550_12.205
2OP214LI/Lid_03C
2OP223/Web_NP8511_320
2OP225/Web_NP8512_13.25
2OP227/Web_NP8541_11
2OP229/Web_NP8512_9.5
2OP230/Web_NP8512_12.205
2OP237LI/Lid_10A
2OP238LI/Lid_04G_560
2OP239/Web_NP8512_310
2OP242LI/Lid_17C
2OP245LI/Lid_04K_.560
2OP252/Web_NP8536_11.875
2OP253LI/Lid_17B
2OP260/Web_NP8535_12.75
2OP262/Web_NP8541_11
2OP278LI/Lid_03D
2OP279/Web_NP8512_12.205
2OP280LI/Lid_17A
2OP284/Web_NP8512_12.205
2OP287LI/Lid_04H_560
2OP300LI/Lid_13E
2OP308/Web_NP8512_12.205
2OP308SLI/Lid_19A
2OP312/Web_NP8541_12.205
2OP328/Web_NP8512_12.205
2OP337/Web_NP8512_310
2OP345/Web_NP8508_12.0
2OP348LI/Lid_18D
2OP350LI/Lid_05A
2OP352LI/Lid_08E
2OP354/Web_NP8508_8.5
2OP365/Web_NP8536_11.875
2OP374/Web_NP8508_10.0
2OP382/Web_NP8530_8.25
2OP383LI/Lid_05H
2OP384LI/Lid_05I
2OP385LI/Lid_18A
2OP392LI/Lid_18E
2OP421/Web_NP8535_8.0
2OP423/Web_NP8508_8.5
2OP435LI/Lid_04C_560
2OP448LI/Lid_18B
2OP456LI/Lid_08C
2OP460/Web_NP8530_12.705
2OP467/Web_NP8508_9.5
2OP485/Web_NP8512_12.205
2OP487/Web_NP8508_9.5
2OP506/Web_NP8508_8.5
2OP509LI/Lid_08B
2OP510LI/Lid_08D
2OP520LI/Lid_13B
2OP523LI/Lid_05B
2OP527LB/Canister_04A
2OP527LI/Lid_10B
2OP560/Web_NP8440_310
2OP563/Web_NP8530_12.205
2OP572/Web_NP8522_12.204
2OP579/Web_NP8508_10.25
2OP585/Web_NP8512_12.205
2OP585SLI/Lid_19A
2OP587/Web_NP8550_12.205
2OP601/Web_NP8512_12.2
2OP617LI/Lid_10C
2OP624/Web_NP8550_11.0
2OP624SLI/Lid_22A
2OP630/Web_NP8440_310
2OP638/Reduced Testing_NY_Can 2OP638
2OP638C/Reduced Testing_NY_Can 2OP638C
2OP638LI/Lid_13D
2OP640/Web_NP8508_10.0
2OP645/Web_NP8440_310
2OP648/Canister_04B
2OP657LI/Lid_05C
2OP674LB/Canister_04A
2OP674LI/Lid_10D
2OP683LI/Lid_03E
2OP705/Web_NP8530_9.0
2OP710/Canister_11A
2OP713/Web_NP8530_8.25
2OP721/Web_NP8512_12.2
2OP728/Web_NP8508_6.88
2OP742SF/Web_NP6360_28.0
2OP755/Web_NP8530_9.0
2OP776LI/Lid_06C
2OP806/Web_NP8430_310
2OP812LB/Canister_05A
2OP832LI/Lid_05D
2OP838LI/Lid_07A
2OP843LI/Lid_13C
2OP870/Web_NP8508_10.25
2OP890/Web_NP8530_11.75
2OP901LI/Lid_05E
2OP906/Web_NP8530_12
2OP906TY/Analytical -Tray_NY_352-122-70-33
2OP913/Web_NP8530_12
2OP920LI/Lid_03F
2OQ001SLI/Lid_15A
2OQ066SLB/Lid_14P
2OQ066TR/Analytical -Tray_NY_Display Tray_01
2OQ111R/Web_NP8535_9.375
2OQ141LI/Lid_05F
2OQ216SLI/Lid_NY_CVS Beauty 360 Micellar MUA
2OQ224TR/Analytical -Tray_NY_14.12-5.87-1.63-36
2OQ258TP/Web_NP8530_11.125
2OQ271TR/Analytical -Tray_NY_16-4-1.40-30
2OQ335R/Web_NP8530_12.5
2OQ387/Web_NP8530_12.125
2OQ449R/Web_NP8535_11.75
2OQ543R/Web_NP8530_11.75
2OQ558TP/Web_NP8530_10.25
2OQ753/Reduced Testing_NY_Can 2OQ753
2OQ753X/Reduced Testing_NY_Can 2OQ753 X
2OQ797/Canister_02A
2OQ8110R/Web_NP8530_12.5
2OQ850/Reduced Testing_NY_Can 2OQ850
2OQ850LI/Lid_11B
2OQ850X/Reduced Testing_NY_Can 2OQ850X
2OQ869LI/Lid_03G
2OQ875LI/Lid_05G
2OQ882LI/Lid_12A
2OQ890LI/Lid_03H
2OQ9027R/Web_NP8550_12.25
2OQ9027SLI/Lid_NY_buybuy BABY Hand & Face
2OQ918LI/Lid_04D_560
2OQ999SLI/Lid_White Mini-Solo Lid
2OS123/Web_NP3870_11.5
2OS293/Web_NP3860_15.5
2OS356/Web_NP3860_12.5
2OS456/Web_NP3870_11.5
2OU127/Web_NP5116_9.0
2OU128/Web_NP5116_9.0
2OU130/Web_NP5115_9.0
2OU131/Web_NP3895_8.5
2OU157/Web_NP5115_7.5
2OU165/Web_NP5116_8.5
2OU171/Web_NP5116_9.0
2OU191/Web_NP5115_7.5
2OU201/Web_NP5115_7.5
2OU211/Web_NP5116_9.0
2OU240/Web_NP5116_9.0
2OU245/Web_NP5115_7.5
2OU252/Web_NP3459_8.0
2OU262/Sani-snap with Hook/waste
2OU265/Web_NP3895_8.5
2OU275/Web_NP5116_8.5
2OU284/Web_NP5116_8.5
2OU295/Web_NP5115_7.5
2OU345/Web_NP5115_7.5
2OU351/Web_NP5210_7.0
2OU354/Web_NP5115_7.5
2OU372W/Web_NP5115_7.5
2OU385/Web_NP5115_7.5
2OU434/Web_NP5115_7.5
2OU444/Web_NP5115_7.5
2OU446/Web_NP5116_9.0
2OU451/Web_NP5116_9.0
2OU475/Web_NP5115_7.5
2OU500/Web_NP5115_7.5
2OU526/Web_NP5115_7.5
2OU555/Web_NP3460_7.5
2OU573/Web_NP5115_7.5
2OU585/Web_NP5116_9.0
2OU591/Web_NP5116_9.0
2OU611/Web_NP5211_7.0
2OU621/Web_NP5116_9.0
2OU710/Web_NP5116_9.0
2OX105B/Bottle, 10.5ml
2OX123BTL/Analytical -canister _32oz White HDPE Spray Bottle
2OX123CAP/Lid_NY_white Polypropylene Cap with Foam Liner
2OX123TRG/Lid_NY_white on/off Trigger Sprayer
2OX131RBTL/Analytical -canister _32oz White HDPE Spray Bottle, 28/400 Ratchet Neck
2OX131RTRG/Lid_NY_white on/off passive vent Trigger Sprayer
2OX131T/Lid_NY_white on/off Passive Vent Trigger Sprayer
2OX465B/Bottle, 26ml
2OX465FS/Foil seal 26 ml
2OX465ROLL/Foil Seal Roll
2OX510B/Bottle, 5.1ml
2OX510FS/Foil seal 5.1 ml
2PB108/Web_NP5114_41.0
2PB114X/Web_NP3860_5.0
2PD351/Web_NP5116_4.5
2PQ594/Web_NP8508_6.88
2PS231/Web_NP3870_11.5
2PS273/Web_NP3860_15.5
2PS324/Web_NP3860_12.5
2PS407/Web_NP3885_13.0
2PS411/Web_NP3870_11.5
2PS419/Web_NP3885_16
2PS428/Web_NP3860_8.0
2PS480/Web_NP3870_11.5
2PU872/Web_NP5115_8.5
2QB108/Web_NP5114_41.0
2QB114X/Web_NP3860_5.0
2QS231/Web_NP3870_11.5
2QS273/Web_NP3860_15.5
2QS411/Web_NP3870_11.5
2QS428/Web_NP3860_8.0
2QS480/Web_NP3870_11.5
2RB108/Web_NP5114_41.0
2RB114X/Web_NP3860_5.0
2RP540/Web_NP9508_7
2RQ119/Web_NP9508_6.375
2RQ890/Web_NP9508_18.125
2RQ890P/Web_NP9508_18.125
2RS324/Web_NP3860_12.5
2RS407/Web_NP3885_13.0
2RS419/Web_NP3885_16
2S3870/Web_NP3870_11.5
2S5115U/Web_NP5115U_7.5
2S5130/Web_NP5130U_11.5
2SB108/Web_NP5114_41.0
2SS231/Web_NP3870_11.5
2SS273/Web_NP3860_15.5
2SS428/Web_NP3860_8.0
2TS231/Web_NP3870_11.5
2X3870/Web_NP3870_9.0
2X5116/Web_NP5116_8.5_UP
2XQ551LI/Lid_03I
2XX999SW/Film_20.0
2ZA407SLI/Solo Lid_02
2ZA863SLI/Solo Lid_01
3AA403/Applicator_NY_Applicator – 6” Unembossed Spunlaced Nonwoven
3AA615/Applicator_NY_Rolls Embossed Spunlaced Nonwoven-58gsm-7.25"
3AA621-SN/Applicator_NY_9.8" Rolls Embossed Spunlaced Nonwoven - 45gsm
3AA621/Applicator_NY_9.8" Rolls Embossed Spunlaced Nonwoven - 45gsm
3AA634/Applicator_31A
3AA640/Applicator_NY_70 gsm Dispersible - 6.96"
3AB103/Applicator_NY_40 gsm Thermobond 100% PP Nonwoven – 30 mm – 1 3/16”
3AC124/Applicator_NY_89.7 gsm Rayon/Polypropylene
3AD03-PS/Applicator_NY_60 gsm Dispersible - 7"
3AD04-N/Applicator_NY_70 gsm Dispersible - 24.25"
3AD04-PS/Applicator_NY_60 gsm Dispersible - 25.25"
3AD05-PS/Applicator_NY_60 gsm Dispersible - 27.5"
3AD06-PS/Applicator_NY_70 gsm Dispersible - 7"
3AD08-PS/Applicator_NY_70gsm Dispersible-25.25"
3AD239/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven6”
3AD504/Applicator_15J
3AD659/Applicator_10A
3AH247/Applicator_10E
3AH500/Applicator_19A
3AM321-G/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 8"
3AM321-J/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 8"
3AM451-G/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 8"
3AM451-J/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 2X8"
3AM452-B/Applicator_NY_Rolls Embossed Spunlaced Nonwoven-32"
3AM452-G/Applicator_NY_Rolls Embossed Spunlaced Nonwoven-32"
3AM452-J/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 32"
3AM466-G/Applicator_NY_Rolls Embossed Spunlaced Nonwoven-40"
3AM466-J/Applicator_NY_Rolls Embossed Spunlaced Nonwoven-40"
3AM652/Applicator_NY_58 gsm Premium Airlaid 35" - Hour Glass Emboss
3AM757/Applicator_NY_55gsm Big Dot Embossed Nonwoven
3AM758/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 52gsm - 7.5"
3AM819/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 55gsm - 2X8"
3AM857/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 55gsm - 8"
3AP135/Applicator_NY_32" Rolls Embossed Spunlaced Nonwoven-50 gsm
3AP148/Applicator_01B
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2"
3AP590/Applicator_NY_35” Rolls Spunlaced Nonwoven – 35 gsm
3AP715/Applicator_02D
3AP811/Applicator_NY_28" Rolls Spunlaced Nonwoven - 35 gsm
3AQ341-G/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 37.5"
3AQ347/Applicator_35A
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm
3AQ470-M/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75" - 50gsm
3AQ717/Applicator_16A
3AS004-AT/Applicator_NY_Rolls Spunlaced Nonwoven - 30"-50gsm
3AS004-MT/Applicator_NY_Rolls Spunlaced Nonwoven - 30"-50gsm
3AS006-AT/Applicator_NY_Rolls Spunlaced Nonwoven - 2X7.5"-57gsm
3AS006-J/Applicator_NY_Rolls Spunlaced Nonwoven - 2X7.6" - 57gsm
3AS008-AT/Applicator_NY_Rolls Spunlaced Nonwoven - 960mm - 57gsm
3AS008-J/Applicator_NY_Rolls spunlaced Nonwoven - 37.75" - 57gsm
3AS008-J/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75" - 57gsm
3AS009-AT/Applicator_NY_Rolls spunlaced Nonwoven - 195mm - 57gsm
3AS010-H/Applicator_NY_Rolls Spunlaced Nonwoven Aperture- 9.8" 60gsm
3AS05-AG/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 42"
3AS305-J/Applicator_NY_Rolls Spunlaced Nonwoven - 30.5" - 57gsm
3AS305-S/Applicator_NY_Rolls Spunlaced Nonwoven - 30.5" 57gsm
3AS411/Applicator_NY_Swabstick Rayon-Double Tipped, Gamma Irradiated-Topical Use
3AS806-A/Applicator_NY_Rolls Spunlaced Nonwoven - 30"-57gsm
3AS806-J/Applicator_NY_Nonwoven Rolls Spunlaced Nonwoven - 30"-57gsm
3AS816-A/Applicator_NY_Rolls Spunlaced Nonwoven - 7.375"-57gsm
3ASPN03U-RUC/Applicator_NY_Spunlace Nonwoven - 35 gsm - 510mm
3AT015-MU/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 58gsm 15"
3AT100/Applicator_NY_Rolls Embossed Spunlaced Nonwoven-12.25"-58gsm
3AU120/Applicator_NY_ 10" Embossed Spunlaced Nonwoven - 45 gsm
3AU125/Applicator_NY_ 6 1/4" Unembossed Spunlaced Nonwoven
3AU132/Applicator_10F
3AU134/Applicator_06B
3AU228/Applicator_06I
3BA523/Applicator_NY_Rolls Spunlaced Nonwoven Exfoliating - 29.5" - 56gsm
3BA527/Applicator_NY_51.2 gsm Airlaid Embossed - 35" Roll
3BA618-J/Applicator_06N
3BA640-2/Applicator_NY_66 gsm Dispersible - 6.96"
3BA640/Applicator_NY_70 gsm Dispersible - 6.96"
3BB164/Applicator_01G
3BB241/Applicator_NY_71.6 gsm Cellulose/Polypropylene
3BB406/Applicator_NY_Applicator Textured Thermal Bond – 52.6 gsm
3BC296/Applicator_NY_53.8 gsm Cellulose/Polypropylene
3BD290/Applicator_20A
3BJ141/Applicator_07K
3BM321-G/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 8"
3BM321-J/Applicator_NY_Rolls embossed Spunlaced Nonwoven - 8" (J)
3BM451-G/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 8"
3BM451-J/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 2X8"
3BM452-G/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 32"
3BM452-J/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 32"
3BM758-F/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 52gsm - 7.5"
3BM819/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 55gsm 2X7.5"
3BQ335-H/Applicator_NY_Rolls spunlaced Nonwoven Exfoliating-30.5"-56gsm
3BQ335/Applicator_NY_Rolls Spunlaced Nonwoven Exfoliating - 30.5" - 56gsm
3BQ472/Applicator_NY_ 32" Embossed Spunlaced Nonwoven
3BT335/Applicator_NY_Rolls spunlaced Nonwoven Exfoliating-12.25"-56gsm
3BU872/Applicator_10B3aa403
3CA620/Applicator_NY_Rolls Spunlaced Nonwoven Exfoliating - 7 3/8" - 56gsm
3CC124/Applicator_NY_89.7 gsm Rayon/Polypropylene
3CC223/Applicator_28C
3CM321-G/Rolls embossed Spunlaced nonwoven-8"
3CQ125/Applicator_06C(_35",_28")
3DA620/Applicator_NY_Rolls Spunlaced Nonwoven Exfoliating - 7.375" - 56gsm
3DM612-S/Applicator_NY_Rolls Spunlaced Nonwoven - 2X8" -52gsm
3DM612-SN/Applicator_NY_Rolls Spunlaced Nonwoven - 2X8" -52gsm
3DQ125/Applicator_06D(_35",_28")
3DQ139/Applicator_06E
3DQ185-H/Applicator_NY_Rolls spunlaced Nonwoven Exfoliating-30.5"-56gsm
3DQ185/Applicator_NY_Rolls Spunlaced Nonwoven Exfoliating-30.5"-56gsm
3DQ595/Applicator_10D
3DQ595N/Applicator_10I
3DQ630/Applicator_10C
3DQ630N/Applicator_01D
3DT185/Applicator_NY_Rolls spunlaced Nonwoven Exfoliating-12.25"-56gsm
3OA236-E/Applicator_NY_Rolls Spunlaced Nonwoven - 7.5" - 58 gsm
3OA276/Applicator_15H
3OA301/Applicator_32B
3OA322/Applicator_NY_Rolls Embossed Spunlaced Nonwoven 40 gsm-11 1/2"
3OA347/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm
3OA362/Applicator_NY_Rolls Spunlaced Nonwoven-35gsm-7.2"
3OA392DH/Applicator_NY_39.2" Rolls Embossed Spunlaced Nonwoven - 49 gsm
3OA403/Applicator_NY_6" Unembossed Spunlaced Nonwoven
3OA407/Applicator_NY_Unembossed Spunlaced Nonwoven -55gsm- 7 3/8”
3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm
3OA435/Applicator_NY_50gsm,Hydroentangled Nonwoven-11 1/4"
3OA462/Applicator_NY_Rolls spunlaced Nonwoven -7.1"-62gsm
3OA565/Applicator_NY_11.5” Rolls Embossed Spunlaced Nonwoven – 50 gsm
3OA580/Applicator_38A
3OA588/Applicator_NY_24.5" Unembossed Spunlaced Nonwoven
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm
3OA613/Applicator_06K
3OA615/Applicator_NY_Rolls Embossed Spunlaced Nonwoven -58gsm - 7.5"
3OA618-A/Applicator_06H
3OA618/Applicator_29A
3OA621/Applicator_06F
3OA623/Applicator_07Q
3OA624/Applicator_NY_8" Rolls Embossed Spunlaced Nonwoven - 49gsm
3OA629/Applicator_NY_52.5 gsm Premium Airlaid 6.375" rolls
3OA630/Applicator_8" Embossed Spunlaced Nonwoven
3OA630A/Applicator_NY_8" Embossed Spunlaced Nonwoven
3OA700/Applicator_NY_Rolls Embossed Spunlaced Nonwoven -61gsm - 7"
3OA725/Applicator_NY_Rolls Spunlaced Nonwoven -55gsm-7.25"
3OA750/Applicator_NY_Rolls Embossed spunlaced Nonwoven -55gsm-7.5"
3OA765/Applicator_NY_Rolls Spunlaced Nonwoven - 7" 65gsm
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven
3OA916-2/Applicator_NY_60 gsm Dispersible - 6.96"
3OA916/Applicator_NY_ 60 gsm Dispersible- 6.96"
3OA933/Applicator_NY_43 gsm Flat Omega 7.1"
3OA934/Applicator_NY_43gsm Flat Omega 35"
3OA945/Applicator_NY_Spunlaced Nonwoven - 50gsm - 7.1" - Unembossed
3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss
3OB108/Applicator_NY_Applicator Textured Thermal Bond - 55.8 gsm - 25 ½"
3OB128/Applicator_14C
3OB465/Applicator_28A
3OB603/Applicator_NY_Applicator Thermal Bonded Nonwoven – 30 mm – 1 3/16”
3OC105/Applicator_28B
3OC455/Applicator_14B
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6”
3OC610/Applicator_NY_89.7 gsm Rayon/Polypropylene 3"
3OC699/Applicator_NY_Applicator Thermal Bonded Nonwoven – 44 mm – 1 3/4”
3OC705/Applicator_NY_Rolls Spunlaced Nonwoven - 3.5" - 50gsm
3OD123/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 6"_01
3OD125/Applicator_21A
3OD135/Applicator_NY_SPC Nonwovens Diamond Embosed 30gsm-6"
3OD179/Applicator_23A
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4"
3OD198/Applicator_22A
3OD239/Applicator_06R
3OD304/Applicator_19B
3OD333/Applicator_NY_6” Rolls Spunlaced Nonwoven – 35 gsm
3OD405/Applicator_NY_Rolls Spunlaced Nonwoven – 75.9 gsm – 5 1/8”
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven
3OD504/Applicator_15C
3OD535/Applicator_34A
3OD724/Applicator_NY_SPC Nonwovens Diamond Embosed 30gsm-7 7/8"
3OG350/Applicator_06M
3OG517/Applicator_NY_45gsm, Hydroentangled Nonwoven 48-1/2"
3OH061/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 7.875"
3OH115/Applicator_NY_7.1" Rolls Spunlaced Nonwoven -35 gsm
3OH126/Applicator_NY_ Rolls Spunlaced Nonwoven-6"-50gsm
3OH126A/Applicator_NY_Rolls Spunlaced Nonwoven - 6" - 50gsm
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7”
3OH155/Applicator_07P
3OH225/Applicator_08C
3OH250/Applicator_NY_Paper Crepe 23 lb., 10 3/4"
3OH334/Applicator_NY_Spunlaced Nonwoven - 50 gsm - 7.1"
3OH334TK/Applicator_NY_Spunlaced Nonwoven - 50gsm - 7.2" - Unembossed
3OH460/Applicator_05A
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7”
3OH592/Applicator_36A
3OH633/Applicator_NY_Applicator Thermal Bonded Nonwoven 25gsm
3OH725/Applicator_15G
3OJ137/Applicator_NY_70 gsm Dispersible - 6"
3OJ171/Applicator_NY_40gsm Hydroentangled Nonwoven 6" with Diamond Emboss
3OJ213/Applicator_08A
3OJ660/Applicator_08B
3OM210/Applicator_NY_52.5 gsm Premium Airlaid 2X8" rolls
3OM211/Applicator_NY_52.5 gsm Premium Airlaid 32"
3OM248/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 2X8"
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9"
3OM353-M/Applicator_NY_Embossed Spunlaced Nonwoven-2X8.125"-50 gsm
3OM353-SN/Applicator_NY_Embossed Spunlaced Nonwoven-2X8"-50 gsm
3OM427-M/Applicator_NY_Embossed Spunlaced Nonwoven-8.125"-50 gsm
3OM460-J/Applicator_NY_Rolls Spunlaced Nonwoven-2X7.5"-60gsm
3OM532-M/Applicator_NY_Embossed Spunlaced Nonwoven - 8.125" - 50gsm
3OM630-ASN/Applicator_NY_Rolls Spunlaced Nonwoven - 7.28" - 50gsm
3OM630-M/Applicator_NY_Spunlaced Nonwoven 50 gsm Flat
3OM630-SN/Applicator_NY_Rolls Spunlaced Nonwoven - 7.35" - 48gsm
3OM630-SN/Applicator_NY_Rolls Spunlaced Nonwoven - 7.35" - 50gsm
3OM652/Applicator_32A
3OM685/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 35" - 50 gsm
3OM757/Applicator_NY_55 gsm Spunlaced CSC Nonwoven 2X8 inch rolls
3OM758/Applicator_NY_55 gsm Spunlaced CSC Nonwoven 8 inch rolls
3OM765/Applicator_NY_Rolls Spunlaced Nonwoven - 42" - 65gsm
3OM814/Applicator_NY_Rolls Spunlaced Nonwoven - 2X8" -45gsm
3OM816/Applicator_NY_Rolls Spunlaced Nonwoven - 8" - 45gsm
3OM819/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 55gsm-2X8"
3OM857/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 55gsm-8.0125"
3OM924/Applicator_01I
3OM954-TJ/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 35"
3OM954/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 35"
3OM975/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 55gsm - 2X7.5"
3OM996/Applicator_NY_58 gsm Premium Airlaid 7" - Hour Glass Emboss
3OP023/Applicator_15B
3OP036/Applicator_NY_40gsm Hydroentangled Nonwoven 36" with Diamond Emboss
3OP043/Applicator_NY_Rolls Wet Laid Nonwoven - 40" - 54gsm
3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm
3OP078/Applicator_NY_37 1/2" Rolls Spunlaced Nonwoven - 35 gsm
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30"
3OP093/Applicator_NY_Embossed Nonwoven - 24"
3OP093PD/Applicator_NY_Die Cut Pads
3OP095/Applicator_06L
3OP113/Applicator_NY_Flat CSC Spunlaced Nonwoven - 45 gsm - 30"
3OP120/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 30”
3OP134/Applicator_NY_35gsm Stora Flat - 30"
3OP135/Applicator_07R
3OP136/Applicator_NY_35 gsm Unembossed Nonwoven
3OP138/Applicator_36D
3OP138N/Applicator_NY_Spunlaced Nonwoven - 40gsm - 30" - Diamond Emboss
3OP142/Applicator_07S
3OP144/Applicator_NY_Unembossed Spunlaced Nonwoven -50gsm 30”
3OP156/Applicator_06A
3OP159/Applicator_NY_36 gsm Unembossed Nonwoven – 36”
3OP175/Applicator_07O
3OP176/Applicator_NY_35” - 889 mm Rolls Embossed Spunlaced Nonwoven – 44 gsm
3OP179/Applicator_NY_28” Rolls Embossed Spunlaced Nonwoven – 45 gsm
3OP195/Applicator_NY_Spunlaced Embossed Nonwoven - 37.5"
3OP197/Applicator_NY_Spunlaced Embossed Nonwoven -37.5"
3OP197A-SU/Applicator_NY_Spunlaced Embossed Nonwoven - 900
3OP197A/Applicator_NY_Spunlaced Embossed Nonwoven -35"
3OP199/Applicator_NY_Spunlaced Nonwoven- Flat - 37.5"
3OP205/Applicator_NY_Rolls spunlaced Nonwoven Exfoliating - 59gsm
3OP211/Applicator_03A
3OP222/Applicator_NY_SPC Nonwovens Diamond Embosed 30gsm-30"
3OP229/Applicator_NY_Irradiated White, non-apertured Spunlace non-woven with 3mm Dot Pattern, 50gsm
3OP230/Applicator_NY_irradiated,white nonwoven 3mm dot 50gsm
3OP237/Applicator_15A
3OP241/Applicator_NY_Spunlaced Nonwoven - 40gsm-34.25” - Unembossed
3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5"
3OP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 36"
3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5"
3OP280/Applicator_10G
3OP284/Applicator_NY_Spunlaced Nonwoven 55 gsm - 30"
3OP308/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2"
3OP333-T/Applicator_NY_Spunlaced Nonwoven-50gsm-28.5"-Flat
3OP333/Applicator_NY_Spunlaced Nonwoven – 50 gsm – 28.5” - Unembossed
3OP333TK/Applicator_NY_Spunlaced Nonwoven - 50gsm - 40" - Unembossed
3OP334/Applicator_NY_Spunlaced Nonwoven – 50 gsm – 34.25”-Unembossed
3OP338/Applicator_07M
3OP345/Analytical -Applicator_NY_Rolls Embossed Spunlaced Nonwoven-63gsm-24.4"
3OP345G/Applicator_NY_ Rolls Spunlaced Nonwoven-63gsm-24.4"
3OP383/Applicator_06O
3OP435/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 30"
3OP444/Applicator_NY_SPC Nonwovens Diamond Embosed 30gsm-36"
3OP445/Applicator_NY_36 gsm Unembossed Nonwoven – 37 1/2”
3OP448/Applicator_08D
3OP453/Applicator_06P
3OP487/Applicator_08E
3OP490/Applicator_07J
3OP495/Applicator_07F
3OP540/Applicator_01F
3OP555/Applicator_NY_21" Rolls Spunlaced Nonwoven - 35gsm
3OP555TK/Applicator_NY_20.08" Rolls Spunlaced Nonwoven - 35gsm
3OP567/Applicator_NY_Rolls Spunlaced Nonwoven - 31" - 50gsm
3OP572/Applicator_NY_Spunlace Nonwoven 50gsm - 37 1/2"
3OP572X/Applicator_NY_Spunlace Nonwoven 50gsm - 30"
3OP580/Applicator_NY_45 gsm, Hydroentangled Nonwoven 28”
3OP590/Applicator_07E
3OP607/Applicator_07B
3OP638/Applicator_15I
3OP657/Applicator_30A(_28", _35")
3OP666/Applicator_NY_SPC Nonwovens Diamond Embosed 30gsm-37.5"
3OP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-31"
3OP705/Applicator_NY_Rolls Spunlaced Nonwoven – 25 1/2” – 52 gsm
3OP710/Applicator_NY_36 gsm Unembossed Nonwoven – 27 ½”
3OP713/Applicator_NY_Rolls Spunlaced Nonwoven - 20 1/4" -52gsm
3OP715/Applicator_02B
3OP728/Applicator_NY_35 gsm Unembossed Nonwoven
3OP784/Applicator_NY_27" Rolls Spunlaced Nonwoven - 35 gsm
3OP811/Applicator_07D
3OP895/Applicator_07N
3OP901/Applicator_30B(_28", _35")
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5"
3OQ028-D/Applicator_NY_Unembossed Spunlaced Nonwoven -50gsm 37 ½”
3OQ028/Applicator_NY_Unembossed Spunlaced Nonwoven -50gsm 37.5”
3OQ084N/Applicator_NY_3OQ084N Applicator Thermal Bonded Nonwoven – 42”
3OQ095-G/Analytical -Applicator_NY_Rolls Embossed Spunlaced Nonwoven-63gsm 30.5"
3OQ095-M/Analytical -Applicator_NY_Rolls Embossed Spunlaced Nonwoven-63gsm 30.5"
3OQ095/Applicator_06J
3OQ103/Applicator_NY_Unembossed Spunlaced Nonwoven- 55gsm 30 1/2"
3OQ133/Applicator_NY_Rolls Embossed Spunlaced Nonwoven-45gsm-37.5"
3OQ169/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 55gsm-37.5"
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven
3OQ228/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 58gsm 22.8"
3OQ255/Applicator_NY_Rolls Spunlaced Nonwoven-35gsm-25.5"
3OQ300-M/Applicator_NY_Embossed Spunlaced Nonwoven-37.75"-50gsm
3OQ300-SN/Applicator_NY_Embossed Spunlaced Nonwoven - 37 3/4" - 50gsm
3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2"
3OQ367/Applicator_NY_ 70 gsm Dispersible-28.5"
3OQ374-M/Applicator_NY_Nonwoven Rolls unembossed - 7.5" - 50gsm
3OQ431/Applicator_NY_Roll embossed spunlaced nonwoven - 37.5"
3OQ432-M/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 58gms - 30.5"
3OQ432-S/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 58gsm 30.5"
3OQ437/Applicator_NY_36 gsm Unembossed Nonwoven – 30"
3OQ437300/Applicator_17A
3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven
3OQ472/Applicator_11A
3OQ478-M/Applicator_NY_Rolls Embossed Spunlaced Nonwoven-58gsm-37.75"
3OQ500/Applicator_09A
3OQ501/Applicator_NY_37.75” rolls Thermal Bonded Nonwoven 50gsm 80PP/20V
3OQ503-MU/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 58gsm 30.5"
3OQ505/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 58 gsm - 7.4"
3OQ541-E/Applicator_NY_Rolls Spunlaced Nonwoven - 30 1/2" - 58 gsm
3OQ551/Applicator_15E
3OQ573-MU/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 58gsm 37.5"
3OQ637/Applicator_NY_ 70 gsm Dispersible-28.5"
3OQ641-2/Applicator_NY_70 gsm Dispersible - 30.5"
3OQ641-PS/Applicator_NY_70 gsm Dispersible - 30.5"
3OQ641/Applicator_NY_70 gsm Dispersible - 30.5"
3OQ645N/Applicator_NY_30" Rolls Thermal Bonded Nonwoven – 25 gsm
3OQ669/Applicator_13A
3OQ705-MU/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 58gsm 37 1/2"
3OQ705/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 58gsm - 37 1/2
3OQ725/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 58gsm 37.5"
3OQ807/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 58gsm - 30.5"
3OQ845/Applicator_01A
3OQ850/Applicator_NY_Thermal Bonded Nonwoven-37 1/2"
3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2"
3OQ869/Applicator_15F
3OQ890/Applicator_01E
3OQ890N/Applicator_01E
3OQ975/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 58gsm 7.4"
3OS214/Applicator_25BB
3OS293/Applicator_NY_Swabstick_3OS293
3OS324/Applicator_25CA
3OS334/Applicator_NY_Swabstick_3OS334
3OS356/Applicator_NY_Swabstick Foam – Double Tipped
3OS407/Applicator_25B
3OS419/Applicator_25C
3OS428/Applicator_NY_Swabstick_3OS428
3OS450/Applicator_37A
3OS484/Applicator_26A
3OS639/Applicator_NY_Swabstick Rayon-Double Tipped, Gamma Irradiated - Oral use
3OT028-D/Applicator_NY_Unembossed Spunlaced Nonwoven-50gsm-15"
3OT103/Applicator_NY_Unembossed Spunlaced Nonwoven-55gsm-12.25"
3OT432-S/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 58gsm 12.25"
3OT501/Applicator_NY_15" Rolls Thermal Bonded Nonwoven 50gsm
3OT705-F/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 58gsm 15"
3OT705/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 58gsm 15"
3OU120/Applicator_12A
3OU125/Applicator_27B
3OU131/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 11.75"
3OU165/Applicator_15D
3OU212/Applicator_NY_6 1/4" Unembossed Spunlaced Nonwoven
3OU214/Applicator_07L
3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm
3OU265/Applicator_10H
3OU275/Applicator_36C
3OU351/Applicator_NY_1.25 oz/sy - 42.4gsm Spunbond Nonwoven
3OU444/Applicator_NY_60 gsm Dispersible - 6"
3OX105F/Applicatpr_NY_Foam 10.5ml
3OX105H/Handle, 10.5 ml
3OX465F/Foam 26 ml
3OX465H/Handle, 26 ml
3OX465P/Pledget, 26 ml
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped
3OX510F/Foam 5.1 ml
3OX510H/Handle, 5.1 ml
3TQ501/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 58gms -37.75"
3TQ541/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 58gms -30.5"
3TT541/Applicator_NY_Rolls Embossed Spunlaced Nonwoven - 58gsm 12.25"
3XP134/Applicator_NY_35gsm Stora Flat - 36"
3XP138/Applicator_36B
3XP138N/Applicator_NY_Spunlaced Nonwoven - 40gsm - 37.5" - Diamond Emboss
3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30"
3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42"
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75"
4AB1110M/Chem_206A
4DT00101/Chem_183A
4OB712OM/Chem_187A
4OS424AS/Chem_190A
4OU25201/Chem_NY_L'Oreal Sublime Bronze Tanning Liquid
4OY11001/Chem_194A
50M35650/
5AA204TR/Box_NY_0.018SBS_300_3.5-4.25-6
5AA284TR/Box_37C
5AA284TR/Box_NY_0.018SBS_300_6-5.875-3.625
5AA302TY/Box_74A
5AA342BT/Box_NY_0.018CCNB_300_3.625-1.25-2.375
5AA347BT/Box_NY_0.020SBS_300_3.375-1.313-2.625
5AA366SL/Box_43E
5AA366TR/Box_78A
5AA404S48/Box_NY_0.018SBS_111_6-5.875-3.625
5AA405TY/Box_74A
5AA411SL/Box_NY_0.020SBS_221_3.6875-3.9375-5.875
5AA466S48/Box_NY_0.024BC_113_7.000-6.000-6.3125
5AA491500/Box_NY_0.020BC_110_6.875-4.375-7.000
5AA527TR/Box_64A
5AA540S48/Box_NY_0.018SBS_100_7.1875-3.6875-6.0625
5AA54100/Box_70A
5AA56448/Box_21A
5AA565TR/Box_NY_0.024WCN_303_5.5-7.125-7.625
5AA592S48/Box_NY_0.018SBS_111_6-5.3125-3.6875
5AA617S96/Box_NY_0.018SBS_47B_3.625-3.875-6
5AA621TY/Box_37D
5AA62596/Box_47B
5AA632S60/Box_NY_0.020SBS_101_6-3.625-7
5AA673SL/Box_NY_0.020SBS_221_3.625-5.25-6
5AA682S96/Box_47B
5AA72148/Box_44A
5AA735TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375
5AA759SL/Box_NY_0.024BC_220_5.8125-3.3125-11.375
5AA759TY/Box_NY_0.024WCN_350_5.75-11.375-3.25-1.50
5AA8065SL/Box_NY_0.020BC_220_6.063-4-6.828
5AA81696/Box_22B
5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375
5AA922TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375
5AA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5
5AB10800/Box_01Q
5AB114IS/Box_61A
5AB123CP/Label_NY_5AB123CP_1.9375-1.875
5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75
5AB29307/Box_79A
5AB34700/Box_01A
5AB41210/Box_01A
5AB43201/Box_01Z
5AB432FA/Box_01Z
5AB51200/Box_01B
5AB687IS/Box_49A
5AC31100/Box_01C
5AC31100/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625
5AC42170/Box_01J
5AC45560/Box_20A
5AC80100/Box_NY_0.018UCSBS_100_5.0625-1.75-3.0625
5AD01524/Box_NY_0.018SBS_107_4-2.03125-4.5
5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5
5AD17748/Box_01O
5AD24031/Box_06A
5AD24200/Box_01D
5AD24200/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625
5AD24301/Box_NY_0.018SBS_107_4-2.03125-4.5
5AD32130/Box_03A
5AD35185/Box_01D
5AD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625
5AD351IS/Box_60B
5AD40100/Box_01D
5AD40100/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625
5AD42240/Box_11A
5AD43600/Box_01D
5AD43600/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625
5AD45330/Box_NY_0.018WCCN_100_2.3125-1.625-6.0625
5AD45430/Box_01N
5AD45460/Box_01K
5AD45460/Box_NY_0.018CCNB_100_3.1875-2.3125-6.0625
5AD47030/Box_NY_0.018SBS_112_2.3125-1.625-6.0625
5AD535BP/Box_01Y
5AD535LC/Box_62A
5AD606TY/Box_37A
5AD62120/Box_48A
5AD62130/Box_03A
5AD70324/Box_26A
5AD71230/Box_NY_0.018SBS_112_2.3125-1.625-6.0625
5AD712TY/Box_NY_0.020WCN_303_4.75-6.6875-6.09375
5AD71830/Box_NY_0.018CCNB_100_2.3125-1.625-6.0625
5AD71860/Box_NY_0.018CCNB_100_3.1875-2.3125-6.0625
5AD71890/Box_NY_0.020SBS_171_4.5-2.1875-8
5AD76630/Box_83A
5AF20400/Box_NY_0.018SBS_101_4.9375-3.375-6.25
5AG567TR/Box_NY_0.024WCN_307_4.4375-5.875-10
5AG60500/Box_65A
5AH06182/Box_NY_0.020SBS_121_6.25-2.5625-5.5
5AH12805/Box_18A
5AH12824/Box_01T
5AH14024/Box_NY_0.018SBS_100_4-2.125-6.125
5AH143PB/
5AH14624/Box_NY_0.018SBS_107_4-2.03125-4.5
5AH22524/Box_24A
5AH22524/Box_NY_0.018SBS_107_4-2.03125-4.5
5AH30116/Box_23A
5AH44116/Box_23A
5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5
5AH46018/Box_01U
5AH565CL/
5AH565PB/
5AH58195/Box_NY_0.020SBS_127_6.25-2.5625-4.625
5AH59200/Box_NY_0.020SBS_127_6.25-2.5625-4.625
5AH72500/Box_19A
5AH72500/Box_NY_0.020SBS_127_6.25-2.5625-4.625
5AJ123BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625
5AJ128BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625
5AJ130BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625
5AJ13724/Box_NY_0.018SBS_47A_3.875-1.3125-5.5
5AJ165SL/Box_NY_0.020BC_223_7.8125-7.1875-5.500
5AJ165TY/Box_NY_0.024WCN_303_5.5-7.125-7.625
5AJ170BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625
5AJ190BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625
5AJ200BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625
5AJ230BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625
5AJ324BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625
5AJ70168/Box_NY_0.020SBS_101_4-4.875-8
5AP0123IS/
5AP0124IS/
5AP027TY/Box_57A
5AP062TY/Box_NY_0.020SBS_105_3.125-10.25-5.375
5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75
5AP07930/Box_NY_0.020BC_110_6.375-3.4375-4.75
5AP080TY/Box_NY_32NPW_342_20.75-14.125-7.4375
5AP123SL/Box_29A
5AP136TY/Box_09A
5AP140TRY/Box_NY_0.020WCNB_300_3.75-12.5-5
5AP144TY/Box_NY_0.020WCN_307_3.75-6.25-4.875
5AP146SL/Box_29A
5AP147TY/Box_NY_0.020WCNB_323_3.375-2-2.625
5AP155SL/Box_NY_0.020BC_110_5.3125-3.750-10.500
5AP162TY/Box_NY_0.020WCN_307_3.75-6.25-4.875
5AP169BT/Box_NY_0.020WCN_306_3.375-2-2.625
5AP173SL/Box_NY_0.018SBS_340_20.25-4.53125-2.75
5AP188SL/Box_35A
5AP188SL/Box_NY_0.018SBS_340_20.25-4.53125-2.75
5AP204BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625
5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625
5AP21195/Box_01M
5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625
5AP229TY/Box_NY_0.024SBS_303_8.228-3.661-6.024
5AP278SL/Box_29A
5AP27995/Box_28A
5AP27995/Box_NY_0.020BC_101_9-6.625-4.5625
5AP308TY/Box_63A
5AP312TY/Box_NY_0.020SBS_300_4.625-5.875-3.75
5AP349SL/Box_81A
5AP354TY/Box_NY_0.020SBS_105_3.125-10.25-5.375
5AP421TY/Box_73A
5AP423TY/Box_NY_0.020SBS_128_5.375-5-5.5
5AP438TY/Box_33A
5AP438TY/Box_NY_0.020SBS_305_2.5-6.25-5.5
5AP449PD/Box_NY_0.020BC_400_7.4375-8.3125
5AP460TY/Box_NY_0.020SBS_105_4.125-7.875-5.5
5AP479SL/Box_35A
5AP479SL/Box_NY_0.018SBS_340_20.25-4.53125-2.75
5AP506TY/Box_NY_0.020SBS_105_3.125-10.25-5.375
5AP563TY/Box_NY_0.020SBS_105_5.375-7.375-4.125
5AP579TY/Box_27A
5AP624BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625
5AP728TY/Box_09A
5AP742SL/Box_81B
5AP87015/Box_NY_0.020BC_100_8.5-3.75-5.75
5AP89024/Box_NY_0.024BC_100_7.5-6.5-4.75
5AQ39130SL/Box_43C
5AQ429TR/Box_NY_0.024CCN_308_9-6.75-6.125
5AQ476TR/Box_55A
5AQ477TR/Box_37E
5AQ499SL/Box_NY_0.020BC_200_6-4.438-4.063
5AQ499TR/Box_NY_0.020CCNB_300_6-4.375-4
5AQ548TR/Box_NY_0.020SBS_300_3.75-8.375-4.4375
5AS17250/Box_01E
5AS334IS/Box_NY_NA50#_402_5-3.25
5AS40750/Box_01R
5AS41950/Box_01S
5AS42050/Box_01E
5AS42325/Box_01E
5AS81650/Box_01E
5AU13195/Box_NY_0.024SBS_127_4.39063-4.39063-12
5AU16500/Box_02A
5AU16500/Box_NY_0.024SBS_102_4.390625-4.390625-12
5AU19101/Box_NY_0.020SBS_129_5.75-3-3.875
5AU239IB/Box_NY_NA_CB_100_5.25-3.9375-5.875
5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6
5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125
5AU26595/Box_02A
5AU26595/Box_NY_0.024SBS_102_4.390625-4.390625-12
5AU27500/Box_02B
5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375
5AU33901/Box_NY_0.020SBS_129_5.75-3-3.875
5AU34510/Box_NY_0.018SBS_131_3.875-1.875-3
5AU37230/Box_NY_0.018SBS_101_3.375-3.375-2.9375
5AU38510/Box_56A
5AU44410/Box_NY_0.016SBS_124_2.9375-1.875-3.8125
5AU451TY/Box_01P
5AU451TY/Box_NY_0.024WCNB_100_7-7-5
5AU47500/Box_NY_0.016SBS_130_3.25-1.5-3.875
5AU50010/Box_NY_0.016SBS_124_2.9375-1.875-3.8125
5AU536D/Box_NY_0.020SBS_129_5.75-3-3.875
5AU54901/Box_NY_0.018SBS_100_101.6-44.45-101.6
5AU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125
5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4
5AU62620/Box_NY_0.016BMPB_122_3.625-3-3.8125
5AU710TY/Box_82A
5AU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125
5BA284TR/Box_NY_0.018SBS_300_6-5.875-3.625
5BA323BX/Box_NY_0.018SBS_47B_3.9375-3.25-5.8125
5BA335TR/Box_37G
5BA347BT/Box_NY_0.020SBS_300_3.375-1.313-2.625
5BA402S96/Box_22A
5BA402TR/Box_37C
5BA402TR/Box_NY_0.018SBS_300_6-5.875-3.625
5BA403SL/Box_36A
5BA404S48/Box_NY_0.018SBS_111_6-5.875-3.625
5BA56448/Box_NY_0.018SBS_122_3.5-3.375-5.625
5BA581TR/Box_66A
5BA617S96/Box_NY_0.018SBS_47B_3.625-3.875-6
5BA62596/Box_NY_0.018SBS_47B_3.9375-3.25-5.8125
5BA682S96/Box_NY_0.018SBS_47B_3.9375-3.25-5.8125
5BA735TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375
5BA922TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375
5BA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5
5BB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375
5BB29307/Box_NY_0.018CCN_100_4.1875-4.1875-2.1875
5BB32071/Box_05A
5BB43201/Box_NY_0.016SBS_100_4-2.25-2.25
5BB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375
5BB973IS/Box_60A
5BC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625
5BD01524/Box_NY_0.018SBS_107_4-2.03125-4.5
5BD02424/Box_NY_0.018SBS_107_4-2.03125-4.5
5BD24624/Box_26A
5BD30412/Box_10A
5BD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625
5BD351IS/Box_49A
5BD38430/Box_03A
5BD43600/Box_01D
5BD45460/Box_01K
5BD45460/Box_NY_0.018CCNB_100_3.1875-2.3125-6.0625
5BD50490/Box_16A
5BD50748/Box_01K
5BD50748/Box_NY_0.018CCNB_100_3.1875-2.3125-6.0625
5BD53530/Box_01I
5BD606TY/Box_37A
5BD65900/Box_01D
5BD65900/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625
5BD71600/Box_01D
5BD71600/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625
5BD71830/Box_NY_0.018CCNB_103_2.3125-1.625-6.0625
5BD71860/Box_NY_0.018WCN_171_3.1875-2.3125-6.0625
5BD71890/Box_NY_0.020SBS_171_4.5-2.1875-8
5BD72400/Box_01D
5BD72400/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625
5BD76630/Box_NY_0.018WCCN_103_2.3125-1.625-6.0625
5BF206R30/Box_NY_0.018SBS_122_3.375-3.125-4.9375
5BH12824/Box_NY_0.018SBS_100_4-2.125-6.125
5BH14024/Box_NY_0.018SBS_100_4-2.125-6.125
5BH21724/Box_24A
5BH21724/Box_NY_0.018SBS_107_4-2.03125-4.5
5BH22018/Box_23A
5BH22524/Box_NY_0.018SBS_107_4-2.03125-4.5
5BH53050/Box_08A
5BH59200/Box_NY_0.020SBS_127_6.25-2.5625-4.625
5BH72500/Box_19A
5BH72500/Box_NY_0.020SBS_127_6.25-2.5625-4.625
5BJ123BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625
5BJ13724/Box_NY_0.018SBS_47A_3.875-1.3125-5.5
5BJ16530TY/Box_NY_0.024WCN_303_5.5-7.125-7.625
5BJ170BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625
5BJ190BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625
5BJ200BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625
5BJ324BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625
5BP115SL/Box_29B
5BP146SL/Box_29A
5BP155SL/Box_NY_0.020BC_223_5.1875-3.5625-10.5
5BP173SL/Box_NY_0.018SBS_340_20.25-4.53125-2.75
5BP37815/Box_27A
5BP438TY/Box_NY_0.020SBS_305_2.5-6.25-5.5
5BP87015/Box_NY_0.020BC_100_8.5-3.75-5.75
5BQ05300/Box_53A
5BQ39130SL/Box_43D
5BQ467SL/Box_70B
5BQ476TR/Box_NY_0.024WCK_351_3.75-12.5-3.9375-2.5625
5BS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125
5BS42325/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125
5BU16500/Box_02A
5BU16500/Box_NY_0.024SBS_102_4.390625-4.390625-12
5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875
5BU21210/Box_56A
5BU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125
5BU26595/Box_NY_0.024SBS_102_4.390625-4.390625-12
5BU27500/Box_NY_0.024SBS_102_4.3906-4.3906-12
5BU37230/Box_NY_0.018SBS_101_3.8125-3.375-2.9375
5BU38510/Box_NY_0.016SBS_124_2.9375-1.875-3.8125
5BU44410/Box_NY_0.016SBS_124_2.9375-1.875-3.8125
5BU451TY/Box_01P
5BU451TY/Box_NY_0.024WCNB_100_7-7-5
5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125
5BU62620/Box_NY_0.016BMPB_122_3.625-3-3.8125
5BU710TY/Box_NY_0.024WCKB_100_6-5-3
5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125
5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875
5CA335TR/Box_37G
5CA402S96/Box_47A
5CA402S96N/Box_47A
5CA403SL/Box_NY_0.020SBS_221_3.625-5.5-6
5CA56448/Box_NY_0.018SBS_122_3.5-3.375-5.625
5CA682S96/Box_NY_0.018SBS_47B_3.9375-3.25-5.8125
5CA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5
5CB10000/Box_01B
5CB10800/Box_01Q
5CB114IS/Box_49B
5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375
5CB270IS/Box_49A
5CB29307/Box_NY_0.018CCN_100_4.1875-4.1875-2.1875
5CB43201/Box_NY_0.016SBS_100_4-2.25-2.25
5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375
5CB81071/Box_05A
5CC58570/Box_01W
5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625
5CD24624/Box_26A
5CD27400/Box_01D
5CD27400/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625
5CD30630/Box_04A
5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625
5CD38430/Box_NY_0.018NB_100_2.3125-1.625-6.0625
5CD43600/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625
5CD45430/Box_01N
5CD53530/Box_01Y
5CD606TY/Box_37A
5CD69500/Box_01D
5CD69500/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625
5CD72400/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625
5CH11524/Box_24A
5CH11524/Box_NY_0.018SBS_107_4-2.03125-4.5
5CH12824/Box_NY_0.018SBS_100_4-2.125-6.125
5CH21130/Box_01L
5CH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25
5CH21724/Box_NY_0.018SBS_107_4-2.03125-4.5
5CH22018/Box_NY_0.018SBS_101_3.875-2-4.75
5CH22524/Box_NY_0.018SBS_107_4-2.03125-4.5
5CH54082/
5CJ13724/Box_NY_0.018SBS_47A_3.875-1.3125-5.5
5CP37815/Box_27A
5CP490SL/Box_81A
5CQ18115/Box_37F
5CS273IS/Box_NY_NA50#_402_5-3.25
5CS324IS/Box_NY_NA50#_402_5-3.25
5CS40750/Box_01R
5CS41950/Box_01S
5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125
5CS42325/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125
5CS428IS/Box_NY_NA50#_402_5-3.25
5CS63925/Box_01E
5CU21210/Box_NY_0.016SBS_124_2.9375-1.875-3.8125
5CU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125
5CU38510/Box_NY_0.016SBS_124_2.9375-1.875-3.8125
5CU451TY/Box_NY_0.024WCNB_100_7-7-5
5DA402S96/Box_47A
5DA402S96/Box_NY_0.018SBS_47A_3.625-3.875-6
5DA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5
5DB00707/Box_01A
5DB01051/Box_07A
5DB01051/Box_NY_0.016UCSBS_123_6.125-1.5-2.375
5DB10800/Box_01Q
5DB12800/Box_01B
5DB29307/Box_NY_0.018CCN_100_4.1875-4.1875-2.1875
5DB43201/Box_NY_0.016SBS_100_4-2.25-2.25
5DB43201/Box_NY_0.016SBS_100_4-2.25-2.25
5DB47000/Box_01B
5DB80351/Box_NY_0.016UCSBS_123_6.125-1.5-2.375
5DC33400/
5DC42500/Box_01C
5DC42500/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625
5DD24624/Box_26A
5DD43600/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625
5DD45430/Box_01N
5DD45460/Box_NY_0.018CCNB_100_3.1875-2.3125-6.0625
5DD72400/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625
5DH04095/Box_19A
5DH04095/Box_NY_0.020SBS_127_6.25-2.5625-4.625
5DH07050/Box_69A
5DH11524/Box_24A
5DH11524/Box_NY_0.018SBS_107_4-2.03125-4.5
5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125
5DH14324/Box_01X
5DH21130/Box_01L
5DH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25
5DH21724/Box_NY_0.018SBS_107_4-2.03125-4.5
5DH22018/Box_NY_0.018SBS_101_3.875-2-4.75
5DH35650/Box_01F
5DP490SL/Box_NY_0.016SBS_340_11.46875-4.5-2.75
5DQ50615/Box_37H
5DS40750/Box_01R
5DS41950/Box_01S
5EB01051/Box_NY_0.016UCSBS_123_6.125-1.5-2.375
5EB10300/Box_01B
5EB114IS/Box_49B
5EB39702/Box_01A
5EB43201/Box_NY_0.016SBS_100_4-2.25-2.25
5EB47000/Box_NY_0.016UCSBS_100_6.125-2.75-2.375
5EB59800/Box_01B
5EB80300/Box_01B
5EC12400/Box_01C
5EC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625
5ED11000/Box_01D
5ED11000/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625
5ED31000/Box_01D
5ED31000/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625
5ED34600/Box_01D
5ED34600/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625
5ED41900/Box_01D
5ED41900/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625
5ED43800/Box_01AA
5ED65500/Box_01D
5ED65500/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625
5EH07050/Box_69B
5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25
5EH21724/Box_NY_0.018SBS_107_4-2.03125-4.5
5EH22018/Box_NY_0.018SBS_101_3.875-2-4.75
5EH35650/Box_01AD
5ES273IS/Box_NY_NA50#_402_5-3.25
5ES324IS/Box_NY_NA50#_402_5-3.25
5ES428IS/Box_NY_NA50#_402_5-3.25
5ES68625/
5EU13295/Box_02A
5EU13295/Box_NY_0.024SBS_102_4.390625-4.390625-12
5FB01051/Box_NY_0.016UCSBS_123_6.125-1.5-2.375
5FB10300/Box_NY_0.016UCSBS_100_6.125-2.750-2.375
5FB10800/Box_01Q
5FB114IS/Box_49B
5FB16400/Box_01B
5FB37900/Box_01V
5FB43201/Box_NY_0.016SBS_100_4-2.25-2.25
5FB47000/Box_NY_0.016UCSBS_100_6.125-2.75-2.375
5FB80300/Box_01B
5FB85602/Box_01A
5FC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625
5FC22324/Box_01J
5FD35100/Box_01D
5FD35100/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625
5FD36330/Box_03A
5FD41900/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625
5FD74800/Box_01D
5FD74800/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625
5FH04095/Box_19A
5FH04095/Box_NY_0.020SBS_127_6.25-2.5625-4.625
5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625
5FH14324/Box_01X
5FH24795/Box_19A
5FH24795/Box_NY_0.020SBS_127_6.25-2.5625-4.625
5FH35650/Box_NY_0.016UCSBS_100_5-6.03125-3.03125
5FS273IS/Box_60C
5FS40750/Box_01R
5FS41950/Box_01S
5FS42250/Box_01E
5FS428IS/Box_60C
5FS46250/Box_01E
5FS48450/Box_NY_0.016UCSBS_100_6.5-3.375-5.8125
5FS84925/Box_01E
5GB10300/Box_NY_0.016UCSBS_100_6.125-2.750-2.375
5GB10800/Box_NY_0.018SBS_100_5.125-2.25-2.09375
5GB114IS/
5GB33902/Box_01A
5GB59800/Box_NY_0.016UCSBS_100_6.125-2.75-2.375
5GB60302/Box_01A
5GB69000/
5GC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625
5GC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625
5GC29600/Box_01C
5GC29600/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625
5GC69900/Box_01C
5GC69900/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625
5GD19831/Box_67A
5GD35100/Box_01AB
5GD40812/Box_68A
5GD45430/Box_NY_0.018NB_100_2.3125-1.625-6.0625
5GD74800/Box_01D
5GD74800/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625
5GH04095/Box_NY_0.020SBS_127_6.25-2.5625-4.625
5GS273IS/
5GS324IS/
5GS41125/Box_01E
5GS428IS/
5GS48450/Box_NY_0.016UCSBS_100_6.5-3.375-5.8125
5HB05502/Box_01A
5HB10800/Box_NY_0.018SBS_100_6.125-2.25-2.5625
5HB114IS/Box_60D
5HB16400/Box_NY_0.016UCSBS_100_6.125-2.75-2.375
5HB33902/Box_01A
5HB34950/
5HB59800/Box_NY_0.016UCSBS_100_6.125-2.75-2.375
5HB68800/Box_05B
5HB71200/Box_01B
5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625
5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625
5HC69900/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625
5HD45430/Box_NY_0.018WCN_103_2.3125-1.625-6.0625
5HD74800/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625
5HH04095/Box_NY_0.020SBS_127_6.25-2.5625-4.625
5HQ594TR/Box_09A
5HS23125/Box_01E
5HS273IS/Box_NY_NA50#_402_5-3.25
5HS324IS/Box_NY_NA50#_402_5-3.25
5HS40750/Box_NY_0.018SBS_100_7.125-3-6.625
5HS41125/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125
5HS41350/Box_01E
5HS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125
5HS42450/Box_01E
5HS428IS/Box_NY_NA50#_402_5-3.25
5HS48050/Box_01E
5HS48450/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125
5HS82125/Box_01E
5HT00128/Box_01H
5HT00272/Box_01G
5JB10800/Box_NY_0.018SBS_100_5.125-2.25-2.09375
5JB114IS/Box_NY_NA50#_402_2.5-3.25
5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375
5JB34950/Box_07A
5JB34950/Box_NY_0.016UCSBS_123_6.125-1.5-2.375
5JB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375
5JB71200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375
5JC69900/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625
5JH04095/Box_NY_0.020SBS_127_6.25-2.5625-4.625
5JS16225/Box_01E
5JS23125/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125
5JS273IS/Box_NY_NA50#_402_5-3.25
5JS40750/Box_NY_0.018SBS_100_7.125-3-6.625
5JS41125/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125
5JS41350/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125
5JS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125
5JS42450/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125
5JS428IS/Box_NY_NA50#_402_5-3.25
5JS48050/Box_NY_0.016UCSBS_100_6.5-3.375-5.8125
5JS82125/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125
5JU87250/Box_02B
5JU87295/Box_02B
5KB10800/Box_NY_0.018SBS_100_5.125-2.25-2.09375
5KB114IS/Box_NY_NA50#_402_2.5-3.25
5KB40600/Box_01B
5KB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375
5KB71200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375
5KC69900/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625
5KH04095/Box_NY_0.020SBS_127_6.25-2.5625-4.625
5KQ594TR/Box_09A
5KS16225/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125
5KS23125/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125
5KS273IS/Box_NY_NA50#_402_5-3.25
5KS324IS/Box_NY_NA50#_402_5-3.25
5KS40750/Box_NY_0.018SBS_100_7.125-3-6.625
5KS41125/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125
5KS41350/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125
5KS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125
5KS42450/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125
5KS428IS/Box_NY_NA50#_402_5-3.25
5KS48050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125
5KS82125/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125
5KU87250/Box_02B
5LB10800/Box_NY_0.018SBS_100_5.125-2.25-2.09375
5LB114IS/Box_NY_NA50#_402_2.5-3.25
5LB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375
5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375
5LQ594TR/Box_NY_0.020SBS_128_5.375-5-5.5
5LS16225/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125
5LS23125/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125
5LS273IS/Box_NY_NA50#_402_5-3.25
5LS40750/Box_NY_0.018SBS_100_7.125-3-6.625
5LS41350/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125
5LS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125
5LS42450/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125
5LS428IS/Box_NY_NA50#_402_5-3.25
5LS48050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125
5LS82125/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125
5LU87250/Box_NY_0.024SBS_127_4.375-4.375-12
5MB10800/Box_NY_0.018SBS_100_5.125-2.25-2.09375
5MB114IS/Box_NY_NA50#_402_2.5-3.25
5MS23125/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125
5MS273IS/Box_NY_NA50#_402_5-3.25
5MS324IS/Box_NY_NA50#_402_5-3.25
5MS40750/Box_NY_0.018SBS_100_7.125-3-6.625
5MS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125
5MS428IS/Box_NY_NA50#_402_5-3.25
5MS48050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125
5MS82125/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125
5MU87250/Box_NY_0.024SBS_127_4.375-4.375-12
5NB10800/Box_NY_0.018SBS_100_5.125-2.25-2.09375
5NB114IS/Box_NY_NA50#_402_2.5-3.25
5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375
5NS23125/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125
5NS273IS/Box_NY_NA50#_402_5-3.25
5NS324IS/Box_NY_NA50#_402_5-3.25
5NS428IS/Box_NY_NA50#_402_5-3.25
5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12
5OA012SL/Box_NY_0.020BC_220_6.375-4.1875-8.125
5OA204SLV/Box_NY_0.020BC_200_3.5-4.313-6.063
5OA335SL/Box_NY_0.024BC_220_7.625-7.0625-8.6875
5OA346BX/Box_NY_0.020CCNB_100_10.25-4.125-7.5
5OA347PB/Box_NY_LDPE_407_8.5-6.6-4-4
5OA366SL/Box_42A
5OA366TR/Box_37B
5OA447SL/Box_NY_0.020BC_220_4.0625-7.0625-6.5
5OA447TR/Box_NY_0.020CCNB_300_4-6.5-7
5OA466S48/Box_84A_0.024
5OA498SL/Box_NY_0.024BC_200_4.25-4.063-6.063
5OA520SL/Box_NY_0.024BC_220_3.4375-5.875-12
5OA520TY/Box_NY_0.024WCN_303_3.375-12-5.75-1.5
5OA567SL/Box_NY_0.020BC_220_5.5625-6.3125-3.875
5OA594SL/Box_NY_0.020BC_200_8-3.438-5.688
5OA594TR/Box_NY_0.020CCNB_300_8-3.375-5.625
5OA599DV1/Box_NY_NA_339_10.25-10.125
5OA599DV2/Box_NY_NA_339_8.50-4.75
5OA621SL/Box_31A
5OA670HT/Box_54A
5OA709BT/Box_NY_0.018SBS_309_3.375-1.75-2.5
5OA710SL/Box_NY_0.020BC_220_6-4.3125-6.750
5OA714TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375
5OA738SL/Box_NY_0.024BC_220_6.25-4.313-6.063
5OA8064SL/Box_NY_0.020BC_220_6.25-4-5.063
5OA8065SL/Box_NY_0.020BC_220_6.25-4-5.813
5OA922TY/Box_NY_0.020SBS_100_4.5-7.9375-7.9375
5OB123CP/Label_NY_NP9300_1.9375-1.875
5OC105PD/Box_50E
5OD125PB/Box_80A
5OD173YS/
5OD239ID/Box_52A
5OD239PD/Box_51A
5OD304WC/
5OD371SL/Box_75A
5OD454SL/Box_71A
5OD718SL/Box_NY_0.020BC_200_9.0625-8.1875-4.500
5OD9033SL100/Box_NY_0.018TBC_200_6.234-3.987-4.594
5OH115TY/Box_43A
5OH143CL/
5OH158PB/Box_72A
5OH211GB/
5OH211PD/Box_50B
5OH269PB/
5OH382PD/
5OH382PD/Box_50D
5OH460IP/Box_39A
5OH617PB/Box_72A
5OJ124BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625
5OJ132TY/Box_NY_0.024WCN_303_5.5-7.125-7.625
5OJ137IN/Box_NY_0.018BC_114_4.1718-4.0468-5.7187
5OJ151BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625
5OJ165SL/Box_NY_0.020BC_222_7.6875-7.1875-5.500
5OJ230BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625
5OJ345SL/Box_NY_0.020BC_200_7.1875-8.125-4.03125
5OK457DR/
5OK457PD/Box_50A
5OM9029BT/Box_NY_0.020CCNB_300_9.625-3-6
5OP037TY/Analytical
5OP043PB/Bag_NY_5OP043PB
5OP065BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375
5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375
5OP073BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375
5OP075BT/Box_NY_0.018SBS_300_4.5-2-1.688
5OP079BT/Box_NY_0.018SBS_320_4.5-2-1.688
5OP081BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375
5OP081CT/Box_NY_0.018SBS_101_4.6875-3.25-10.875
5OP081PB/Box_NY_LDPE_407_13-5.5-6-4
5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875
5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4
5OP140SL/Box_NY_0.020BC_220_3.8125-5.5-13.5
5OP140SLV/Box_NY_0.020BC_200_3.8125-5.125-12.5
5OP140TY/Box_NY_0.020WCNB_303_3.75-13.5-5.4375
5OP144SL/Box_NY_0.020BC_360_3.8125-4.9375-6.250
5OP149SL/Box_NY_0.020BC_220_3.875-5.125-7.875
5OP149TY/Box_NY_0.020BC_301_6.125-4-2.375
5OP155SL/Box_NY_0.020BC_5.3125-3.750-10.500
5OP169BT/Box_NY_0.020WCN_306_3.375-2-2.625
5OP189IN/Box_NY_0.020BC_100_4.875-3.625-6.375
5OP195TY/Analytical
5OP197TY/Analytical
5OP205BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375
5OP209SL/Box_NY_0.020BC_220_4.9375-3.6875-4.625
5OP210TY/Box_NY_0.020SBS_300_4.625-5-3.625
5OP214GL/
5OP227TY/Box_NY_0.020SBS_300_4.625-4.875-3.625
5OP249TW/Box_NY_0.020SBS_115_3.25-6.5-4.625
5OP252BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375
5OP308PD/Box_50C
5OP312SL/Box_NY_0.020BC_220_5.9375-3.750-4.625
5OP365BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375
5OP378PD/Box_58A
5OP438SL/Box_34A
5OP45520TK/Box_NY_0.020BC_100_12.375-3.15625-5
5OP460PD/Box_NY_NA_400_11.875-4.625
5OP460TY/Box_NY_0.020SBS_105_4.125-7.375-5.375
5OP467SL/Box_NY_0.020BC_200_3.1875-6.6875-10.375
5OP467TR/Box_NY_0.020SBS_201_3.125-10.375-6.625
5OP624PB/Box_NY_0.0018LDPE_407_12-5.5-5-4
5OP705PB/Box_NY_0.0018LDPE_407_12.5-6.5-5-4
5OP710SL/Box_38A
5OP911TY/Box_NY_0.020SBS_100_4.5-7.9375-7.9375
5OP913TY/Analytical -Tray_NY_Display Tray_02
5OQ008BNSL/Box_NY_0.020BC_200_6.75-4.063-4.563
5OQ008BNTR/Box_NY_0.020CCNB_300_6.75-4-4.50
5OQ011SL/Box_NY_0.020BC_200_11.125-3.813-4.438
5OQ021TR/Box_NY_0.024CRB_300_7.375-4.50-8.00
5OQ028SL/Box_NY_0.020BC_220_10-5.063-4.313
5OQ028SL/Box_NY_0.020BC_220_5.0625-4.3125-10
5OQ028TR/Box_NY_0.020CCNB_300_10-5-4.25
5OQ028TR/Box_NY_0.020CCNB_300_5-10-4.25
5OQ137SL/Box_NY_0.024BC_200_12.5-4.063-4.438
5OQ204SL/Box_NY_0.020BC_220_8.094-4.594-7.375
5OQ231SL/Box_NY_0.024BC_220_13.75-4.063-4.563
5OQ231TR/Box_NY_0.020CCNB_300_13.75-4-4.50
5OQ252SL/Box_NY_0.018BC_224_4.392-3.8125-14.250
5OQ401SL/Box_NY_0.024BC_341_6.875-6.250-3.000
5OQ409SL/Box_43B
5OQ445SL/Box_NY_0.020BC_200_5.75-4.063-4.563
5OQ492SL/Box_NY_0.020BC_220_8.375-3.875-4.5
5OQ5019SL/Box_NY_0.018BC_220_6.59375-3.53125-7.5
5OQ5042TR/Box_NY_0.020CCNB_300_10.125-3.750-4.375
5OQ548SL/Box_NY_0.020BC_200_8.375-3.875-4.5
5OQ558TR/
5OQ594SL/Box_77A
5OQ635BX/Box_NY_0.020BC_100_13.75-3.875-4.625
5OQ637SL/Box_NY_0.024BC_200_10.75-3.688-5.063
5OQ778S25/Box_NY_0.020BC_100_12-3.875-4.375
5OQ793TR/Box_NY_0.020CCNB_300_3.5-4.875-4.75
5OQ797IN/Box_59A
5OQ808SL/Box_NY_0.020BC_220_4.4375-3.375-9
5OQ808TR/Box_NY_0.020WCN_300_4.375-9-3.25
5OQ999TR/Box_NY_0.020BC_301_3.7500-8.375-4.4375
5OU295PB/Box_76A
5OU295PB/Box_NY_0.002LDPE_407_4-5
5OU295PD/
5OU351LP/SHIP_NY_32K_NA_19.500X11.750
5OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500
5OU475IS/Box_NY_NA50#_404_4.375-3.625
5OU475SL/Box_NY_BrownPaper_2.624-1.9680
5OU526PD/Box_NY_0.020BC_400_14.750-9.500
5OX10500/Box_NY_40W_14.875-5.625-7.125
5OX465LB/Box_NY_HDPE_407_24-20-16
5PB10800/Box_NY_0.018SBS_100_5.125-2.25-2.09375
5PB114IS/Box_NY_NA50#_402_2.5-3.25
5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625
5PS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125
5QB114IS/Box_NY_NA50#_402_2.5-3.25
5QS273IS/Box_NY_NA50#_402_5-3.25
5QS324IS/Box_NY_NA50#_402_5-3.25
5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125
5QS428IS/Box_NY_NA50#_402_5-3.25
5XR11689/
5YL619PB/
6AA119S6/SHIP_32_27A
6AA119S9/SHIP_32_27B
6AA250S56/SHIP_NY_40K_100_14-8-9
6AA34724/SHIP_NY_32K_100_8.25-7-6.25
6AA350S64/SHIP_NY_40K_100_14-8-9.5
6AA366R48/SHIP_44_2AQ
6AA402S48/SHIP_44_2AF
6AA402S96/SHIP_44_2E
6AA436XF2B/SHIP_40_1T
6AA437FWB/SHIP_NY_44K_100_14.125-7.6875-10.625
6AA445S42/SHIP_NY_32K_100_13-7-7.875
6AA507S40/SHIP_NY_44K_100_13.75-8.5-6.75
6AA540S40/SHIP_NY_32K_100_16.5625-6.3125-5.9375
6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875
6AA630FWB/SHIP_NY_44K_100_14.125-7.6875-10.625
6AA632TY/SHIP_NY_44K_231_18.5-7.375-14
6AA641TC/SHIP_NY_32K_201_47.25-34.75-3.000
6AA673S30/SHIP_NY_44K_150_20.0625-9.1875-10.9375
6AA673TY/SHIP_NY_44K_10A_18.75-7.875-10.8125
6AA675S632/SHIP_NY_44SBS_100_9-6.5-11
6AA709S48/SHIP_NY_40K_100_15-11-6.125
6AA709SWN/SHIP_NY_40K_100_14.75-9.375-5.5
6AA73530/SHIP_NY_32K_100_10-8.375-8.25
6AA76930/SHIP_NY_44K_100_9-9-4.75
6AA8054HSC/SHIP_NY_32K_150_9.188-8.188-9.25
6AA8054TR/SHIP_NY_32W_250_8.875-7.875-6.5
6AA810R48/SHIP_40_1B
6AA88230/SHIP_NY_44K_100_9-9-4.75
6AA91130/SHIP_NY_32K_100_10-8.375-8.25
6AA92230/SHIP_NY_32K_100_10-8.375-8.25
6AB16400/SHIP_NY_32K_100_16-4.25-5.125
6AB34700/SHIP_26_1A
6AB41210/SHIP_29_1I
6AB89435/SHIP_29_1E
6AD11055/SHIP_NY_32K_100_10.625-6.375-11.25
6AD17385/SHIP_44_2AB
6AD249JR/SHIP_32_1AW
6AD33333/SHIP_NY_32K_100_21-12.5-14
6AD38731/SHIP_32_1AR
6AD40100/SHIP_32_1AD
6AD41900/SHIP_NY_32K_100_13.25-12-6.125
6AD43600/SHIP_NY_32K_100_13.25-12-6.125
6AD454KCK/SHIP_32_9A
6AD50490/SHIP_32_1A
6AD504JR/SHIP_32_1AX
6AD535KC/SHIP_32_9A
6AD606SK/SHIP_32_9A
6AD71600/SHIP_32_1AD
6AG11284/SHIP_NY_32K_100_9.75-8.5-11
6AG11584/SHIP_NY_32K_100_9.75-8.5-11
6AH06182/SHIP_NY_32K_100_13.6875-6.4375-11
6AH14324/SHIP_NY_32K_100_12.4375-9-5.25
6AH58195/SHIP_NY_32K_100_13.4375-6.25-9.6875
6AH59200/SHIP_32_1AF
6AH72500/SHIP_32_1AF
6AJ112CBX/SHIP_NY_32K_100_8.75-5.625 -9
6AJ13724/SHIP_NY_32K_100_9-8.5-5.875
6AJ14543/SHIP_44_4A
6AJ17025/SHIP_NY_32K_100_12-7.25-6.75
6AJ19025/SHIP_NY_32K_100_12-7.25-6.75
6AJ21396/SHIP_44_2U
6AJ22750/SHIP_44_2U
6AM076S6/SHIP_32_28A
6AM1500S72J/SHIP_44_2AP
6AM233XT/SHIP_44_2H
6AM30472/SHIP_NY_44K_100_15.5-9-12
6AM363XF6B/SHIP_44_2AH
6AM420SF3B/SHIP_44_2AI
6AM521S64/SHIP_NY_44K_100_13.75-8.5-9.5
6AM525S72/SHIP_NY_44K_100_13.75-8.5-10.75
6AM70172/SHIP_NY_44K_100_15.5-9-12
6AM8801CB/Ship_NA_2A
6AM919S9/SHIP_200_7A
6AM928S80/SHIP_NY_44K_100_15.5-9-10.5
6AM929SB/SHIP_200_5E
6AM960TR/SHIP_200_5A
6AM998DV/SHIP_200_5B
6AMC9603BDL/SHIP_44_2AM
6AP0112BS/SHIP_32_10A
6AP0112HC/SHIP_32_6A
6AP0112TR1/SHIP_32_5A
6AP0112TR2/SHIP_32_5A
6AP0123BS/SHIP_200_3A
6AP0123HD/SHIP_200_2A
6AP0123TY/SHIP_200_1B
6AP0123TY1/SHIP_200_1B
6AP0123TY2/SHIP_200_1B
6AP0124TY/SHIP_200_1B
6AP02372/SHIP_32_1AU
6AP02810/SHIP_NY_32K_100_15.375-11.5625-10.5625
6AP03410/SHIP_NY_32K_100_15.375-11.5625-10.5625
6AP03672/SHIP_NY_32K_100_16.3125-12.25-7.5625
6AP03707/SHIP_NY_32K_100_13-7.3125-4
6AP06220/SHIP_NY_32K_100_10.625-10-5.625
6AP06325/SHIP_NY_32K_100_12.4375-5.4375-4.1875
6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75
6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5
6AP06728/SHIP_NY_32K_100_7.125-7-6.0625
6AP07328/SHIP_NY_32K_100_13.25-7.75-5.75
6AP08084/SHIP_NY_40NPW_100_10-6.75-7
6AP08128/SHIP_NY_32K_100_13.25-7.75-5.75
6AP08150/SHIP_NY_32K_100_17.375-12.0625-10.5
6AP081CT/SHIP_NY_44K_150_20.3125-14.1875-11.0625
6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75
6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875
6AP10110/SHIP_NY_32K_100_10.250-6.250-4.750
6AP11325/SHIP_NY_44K_100_10.625-9.75-7.75
6AP115TY/SHIP_32_7A
6AP12072/SHIP_40_1K
6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625
6AP13872/SHIP_32_1BL
6AP15125/SHIP_NY_32K_100_12-7.25-6.75
6AP15472/SHIP_44_2L
6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875
6AP17080/SHIP_NY_32K_100_14.625-6.375-5
6AP173TY/SHIP_32_3A
6AP18125/SHIP_NY_44K_100_11.5-8.75-5.75
6AP18905/SHIP_NY_32K_100_7.625-5.1875-6.625
6AP19525/SHIP_NY_32K_100_12.4375-5.4375-4.1875
6AP19725/SHIP_NY_32K_100_12.4375-5.4375-4.1875
6AP20425/SHIP_NY_32K_100_12-7.25-6.75
6AP20528/SHIP_NY_32K_100_13.25-7.75-5.75
6AP21195/SHIP_44_2I
6AP214TY/SHIP_32_7A
6AP22325/SHIP_NY_32K_100_12.4375-5.4375-4.1875
6AP22884/SHIP_40_1H
6AP22907/SHIP_NY_32K_100_8.5625-7.8125-6.4375
6AP23025/SHIP_NY_32K_100_11.25-5.8125-4.3125
6AP23772/SHIP_32_1AC
6AP23872/SHIP_NY_32K_100_16.3125-12.25-7.5625
6AP23925/SHIP_NY_32K_100_11.25-5.8125-4.3125
6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625
6AP24372/SHIP_NY_44K_100_15.25-10.375-15.4375
6AP24572/SHIP_NY_32K_100_16.3125-12.25-7.5625
6AP24684/SHIP_NY_40K_100_15.375-10.5625-9.625
6AP249TW/SHIP_NY_32K_100_13.75-10.375-9.75
6AP25228/SHIP_NY_32K_100_7.125-7-6.0625
6AP25372/SHIP_NY_44K_100_15.25-10.375-15.4375
6AP25784/SHIP_NY_40K_100_15.375-10.5625-9.625
6AP26384/SHIP_NY_32K_100_15.1875-10.1875-9.625
6AP27284/SHIP_NY_44K_100_15.375-10.5625-9.625
6AP278TY/SHIP_NY_32S_240_33.625-34.625
6AP27995/SHIP_32_1BC
6AP29484/SHIP_NY_32K_100_15.1875-10.1875-9.625
6AP30895/SHIP_32_1BI
6AP31230/SHIP_NY_32K_100_12.3125-9.6875-7.875
6AP32472/SHIP_32_1AU
6AP32825/SHIP_32_1P
6AP352TR/SHIP_32_2A
6AP36528/SHIP_NY_32K_100_7.125-7-6.0625
6AP37072/SHIP_32_1BN
6AP43840/SHIP_32_1K
6AP44172/SHIP_32_1AU
6AP44584/SHIP_NY_40K_100_15.375-10.5625-9.625
6AP45520TK/SHIP_NY_44K_100_12.625-9.875-5.125
6AP46025/SHIP_NY_32CW_100_12.875-5.625-7.625
6AP471TY/SHIP_NY_32SBS_240_11.0625x15.0625x2
6AP479TY/SHIP_32_3A
6AP48595/SHIP_32_1BC
6AP50620/SHIP_NY_32K_100_10.625-10-5.625
6AP55084/SHIP_29_1B
6AP56025/SHIP_NY_32K_100_11.25-5.8125-8.625
6AP56325/SHIP_NY_32CW_100_12.875-5.625-7.625
6AP56784/SHIP_NY_44K_100_15.375-10.5625-9.625
6AP57225/SHIP_NY_32K_100_11.25-5.8125-4.3125
6AP57295/SHIP_NY_44K_100_15.1875-11-7.875
6AP57915/SHIP_44_2X
6AP58595/SHIP_32_1BI
6AP590DV/SHIP_44_3A
6AP590MDV/SHIP_44_3B
6AP590MTY/SHIP_44_8B
6AP590TY/SHIP_44_8A
6AP60280/SHIP_NY_32K_100_14.625-6.375-5
6AP60784/SHIP_NY_44K_100_15.375-10.5625-9.625
6AP61572/SHIP_44_2L
6AP62425/SHIP_NY_32K_100_12-7.25-6.75
6AP63025/SHIP_NY_32K_100_11.25-5.8125-4.3125
6AP63884/SHIP_40_1I
6AP64525/SHIP_NY_32K_100_11.25-5.8125-4.3125
6AP66125/SHIP_NY_44K_100_15.125-11.0625-6.9375
6AP66784/SHIP_NY_44K_100_15.375-10.5625-9.625
6AP70510/SHIP_NY_32K_100_18.875-11.625-9
6AP71084B/SHIP_32_1BN
6AP73972/SHIP_32_4A
6AP78872/SHIP_40_1D
6AP80625/SHIP_NY_32K_100_11.25-5.8125-8.625
6AP81284/SHIP_29_1B
6AP83284/SHIP_32_1G
6AP84372/SHIP_44_2L
6AP86772/SHIP_NY_32K_100_16.3125-12.25-7.5625
6AP87015/SHIP_NY_44K_100_17.9375-9-7.75
6AP89024/SHIP_NY_44K_100_20.375-15.625-6
6AP90184/SHIP_32_1G
6AP92084/SHIP_32_1AQ
6AP97584/SHIP_32_1AQ
6AQ006S15/SHIP_NY_40K_100_10.25-9.375-8.625
6AQ006S18/SHIP_NY_40K_100_10.25-9.25-10.75
6AQ006S30/SHIP_NY_44K_100_10.50-9.50-10.50
6AQ011S25/SHIP_NY_32K_100_11.375-8.125-4.500
6AQ019S36X/SHIP_NY_32K_100_15.50-9-10.50
6AQ028R25/SHIP_NY_44K_100_10.50-10.50-4.5
6AQ20372/SHIP_32_1AU
6AQ205R30/SHIP_NY_32K_100_8.25-5.25-5.25
6AQ235R25/SHIP_NY_32K_100_14.50-6-7.50
6AQ335R30/SHIP_NY_40K_100_10.75-8.75-6.00
6AQ499S25/SHIP_NY_40K_100_11.50-8.75-4
6AQ50720/SHIP_NY_32K_100_12.00-9.00-12.50
6AQ588R30/SHIP_NY_40K_100_10.25-9.625-5
6AQ59448/SHIP_32_1I
6AQ707R10/SHIP_NY_32K_100_10-5.25-6.25
6AQ76584/SHIP_NY_44K_100_15.25-10.375-7.6875
6AQ797HD/SHIP_32_8A
6AQ90172/SHIP_44_2L
6AQ95084/SHIP_44_2AK
6AS27350/SHIP_44_2O
6AS40750/SHIP_32_1AP
6AS41950/SHIP_40_1O
6AU075JR/SHIP_32_1AX
6AU12310/SHIP_32_1U
6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5
6AU13195/SHIP_NY_32K_100_13.4375-6.25-9.6875
6AU16500/SHIP_NY_32K_100_13.625-4.5-12.25
6AU20100/SHIP_NY_32K_100_14.875-9.625-7.625
6AU25101/SHIP_NY_32K_100_14.875-9.625-7.625
6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625
6AU26595/SHIP_NY_32K_100_13.625-4.5-12.25
6AU27500/SHIP_NY_32K_100_13.625-4.5-12.25
6AU33901/SHIP_NY_32K_100_11.8125-9.375-8.5
6AU34510/SHIP_NY_32K_100_8.125-8.125-9
6AU52601/SHIP_NY_32K_100_14.875-9.625-7.625
6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5
6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688
6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25
6AU71020/SHIP_40_1Y
6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75
6BA403S3/SHIP_44_14B
6BA403TR/SHIP_44_10A
6BA436XF6/SHIP_32_25A
6BA437FWB/SHIP_NY_44WP_100_14.125-7.6875-10.625
6BA580FW/SHIP_44_2W
6BA630FWB/SHIP_NY_44WP_100_14.125-7.6875-10.625
6BA67040/SHIP_32_1BR
6BA673S30/SHIP_NY_44K_150_20.5-8.3125-11.3125
6BA673TY/SHIP_NY_44SBS_10A_19.5-7.4375-11.125
6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125
6BA972CK/SHIP_NY_44K_100_18-8.25-10.25
6BB10800/SHIP_32_1S
6BC12400/SHIP_29_1H
6BC22370/SHIP_32_1N
6BD11000/SHIP_32_1AD
6BD17385/SHIP_44_2AB
6BD24780/SHIP_44_2Q
6BD27400/SHIP_32_1AD
6BD33333/SHIP_NY_32K_100_21-12.5-10.5
6BD35100/SHIP_32_1AD
6BD35185/SHIP_44_2P
6BD35187/SHIP_44_2P
6BD40812/SHIP_32_1E
6BD72400/SHIP_32_1AD
6BH46018/SHIP_29_1K
6BH59200/SHIP_NY_32K_100_13.4375-6.25-9.6875
6BH72500/SHIP_32_1AF
6BJ21396/SHIP_44_2V
6BJ66072/SHIP_32_1AB
6BM225XT/SHIP_44_2H
6BM233XT/SHIP_44_2H
6BM321S9/SHIP_32_24A
6BM420S9/SHIP_40_6A
6BMC667S4J/SHIP_32_26A
6BP02724/SHIP_32_1M
6BP03410/SHIP_NY_44K_100_18.875-11.625-9
6BP043P/SHIP_NY_32K_100_15.625-8-7
6BP131TW/SHIP_32_1AK
6BP13872/SHIP_NY_44K_100_15.25-10.375-15.4375
6BP14884/SHIP_40_1I
6BP21195/SHIP_NY_44K_100_14.375-12.1875-9.25
6BP22325/SHIP_NY_32K_100_12.4375-5.4375-4.1875
6BP22884/SHIP_NY_44K_100_15.375-10.5625-9.625
6BP23872/SHIP_NY_32K_100_16.3125-12.25-7.5625
6BP24284/SHIP_NY_32K_100_15.1875-10.1875-9.625
6BP24684/SHIP_NY_32K_100_15.1875-10.1875-9.625
6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875
6BP25784/SHIP_NY_32K_100_15.1875-10.1875-9.625
6BP27284/SHIP_NY_32K_100_15.1875-10.1875-9.625
6BP28425/SHIP_32_1P
6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875
6BP33724/SHIP_32_1M
6BP37815/SHIP_44_2X
6BP4004P/SHIP_44_2D
6BP43572/SHIP_32_1AU
6BP44584/SHIP_NY_40K_100_15.375-10.5625-9.625
6BP46025/SHIP_NY_32CW_100_13.5-6.375-8
6BP471TY/SHIP_NY_32SBS_240_11.0625x15.0625x2
6BP49584/SHIP_44_2K
6BP5005P/SHIP_44_2D
6BP54072/SHIP_44_2L
6BP63884/SHIP_NY_44K_100_15.375-10.5625-9.625
6BP71075/SHIP_32_1AC
6BP72125/SHIP_32_1P
6BP73684/SHIP_40_1I
6BP87015/SHIP_NY_44K_100_17.9375-9-7.75
6BP92084/SHIP_32_1AQ
6BP97584/SHIP_32_1AQ
6BQ59448/SHIP_32_1I
6BQ79724/SHIP_44_2AE
6BQ79772/SHIP_40_1D
6BQ95084/SHIP_NY_44K_100_15.375-10.5625-9.625
6BS42450/SHIP_32_1AY
6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25
6BU23949/SHIP_32_1L
6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625
6BU26595/SHIP_NY_32K_100_13.625-4.5-12.25
6BU27500/SHIP_NY_32K_100_13.625-4.5-12.25
6BU71020/SHIP_NY_40K_100_15.625-6.375-6.5
6CA403TR/SHIP_NY_44K_10A_19.75-7.875-11.125
6CA580FW/SHIP_NY_44K_100_16.75-8.125-10
6CA972CK/SHIP_NY_44K_100_18-8.25-10.25
6CB12800/SHIP_NY_29K_100_14.625-6.375-5
6CC12400/SHIP_NY_32K_100_15.75-8.625-5.1875
6CC22370/SHIP_NY_32K_100_11.125-6.375-10.375
6CC42500/SHIP_NY_32K_100_15.75-8.625-5.1875
6CD33333/SHIP_NY_44K_100_21-12.5-10.5
6CD35100/SHIP_32_1AD
6CD35185/SHIP_44_2P
6CD69500/SHIP_32_1AD
6CD72400/SHIP_NY_32K_100_13.25-12-6.125
6CD74800/SHIP_32_1AD
6CJ21396/SHIP_NY_44K_100_16.625-9.375-7
6CJ22750/SHIP_NY_44K_100_16.625-8.625-7
6CM233XT/SHIP_NY_44K_100_14.125-7.6875-10.625
6CP02724/SHIP_NY_32K_100_11-5.8125-4.0625
6CP03410/SHIP_NY_44K_100_18.875-11.625-7.6875
6CP13872/SHIP_NY_44K_100_15.25-10.375-15.4375
6CP14884/SHIP_NY_40K_100_15.375-10.5625-9.625
6CP214TY/SHIP_32_7A
6CP21784/SHIP_32_1AO
6CP22884/SHIP_NY_44K_100_15.375-10.5625-9.625
6CP23872/SHIP_NY_32K_100_16.3125-12.25-7.5625
6CP24572/SHIP_NY_32K_100_16.3125-12.25-7.5625
6CP33724/SHIP_NY_32K_100_11-5.8125-4.0625
6CP37815/SHIP_44_2X
6CP43572/SHIP_32_1AU
6CP44584/SHIP_NY_40K_100_15.375-10.5625-9.625
6CP46025/SHIP_NY_32CW_100_12.875-5.625-8
6CP60125/SHIP_32_1P
6CP63884/SHIP_NY_44K_100_15.375-10.5625-9.625
6CP71084/SHIP_29_1B
6CP92084/SHIP_32_1AQ
6CP97584/SHIP_NY_40K_100_15.375-10.5625-9.625
6CQ79772/SHIP_40_1D
6CQ95084/SHIP_NY_44K_100_15.375-10.5625-9.625
6CS42450/SHIP_NY_29K_100_17.125-6.6875-11.75
6CT00229/SHIP_32_1AA
6DA580FW/SHIP_NY_44K_100_16.75-8.125-10.75
6DB60307/SHIP_40_1L
6DB71200/SHIP_32_1AL
6DB77801/SHIP_29_1L
6DC12400/SHIP_NY_32K_100_15.75-8.625-5.1875
6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375
6DD33333/SHIP_NY_44K_100_21-12.5-10.5
6DD35185/SHIP_32_1BS
6DD65500/SHIP_32_1AD
6DD72400/SHIP_NY_32K_100_13.25-12-6.125
6DD74800/SHIP_NY_32K_100_13.25-12-6.125
6DH04082/SHIP_32_1AF
6DJ14108/SHIP_44_2U
6DJ14143/SHIP_44_2S
6DM225XT/SHIP_NY_44K_100_14.125-7.6875-10.625
6DM233XT/SHIP_NY_44K_100_14.125-7.6875-10.625
6DP13184/SHIP_32_30A
6DP13872/SHIP_NY_44K_100_15.1875-10.1875-15.1875
6DP22884/SHIP_NY_32K_100_15.1875-10.1875-9.625
6DP43572/SHIP_32_1AU
6DP44584/SHIP_NY_40K_100_15.375-10.5625-9.625
6DP63884/SHIP_NY_44K_100_15.375-10.5625-9.625
6DP71084/SHIP_29_1B
6DP97584/SHIP_NY_40K_100_15.375-10.5625-9.625
6DQ387S40/SHIP_NY_40K_100_14.125-7.6875-5
6DQ59448/SHIP_40_1W
6DQ85484/N/A
6DS42450/SHIP_NY_29K_100_17.125-6.6875-11.75
6DS68625/SHIP_29_1J
6EA500F48/SHIP_40_1U
6EA580FW/SHIP_NY_44K_100_16.750-8.125-10.750
6EB10300/SHIP_NY_29K_100_14.625-6.375-5
6EB37900/SHIP_32_1BE
6EB71200/SHIP_NY_32K_102_14.625-6.375-5
6EC12400/SHIP_NY_32K_100_15.75-8.625-5.1875
6EC69900/SHIP_NY_32K_100_15.75-8.625-5.1875
6ED12555/SHIP_NY_32K_100_10.625-6.375-11.25
6ED35185/SHIP_32_1BU
6EH04082/SHIP_NY_32K_100_13.4375-6.25-9.6875
6EJ14108/SHIP_NY_44K_100_16.625-8.625-7
6EJ14143/SHIP_NY_44K_100_16.5-10.875-5.5
6EM225XT/SHIP_NY_44K_100_14.125-7.6875-10.625
6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625
6EP63884/SHIP_NY_32K_100_15.1875-10.1875-9.625
6EP71084/SHIP_29_1B
6EQ387S40/SHIP_NY_40K_100_14.125-7.6875-5
6EQ59470/SHIP_32_1Q
6EQ85484/SHIP_40_1I
6EQ87084/SHIP_NY_44K_100_15.375-10.5625-9.625
6EQ88272/SHIP_44_2L
6ES42450/SHIP_NY_29K_100_17.125-6.6875-11.75
6ES48450/SHIP_NY_29K_100_17.125-6.6875-11.75
6EU13295/SHIP_NY_32K_100_13.625-4.5-12.25
6FB10300/SHIP_NY_32K_102_14.625-6.375-5
6FB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875
6FB60335/SHIP_29_1L
6FB71200/SHIP_NY_32K_102_14.625-6.375-5
6FC69900/SHIP_NY_32K_100_15.75-8.625-5.1875
6FD35185/SHIP_NY_32K_100_13.25-12-12.4375
6FH04082/SHIP_NY_32K_100_13.4375-6.25-9.6875
6FH24795/SHIP_32_1AF
6FJ14108/SHIP_NY_44K_100_16.625-8.625-7
6FJ14143/SHIP_NY_44K_100_16.5-10.875-5.5
6FM225XT/SHIP_NY_44K_100_14.125-7.6875-10.625
6FP43572/SHIP_NY_32K_100_16.3125-12.25-7.5625
6FP92084/SHIP_NY_40K_100_15.375-10.5625-9.625
6FQ387S40/SHIP_NY_40K_100_14.125-7.6875-5
6FQ59470/SHIP_32_1Q
6FQ87084/SHIP_NY_44K_100_15.375-10.5625-9.625
6FQ88272/SHIP_NY_44K_100_15.25-10.375-15.4375
6FQ94384/SHIP_NY_44K_100_15.375-10.5625-9.625
6FS48450/SHIP_NY_29K_100_17.125-6.6875-11.75
6GA580FW/SHIP_NY_44K_100_16.750-8.125-10.750
6GB10300/SHIP_NY_32K_100_14.625-6.375-5
6GB33935/N/A
6GB33937/SHIP_29_1C
6GB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875
6GC69900/SHIP_NY_32K_100_15.75-8.625-5.1875
6GD35185/SHIP_NY_32K_100_13.25-12-12.4375
6GH04082/SHIP_NY_32K_100_13.4375-6.25-9.6875
6GQ387S40/SHIP_NY_40K_100_14.125-7.6875-5
6GQ70040/SHIP_40_1P
6GQ87084/SHIP_NY_44K_100_15.375-10.5625-9.625
6GQ88272/SHIP_NY_44K_100_15.25-10.375-15.4375
6GQ94384/SHIP_NY_44K_100_15.375-10.5625-9.625
6GS41125/SHIP_29_1J
6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75
6GT00128/SHIP_32_1AZ
6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875
6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875
6HH04082/SHIP_NY_32K_100_13.4375-6.25-9.6875
6HQ08472/SHIP_44_2L
6HQ34540/SHIP_40_1P
6HQ50772/SHIP_44_2L
6HQ59470/SHIP_44_2AO
6HQ70040/SHIP_40_1P
6HQ85084/SHIP_40_1I
6HQ87084/SHIP_NY_32K_100_15.1875-10.1875-9.625
6HQ88272/SHIP_NY_44K_100_15.1875-10.1875-15.1875
6HS41125/SHIP_NY_29K_100_17.125-6.6875-11.75
6HS41350/SHIP_26_1C
6HS48050/SHIP_29_1J
6HS82125/SHIP_29_1J
6HU87295/SHIP_NY_32K_100_13.625-4.5-12.25
6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750
6JB40600/SHIP_NY_29K_100_14.625-6.375-5
6JC29600/SHIP_29_1H
6JQ08472/SHIP_NY_44K_100_15.1875-10.1875-15.1875
6JQ34540/SHIP_NY_40K_100_14.125-7.6875-5
6JQ50772/SHIP_NY_44K_100_15.25-10.375-15.4375
6JQ55172/SHIP_44_2L
6JQ70040/SHIP_NY_40K_100_14.125-7.6875-5
6JQ85084/SHIP_NY_40K_100_15.375-10.5625-9.625
6JQ86984/SHIP_40_1I
6JS16225/SHIP_26_1C
6JS23125/SHIP_32_1AY
6JS41125/SHIP_NY_29K_100_17.125-6.6875-11.75
6JS41350/SHIP_NY_29K_100_17.125-6.6875-11.75
6JS48050/SHIP_NY_29K_100_17.125-6.6875-11.75
6JS82125/SHIP_NY_29K_100_17.125-6.6875-11.75
6JU87295/SHIP_NY_32K_100_13.625-4.5-12.25
6KB40600/SHIP_NY_29K_100_12.375-5.5-4.625
6KQ34540/SHIP_NY_40K_100_14.125-7.6875-5
6KQ50772/SHIP_NY_44K_100_15.25-10.375-15.4375
6KQ70040/SHIP_NY_40K_100_14.125-7.6875-5
6KQ85084/SHIP_NY_40K_100_15.375-10.5625-9.625
6KQ86984/SHIP_NY_40K_100_15.375-10.5625-9.625
6KS16225/SHIP_NY_29K_100_17.125-6.6875-11.75
6KS23125/SHIP_NY_29K_100_17.125-6.6875-11.75
6KS41125/SHIP_NY_29K_100_17.125-6.6875-11.75
6KS41350/SHIP_NY_29K_100_17.125-6.6875-11.75
6KS48050/SHIP_NY_29K_100_17.125-6.6875-11.75
6KS82125/SHIP_NY_29K_100_17.125-6.6875-11.75
6KU87295/SHIP_NY_32K_100_13.625-4.5-12.25
6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625
6LQ50772/SHIP_NY_44K_100_15.1875-10.1875-15.1875
6LQ55172/SHIP_44_2L
6LQ70040/SHIP_NY_40K_100_14.125-7.6875-5
6LQ85084/SHIP_NY_40K_100_15.375-10.5625-9.625
6LQ86984/SHIP_NY_40K_100_15.375-10.5625-9.625
6LQ89072/SHIP_44_2L
6LS16225/SHIP_NY_29K_100_17.125-6.6875-11.75
6LS23125/SHIP_NY_29K_100_17.125-6.6875-11.75
6LS41350/SHIP_NY_29K_100_17.125-6.6875-11.75
6LS48050/SHIP_NY_29K_100_17.125-6.6875-11.75
6LS82125/SHIP_NY_29K_100_17.125-6.6875-11.75
6LU87295/SHIP_NY_32K_100_13.625-4.5-12.25
6MQ55172/SHIP_NY_44K_100_15.25-10.375-15.4375
6MQ85084/SHIP_NY_32K_100_15.1875-10.1875-9.625
6MQ86984/SHIP_NY_40K_100_15.375-10.5625-9.625
6MQ89072/SHIP_NY_44K_100_15.25-10.375-15.4375
6MS23125/SHIP_NY_29K_100_17.125-6.6875-11.75
6MS48050/SHIP_NY_29K_100_17.125-6.6875-11.75
6MU87295/SHIP_NY_32K_100_13.625-4.5-12.25
6NQ55172/SHIP_NY_44K_100_15.25-10.375-15.4375
6NQ86984/SHIP_NY_44K_100_15.375-10.5625-9.625
6NQ89072/SHIP_NY_44K_100_15.25-10.375-15.4375
6NS23125/SHIP_NY_29K_100_17.125-6.6875-11.75
6OA119S3B/SHIP_40_1S
6OA119S48/SHIP_32_1BJ
6OA144S30/SHIP_NY_32NA_200_11.375-1.250-10.625
6OA203R42/SHIP_NY_32K_100_13-7-6
6OA203T48/SHIP_NY_40K_100_21.5-7.0-10.0
6OA204R18/SHIP_NY_32K_100_9.375-4.75-7.125
6OA250S56/SHIP_NY_40K_100_14-8-8.5
6OA30248/SHIP_32_1BP
6OA343S16/SHIP_NY_40K_100_12.50-9.50-10.75
6OA345S40/SHIP_NY_32K_100_13.0-8.0-8.75
6OA346S48/SHIP_NY_32K_100_13.0-8.0-10.5
6OA350S64/SHIP_NY_40K_100_14-8-9.25
6OA366F48/SHIP_40_1C
6OA366R48/SHIP_44_2A
6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875
6OA42640/SHIP_40_1F
6OA42640A/SHIP_40_1X
6OA436TC/SHIP_NY_32NA_201_47.25-37.25-3
6OA437FW/SHIP_44_2H
6OA44872/SHIP_NY_32K_100_13-10-7
6OA466DR/SHIP_32_29A
6OA466S48/SHIP_44_2AN
6OA494S48/SHIP_NY_32K_100_12.25-7.5-7.125
6OA498S30/SHIP_NY_32K_100_9.5-8.75-7
6OA507S40/SHIP_NY_32K_100_13.75-8.5-6.75
6OA52048/SHIP_NY_40K_100_12.5-7.25-6.25
6OA527R40/SHIP_40_1Q
6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625
6OA54130AZ/SHIP_32_1BK
6OA543TR/SHIP_200_4A
6OA554CST/SHIP_NY_NAK_231_29-8.50-3.375
6OA557TY/SHIP_32_19A
6OA560S4B/SHIP_NY_40K_100_13-7-8.25
6OA56448/SHIP_44_2B
6OA580CB/Ship_NA_1A
6OA582CB/SHIP_NY_NA_200_2.5-2.5-32.625
6OA586R42/SHIP_NY_44K_100_13-7-6
6OA587S40/SHIP_NY_32K_100_15.75-8-7
6OA599CB/SHIP_NY_NAFB_251_2.5-2.5-37.0
6OA599PD/SHIP_NY_32K_219_41.125-35.5
6OA599TC/SHIP_NY_32K_201_40.75-36-3
6OA605CB/Ship_NY_NA_230_3-3-38_0.225
6OA612R3B/SHIP_NY_32K_100_9-7-9
6OA62140/SHIP_40_1A
6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125
6OA63748/SHIP_NY_44K_100_13-10.625-8.75
6OA65748/SHIP_44_2C
6OA67040/SHIP_32_1BH
6OA682S96/SHIP_NY_44K_100_10.375-6.125-8
6OA709PD/SHIP_NY_40NA_250_14.75-7
6OA717F604/SHIP_NY_32K_100_12-7.25-6.75
6OA72148/SHIP_44_2J
6OA732S60/SHIP_NY_32K_100_17.25-7-11
6OA737S48/SHIP_NY_40K_100_10.5-9.25-6.5
6OA738S3BL/SHIP_NY_44K_100_17.5-6.75-5
6OA738S48/SHIP_NY_40K_100_13.5-6.25-6.5
6OA75948/SHIP_NY_44K_100_11.6875-6.25-6.75
6OA766TR/SHIP_32_22C
6OA796R48/SHIP_NY_32K_100_10.375-6.125-8
6OA8003TC/SHIP_NY_32K_201_43.25-38.00-3.00
6OA8009S110/SHIP_NY_32K_100_9-7.625-8.25
6OA8026R5BL/SHIP_NY_44K_100_17.25-7.00-6.00
6OA8033S48/SHIP_NY_32K_100_9.00-7.50-6.00
6OA8054INS/SHIP_NY_32K_231_8.688-7.625-2.25
6OA8054PD/SHIP_NY_32K_201_39.75-36.00
6OA8054TC/SHIP_NY_32K_201_40.75-37.00-3.00
6OA8071S3BL/SHIP_NY_40K_100_17.25-7.00-8.00
6OA808F2B/SHIP_32_1BF
6OA8092TC/SHIP_NY_32K_201_47.75-37.25-3
6OA8109S24/SHIP_NY_32K_100_9.375-9.188-8.125
6OA8121S42/SHIP_NY_32K_100_9-6.75-5.50
6OA8125S4/SHIP_NY_44K_100_19.25-15.25-6.25
6OA9008GT/SHIP_NY_600#K_150_40.25-40.25-34
6OA9009S32/SHIP_NY_32K_100_12.5-6.5-7
6OA9010S32/SHIP_NY_32K_100_12.5-6.5-7
6OA9011S32/SHIP_NY_32K_100_12.5-6.5-7
6OA9012S32/SHIP_NY_32K_100_12.5-6.5-7
6OA9013S32/SHIP_NY_32K_100_12.5-6.5-7
6OA9016S32/SHIP_NY_32K_100_12.5-6.5-7
6OA9048R40/SHIP_NY_32K_100_12-8-8
6OA917S42/SHIP_NY_32K_100_13.75-7.50-6.00
6OA935S60/SHIP_NY_32K_100_9.5-7.0-7.75
6OA936TWBL/SHIP_NY_32K_100_14-7-10.25
6OA999DR/SHIP_200_5C
6OA999FP/SHIP_44_13A
6OA999HS/SHIP_44_11A
6OA999PS/SHIP_200_5D
6OA999SH/SHIP_44_12A
6OA999TC/SHIP_32_22B
6OA999TR/SHIP_200_6A
6OA999VS/SHIP_44_11B
6OB01148/SHIP_32_1E
6OB01148A/SHIP_32_1E
6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75
6OB29307/SHIP_NY_32K_100_8.625-8.625-11.1875
6OB31400/SHIP_32_1AL
6OB33973/SHIP_44_2F
6OB389PD/Ship_NA_3A
6OB43201/SHIP_NY_32K_100_8.25-4.75-7.125
6OB48903/SHIP_40_1L
6OC10560/SHIP_32_1BV
6OC22324/SHIP_32_1N
6OD00473/SHIP_32_1BD
6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875
6OD02400/SHIP_NY_32K_100_14.875-9.625-7.625
6OD11087/SHIP_44_2AB
6OD13500A/SHIP_32_1AD
6OD17748/SHIP_29_1D
6OD23980/SHIP_44_2Q
6OD239DR/SHIP_44_1A
6OD239DV/SHIP_44_1D
6OD30495/SHIP_40_1N
6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25
6OD42240/SHIP_32_1D
6OD438SS/SHIP_32_12A
6OD45460/SHIP_32_1BM
6OD51048/SHIP_32_1AE
6OD535WD/SHIP_40_2A
6OD57630/SHIP_44_2T
6OD606CV/SHIP_29_2A
6OD71230/SHIP_NY_32K_100_9.75-6.4375-7.125
6OD71248/SHIP_NY_32K_100_13-10-10.375
6OD71890/SHIP_NY_32K_100_9.375-9.1875-8.125
6OD77440/SHIP_29_1F
6OG202D2/SHIP_40_1R
6OH01224/SHIP_32_1W
6OH11506/SHIP_32_1T
6OH11524/SHIP_NY_32K_100_12.4375-9-5.25
6OH11524A/SHIP_32_1AJ
6OH124CK/SHIP_NY_44K_100_14.5-10.75-6.5
6OH14324/SHIP_32_1Y
6OH143WM/SHIP_32_23A
6OH18436/SHIP_32_1AM
6OH21112/SHIP_32_1F
6OH21130/SHIP_32_1AN
6OH211GB/SHIP_32_1AV
6OH22018/SHIP_32_1B
6OH22524/SHIP_NY_32K_100_9.875-6.1875-6.875
6OH36328/SHIP_32_1Z
6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375
6OH51000/SHIP_44_2R
6OH53022/SHIP_32_1AS
6OJ112CB/SHIP_NY_NA_200_2.5-2.5-19.375"
6OJ16530/SHIP_NY_32K_100_11.625-7.875-7.5
6OJ17142/SHIP_NY_32K_100_15.375-11.5625-10.5625
6OJ29143/SHIP_44_2S
6OJ45143/SHIP_275_1A
6OJ66072/SHIP_32_1C
6OJ701HSC/SHIP_NY_32K_150_20.9375-16.8125-8.3125
6OK45781/SHIP_44_2Z
6OK75595/SHIP_32_1R
6OM105CB/SHIP_NY_NAFB_251_2.5-2.5-43.0
6OM1288S88/SHIP_44_2Y
6OM217CB/SHIP_NY_NAFB_251_2.5-2.5-34.5
6OM245S56/SHIP_NY_40K_100_18.5-8.5-7.0
6OM260S728/SHIP_NY_32K_100_9.375-8.188-9.625
6OM311S72/SHIP_NY_40K_100_18.688-9.188-8.50
6OM345S72/SHIP_NY_44K_100_18.5-8.5-8.625
6OM345T72/SHIP_NY_32K_100_21.75-9-12.5
6OM35650/SHIP_NY_32K_100_15.625-6.125-10.25
6OM416SF3B/SHIP_44_2AJ
6OM426CB/SHIP_NY_NAFB_251_2.50-2.50-41.00
6OM445S64/SHIP_NY_44K_100_18.50-8.50-8.00
6OM521S3B/SHIP_NY_44K_100_22.50-9.25-7.75
6OM521S64/SHIP_NY_32K_100_13.75-8.5-9.5
6OM525S3B/SHIP_NY_44K_100_22.50-9.25-8.75
6OM525S72/SHIP_NY_32K_100_13.75-8.5-10.75
6OM608TR/SHIP_32_22A
6OM616TC/SHIP_NY_32K_201_42.50-33.25-3.000
6OM632TR/SHIP_NY_32K_201_47-34.500-3.000
6OM653TR/SHIP_NY_29K_202_15.375-13.875-4.000
6OM654EZTR/SHIP_NY_29K_201_42.3125-39.500-3.000
6OM752EZCB/Analyti+C455:X455cal - SHIP_NY_NA_200_2.5-2.5-37.375
6OM8051S72/SHIP_NY_40K_100_18.50-8.50-5.75
6OM8074S60/SHIP_NY_40K_100_19.50-9.00-7.00
6OM8095CB/SHIP_NY_NAFB_251_2.50-2.50-38.25
6OM8096CB/SHIP_NY_NAFB_251_2.50-2.50-40.5
6OM814EZTR/SHIP_NY_29K_201_47.3125-37.9375-3.000
6OM816EZTR/SHIP_NY_29K_201_44.8125-39.500-3.000
6OM819EZCB/SHIP_NY_NA_200_2.5-2.5-37.875
6OM8230S886/SHIP_NY_32K_100_9.375-8.188-8
6OM911S3BL/SHIP_NY_44K_100_17.75-8.75-8.00
6OM924CB/SHIP_NY_NA_200_2.5-2.5-43.75
6OM924TY/SHIP_29_3A
6OM929HSC/SHIP_44_14A
6OM936S80/SHIP_NY_40K_100_10.50-9.00-13.625
6OM956TR/SHIP_NY_32_201_45.625-39.25-3
6OM960CB/SHIP_NY_NA_200_2.5-2.5-33.75
6OM960TY/SHIP_32_21A
6OM961CB/SHIP_NY_NAFB_251_2.50-2.50-43.50
6OM967S3BL/SHIP_NY_44K_100_17.75-8.75-7.00
6OM968CB/SHIP_NY_NAFB_251_2.5-2.5-37.25
6OM970BNS/SHIP_NY_32K_100_14-8.188-8
6OM975S72/SHIP_NY_32K_100_9.5-8.75-9.5
6OMC916T80/SHIP_NY_32K_100_11.25-8.75-8.5
6OP011201/SHIP_32_1BB
6OP0112FP/SHIP_32_16A
6OP012300/SHIP_44_2AG
6OP0123SF/SHIP_200_1A
6OP0123TF/SHIP_44_9A
6OP0124BF/SHIP_200_1C
6OP01784/SHIP_NY_29K_100_10-6.750-7
6OP027LP/SHIP_32_13A
6OP07530/SHIP_NY_44K_100_14.5-6.75-9.75
6OP080TC/SHIP_NY_29K_201_44.250-36.750-3.000
6OP081LP/SHIP_NY_32NA_250_17.250-11.938
6OP081TY/SHIP_NY_32W_202_19.375-13.4375-10.875
6OP08725/SHIP_NY_44K_100_11.5-9.5-6
6OP09505/SHIP_44_2AA
6OP11325/SHIP_NY_44K_100_10.25-9.75-7.75
6OP12684/SHIP_44_2AK
6OP126DV/SHIP_32_18A
6OP12984/SHIP_32_1V
6OP14025/SHIP_NY_40K_100_13-8-5.25
6OP14205/SHIP_32_1O
6OP14725/SHIP_NY_44K_100_14.5-10.75-6.5
6OP14930/SHIP_NY_44K_100_15.5-8.375-10.125
6OP15540/SHIP_NY_32K_100_11-11-7.75
6OP214CB/Ship_NA_1C
6OP214CP/SHIP_32_20A
6OP214FP/SHIP_32_4B
6OP214SP/SHIP_44_6A
6OP34524/SHIP_44_2G
6OP349TW/SHIP_NY_40K_100_14-9.1875-9.5
6OP352PQ/SHIP_32_15A
6OP35420/SHIP_NY_32K_100_10.625-7-5.625
6OP38372/SHIP_NY_40K_100_13.625-10.1875-8.3125
6OP400DV/SHIP_44_1B
6OP40384/SHIP_NY_32K_100_13.5625-9.0625-9.5
6OP42195/SHIP_32_1BQ
6OP46515/SHIP_NY_32K_100_10.75-6.5-5.75
6OP467SH/SHIP_NY_32K_100_13.75-11.125-7
6OP479CP/SHIP_32_19A
6OP479FP/SHIP_32_1BA
6OP479PS/SHIP_26_2A
6OP48720/SHIP_NY_32K_100_14-8.75-4.6875
6OP52784/SHIP_29_1B
6OP52784A/SHIP_NY_32K_100_10-6.75-5.75
6OP572LP/SHIP_NY_32NA_250_15-10.875
6OP59084/SHIP_32_1H
6OP590CB/Ship_NA_1B
6OP590CP/SHIP_44_7A
6OP590SS/SHIP_32_11A
6OP65772/SHIP_NY_40K_100_13.625-10.1875-8.625
6OP705PD/SHIP_NY_32NA_250_18.6875-11.4375
6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625
6OP823TC/SHIP_NY_32K_201_47.25-37.500-11.063
6OP83872/SHIP_32_1AH
6OP90025/SHIP_NY_40K_100_14.125-10-9.625
6OP90625/SHIP_NY_32K_100_14.125-10-10.875
6OP906LP/SHIP_NY_32NA_250_13.875-9.250
6OP91325/SHIP_NY_44K_100_13.75-9-5.5
6OQ007STW/SHIP_NY_32K_100_12-8.625-9.188
6OQ008S25/SHIP_NY_44K_100_10.375-10.375-4.5
6OQ011S25/SHIP_NY_32K_100_11.375-8.125-4.50
6OQ070S150HC/SHIP_NY_55K_100_21.625-12.125-12.625
6OQ137S25/SHIP_NY_32K_100_13-8.50-5.00
6OQ14372/SHIP_32_1AI
6OQ153S25/SHIP_NY_44K_100_12-9.25-8
6OQ181R15/SHIP_32_1BO
6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188
6OQ224R25/SHIP_NY_40K_100_11.75-8.125-4.813
6OQ231R25/SHIP_NY_40K_100_14.25-8.75-5.0
6OQ235R25/SHIP_NY_32K_100_14.5-6-7.5
6OQ301S30/SHIP_44_2AL
6OQ335R30/SHIP_40_1V
6OQ346R20/SHIP_NY_40K_100_12.5-9.5-10.875
6OQ349S25/SHIP_NY_32K_100_11.875-5-8
6OQ4171R12/SHIP_NY_44K_100_15.188-12.375-5.125
6OQ445S25/SHIP_NY_32K_100_11.625-4.5-4.875
6OQ451R05/SHIP_NY_32K_100_8-5-6.25
6OQ467S30/SHIP_NY_40K_100_17-7.50-9.25
6OQ5014R25/SHIP_NY_32K_100_12.00-5.00-7.50
6OQ5018S25/SHIP_NY_32K_100_9.25-6.00-8.00
6OQ5020S15/SHIP_NY_40K_100_9.875-9.6875-9.875
6OQ5032S30/SHIP_NY_32K_100_10.438-9.875-9.875
6OQ5042S36/SHIP_NY_40K_100_15.50-9-10.50
6OQ50720/SHIP_NY_32K_100_12.00-9.00-12.50
6OQ516S25/SHIP_NY_44K_100_11.875-5.0-7.0
6OQ543R25/SHIP_NY_32K_100_10.50-5.50-9.00
6OQ548R15/SHIP_NY_32K_100_12.375-8.875-9.375
6OQ59448/SHIP_32_1I
6OQ594LP/SHIP_32_17A
6OQ60584/SHIP_44_2K
6OQ607R30/SHIP_NY_40K_100_10.875-5.75-10.25
6OQ613R700/SHIP_NY_44K_100_17-11.50-9
6OQ635S30/SHIP_NY_40K_100_14.25-8.25-5
6OQ637S20/SHIP_NY_32K_100_11.25-8.00-6.00
6OQ701BK/SHIP_NY_44K_100_19.625-9.25-10.75
6OQ707R10/SHIP_NY_32K_100_10-5.25-6.50
6OQ778S25/SHIP_NY_32K_100_12.375-8.125-5.375
6OQ797BS/SHIP_40_4A
6OQ797FL/SHIP_44_5A
6OQ797TY/SHIP_40_5A
6OQ8008R30/SHIP_NY_44K_100_14.50-8.50-8.00
6OQ8050R2BL/SHIP_NY_44K_100_12.00-9.00-6.00
6OQ826S25/SHIP_NY_32K_200_14.5-6-4
6OQ84584/SHIP_40_1G
6OQ850DV/SHIP_44_1C
6OQ85984/SHIP_NY_44K_100_15.375-10.5625-9.625
6OQ87170/SHIP_32_1Q
6OQ890PD/SHIP_32_13B
6OQ916S40/SHIP_NY_32K_100_12.25-5.25-6.5
6OQ972R15/SHIP_NY_32K_100_12.00-9.00-9.25
6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938
6OS273IN/SHIP_40_3A
6OS42850/SHIP_44_2N
6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938
6OS46250/SHIP_29_1J
6OU12066/SHIP_40_1M
6OU129TY/SHIP_NY_32White_203_6.5-9.625-4.5
6OU133TY/SHIP_NY_32White_203_6.5-9.625-4.5
6OU141TY/SHIP_NY_32White_203_6.5-9.625-4.5
6OU21470/SHIP_40_1J
6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125
6OU35199/SHIP_NY_32K_100_19.75-12-11.375
6OU351LP/SHIP_NY_32K_NA_19.500X11.750
6OU351SS/SHIP_NY_32K_NA_44.250X39.750
6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500
6OU37230/SHIP_NY_32K_100_10.75-8-6.125
6OU38510/SHIP_32_1BG
6OU45160/SHIP_29_1A
6OU47500/SHIP_NY_29K_100_10-9.5625-6.75
6OU55510/SHIP_NY_32K_100_11.75-6.25-8
6OU62620/SHIP_NY_32K_100_15-9.5-8
6OX12309/SHIP_NY_40K_100_14.5-9-9.750
6OX13109/SHIP_NY_44K_100_13.875-9.375-12.625
6OX131DVD/SHIP_NY_44K_250_13.625-2.9375-10.875
6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500
6PQ55172/SHIP_NY_44K_100_15.25-10.375-15.4375
6PQ86984/SHIP_NY_32K_100_15.1875-10.1875-9.625
6PQ89072/SHIP_NY_44K_100_15.1875-10.1875-15.1875
6PS23125/SHIP_NY_29K_100_17.125-6.6875-11.75
6QQ55172/SHIP_NY_44K_100_15.25-10.375-15.4375
6RQ55172/SHIP_NY_44K_100_15.1875-10.1875-15.1875
6UK88381/SHIP_44_2AD
6YA150TR/SHIP_NY_29K_201_47.625-37.3125-3.000
6YM4423TR/SHIP_NY_29K_201_40.4375-39.500-3.000
6YM4426CB/SHIP_NY_NA_200_2.5-2.5-37.125
6YM4426TR/SHIP_NY_29K_201_40.4375-38.875-3.000
6YM705CP/SHIP_NY_32K_221_47.50-37.00-4.00
6YM706EZTR/SHIP_NY_29K_201_40.750-37.750-3.000
6YM934CP/SHIP_32_1BT
6YMC667TR/SHIP_NY_32NA_201_46-36.375-3
6YQ705CP/SHIP_NY_32K_221_47.50-37.00-4.00
6YQ705FIL/SHIP_NY_32K_200_14.813-4.438-18.0
6ZS505PD/SHIP_32_14A
7A4AA580/Chem_NY_Disodium Cocoamphodipropinate
7A4AL930/Chem_08A
7A4AS162/Chem_09A
7A4BF774/Chem_10A
7A4BL625/Chem_11A
7A4BQ164/Chem_12A
7A4CD419/Chem_13A
7A4FD110/Chem_Sodium Chloride USP/NF
7A4OA019/Chem_NY_Mackam 2C-LV
7A4OA234/Chem_15A
7A4OA347/Chem_NY_Sodium-L(+) Lactate 60%
7A4OA443/Chem_16A
7A4OA709/Chem_NY_Chamomile Extract # 11-19-2
7A4OA863/Chem_NY_Lavender Silk UAC0285/00
7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1
7A4OA940/Chem_NY_Carbopol 940 RI - 0231
7A4OB223/Chem_189A
7A4OB314/Chem_NY_Isopropyl alcohol 99% USP
7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline
7A4OB397/Chem_04A
7A4OB401/Chem_NY_LEXFEEL D4_01
7A4OB690/Chem_01A
7A4OC109/Chem_NY_Sodium Xylene Sulfonate
7A4OD310/Chem_17A
7A4OD791/Chem_CB_Fleet
7A4OF292/Chem_18A
7A4OF602/Chem_19A
7A4OF774/Chem_01A
7A4OG169/Chem_20A
7A4OH247/Chem_21A
7A4OH460/Chem_22A
7A4OJ137/Chem_NY_Glycerine, USP Kosher
7A4OK459/Chem_NY_Salicylic Acid USP
7A4OM002/Chem_NY_ Dissolvine GL-47-S
7A4OM799/Chem_NY_Camillia Sinensis Leaf Extract in Glycerin
7A4OM962/Chem_NY_Plantacare 818 UP
7A4OP027/Chem_NY_Ceraphyl 41
7A4OP054/Chem_NY_Mango Extract 217552 PF
7A4OP065/Chem_NY_ Fragrance Plaisir Gourmand 23
7A4OP073/Chem_NY_Fragrance Bali Leia 5
7A4OP081/Chem_NY_Fragrance SC 2705 LF
7A4OP093/Chem_NY_Essential Oil 27A Fragrance
7A4OP113/Chem_NY_Vitis Vinifera (Grape) Fruit Water_01
7A4OP197/Chem_NY_D-Panthenol 75W_01
7A4OP203/Chem_NY_Hydroxyethyl Cellulose - RI-0427
7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract
7A4OP207/Chem_NY_Phytosphingosine HCL_01
7A4OP211/Chem_188A
7A4OP230/Chem_NY_Tegosoft CT
7A4OP284/Chem_NY_Grape Fruit Extract
7A4OP345/Chem_195A
7A4OP448/Chem_208A
7A4OP485/Chem_NY_Panthenol
7A4OP567/Chem_NY_Surfonic L12-6_Surfactant
7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft
7A4OP630/Chem_NY_Fragrance Orchid Dew
7A4OP661/Chem_NY_Atopalm Cream
7A4OP705/Chem_NY_Allantoin
7A4OP806/Chem_205A
7A4OP862/Chem_191A
7A4OQ028/Chem_NY_Emulgin VL-75
7A4OQ084/Chem_23A
7A4OQ100/Chem_24A
7A4OQ132/Chem_25A
7A4OQ203/Chem_NY_Carbopol Ultrez-10_7A4OQ203
7A4OQ214/Chem_27A
7A4OQ229/Chem_Trisodium Citrate Dihydrate USP
7A4OQ260/Chem_29A
7A4OQ353/Chem_30A
7A4OQ518/Chem_31A
7A4OQ595/Chem_32A
7A4OQ758/Chem_NY_Protachem HCO-40
7A4OQ792/Chem_33A
7A4OQ826/Chem_NY_Calming Facial Fragrance
7A4OQ855/Chem_NY_Naturals Fragrance
7A4OQ890/Chem_34A
7A4OQ904/Chem_NY_Lemon DUP Fragrance for wet wipes
7A4OQ992/Chem_NY_Actiphyte Sea Kelp AQ PH
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP
7A4OS411/Chem_36A
7A4OS639/Chem_37A
7A4OU351/Chem_NY_Dehydrated Alcohol USP
7A4OU555/Chem_NY_Phenylephrine HCL
7A4OU870/Chem_38A
7A4OX110/Chem_39A
7A4OX131/Chem_NY_Hydrogen Peroxide
7A4OX231/Chem_NY_D&C Yellow 10
7A4OX232/Chem_40A
7A4OY806/Chem_41A
7ACHE01/Chem_NY_Lamesoft PO 65
7ACHE02/Chem_NY_Caprylic/Capric Triglyceride
7ACHE03/Chem_NY_Bahama Wave 12-15-66665
7ACHE06/Chem_NY_Apricot FCL Fragrance 12-15-67645
7ACHE07/Chem_NY_Night Calming FCL Fragrance Symrise 734732
7ACHE08/Chem_NY_Detoxifying FCL Fragrance 12-15-68055
7ACHE09/Chem_NY_Baby Fresh Fragrance F-170564
7ACHE10/Chem_NY_Vivapur COS 8
7ACHE11/Chem_NY_Radiant 2 RTC-000069
7ACHE12/Chem_NY_Actiphyte Soybean GL
7ACHE13/Chem_NY_Plantasil Micro
7ACHE14/Chem_NY_Actiplex 5413 GL
7ACHE15/Chem_NY_Bell Botanical Blend_01
7ACHE20P/Chem_NY_Powerman
7ACHE22P/Chem_NY_Lactic Acid
7ACHE24P/Chem_NY_Coconut Milk Extract_01
7ACHE29P/Chem_NY_Orange Flower Extract_01
7ACHE30P/Chem_NY_Grape Fruit Extract_01
7ACHE35/Chem_NY_White Tea FCL Fragrance
7ACHE37/Chem_NY_Hexyl Cellosolve
7ACHE38/Chem_NY_Plurafac SL-62
7ACHE40/Chem_NY_Neossance Hemisqualane
7ACHE41/Chem_NY_Stabylen-30
7ACHE43/Chem_NY_Jarplex SB-12
7ACHE44/Chem_NY_Natural Oil Blend
7ACHE46/Chem_NY_Burt's Bees cucumber and sage Fragrance
7ACHE47/Chem_NY_Green Tea NP
7ACHE47/Chem_NY_Green Tea NP
7ACHE48/Chem_NY_Night Calming Fragrance
7ACHE49/Chem_NY_Tahitian Island Fragrance
7ACHE50/Chem_NY_Green Tea Melon Frag
7ACHE51/Chem_NY_Peaches & Honey Frag
7ACHE53/Chem_NY_Gluconic acid 50%
7ACHE54/Chem_NY_Cucumber & Green Tea Fragrance
7ACHE56/Chem_NY_STEOL CS-130
7ACHE57P/Chem_NY_Aromatic Chamomile Extract Blend
7ACHE60P/Chem_NY_Festive Cranberry MOD1
7ACHE61P/Chem_NY_Sodium Benzoate
7ACHE67/Chem_NY_Age Defying WS Fragrance
7ACHE68/Chem_NY_Evening Soothe WS Fragrance
7ACHE69/Chem_NY_Ultra Calm WS Fragrance
7ACHE70/Chem_NY_Radiant WS Fragrance
7ACHE71/Chem_NY_Fresh Spring Fragrance
7B4AA580/Chem_NY_Acetic Acid - USP
7B4AL930/Chem_42A
7B4AS318/Chem_43A
7B4FD110/Chem_44A
7B4GB598/Chem_45A
7B4OA160/Chem_46A
7B4OA234/Chem_47A
7B4OA242/Chem_48A
7B4OA347/Chem_NY_Calming Waters INCI w/symfresh
7B4OA709/Chem_NY_Bell fragrance
7B4OA863/Chem_NY_Lavender Extract
7B4OB521/Chem_49A
7B4OB690/Chem_50A
7B4OC109/Chem_51A
7B4OD232/Chem_52A
7B4OJ137/Chem_NY_Butylated Hydroxyanisole
7B4OM799/Chem_NY_Bell_Cucumber Extract in Glycerin
7B4OP027/Chem_NY_Abil B88183 PH
7B4OP054/Chem_NY_Symrelief 100_01
7B4OP065/Chem_NY_Chamomile Flower Extract
7B4OP093/Chem_NY_PEG-4
7B4OP113/Chem_NY_Mentha Piperita Extract_01
7B4OP118/Chem_NY_Polyethylene Glycol 1450
7B4OP203/Chem_NY_Diazolidinyl Urea – RI-1157
7B4OP205/Chem_NY_Fragrance Emma 8
7B4OP209/Chem_NY_Tinogard HS
7B4OP230/Chem_NY_Cosvat (Elestab CPN Ultra Pure)
7B4OP279/Chem_NY_Potassium Sorbate_7B4OP279
7B4OP284/Chem_NY_Sodium Citrate USP
7B4OP601/Chem_58A
7B4OP661/Chem_NY_Pemulen TR-2
7B4OP705/Chem_NY_Zinc Gluconate
7B4OP862/Chem_59A
7B4OQ028/Chem_NY_Sorbitan Monolaurate
7B4OQ084/Chem_60A
7B4OQ209/Chem_61A
7B4OQ214/Chem_62A
7B4OQ595/Chem_63A
7B4OQ794/Chem_NY_Disodium EDTA-USP
7B4OQ855/Chem_NY_Hydromanil H. GL
7B4OQ992/Chem_NY_Refined Sea Salt
7B4OS639/Chem_65A
7B4OU351/Chem_NY_Citric Acid Anhydrous - USP/NF_01
7B4OX110/Chem_66A
7B4OX131/Chem_NY_Dequest 2010
7B4OX232/Chem_40A
7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange
7C4AS318/Chem_68A
7C4FD110/Chem_69A
7C4OA160/Chem_70A
7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides (Lipex 102E-75 100%)
7C4OA863/Che_NY_Chamomile Extract
7C4OA940/Chem_NY_Propylene Glycol-RI-0570
7C4OB690/Chem_71A
7C4OB853/Chem_72A
7C4OD232/Chem_NY_Methylparaben NF/USP
7C4OD423/Chem_74A
7C4OF250/Chem_181A
7C4OG169/Chem_75A
7C4OG735/Chem_76A
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract
7C4OP054/Chem_NY_Betaine_01
7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30%
7C4OP081/Chem_NY_Extrapone Green Tea GW 660403
7C4OP203/Chem_NY_Leontopodium Alpinum Flower/Leaf Extract – RI-119235
7C4OP230/Chem_NY_Cetiol CC
7C4OP279/Chem_77A
7C4OP284/Chem_NY_Grape Fruit Extract
7C4OP338/Chem_78A
7C4OP601/Chem_79A
7C4OP661/Chem_NY_Euxyl K 701
7C4OP862/Chem_80A
7C4OQ028/Chem_NY_Isohexadecane
7C4OQ156/Chem_NY_Ascorbyl PM
7C4OQ595/Chem_81A
7C4OS639/Chem_82A
7C4OU218/Chem_83A
7C4OU351/Chem_NY_Trisodium Citrate Dihydrate USP_01
7C4OX110/Chem_84A
7C4OX131/Chem_NY_Fresh Apple 607 Fragrance
7D4AA580/Chem_NY_Monoethanolamine
7D4AB636/Chem_NY_Benzocaine USP
7D4FD110/Chem_NY_Purified Water USP
7D4OA863/Chem_NY_Lavender Silk Encapsulated Fragrance (myrcene free)
7D4OA940/Chem_NY_Sodium Benzoate RI-0660
7D4OB853/Chem_86A
7D4OD232/Chem_NY_Propylparaben USP
7D4OD376/Chem_NY_Vitamin A & D3
7D4OG169/Chem_88A
7D4OP065/Chem_NY_ Glydant Plus Liquid
7D4Op081/Chem_NY_Honeybush Tea Extract 173705 PF
7D4OP118/Chem_NY_Grapefruit Extract_7D4OP118
7D4OP203/Chem_NY_Dimethicone Copolyol RI-2322
7D4OP230/Chem_NY_Zemea_01
7D4OP338/Chem_89A
7D4OP485/Chem_90A
7D4OP601/Chem_91A
7D4OP862/Chem_02A
7D4OQ028/Chem_NY_Dicaprylyl Ether
7D4OQ209/Chem_92A
7D4OQ214/Chem_93A
7D4OQ967/Chem_94A
7D4OS639/Chem_95A
7D4OUXP1/Chem_NY_Glycopyrronium Tosylate Monohydrate_02
7E4AS318/Chem_96A
7E4OA863/Chem_NY_Lav Silk Stab Mod 2 (myrcene free)
7E4OA940/Chem_NY_Tocopheryl Acetate-SYN RI-1168
7E4OD232/Chem_97A
7E4OP027/Chem_NY_Benzalkonium Chloride 50% NF
7E4OP065/Chem_NY_Benzyl Alcohol
7E4OP118/Chem_99A
7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059
7E4OP230/Chem_NY_Lexfeel Natural
7E4OP279/Chem_100A
7E4OP338/Chem_101A
7E4OP601/Chem_102A
7E4OP806/Chem_58B
7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate
7E4OS639/Chem_103A
7F4BP869/Chem_NY_Citric Acid Anhydrous - USP
7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil
7F4OA940/Chem_NY_Decyl Polyglycoside RI-3274
7F4OD110/Chem_106A
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC
7F4OP065/Chem_NY_Mineral Oil 70
7F4OP203/Chem_NY_PEG-11 Methyl Ether Dimethicone - RI-24039
7F4OP230/Chem_NY_Diva Mod 4 RY-007326
7F4OP279/Chem_NY_Takasago Pomegranate 4865/3
7F4OP308/Chem_108A
7F4OP601/Chem_109A
7F4OP806/Chem_NY_Phenoxyethanol_7F4OP806
7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate
7G4OA160/Chem_NY_Floresta 443115K
7G4OA879/Chem_02A
7G4OA940/Chem_NY_Aloe Vera Gel Powder (200X Conc) RI-3764
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM
7G4OP203/Chem_NY_Solutions Destination Cleansing 5351254 – RI-120765
7G4OP308/Chem_111A
7G4OP485/Chem_NY_Bisabolol
7G4OP601/Chem_113A
7G4OP806/Chem_198A
7G4OQ028/Chem_NY_Aminomethylpropanol
7H4OA940/Chem_NY_Phenoxyethanol 98% Min RI-11921
7H4OP027/Chem_07A
7H4OP203/Chem_NY_Hexylene Glycol - RI-0425
7H4OP308/Chem_NY_Feverfew CMP-02A
7H4OP485/Chem_06B
7H4OP601/Chem_NY_DUB PTO
7H4OP806/Chem_NY_Kelp & Sea Salt Fragrance
7H4OQ028/Chem_NY_Fragrance IA 01-13-52229
7J4OA940/Chem_NY_Natural Kids Apple 485442D RI-125366
7J4OP027/Chem_182A
7J4OP203/Chem_NY_PEG-90M - RI-1183
7J4OP485/Chem_NY_Butylene Glycol_7J4OP485
7J4OP601/Chem_NY_Tegosoft CO
7K4OA940/Chem_NY_Sodium Hydroxide Solution 50% – RI-0659
7K4OP027/Chem_116A
7K4OP485/Chem_NY_Glycerin
7K4OP601/Chem_NY_Actiphyte of Aloe Vera 10
7L4OP027/Chem_119A
7L4OP203/Chem_NY_Glycerin - RI-0420
7L4OP485/Chem_NY_Polysorbate 20_7L4OP485
7L4OP601/Chem_121A
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP
7M4OP203/Chem_NY_Disodium EDTA - RI-3736
7M4OP485/Chem_NY_Emulgade CM_7M4OP485
7M4OP601/Chem_NY_Isononyl Isononanoste
7N4OP203/Chem_NY_Citric Acid Anhydrous - RI-0220
7N4OP485/Chem_NY_Citric Acid USP_7N4OP485
7O4AA580/Chem_NY_Ethylene Glycol Monohexyl Ether
7O4AJ131/Chem_NY_Acticide LG
7O4AP135/Chem_NY_Orange Fragrance
7O4AQ084/Chem_125A
7O4AQ467/Chem_NY_Fragrance CE-154785
7O4AQ500/Chem_04A
7O4AQ516/Chem_NY_Fragrance F-161194
7O4AQ763/Chem_NY_GB Hand Fragrance F-161684
7O4AQ904/Chem_NY_Lemon Fragrance_IA
7O4AQ918/Chem_NY_Green Apple Fresh Scent_IA
7O4BG500/Chem_NY_BTC 1010-50%
7O4BH999/Chem_NY_Procetyl AWS_7O4BH999
7O4BJ131/Chem_NY_Acticide CBM-2
7O4BM301/Chem_126A
7O4BM422/Chem_NY_Shea Butter Fragrance
7O4BQ164/Chem_127A
7O4BU125/Chem_09A
7O4CQ471/Chem_129A
7O4DH247/Chem_201A
7O4DM102/Chem_46A
7O4DX220/Chem_NY_Neutralizing Agent
7O4GA493/Chem_NY_Flowering Herbs Rev 10/08_01-08-38037
7O4OA012/Chem_NY_Feminine Hygiene Fragrance
7O4OA043/Chem_NY_Creamy Soy Milk Fragrance G9810 WS
7O4OA160/Chem_130A
7O4OA234/Chem_131A
7O4OA347/Chem_NY_Lauryl Glucoside
7O4OA356/Chem_NY_Fermented Vegetable Extract
7O4OA401/Chem_132A
7O4OA407/Chem_NY_Gentle Floral 79904B Fragrance
7O4OA490/Chem_133A
7O4OA517/Chem_196A
7O4OA526/Chem_NY_Island Splash Fragrance
7O4OA580/Chem_NY_Glycacil SG
7O4OA617/Chem_NY_Water Soluble Baby Fragrance
7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance
7O4OA709/Chem_NY_Cucumber Extract # 11-19-3
7O4OA862/Chem_NY_Delicate Aloe – Givaudan UAC03417/00
7O4OA863/Chem_NY_Lavender Silk Encapsulated Fragrance
7O4OA879/Chem_NY_Tomadol 25-7
7O4OA911/Chem_NY_Zinc Sulfate Heptahydrate
7O4OA940/Chem_NY_Simethicone Emulsion AF RI-0167
7O4OB103/Chem_NY_20% Chlorhexidine Gluconate (CHG) Concentrate
7O4OB401/Chem_NY_LEXFEEL D7_01
7O4OB521/Chem_135A
7O4OD302/Chem_136A
7O4OD351/Chem_NY_Benzalkonium Chloride 50%
7O4OD363/Chem_137A
7O4OD376/Chem_NY_Pramoxine HCL (USP Grade)
7O4OD419/Chem_138A
7O4OD748/Chem_139A
7O4OD791/Chem_CB_Fleet
7O4OF050/Chem_140A
7O4OF575/Che_NY_Plantaren 2000N - UP
7O4OF602/Chem_141A
7O4OG169/Chem_142A
7O4OG455/Chem_NY_Citrus Floor Fragrance
7O4OG510/Chem_NY_Wet Floor Fragrance.
7O4OH247/Chem_143A
7O4OH460/Chem_144A
7O4OH889/Chem_145A
7O4OH999/Chem_NY_Pink Grapefruit - CE-145431
7O4OJ123/Chem_NY_NET-KOB
7O4OJ131/Chem_NY_Fragrance-Geranium & Violet
7O4OJ137/Chem_NY_Phenylephrine Hydrochloride
7O4OM002/Chem_NY_Plantapon TF
7O4OM165/Chem_146A
7O4OM212/Chem_214A
7O4OM301/Chem_147A
7O4OM434/Chem_NY_Lavender Fragrance
7O4OM502/Chem_NY_Wet Floor Fragrance Lavender Vanilla 01-12-49043
7O4OM601/Chem_NY_PEG-50 Shea Butter
7O4OM657/Chem_NY_PC Cucumber Green Tea Type
7O4OM666/Chem_NY_Sodium Benzoate EDF NF/FCC
7O4OM799/Chem_148A
7O4OM857/Chem_NY_Refreshing Cucumber Water Soluble Fragrance
7O4OM945/Chem_NY_Lemon Fragrance
7O4OM962/Chem_NY_Emulgade NICE
7O4OP027/Chem_149A
7O4OP054/Chem_NY_Tego Cosmo C 250_01
7O4OP065/Chem_NY_Actipone White Tea
7O4OP073/Chem_NY_Minerals/Spring Water
7O4OP075/Chem_NY_Magique Lift 02LF
7O4OP081/Chem_NY_Triple Tea Extract
7O4OP087/Chem_NY_Actiplex 5327
7O4OP093/Chem_NY_PEG-11
7O4OP113/Chem_NY_Robertet The Matcha I
7O4OP118/Chem_150A
7O4OP135/Chem_NY_Orange Disinfecting Fragrance
7O4OP167/Chem_NY_Clearasil NBE International Aromatic 01-13-5144
7O4OP183/Chem_NY_Rose Flower Extract_01
7O4OP197/Chem_NY_Emulgade Smart/MB_01
7O4OP203/Chem_NY_Polysaccharide Glucose – RI-3349
7O4OP205/Chem_NY_Citrus Peach Extract
7O4OP207/Chem_NY_Zinc PCA_01
7O4OP209/Chem_NY_Ceraphyl 28
7O4OP211/Chem_153A
7O4OP228/Chem_NY_Ecosurf EH-9 Surfactant
7O4OP230/Chem_NY_Hydromanil
7O4OP241/Chem_NY_Citrus Scent – CE-144477
7O4OP253/Chem_NY_BTC 1010
7O4OP279/Chem_NY_Tocopheryl Acetate_7O4OP279
7O4OP284/Chem_155A
7O4OP338/Chem_156A
7O4OP345/Chem_157A
7O4OP352/Chem_158A
7O4OP421/Chem_131A
7O4OP448/Chem_197A
7O4OP527/Chem_159A
7O4OP567/Chem_NY_BTC 1210-80%
7O4OP601/Chem_107A
7O4OP630/Chem_202A
7O4OP638/Chem_192A
7O4OP661/Chem_NY_Xanthan Gum
7O4OP705/Chem_NY_Cetaphil Concentrate
7O4OP806/Chem_NY_Carbopol 2020
7O4OP862/Chem_NY_Ethanol-SDA 3C, 190 Proof-Corn Based
7O4OQ028/Chem_NY_Hexylene Glycol
7O4OQ030/Chem_NY_Detoxifying Fragrance
7O4OQ038/Chem_NY_Collasol M_01
7O4OQ084/Chem_162A
7O4OQ181/Chem_163A
7O4OQ214/Chem_NY_Deep Action Fragrance
7O4OQ229/Chem_NY_Witch Hazel
7O4OQ243/Chem_165A
7O4OQ268/Chem_NY_Euxyl K703
7O4OQ318/Chem_209A
7O4OQ342/Chem_166A
7O4OQ409/Chem_167A
7O4OQ413/Chem_NY_Juicy Fruits/Berry Fragrance
7O4OQ437/Chem_168A
7O4OQ467/Chem_NY_Night Calming Fragrance CE-122551
7O4OQ504/Chem_NY_Stuffy Nose Grape
7O4OQ516/Chem_NY_Hydrating Fragrance
7O4OQ763/Chem_169A
7O4OQ826/Chem_NY_Actiphyte of Fever Few GL
7O4OQ855/Chem_NY_Actiphyte of Willow Bark GL
7O4OQ890/Chem_170A
7O4OQ904/Chem_171A_IA
7O4OQ918/Chem_172A_IA
7O4OQ992/Chem_NY_Fragrance F-163401
7O4OS407/Chem_173A
7O4OS639/Chem_174A
7O4OU181/Chem_NY_Isopropanol 70%
7O4OU225/Chem_NY_Sodium Hypochlorite 12.5%
7O4OU351/Chem_NY_Glycopyrronium Tosylate Monohydrate_01
7O4OU555/Chem_NY_Lidocaine Base USP
7O4OX131/Chem_NY_Sodium Acetate Anhydrous
7O4OX231/Chem_NY_FD&C Red 40
7P4OP203/Chem_NY_Sodium Citrate - RI-0662
7P4OP485/Chem_176A
7P4OQ214/Chem_177A
7R4OP485/Chem_NY_Phenoxyethanol_7R4OP485
7S4OP485/Chem_NY_Sodium Benzoate NF_7S4OP485
7T4OP485/Chem_NY_Sodium Citrate Dihydrate USP
8AA203CS/Label_NY_NP7030_9-2.625
8AA405CS/Label_A12
8AA4432KC/Label_NY_NP7030_9-2.625
8AA564CK/Label_A12
8AA613KS/Label_NY_NP7030_9-2.625
8AA613SU/Label_NY_NP7030_9-2.625
8AA619SU/Label_NY_NP7030_9-2.625
8AA623RS/Label_NY_NP9850_3.5-3
8AA680KCNY/Label_A12
8AA680RF/Label_A70
8AA682CK/Label_NY_NP7030_9-2.625
8AA682KSNY/Label_NY_NP7030_9-2.625
8AA70900/Label_NY_NP9990_2.5-4.625
8AA711CK/Label_NY_NP7030_9-2.625
8AA739CK/Label_A12
8AA759CK/Label_A12
8AA766CPL/Label_A34
8AA8032BBL/Label_NY_NP9851_3.5-3.5
8AA941CK/Label_NY_NP7030_9-2.625
8AB010CK/Label_A23
8AB100CK/Label_A32
8AB196CK/Label_NY_NP7030_6-5
8AB293CK/Label_NY_NP7030_10.5-5.25
8AB432CK/Label_A64
8AB470CK/Label_NY_NP7030_6-4
8AB512CL/Label_NY_NP7030_6-4
8AC311SL/Label_A01
8AC585CK/Label_A32
8AD177CK/Label_A12
8AD304CK/Label_A12
8AD363CK/Label_A12
8AD427CK/Label_A12
8AD436KS/Label_A23
8AD454CK/Label_A12
8AD507CK/Label_A23
8AD774SH/Label_A23
8AH128CK/Label_NY_NP7030_9-2.625
8AH140CK/Label_NY_NP7030_9-2.625
8AH143KL/Label_A34
8AH143LC/Label_NY_NP7030_10-4.25
8AH211CK/Label_A12
8AH220CK/Label_NY_NP7030_9-2.625
8AH225CK/Label_A12
8AH250SH/Label_A23
8AJ165CK/Label_NY_NP7030_9-2.625
8AJ660CK/Label_A12
8AM071EZSU/Label_A12
8AM421PB/Label_A54
8AP01701/Label_NY_NP7100/NP7200_3.9659-4.3336
8AP02300/Label_A22
8AP02301/Label_A22
8AP02501/Label_NY_NP7100/NP7200_3.9659-4.3336
8AP02601/Label_NY_NP7100/NP7200_3.9659-4.3336
8AP02700/Label_NY_NP9815_01A
8AP02700RF/Label_NY_NP9815_73-85
8AP034RSL/Label_NY_NP9800_3.375-2.3437
8AP03501/Label_A35
8AP03600/Label_NY_NP7200_4.5-4.5
8AP03601/Label_NY_NP7200_4.5-4.5
8AP09700/Label_NY_NP9800_2.3437-3.375
8AP102SI/Label_A06
8AP10400/Label_NY_NP9800_2.3437-3.375
8AP11601/Label_NY_NP7100/NP7200_3.9659-4.3336
8AP12400/Label_NY_NP7200_5.5625-6.25
8AP12600/Label_NY_NP7200_5.8125-6.5
8AP12601/Label_NY_NP7200_5.8125-6.5
8AP12800/Label_A22
8AP12801/Label_NY_NP7100/NP7200_4.5-4
8AP128CK/Label_NY_NP7030_9-2.625
8AP13400/Label_A22
8AP13401/Label_A22
8AP134CS/Label_A23
8AP136CK/Label_A12
8AP13800/Label_NY_NP7200_6.25-4
8AP13801/Label_NY_NP7200_6-3.75
8AP14000/Label_NY_NP9800_2.375-3.5
8AP140RF/Label_NY_NP9800_2.375-3.5
8AP14700/Label_A12
8AP14800/Label_A41
8AP14801/Label_A41
8AP15300/Label_A40
8AP156CS/Label_A34
8AP15800/Label_NY_NP9800_2.375-3.5
8AP158RF/Label_NY_NP9800_3.5-2.375
8AP159CS/Label_A23
8AP16100/Label_NY_NP9800_2.375-3.5
8AP169CK/Label_NY_NP7030_9-2.625
8AP172CK/Label_NY_NP7030_9-2.625
8AP17301/Label_NY_NP7200_5.25-6.25
8AP176RSL/Label_NY_NP9800_2.375-3.5
8AP17800/Label_NY_NP9800_3.5-2.375
8AP181RSL/Label_NY_NP9800_2.375-3.5
8AP20300/Label_A17
8AP20300A/Label_NY_NP7700_3.375-2.3437
8AP21100/Label_NY_NP9890_85-73
8AP21100A/Label_NY_NP9890_85-73
8AP21700/Label_A20
8AP21900/Label_NY_NP7200_3.9659-4.3336
8AP21901/Label_NY_NP7200_3.9659-4.3336
8AP219CK/Label_NY_NP7030_9-2.625
8AP225CK/Label_NY_NP7030_9-2.625
8AP22800/Label_NY_NP7200_5.8125-6.5
8AP22801/Label_NY_NP7200_5.5625-6.25
8AP23800/Label_NY_NP7100/NP7200_4.5-4
8AP23801/Label_NY_NP7100/NP7200_4.5-4
8AP24200/Label_NY_NP7200_5.8125-6.5
8AP24300/Label_NY_NP7200_6.25-4.5
8AP24301/Label_NY_NP7200_6.25-4.5
8AP25300/Label_NY_NP7200_6.25-4.5
8AP25301/Label_NY_NP7200_6.25-4.5
8AP28000/Label_A20
8AP28400/Label_NY_NP9890_73-85
8AP28400A/Label_NY_NP9890_85-73
8AP32000/Label_A22
8AP32001/Label_A22
8AP320CK/Label_A34
8AP32400/Label_A22
8AP32401/Label_A22
8AP32800/Label_NY_NP98372_73-85
8AP33700/Label_NY_NP9815_73-85
8AP34201/Label_A35
8AP34500A/Label_NY_NP9800_3.375-2.3437
8AP35200/Label_A22
8AP35201/Label_A22
8AP37800/Label_A18
8AP38301/Label_NY_NP7200_4.25-5.5
8AP383CK/Label_NY_NP7030_9-2.625
8AP43500/Label_A22
8AP43501/Label_A22
8AP44300/Label_A40
8AP44500/Label_A41
8AP44501/Label_A41
8AP45101/Label_A35
8AP45700/Label_NY_NP7200_3.9659-4.3336
8AP45701/Label_NY_NP7200_3.9659-4.3336
8AP457CK/Label_NY_NP7030_9-2.625
8AP460RL/Label_NY_NP9800_3.3465-2.874
8AP46801/Label_A35
8AP47501/Label_NY_NP7200_4.25-5.5
8AP487CK/Label_NY_NP7030_9-2.625
8AP490UP/Label_A12
8AP49500/Label_A20
8AP49501/Label_A20
8AP52000/Label_NY_NP7200_5.5625-6.25
8AP52001/Label_NY_NP7200_5.5625-6.25
8AP520CK/Label_NY_NP7030_9-7
8AP53500/Label_A35
8AP53501/Label_NY_NP7200_4.25-5.5
8AP563RL/Label_NY_NP9800_3.3465-2.874
8AP56701/Label_NY_NP7200_5.8125-6.5
8AP57000/Label_NY_NP7100/NP7200_4.5-4
8AP57001/Label_A22
8AP570CK/Label_A12
8AP59000/Label_A35
8AP59001/Label_A35
8AP590CK/Label_A34
8AP60100/Label_NY_NP9815_01A
8AP60100RF/Label_NY_NP9890_73-85
8AP61500/Label_A20
8AP63500/Label_A35
8AP63501/Label_NY_NP7200_4.25-5.5
8AP63800/Label_NY_NP7200_5.8125-6.5
8AP63801/Label_NY_NP7200_5.5625-6.25
8AP656F0/Label_A33
8AP66701/Label_NY_NP7200_5.8125-6.5
8AP69100/Label_A28
8AP713RSL/Label_NY_NP9800_3.375-2.3437
8AP71500/Label_A63
8AP72100/Label_NY_NP9890_85-73
8AP73400/Label_A28
8AP73600/Label_A41
8AP765CK/Label_A12
8AP794CK/Label_A34
8AP87000/Label_NY_NP9800_2.375-3.5
8AP909BL/Label_B03
8AP92000/Label_A41
8AP92001/Label_A41
8AP97500/Label_A41
8AP97501/Label_A41
8AQ10501/Label_A20
8AQ387RL/Label_NY_NP9800_2.625-5.375
8AQ508TPA/Label_A18
8AQ527TPA/Label_A18
8AQ594CK/Label_A12
8AQ594CS/Label_A12
8AQ76500/Label_A39
8AQ76501/Label_A39
8AQ765CK/Label_A24
8AQ769CK/Label_A36
8AQ8032RSL/Label_NY_NP9802_2.874-3.3465
8AQ8045BBL/Label_NY_NP9850_3-3
8AQ8045FBL/Label_NY_NP9850_3-3
8AQ89901/Label_A38
8AQ95000/Label_A37
8AQ95001/Label_A41
8AS043AL/Label_A08
8AS043CK/Label_A31
8AS420CL/Label_NY_NP7030_6-5
8AS423CL/Label_NY_NP7030_6-5
8AS724AL/Label_A08
8AS724CK/Label_A31
8AU002CK/Label_A45
8AU120CK/Label_NY_NP7030_6-4
8AU214UP/Label_NY_NP7030_9-2.625
8AU385CK/Label_NY_NP7030_9-2.625
8AU451CS/Label_A12
8BA366CS/Label_A12
8BA366CSNY/Label_A12
8BA366KCNY/Label_A12
8BA436P0/Label_A05
8BA436P1/Label_A05
8BA437CK/Label_A12
8BA630CK/Label_A12
8BA680KC/Label_A12
8BB010CK/Label_NY_NP7030_6-4
8BB10800/Label_A31
8BB164CK/Label_NY_NP7030_6-4
8BB196CK/Label_NY_NP7030_6-5
8BB432CK/Label_NY_NP7030_9-2.625
8BB470CK/Label_NY_NP7030_6-4
8BB512CL/Label_NY_NP7030_6-4
8BB598CK/Label_NY_NP7030_6-4
8BD436CS/Label_A23
8BD436KS/Label_NY_NP7030_6-4
8BD454CK/Label_NY_NP7030_9-2.625
8BG52400/Label_A21
8BH115CK/Label_A34
8BH115KS/Label_A34
8BH128CK/Label_NY_NP7030_9-2.625
8BH211CK/Label_A12
8BP02301/Label_A22
8BP115F1/Label_A37
8BP115F1/Label_A37
8BP115L1/Label_A37
8BP12000/Label_A22
8BP12800/Label_NY_NP7100/NP7200_4.5-4
8BP12801/Label_NY_NP7200_4.5-4
8BP128CK/Label_NY_NP7030_9-2.625
8BP13400/Label_A22
8BP13401/Label_A22
8BP134CS/Label_A23
8BP13500/Label_A37
8BP13800/Label_NY_NP7200_6.25-4.5
8BP13801/Label_NY_NP7200_6.25-4
8BP14800/Label_A41
8BP14801/Label_NY_NP7200_5.5625-6.25
8BP15600/Label_A15
8BP156CS/Label_A34
8BP156RF/Label_NY_NP9510_3.3125-1.5625
8BP15900/Label_A20
8BP159CS/Label_A23
8BP214L0/Label_A37
8BP21700/Label_A20
8BP21701/Label_A20
8BP22801/Label_NY_NP7200_5.8125-6.5
8BP24500/Label_NY_NP7200_4-4.5
8BP24501/Label_NY_NP7200_4-4.5
8BP25300/Label_NY_NP7200_6.25-4.5
8BP28400/Label_NY_NP98372_73-85
8BP34200/Label_A35
8BP43500/Label_A22
8BP43501/Label_A22
8BP44301/Label_NY_NP7200_5.25-6.25
8BP44500/Label_NY_NP7200_5.8125-6.5
8BP44501/Label_NY_NP7200_5.5625-6.25
8BP45100/Label_A35
8BP45700/Label_NY_NP7200_3.9659-4.3336
8BP45701/Label_NY_NP7200_3.9659-4.3336
8BP457CK/Label_NY_NP7030_9-2.625
8BP54001/Label_NY_NP7100/NP7200_6-3.75
8BP570CK/Label_NY_NP7030_9-2.625
8BP57500/Label_A40
8BP57501/Label_A40
8BP63801/Label_NY_NP7200_5.5625-6.25
8BP92001/Label_A41
8BP97500/Label_A41
8BP97501/Label_A41
8BQ105CK/Label_A46
8BQ43800/Label_A19
8BQ506CSA/Label_A12
8BQ594CK/Label_NY_NP7030_9-2.625
8BQ594CS/Label_NY_NP7030_9-2.625
8BQ74900/Label_A41
8BQ750BL/Label_B02
8BQ76500/Label_NY_NP7200_6-3.75
8BQ76501/Label_NY_NP7200_3.75-6
8BQ85400/Label_A41
8BQ85401/Label_A41
8BQ85901/Label_A41
8BQ87000/Label_A41
8BQ87001/Label_A41
8BQ89900/Label_A38
8BQ899CK/Label_A48
8BQ901CK/Label_A48
8BQ91401/Label_A20
8BQ914SL/Label_A48
8BS40750/Label_A31
8BS41950/Label_A31
8BS420CL/Label_NY_NP7030_6-5
8BS423CL/Label_NY_NP7030_6-5
8BU451CS/Label_A12
8CB007PS/Label_A23
8CB055CK/Label_A23
8CB10800/Label_A31
8CB11400/Label_A31
8CB432CK/Label_NY_NP7030_9-2.625
8CB598CK/Label_NY_NP7030_6-4
8CD438UP/Label_A30
8CH128CK/Label_NY_NP7030_9-2.625
8CH211CK/Label_A12
8CP02101/Label_A37
8CP115F1/Label_A41
8CP115L1/Label_A41
8CP12001/Label_A22
8CP12800/Label_NY_NP7200_4.5-4
8CP13400/Label_NY_NP7100_4.5-4
8CP13401/Label_NY_NP7100/NP7200_4.5-4
8CP134CS/Label_NY_NP7030_6-4
8CP13801/Label_NY_NP7200_6.25-4.5
8CP15900/Label_NY_NP7100/NP7200_6-3.75
8CP15901/Label_A20
8CP159CS/Label_NY_NP7030_6-4
8CP214F0/Label_A37
8CP43501/Label_A22
8CP44500/Label_NY_NP7200_5.8125-6.5
8CP44501/Label_NY_NP7200_5.8125-6.5
8CP57500/Label_A40
8CP57501/Label_NY_NP7200_5.25-6.25
8CP63801/Label_NY_NP7200_5.8125-6.5
8CP68300/Label_A20
8CP68301/Label_A20
8CP683CK/Label_A48
8CP97501/Label_A41
8CQ34500/Label_A11
8CQ43801/Label_A13
8CQ70000/Label_A11
8CQ74901/Label_A41
8CQ749CK/Label_A32
8CQ84500/Label_A38
8CQ84501/Label_A38
8CQ85900/Label_A41
8CQ85901/Label_NY_NP7100/NP7200_5.5625-6.25
8CQ859CK/Label_A45
8CQ87000/Label_A41
8CQ87001/Label_A41
8CQ88200/Label_A20
8CQ88201/Label_A20
8CQ89900/Label_NY_NP7200_5.25-5.625
8CQ899CK/Label_NY_NP7030_9-7
8CQ90100/Label_A20
8CQ901CK/Label_NY_NP7030_9-7
8CQ91400/Label_A20
8CQ91401/Label_NY_NP7200_6-3.75
8CQ914SL/Label_NY_NP7030_9-7
8CQ94300/Label_A41
8CQ94301/Label_A41
8CS32450/Label_A31
8CS40750/Label_A31
8CS41950/Label_A31
8CS428CK/Label_A31
8DD438UP/Label_A29
8DH115CK/Label_A29
8DP02101/Label_A41
8DP13401/Label_NY_NP7200_4.5-4
8DP134CS/Label_NY_NP7030_6-4
8DP15901/Label_NY_NP7100/NP7200_6-3.75
8DP159CS/Label_NY_NP7030_6-4
8DP43500/Label_NY_NP7200_4-4.5
8DP44501/Label_NY_NP7200_5.8125-6.5
8DP54000/Label_NY_NP7100/NP7200_6-3.75
8DP68300/Label_NY_NP7200_6-3.75
8DP68301/Label_NY_NP7200_6-3.75
8DP683CK/Label_NY_NP7030_9-7
8DP92000/Label_NY_NP7200_6.5-5.8125
8DQ10500/Label_A20
8DQ34500/Label_NY_NP9800_2.625-5.375
8DQ70000/Label_NY_NP9800_2.625-5.375
8DQ84500/Label_A38
8DQ84501/Label_A38
8DQ845SK/Label_A45
8DQ85000/Label_A41
8DQ85400/Label_A41
8DQ85401/Label_A41
8DQ85900/Label_NY_NP7200_5.5625-6.25
8DQ85901/Label_NY_NP7200_5.5625-6.25
8DQ859CK/Label_NY_NP7030_9-7
8DQ87000/Label_NY_NP7200_5.5625-6.25
8DQ87001/Label_NY_NP7200_5.5625-6.25
8DQ88200/Label_A20
8DQ88201/Label_A20
8DQ90100/Label_NY_NP7200_6-3.75
8DQ90101/Label_A20
8DQ901CK/Label_NY_NP7030_9-7
8DQ91400/Label_NY_NP7200_6-3.75
8DQ91401/Label_NY_NP7200_6.25-4.5
8DQ914SL/Label_NY_NP7030_9-7
8DQ94300/Label_A41
8DQ94301/Label_NY_NP7200_5.5625-6.25
8DS27350/Label_A31
8EB10800/Label_A31
8EB11400/Label_A31
8EP15901/Label_NY_NP7200_6.25-4
8EP43501/Label_NY_NP7200_4-4.5
8EP44501/Label_NY_NP7200_5.8125-6.5
8EP68300/Label_NY_NP7200_6.25-4.5
8EP68301/Label_NY_NP7200_6.25-4.5
8EP92001/Label_NY_NP7200_6.5-5.8125
8EQ08400/Label_A20
8EQ08401/Label_A20
8EQ50701/Label_A20
8EQ84500/Label_NY_NP7200_5.25-5.625
8EQ84501/Label_NY_NP7200_5.25-5.625
8EQ845SK/Label_NY_NP7030_9-7
8EQ85001/Label_A41
8EQ859CK/Label_NY_NP7030_9-7
8EQ86900/Label_A41
8EQ86901/Label_A41
8EQ87000/Label_NY_NP7200_5.8125-6.5
8EQ87001/Label_NY_NP7200_5.8125-6.5
8EQ88200/Label_NY_NP7200_6-3.75
8EQ88201/Label_NY_NP7200_6-3.75
8EQ90100/Label_NY_NP7200_6.25-4.5
8EQ90101/Label_NY_NP7100/NP7200_6-3.75
8EQ901CK/Label_NY_NP7030_9-7
8EQ91400/Label_NY_NP7200_6-3.75
8ES32450/Label_A31
8ES40750/Label_A31
8ES41950/Label_A31
8ES428CK/Label_A31
8FB11400/Label_A31
8FQ08401/Label_NY_NP7200_6.25-4.5
8FQ506TPA/Label_A68
8FQ50700/Label_A20
8FQ50701/Label_NY_NP7200_6-3.75
8FQ55101/Label_A20
8FQ85000/Label_NY_NP7200_5.5625-6.25
8FQ85001/Label_A41
8FQ859CK/Label_NY_NP7030_9-7
8FQ86900/Label_A41
8FQ86901/Label_A41
8FQ88200/Label_NY_NP7200_6.25-4
8FQ88201/Label_NY_NP7200_6.25-4
8FQ90101/Label_NY_NP7200_6-3.75
8FQ91400/Label_NY_NP7200_6.25-4.5
8FS27350/Label_A31
8FS40750/Label_A31
8FS428CK/Label_A31
8GB10800/Label_NY_NP7030_6-5
8GQ50700/Label_NY_NP7200_6-3.75
8GQ50701/Label_NY_NP7200_6.25-4
8GQ55100/Label_A20
8GQ55101/Label_A20
8GQ772SK/Label_B03
8GQ85001/Label_NY_NP7200_5.5625-6.25
8GQ86900/Label_NY_NP7200_5.5625-6.25
8GQ86901/Label_NY_NP7100/NP7200_5.5625-6.25
8GQ88200/Label_NY_NP7200_6.25-4.5
8GQ88201/Label_NY_NP7200_6.25-4.5
8GQ89000/Label_A20
8GQ89001/Label_A20
8GQ90101/Label_NY_NP7200_6.25-4.5
8GS27350/Label_A31
8GS40750/Label_NY_NP7030_6-5
8GS41950/Label_NY_NP7030_6-5
8HB10800/Label_NY_NP7030_6-5
8HB11400/Label_A31
8HQ50700/Label_NY_NP7200_6.25-4
8HQ50701/Label_NY_NP7200_6.25-4.5
8HQ55100/Label_A20
8HQ55101/Label_A20
8HQ59400/Label_A15
8HQ85001/Label_NY_NP7200_5.5625-6.25
8HQ86900/Label_NY_NP7200_6.5-5.8125
8HQ86901/Label_NY_NP7200_6.5-5.8125
8HQ89000/Label_NY_NP7200_6-3.75
8HQ89001/Label_A20
8HS32450/Label_A31
8HS40750/Label_NY_NP7030_6-5
8HS41950/Label_NY_NP7030_6-5
8HS428CK/Label_A31
8IQ50700/Label_NY_NP7200_6.25-4.5
8JB10800/Label_NY_NP7030_6-5
8JB11400/Label_A31
8JQ55100/Label_A20
8JQ55101/Label_NY_NP7100/NP7200_6-3.75
8JQ86901/Label_NY_NP7200_6.5-5.8125
8JQ89000/Label_NY_NP7200_6.25-4.5
8JQ89001/Label_NY_NP7200_6-3.75
8JS27350/Label_A31
8JS40750/Label_NY_NP7030_6-5
8JS41950/Label_NY_NP7030_6-5
8JS428CK/Label_A31
8KB10800/Label_NY_NP7030_6-5
8KB11400/Label_NY_NP7030_6-5
8KQ55100/Label_NY_NP7200_6-3.75
8KQ55101/Label_NY_NP7200_6.25-4
8KQ59400/Label_A15
8KQ594RF/Label_NY_NP9510_3.3125-1.5625
8KQ89001/Label_NY_NP7200_6.25-4.5
8KS32450/Label_NY_NP7030_6-5
8KS40750/Label_NY_NP7030_6-5
8KS41950/Label_NY_NP7030_6-5
8KS428CK/Label_NY_NP7030_6-5
8LB10800/Label_NY_NP7030_6-5
8LB11400/Label_NY_NP7030_6-5
8LQ49600/Label_A17
8LQ55100/Label_NY_NP7200_6.25-4
8LQ55101/Label_NY_NP7200_6.25-4
8LS27350/Label_NY_NP7030_6-5
8LS32450/Label_NY_NP7030_6-5
8LS40750/Label_NY_NP7030_6-5
8LS41950/Label_NY_NP7030_6-5
8LS428CK/Label_NY_NP7030_6-5
8MB10800/Label_NY_NP7030_6-5
8MB11400/Label_NY_NP7030_6-5
8MQ55100/Label_NY_NP7200_6.25-4.5
8MQ55101/Label_NY_NP7200_6.25-4.5
8MS32450/Label_NY_NP7030_6-5
8MS428CK/Label_NY_NP7030_6-5
8NB10800/Label_NY_NP7030_6-5
8NB11400/Label_NY_NP7030_6-5
8NS27350/Label_NY_NP7030_6-5
8NS32450/Label_NY_NP7030_6-5
8NS40750/Label_NY_NP7030_6-5
8NS41950/Label_NY_NP7030_6-5
8NS428CK/Label_NY_NP7030_6-5
8OA119CPL/Label_NY_NP7030_10.5-5.25
8OA119KS/Label_A12
8OA127UC/Label_NY_NP7030_9-2.625
8OA226KS/Label_A12
8OA302CS/Label_A12
8OA366P0/Label_A05
8OA366P1/Label_A05
8OA366RF/Label_A27
8OA366TL/Label_A14
8OA366TLNY/Label_A14
8OA407SI/Label_A10
8OA466KSNY/Label_A12
8OA520CK/Label_A12
8OA527P0/Label_A05
8OA527P1/Label_A05
8OA54130AZ/Label_A12
8OA544CK/Label_A12
8OA544PI/Label_A52
8OA544RW/Label_A12
8OA544UC/Label_A61
8OA564CK/Label_A12
8OA565CK/Label_NY_NP7030_9-2.625
8OA565CLF/Label_NY_NP7700_2.625-5.375
8OA582CK/Label_NY_NP7030_9-2.625
8OA619CLF/Label_NY_NP7700_2.877-2.489
8OA621CK/Label_A12
8OA625CK/Label_A12
8OA641CPL/Label_NY_NP7030_10.5-5.25
8OA657CK/Label_A12
8OA660CS/Label_NY_NP7030_9-2.625
8OA664CK/Label_NY_NP7030_9-2.625
8OA670CK/Label_A12
8OA678CK/Label_NY_NP7030_9-2.625
8OA680RF/Label_A07
8OA682CK/Label_A12
8OA682KSNY/Label_A12
8OA70900/Label_NY_NP9990_2.5-4.625
8OA71100/Label_A27
8OA711CK/Label_A12
8OA712CK/Label_NY_NP7030_9-2.625
8OA721CK/Label_A12
8OA775CS/Label_NY_NP7030_9-2.625
8OA796CLF/Label_NY_NP7700_1.9375-4.375
8OA8026BBL/Label_NY_NP9850_3.5-1.625
8OA8026CS/Label_NY_NP7030_9-2.625
8OA8026F5L/Label_NY_NP9850_3.5-5
8OA8026FBL/Label_NY_NP9850_3.5-1.625
8OA8026KS/Label_NY_NP7030_9-2.625
8OA8026KS/Label_NY_NP7030_9-2.625
8OA80800/Label_A50
8OA8121CS/Label_NY_NP7030_9-2.625
8OA8125CS/Label_NY_NP7030_9-2.625
8OA816CK/Label_A12
8OA849CK/Label_A12
8OA9022CS/Label_NY_NP7030_9-2.625
8OA940RL/Label_NY_NP9800_2.25-4.25
8OA941CK/Label_NY_NP7030_9-2.625
8OA962P1/Label_NY_NP7200_4.5-1.5
8OA97200/Label_NY_NP9990_2.5-4.625
8OA999PL/Label_A34
8OB010CK/Label_A23
8OB10800/Label_A31
8OB123CK/Label_NY_NP7030_6-5
8OB164CK/Label_A23
8OB196CK/Label_A12
8OB270CL/Label_A23
8OB293CK/Label_A34
8OB320CK/Label_A25
8OB339CK/Label_A23
8OB432CL/Label_A12
8OB470CK/Label_A23
8OB512CL/Label_A23
8OB598CK/Label_A23
8OB690CK/Label_A23
8OB714CKB/Label_NY_NP7030_9-2.625
8OB803CK/Label_A23
8OB803CK/Label_A23
8OB803CL/Label_A23
8OB803CL/Label_NY_NP7030_6-4
8OB803CS/Label_A23
8OB810CK/Label_A25
8OC421CK/Label_A12
8OC610CK/Label_NY_NP7030_9-2.625
8OD015CK/Label_NY_NP7030_9-2.625
8OD015PQ/Label_NY_NP7030_9-2.625
8OD02400/Label_NY_NP7030_9-2.625
8OD024CK/Label_NY_NP7030_9-2.625
8OD024PQ/Label_NY_NP7030_9-2.625
8OD177CK/Label_A12
8OD185OL/Label_A69
8OD198CK/Label_A23
8OD240CK/Label_A45
8OD242CK/Label_A31
8OD243CK/Label_NY_NP7030_9-2.625
8OD245CK/Label_NY_NP7030_9-2.625
8OD310CK/Label_A23
8OD321CK/Label_A12
8OD346CK/Label_A23
8OD384CK/Label_A12
8OD417CK/Label_NY_NP7030_9-2.625
8OD419CK/Label_A23
8OD422UC/Label_A12
8OD4360B/Label_NY_NP7030_6-4
8OD436CS/Label_A23
8OD436KS/Label_A23
8OD453CK/Label_NY_NP7030_9-2.625
8OD45480CK/Label_NY_NP7030_9-2.625
8OD454CK/Label_A12
8OD454CL/Label_A12
8OD454CS/Label_A12
8OD470CK/Label_NY_NP7030_9-2.625
8OD53536/Label_A12
8OD535CK/Label_A12
8OD535CL/Label_A12
8OD535CS/Label_A12
8OD576CK/Label_A34
8OD60624/Label_A12
8OD606CK/Label_A12
8OD606CS/Label_A12
8OD621CK/Label_A12
8OD621CS/Label_A12
8OD712CK/Label_NY_NP7030_9-2.625
8OD712CS/Label_NY_NP7030_9-2.625
8OD71880CK/Label_NY_NP7030_9-2.625
8OD718CK/Label_A12
8OD718CKA/Label_NY_NP7030_9-2.625
8OD718CL/Label_A12
8OD718CLA/Label_NY_NP7030_9-2.625
8OD718CS/Label_A12
8OD718CSA/Label_NY_NP7030_9-2.625
8OD766CK/Label_A12
8OG11200/Label_NY_NP7200_3.80-4.82
8OG11201/Label_NY_NP7200_3.80-4.82
8OG112CK/Label_NY_NP7030_7.750-5.250
8OG126CS/Label_NY_NP7030_9-2.625
8OG127CS/Label_NY_NP7030_9-2.625
8OG129CS/Label_NY_NP7030_9-2.625
8OG133CS/Label_NY_NP7030_9-2.625
8OG135IR/Label_NY_NP7950_2.125-2.0625
8OG201KS/Label_A12
8OG202CS/Label_A12
8OG301CS/Label_NY_NP7030_9-2.625
8OG302CS/Label_NY_NP7030_9-2.625
8OG30600/Label_A49
8OG30601/Label_A49
8OG306KS/Label_A12
8OG350CP/Label_A14
8OG451KS/Label_A12
8OG605KS/Label_A12
8OH115KS/Label_A34
8OH125CK/Label_NY_NP7030_9-2.625
8OH126CK/Label_NY_NP7030_9-2.625
8OH128CK/Label_A12
8OH128CL/Label_A26
8OH140CK/Label_NY_NP7030_9-2.625
8OH14348/Label_A12
8OH143LC/Label_A29
8OH143WM/Label_A12
8OH146CK/Label_NY_NP7030_9-2.625
8OH158CK/Label_A12
8OH211CK/Label_A12
8OH211GB/Label_A12
8OH217UP/Label_A25
8OH220CK/Label_A12
8OH225CK/Label_A12
8OH241UP/Label_A34
8OH260CK/Label_A12
8OH269CK/Label_A12
8OH301CK/Label_A12
8OH450CK/Label_NY_NP7030_9-2.625
8OH565CK/Label_A12
8OH617CK/Label_A12
8OH757CK/Label_A23
8OH889CK/Label_A23
8OJ112PL/Label_NY_NP7030_10.5-5.25
8OJ131CK/Label_NY_NP7030_9-2.625
8OJ132CK/Label_NY_NP7030_9-2.625
8OJ14500/Label_A42
8OJ145TL/Label_A43
8OJ165CK/Label_NY_NP7030_9-2.625
8OJ171CK/Label_NY_NP7030_9-2.625
8OJ22750/Label_A11
8OJ291CK/Label_A12
8OJ315CK/Label_A12
8OJ660CK/Label_A12
8OJ701CL/Label_NY_NP7030_9-2.625
8OK457CS/Label_A31
8OM005PL/Label_NY_11-8.5
8OM071BND/Label_C01
8OM120BL/Label_A58
8OM120EZRF/Label_A11
8OM248EZCS/Label_NY_NP7030_9-2.625
8OM420CPL/Label_A34
8OM425CPL/Label_NY_NP7030_10.5-5.25
8OM439LUL/Label_NY_NP9920_2.5625-1.1875
8OM492CPL/Label_NY_NP7030_10.5-5.25
8OM525CS/Label_NY_NP7030_9-2.625
8OM9007KS/Label_NY_NP7030_9-2.625
8OM929CPL/Label_A34
8OM930PB/Label_A59
8OM956BS/Label_A51
8OM962KA/Label_NY_NP7030_9-2.625
8OM963KA/Label_NY_NP7030_9-2.625
8OMC192L0/Label_A55
8OMC192L1/Label_A55
8OMC192SU/Label_A12
8OMC690PL/Label_A34
8OMC799CS/Label_A12
8OMC850UP/Label_A60
8OMC960UP/Label_A60
8OP0111SL/Label_A34
8OP0112CK/Label_A12
8OP012300/Label_A12
8OP012400/Label_A12
8OP01700/Label_NY_NP7100/NP7200_3.9659-4.3336
8OP01701/Label_NY_NP7100/NP7200_3.9659-4.3336
8OP017CK/Label_NY_NP7030_9-2.625
8OP017PQ/Label_NY_NP7030_9-2.625
8OP02100/Label_A37
8OP02300/Label_A22
8OP02501/Label_NY_NP7100/NP7200_3.9659-4.3336
8OP025CK/Label_NY_NP7030_9-2.625
8OP02601/Label_NY_NP7100/NP7200_3.9659-4.3336
8OP026CK/Label_NY_NP7030_9-2.625
8OP02700/Label_NP9815_
8OP028RSL/Label_NY_NP9800_3.375-2.3437
8OP02900/Label_NY_NP7100/NP7200_5.25-6.25
8OP02901/Label_NY_NP7100/NP7200_5.25-6.25
8OP029CK/Label_NY_NP7030_9-2.625
8OP034RSL/Label_NY_NP9800_3.375-2.3437
8OP034RSLB/Label_NY_NP9800_2.3437-3.375
8OP03500/Label_A28
8OP03501/Label_A28
8OP035CK/Label_A12
8OP03600/Label_NY_NP7200_4-4.5
8OP03601/Label_NY_NP7200_4-4.5
8OP03700/Label_NY_NP9840_87-75
8OP04500/Label_NP9350_
8OP06300/Label_NY_NP9840_87-75
8OP065RSL/Label_NY_NP9995_2.625-3.375
8OP067RSL/Label_NY_NP9995_2.625-3.375
8OP07200/Label_NY_NP9840_87-75
8OP073RSL/Label_NY_NP9995_2.625-3.375
8OP081RSL/Label_NY_NP9995_2.625-3.375
8OP081TS/Label_NY_NP7117_1.500-1.000
8OP08700/Label_NY_NP9800_3.375-2.625
8OP09100/Label_NY_NP9991_3.375-2.625
8OP09300/Label_NY_NP7110_2.55-3
8OP09301/Label_NY_NP7110_2.55-5
8OP09700/Label_A17
8OP10400/Label_A17
8OP109RSL/Label_NY_NP9800_3.375-2.3437
8OP11300/Label_NY_NP9995_2.625-3.375
8OP115F0/Label_A37
8OP115L0/Label_A37
8OP11600/Label_NY_NP7100/NP7200_3.9659-4.3336
8OP11601/Label_NY_NP7100/NP7200_3.9659-4.3336
8OP116CK/Label_NY_NP7030_9-2.625
8OP116PQ/Label_NY_NP7030_9-2.625
8OP123F1/Label_A53
8OP123L1/Label_A53
8OP12400/Label_A37
8OP12401/Label_A37
8OP12401/Label_NY_NP7200_5.5625-6.25
8OP12600/Label_A37
8OP12601/Label_A37
8OP126CK/Label_A47
8OP12801/Label_A22
8OP128CK/Label_A12
8OP13001/Label_A37
8OP136RL/Label_A15
8OP13800/Label_NY_NP7100/NP7200_6-3.75
8OP13801/Label_A20
8OP14000/Label_A18
8OP140CK/Label_A12
8OP140RF/Label_NY_NP9800_2.375-3.5
8OP144CK/Label_NY_NP7030_9-2.625
8OP146F0/Label_A37
8OP146F0/Label_A37
8OP146F0/Label_NY_NP7200_5.5625-6.25
8OP146F1/Label_A37
8OP146F1/Label_A37
8OP146F1/Label_NY_NP7200_5.5625-6.25
8OP146L0/Label_NY_NP7200_5.5625-6.25
8OP146L1/Label_NY_NP7200_5.5625-6.25
8OP149CK/Label_A12
8OP150CK/Label_A12
8OP15300/Label_A40
8OP15301/Label_A40
8OP153CK/Label_A12
8OP15400/Label_A20
8OP15401/Label_A20
8OP155CK/Label_NY_NP7030_9-2.625
8OP155SI/Label_NY_NP9800_2-1.375
8OP158CK/Label_NY_NP7030_9-2.625
8OP15900/Label_A20
8OP16000/Label_NY_NP9800_2.375-3.5
8OP160CK/Label_NY_NP7030_9-2.625
8OP160RF/Label_NY_NP9800_3.5-2.375
8OP161CK/Label_NY_NP7030_9-2.625
8OP162CK/Label_NY_NP7030_9-2.625
8OP16300/Label_A67
8OP16301/Label_A67
8OP163CK/Label_A12
8OP16400/Label_A16
8OP164CK/Label_A12
8OP169CK/Label_NY_NP7030_9-2.625
8OP172CK/Label_NY_NP7030_9-2.625
8OP17300/Label_A40
8OP17301/Label_A40
8OP17500/Label_NP9350_
8OP176CK/Label_NY_NP7030_9-2.625
8OP176RSL/Label_NY_NP9800_2.375-3.5
8OP17700/Label_NY_NP9800_3.5-2.375
8OP177CK/Label_NY_NP7030_9-2.625
8OP178CK/Label_NY_NP7030_9-2.625
8OP181RSL/Label_NY_NP9800_2.375-3.5
8OP18800/Label_A40
8OP18801/Label_A40
8OP188BL/Label_B01
8OP19000/Label_NY_NP9800_2.375-3.5
8OP190CK/Label_NY_NP7030_9-2.625
8OP19500/Label_NY_NP9840_87-75
8OP19700/Label_NY_NP9840_87-75
8OP20300/Label_A17
8OP205RSL/Label_NY_NP9995_2.625-3.375
8OP21100/Label_NY_NP9815_01A
8OP21100A/Label_NY_NP9890_85-73
8OP214RSL/Label_NY_NP9991_3.375-2.625
8OP21701/Label_A20
8OP21900/Label_NY_NP7200_3.9659-4.3336
8OP21901/Label_NY_NP7200_3.9659-4.3336
8OP219CK/Label_NY_NP7030_9-2.625
8OP223RSL/Label_NY_NP9840_87-75
8OP225CK/Label_NY_NP7030_9-2.625
8OP227RSL/Label_NY_NP9997_2.625-3.375
8OP22800/Label_A41
8OP22801/Label_A41
8OP22900/Label_NY_NP9837_78-61
8OP22900RF/Label_NY_NP98371_78-61
8OP23000/Label_NY_NP9837_73-85
8OP23000RF/Label_NY_NP9837_73-85
8OP23800/Label_NY_NP7100/NP7200_4.5-4
8OP23801/Label_NY_NP7100/NP7200_4.5-4
8OP23900/Label_NY_NP9837_73-85
8OP23900RF/Label_NY_NP9837_73-85
8OP24100/Label_A67
8OP24101/Label_A67
8OP241CK/Label_A12
8OP24200/Label_NY_NP7200_5.8125-6.5
8OP24201/Label_NY_NP7200_5.8125-6.5
8OP24300/Label_NY_NP7200_6.25-4
8OP24301/Label_NY_NP7200_6.25-4
8OP24500/Label_NY_NP7100/NP7200_4-4.5
8OP24501/Label_NY_NP7100/NP7200_4-4.5
8OP245CK/Label_A12
8OP24600/Label_NY_NP7200_5.8125-6.5
8OP252RSL/Label_NY_NP9995_2.625-3.375
8OP25300/Label_NY_NP7200_6.25-4
8OP25301/Label_NY_NP7200_6.25-4
8OP25700/Label_A41
8OP25701/Label_A41
8OP260RSL/Label_NY_NP7700_3.375-2.3437
8OP26100/Label_A37
8OP26101/Label_A37
8OP26300/Label_NY_NP7200_6.5-5.8125
8OP26301/Label_NY_NP7200_6.5-5.8125
8OP263CK/Label_NY_NP7030_8.5-6.5
8OP27200/Label_NY_NP7200_5.8125-6.5
8OP27201/Label_NY_NP7200_5.8125-6.5
8OP278F0/Label_A37
8OP278F1/Label_A37
8OP278L0/Label_A37
8OP278L1/Label_A37
8OP279RL/Label_NY_NP9815_01A
8OP28001/Label_A20
8OP28400/Label_NY_NP9815_01A
8OP29400/Label_NY_NP7200_6.5-5.8125
8OP29401/Label_NY_NP7200_6.5-5.8125
8OP294CK/Label_NY_NP7030_8.5-6.5
8OP308SI/Label_A06
8OP308SIRF/Label_NY_NP9300_2-1.375
8OP31200/Label_NY_NP9996_2.625-3.375
8OP32000/Label_A22
8OP32001/Label_A22
8OP320CK/Label_A34
8OP32800/Label_NY_NP9815_01A
8OP33700/Label_NY_NP9815_01A
8OP33800/Label_NP9350_
8OP338CL/Label_A04
8OP342CK/Label_A12
8OP34500/Label_A09
8OP34500A/Label_NY_NP9800_2.3437-3.375
8OP34600/Label_NY_NP7200_3.9659-4.3336
8OP34601/Label_NY_NP7200_3.9659-4.3336
8OP346CK/Label_NY_NP7030_9-2.625
8OP34800/Label_NY_NP7200_3.9659-4.3336
8OP34801/Label_NY_NP7200_3.9659-4.3336
8OP348CK/Label_NY_NP7030_9-2.625
8OP349CK/Label_A12
8OP35000/Label_A28
8OP35001/Label_A28
8OP350CK/Label_A12
8OP354CK/Label_NY_NP7030_9-2.625
8OP365RSL/Label_NY_NP9995_2.625-3.375
8OP37000/Label_A67
8OP37001/Label_A67
8OP37400/Label_A17
8OP37800/Label_A18
8OP38300/Label_A35
8OP38301/Label_A35
8OP383CK/Label_A12
8OP39200/Label_NY_NP7200_3.9659-4.3336
8OP39201/Label_NY_NP7200_3.9659-4.3336
8OP392CK/Label_NY_NP7030_9-2.625
8OP40300/Label_A40
8OP40301/Label_A40
8OP41000/Label_NP9350_
8OP42100/Label_A17
8OP42300/Label_NY_NP7700_1.5625-3.3125
8OP423BCS/Label_NY_NP7030_9-2.625
8OP423CK/Label_NY_NP7030_9-2.625
8OP43800/Label_A15
8OP44100/Label_A22
8OP44300/Label_A40
8OP44301/Label_A40
8OP443UP/Label_A12
8OP44500/Label_A41
8OP44501/Label_A41
8OP44800/Label_A67
8OP44801/Label_A67
8OP448CK/Label_A12
8OP451CK/Label_A12
8OP45200/Label_NY_NP7100/NP7200_3.9659-4.3336
8OP45201/Label_NY_NP7100/NP7200_3.9659-4.3336
8OP452CK/Label_NY_NP7030_9-2.625
8OP45300/Label_A35
8OP45301/Label_A35
8OP453CK/Label_A12
8OP45600/Label_A22
8OP45601/Label_A22
8OP45700/Label_NY_NP7200_3.9659-4.3336
8OP45701/Label_NY_NP7200_3.9659-4.3336
8OP457CK/Label_NY_NP7030_9-2.625
8OP460RL/Label_NY_NP9800_2.874-3.3465
8OP46100/Label_A35
8OP46101/Label_A35
8OP465CK/Label_A12
8OP465RL/Label_A15
8OP46700/Label_NY_NP9510_2.25-3.875
8OP467CK/Label_NY_NP7030_9-2.625
8OP46800/Label_A28
8OP46801/Label_A28
8OP468CK/Label_A12
8OP47101/Label_NY_NP7200_6.5-5.8125
8OP47300/Label_NY_NP7200_3.9659-4.3336
8OP47301/Label_NY_NP7200_3.9659-4.3336
8OP473CK/Label_NY_NP7030_9-2.625
8OP47400/Label_NY_NP7100/NP7200_5.5625-6.25
8OP47401/Label_NY_NP7100/NP7200_5.5625-6.25
8OP47500/Label_A35
8OP47501/Label_A35
8OP475UP/Label_A12
8OP48500/Label_NY_NP9815_01A
8OP48700/Label_A63
8OP487CK/Label_A12
8OP49500/Label_A20
8OP49501/Label_A20
8OP50900/Label_A40
8OP50901/Label_A40
8OP52000/Label_A41
8OP52001/Label_A41
8OP520CK/Label_A48
8OP52300/Label_A35
8OP52301/Label_A35
8OP52400/Label_A16
8OP527CK/Label_A12
8OP53500/Label_A35
8OP53501/Label_A35
8OP535UP/Label_A12
8OP54800/Label_A41
8OP54801/Label_A41
8OP548CK/Label_A44
8OP56000/Label_NY_NP9837_73-85
8OP56000RF/Label_NY_NP9837_73-85
8OP563RL-IN/Label_NY_NP9800_2.874-3.3465
8OP563RL/Label_NY_NP9800_2.874-3.3465
8OP56700/Label_NY_NP7200_5.8125-6.5
8OP56701/Label_NY_NP7200_5.8125-6.5
8OP57000/Label_A22
8OP57001/Label_A22
8OP570CK/Label_A12
8OP57200/Label_NY_NP98372_73-85
8OP57200RF/Label_NY_NP98372_73-85
8OP575UP/Label_A12
8OP57900/Label_A18
8OP579CK/Label_A57
8OP585SI/Label_A06
8OP585SIRF/Label_NY_NP9300_1.375-2
8OP587RSL/Label_NY_NP9991_3.375-2.625
8OP590ST1/Label_B04
8OP590ST2/Label_B04
8OP590ST3/Label_B04
8OP590ST4/Label_B04
8OP60100/Label_NP9890_
8OP60700/Label_A41
8OP60701/Label_A41
8OP61500/Label_A20
8OP61501/Label_A20
8OP63000/Label_NY_NP9837_01A
8OP63000RF/Label_NY_NP9837_85-73
8OP63500/Label_A35
8OP63501/Label_A35
8OP635UP/Label_A12
8OP63800/Label_A41
8OP63801/Label_A41
8OP640CK/Label_A12
8OP640RL/Label_A17
8OP64200/Label_A22
8OP64201/Label_A22
8OP642CK/Label_A12
8OP64500/Label_NY_NP9837_73-85
8OP64500RF/Label_NY_NP9837_73-85
8OP648CK/Label_A12
8OP65700/Label_A35
8OP65701/Label_A35
8OP657UP/Label_A12
8OP66700/Label_NY_NP7200_5.8125-6.5
8OP66701/Label_NY_NP7200_5.8125-6.5
8OP674CK/Label_A12
8OP69700/Label_A17
8OP697CK/Label_A12
8OP700CK/Label_NY_NP7030_7.750-5.250
8OP700CS/Label_NY_NP7030_7.750-5.250
8OP70500/Label_A17
8OP713RSL/Label_NY_NP9800_3.375-2.3437
8OP72100/Label_NY_NP9815_01A
8OP72800/Label_A15
8OP728CK/Label_A12
8OP734CK/Label_A12
8OP73601/Label_A41
8OP742CK/Label_A12
8OP75000/Label_A22
8OP75001/Label_A22
8OP765CK/Label_A12
8OP765CLA/Label_NY_NP7030_9-2.625
8OP765RL/Label_A62
8OP78800/Label_A35
8OP78801/Label_A35
8OP794RL/Label_A15
8OP80600/Label_NY_NP9837_01A
8OP80600RF/Label_NY_NP9837_85-73
8OP80800/Label_A63
8OP808500/Label_NY_NP7200_3.9659-4.3336
8OP808501/Label_NY_NP7200_3.9659-4.3336
8OP81000/Label_NP9350_
8OP83200/Label_A28
8OP83201/Label_A28
8OP838CK/Label_A34
8OP84200/Label_NP9350_
8OP84300/Label_A20
8OP84301/Label_A20
8OP86700/Label_NY_NP7200_4.5-4
8OP86701/Label_NY_NP7200_4.5-4
8OP87000/Label_NY_NP9800_2.375-3.5
8OP890RSL/Label_NY_NP9800_3.5-2.375
8OP90100/Label_A28
8OP90101/Label_A28
8OP906RL/Label_A18
8OP909BL/Label_B03
8OP909CU/Label_A03
8OP909CUSL/Label_A02
8OP912TP/Label_NP9350_
8OP91300/Label_NY_NP9800_3.5-2.375
8OP92000/Label_A37
8OP97501/Label_A41
8OQ095TS/Label_NY_NP7835_1.5
8OQ095UL/Label_NY_NP9800_01
8OQ113PL/Label_A45
8OQ215CS/Label_A12
8OQ25800/Label_A17
8OQ301TP/Label_NY_NP9802_3.5-2.375
8OQ335CLF/Label_NY_NP7700_2.877-2.489
8OQ349CS/Label_A12
8OQ387RL/Label_NY_NP9800_2.625-5.375
8OQ424CSA/Label_A12
8OQ424TPA/Label_A56
8OQ430KS/Label_NY_NP7030_9-2.625
8OQ432DSP/Label_NY_NP7030_9-2.625
8OQ432KS4/Label_NY_NP7030_9-2.625
8OQ449CS/Label_A12
8OQ5058CS/Label_NY_NP7030_9-2.625
8OQ508CSA/Label_A12
8OQ527CSA/Label_A12
8OQ543CK/Label_A12
8OQ548TP/Label_NY_NP9800_2.874-3.3465
8OQ558CS/Label_A12
8OQ558TP/Label_A56
8OQ601RSL/Label_NY_NP9802_2.874-3.3465
8OQ761CK/Label_A12
8OQ76900/Label_NY_NP7100_6-3.75
8OQ76901/Label_NY_NP7100_6-3.75
8OQ8111CS/Label_NY_NP7030_9-2.625
8OQ9027CS/Label_NY_NP7030_9-2.625
8OQ916KS/Label_NY_NP7030_9-2.625
8OQ95001/Label_A37
8OQ972TP/Label_NY_NP9802_3.5-2.375
8OS172CK/Label_A23
8OS29300/Label_A08
8OS293CK/Label_A31
8OS356CK/Label_A31
8OS40750/Label_A31
8OS41950/Label_A31
8OS420CL/Label_A31
8OS422CK/Label_A23
8OS816CK/Label_A23
8OS849CK/Label_A23
8OU120CK/Label_A23
8OU127CK/Label_NY_NP7030_9-2.625
8OU128CK/Label_NY_NP7030_9-2.625
8OU130CK/Label_NY_NP7030_9-2.625
8OU171CK/Label_NY_NP7030_9-2.625
8OU211CK/Label_NY_NP7030_9-2.625
8OU214UP/Label_A12
8OU240CK/Label_NY_NP7030_9-2.625
8OU372CK/Label_NY_NP7030_9-2.625
8OU385CK/Label_A12
8OU444ULS/Label_NY_NA_12-3
8OU446CK/Label_NY_NP7030_9-2.625
8OU451CK/Label_NY_NP7030_9-2.625
8OU451CS/Label_A12
8OU500CK/Label_NY_NP7030_9-2.625
8OU591CK/Label_NY_NP7030_9-2.625
8OU621CK/Label_NY_NP7030_9-2.625
8OX13100/Label_NY_NP7119_3.80-4.82
8OX13101/Label_NY_NP7118_3.80-4.82
8OX131CK/Label_NY_NP7200_7.75-5.25
8OX20635DHCS/Label_NY_NP7030_9-2.625
8OX20735DHCS/Label_NY_NP7030_9-2.625
8OX20875DHCS/Label_NY_NP7030_9-2.625
8OX20975DHCS/Label_NY_NP7030_9-2.625
8OX351HL/Label_NY_NP9005_2.375-2.25
8PB11400/Label_NY_NP7030_6-5
8PS27350/Label_NY_NP7030_6-5
8PS32450/Label_NY_NP7030_6-5
8PS428CK/Label_NY_NP7030_6-5
8PS428CK/Label_NY_NP7030_6-5
8QB11400/Label_NY_NP7030_6-5
8QS27350/Label_NY_NP7030_6-5
8RB11400/Label_NY_NP7030_6-5
8RS32450/Label_NY_NP7030_6-5
8RS428CK/Label_NY_NP7030_6-5
8SS27350/Label_NY_NP7030_6-5
8XB432CK/Label_NY_NP7900_9-2.625
8XM037CPL/Label_A34
8XM073CPL/Label_A34
8XM076CPL/Label_A34
8ZP57001/Label_NY_NP7900_01A

## CONVERT
METHOD | METHOD_C | Res_TestName | METHOD
AOCS Cd 8-53 | AOCS_Cd_8_53 | Peroxide Value | AOCS Cd 8-53
AOCS Td 1a-64) | AOCS_Td_1a_64 | color (Gardner) | AOCS Td 1a-64)
| APAM_6_018 | % Water | APAM 6.018
| ASTMD_646 | Basis Weight (lbs/ream) | ASTM D 646
|  | Box Inspection | Box Inspection
|  | Applicator_Inspection | #REF!
| Dyne_Test_Kit | Chemical Inspection | #REF!
| IUPAC_2_201M | Dyne Level | Dyne Level
| IUPAC_2_202 | Canister & Lid Inspection | Canister & Lid Inspection
| J&J | Chemical Inspection | TM 208
| J&J | Applicator_Inspection | Odor (TM 208)
| J&J | pH | TM200
| J&J | Canister & Lid Inspection | Print
| J&J | Moisture Vapor Transmission rate | #REF!
| J&J | Residual Solvents | #REF!
| J&J | Temptale Reading | #REF!
| J&J | Shipper Inspection | #REF!
| J&J | Certificate of Analysis | COA
| J&J | Canister Release | COA
| J&J | Caliper | Caliper
| J&J | Swab Measurements | Swab Measurements
| J&J | Applicator Measurements | Applicator Measurements
| J&J | Box Inspection | Certificate of Analysis
| TM7252 | Certificate of Analysis | Certificate of Analysis
| J&J | Certificate of Analysis | Certificate of Compliance
| J&J | Certificate of Compliance | Certificate of Compliance
| J&J | Certificate of Compliance | Certificate of Conformance
| J&J | Certificate of Conformance | Certificate of Conformance
| J&J | Certificate of Conformance | Box Inspection
| J&J | Box Inspection | Dyne Level
| J&J | Dyne Level | Dyne Test Kit
|  | Acid Value | IUPAC 2.201(M)
| JJTMWW78010__JJTM7252 | Saponification Value | IUPAC 2.202
Measurement | Measurement | GC Profile | J&J TM-001862
NPTM 1019 | TM1019 | Assay | J&J TM-002528
ORGANOLEPTIC | TM201 | Absorption | J&J TM-002747
See Analyte for Method # | See_Analyte_for_Method | Iodine Value | J&J TM-003029
STM 101 | STM101 | GC Profile | J&J TM-003030
TC276/WG2N30 930809C | TC276_WG2N30_930809C | Chemical Inspection | J&J TM-003188
TM 101 | TM101 | Odor | J&J TM-003188
TM 102 | TM102 | Acidity | J&J TM-003589
TM 103 | TM103 | Identity of Denatonium Benzoate | J&J TM4028
TM 107 | TM107 | Assay | J&J TM4075
TM 109 | TM109 | color (Gardner) | J&J TM701
TM 132 | TM132 | Loss on Drying | J&J TM7221
TM 141 / Visual | TM141 | Viscosity | J&J TM7423
TM 143 | TM143 | Hydroxyl Value | J&J TMWW7671
TM143 | TM143 | Chemical Inspection | J&J TMWW78010
TM 148 | TM148 | Chemical Inspection | J&J TMWW78010/J&J TM7252
TM 200 | TM200 | Visual Inspection | J&J TMWW78010/J&J TM7252
TM 202 | TM202 | Acid Value | J&J TMWW78011
TM 203 | TM203 | Density @ 20° C | J&J TMWW78012
TM 204 | TM204 | Specific Gravity at 20�C | J&J TMWW78012
TM 206 | TM206 | Specific Gravity at 25�C | J&J TMWW78012
TM 208 | TM208 | Specific Gravity @ 15.56°C/ 60° F | J&J TMWW78012
TM 222 | TM222 | Identification | J&J TMWW78013
TM 226 | TM226 | pH | J&J TMWW78014
TM 3005 | TM3005 | Refractive Index at 25�C | J&J TMWW78015
TM 301/434 | TM301_TM434 | Refractive Index at 20°C | J&J TMWW78015
TM 3014 | TM3014 | Saponification Value | J&J TMWW78016
TM 323 | TM323 | % Water | J&J TMWW78017
TM 366 | TM366 | Chemical Inspection | JJTM7252
TM 375 | TM375 | Chemical Inspection | JJTMWW78010 / JJTM7252
TM 376 | TM376 | Solo Lid Inspection | Solo Lid Inspection
TM 395 | TM395 | Eyespot Repeat | TM 208
TM 399 | TM399 | Web Measurement | Measurement
TM 4001 | TM4001 | Solo Lid Inspection | Measurement
TM 411 | TM411 | Canister & Lid Inspection | Measurement
TM 434 | TM434 | Top Load Compression Strength | NPTM 1019
TM 602 | TM602 | Chemical Inspection | ORGANOLEPTIC
TM 603 | TM603 | Bulk Liquid Inspection | See Analyte for Method #
TM 610 | TM610 | Applicator_Inspection | See Analyte for Method #
TM 613 | TM613 | Chemical Inspection | See Analyte for Method #
TM 869 | TM869 | Finished Good Inspection | See Analyte for Method #
TM 882 | TM882 | Apparent proof | STM 101
TM WW78014 | TMWW78014 | Cloudpoint | TC276/WG2N30 930809C
TM-002291 | TM_002291 | Basis Weight (lbs/ream) | TM 101
TM-002568 | TM_002568 | Basis Weight | TM 101
TM-002569 | TM_002569 | Caliper | TM 102
TM101/TM107/TM132/Measure | TM101_TM107_TM132 | Tensile Strength | TM 103
TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | Seal Strength | TM 107
TM400/408b/421/422/USP | TM400_TM408b_TM421_TM422_USP | Seal strength (NP8536) | TM 107
USP | USP | Seal strength (NP8511) | TM 107
USP <197F> | USP_197F | Swab Measurements | TM 109
USP <197M> | USP_197M | Web Thickness | TM 132
USP <541> | USP_541 | Shipper Inspection | TM 141 / Visual
USP <621> | USP_621 | Web Measurement | TM 143
USP <791> | USP_791 | Flip Lid Force | TM 143
USP-HPLC | USP_HPLC | Applicator Measurements | TM 143
USP/NF | USP_NF | Canister & Lid Inspection | TM 143 / Visual
USP/NF <841> | USP_NF_841 | Label Inspection | TM 143 / Visual
Vatten LB | Vatten_LB | Flip Lid Force | TM 148
Visual | Visual | Lid Pull Force | TM 148
Visual Observation | Visual | pH as is | TM 200
WI QAS 076 | WI_QAS_076 | pH | TM 200
WI QAS 090 | WI_QAS_090 | Density @ 20° C | TM 202
Box Inspection | Box Inspection | Specific Gravity at 25�C | TM 202
Canister & Lid Inspection | TM143 | Specific Gravity at 20�C | TM 202
Odor (TM 208) | TM208 | Specific Gravity @ 15.56°C/ 60° F | TM 202
Color (TM 208) | TM208 | Viscosity | TM 203
Certificate of Conformance | COA | Refractive Index at 25�C | TM 204
Dyne Level | Dyne Level | Refractive Index at 20°C | TM 204
TM200 | TM200 | Total Solids | TM 206
Report result from COA |  | Appearance | TM 208
Residual Solvents |  | Visual Inspection | TM 208
Shipment Temperatutre | NA | Color | TM 208
Shipper Style | TM141 | Solutions and Solubility | TM 222
Temptale Reading |  | Nonvolatile Residue | TM 226
Caliper | TM102 | Active (BTC 1210-80%) | TM 3005
Swab Measurements | TM109 | Assay | TM 301/434
Applicator Measurements | TM142 | Identification | TM 3014
Certificate of Analysis | COA | Total Color | TM 3014
Certificate of Compliance | COA | Ethanol | TM 323
Solo Lid Inspection | SOLO | % NaOCl | TM 366
|  | Acidity | TM 375
|  | Assay | TM 376
|  | %Phenoxyethanol & Sodium Benzoate | TM 395
|  | % Sodium Benzoate | TM 395
|  | % Phenoxyethanol & IPBC | TM 399
|  | % Salicylic Acid | TM 399
|  | Identification A | TM 4001
|  | Identification B | TM 4001
|  | Identification | TM 411
|  | Total % ADBAC | TM 434
|  | % Dual Quat | TM 434
|  | B.cepacia | TM 602
|  | Bacteria | TM 602
|  | Blood Agar | TM 602
|  | E.coli | TM 602
|  | Enrichment_E.coli | TM 602
|  | Mold | TM 602
|  | P.aeruginosa | TM 602
|  | S.aureus | TM 602
|  | Enrichment | TM 602
|  | Celsis_Rapid | TM 603
|  | Bioburden_TAMC | TM 610
|  | Bioburden_TMYC | TM 610
|  | Sterility Controls Raw Materials | TM 610
|  | GramStain_BloodAgar | TM 613
|  | J&J_Pathogens | TM 869
|  | J&J Fungal | TM 869
|  | J&J SDA Controls | TM 869
|  | J&J TSB Controls - TSA | TM 869
|  | J&J TSA Controls | TM 869
|  | J&J TSB Controls - SAB | TM 869
|  | J&J_Bacteria | TM 869
|  | J&J_Enrichment | TM 869
|  | % Benzoic Acid Phenoxyethanol & Dehydroacetic Acid | TM 882
|  | pH as is | TM WW78014
|  | Assay | TM-002291
|  | Total Sugar | TM-002568
|  | Reducing Sugars | TM-002569
|  | Web Inspection | TM101/TM107/TM132/Measure
|  | Chemical Inspection | TM208/220/230/USP/Kao LM
|  | Identification | TM400/408b/421/422/USP
|  | Acidity | USP
|  | Organic Impurites | USP
|  | Organic Impurities Part II | USP
|  | Assay | USP
|  | Calcium | USP
|  | Chloride | USP
|  | Limit of Ethylene Glycol and Diethylene Glycol | USP
|  | Nonvolatile Residue | USP
|  | Heavy Metals | USP
|  | Identification - USP | USP
|  | Identification | USP
|  | Related Compounds | USP
|  | Loss on Drying | USP
|  | Melting Range | USP
|  | Nitrilotriacetic Acid | USP
|  | Optical Rotation | USP
|  | Refractive Index at 20°C | USP
|  | Residue on Ignition | USP
|  | Specific Rotation | USP
|  | Sulfate Content | USP
|  | Identification | USP <197F>
|  | Identification | USP <197M>
|  | Assay | USP <541>
|  | Related Compounds | USP <621>
|  | pH as is | USP <791>
|  | pH | USP <791>
|  | Assay | USP-HPLC
|  | Assay | USP/NF
|  | Ratio of Alkyl Compounds | USP/NF
|  | Identification A | USP/NF
|  | Identification | USP/NF
|  | Identification B | USP/NF
|  | Individual Impurity | USP/NF
|  | Phenol | USP/NF
|  | Limit of Phenol | USP/NF
|  | Phenoxyethanol Content | USP/NF
|  | Purity | USP/NF
|  | Specific Gravity at 20�C | USP/NF
|  | Total Impurities | USP/NF
|  | Specific Gravity at 25�C | USP/NF <841>
|  | Water Content | Vatten LB
|  | 2D Matrix (Barcode) | Visual
|  | Visual Inspection | Visual
|  | Appearance | Visual
|  | Cavity | Visual
|  | Certificate of Analysis | Visual
|  | Color | Visual
|  | Functionality | Visual
|  | Visual Inspection | Visual Observation
|  | Web Inspection | Visual Observation
|  | GramStain_Candida | WI QAS 076
|  | GramStain_Bioburden_TAMC | WI QAS 076
|  | GramStain_Bioburden_TMYC | WI QAS 076
|  | GramStain_J&J_TSA | WI QAS 076
|  | IdCoag_BloodAgar | WI QAS 090
KC Part Number

## analyte
Analyte | Analyte.C
% BZK and Dual-Quat | % BZK and Dual-Quat
% NaOCl by weight (calc) | % NaOCl by weight (calc)
% NaOCl by weight Final (calc) | % NaOCl by weight Final (calc)
2D Matrix (Barcode) | 2D Matrix (Barcode)
Absorption (1%/1cm) | Absorption (1%/1cm)
Acid Value | Acid Value
Acid Value | Acid Value
Acid Value | Acid Value
Acid Value Final | Acid Value Final
Acid Value Final | Acid Value Final
Acid Value Final | Acid Value Final
Acidity (calc) | Acidity (calc)
Acidity (calc) | Acidity (calc)
Acidity (calc) | Acidity (calc)
Acidity as Acetic Acid(10ml burette 0.05ml subdiv) | Acidity as Acetic Acid(10ml burette 0.05ml subdiv)
Acidity Final (Calc) | Acidity Final (Calc)
Acidity Final (Calc) | Acidity Final (Calc)
Acidity Final (Calc) | Acidity Final (Calc)
Active (BTC 1210-80%) | Active (BTC 1210-80%)
Any individual unspecified impuritv | Any individual unspecified impuritv
Any individual unspecified impuritv | Any individual unspecified impuritv
Apparent proof | Apparent proof
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance | Appearance
Appearance (J&J TMWW78010) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance (TM 208) | Appearance
Appearance of Packaging | Appearance of Packaging
Appearance/Matches Embossing Pattern | Appearance/Matches Embossing Pattern
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
Applicator Weight | Applicator Weight
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
AQL Sample Size | AQL Sample Size
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork / Copy | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Artwork_Copy (Color & Text) | Artwork / Copy
Assay | Assay
Assay | Assay
Assay | Assay
Assay | Assay
Assay | Assay
Assay | Assay
Assay | Assay
Assay | Assay
Assay | Assay
Assay | Assay
Assay | Assay
Assay | Assay
Assay | Assay
Assay | Assay
Assay | Assay
Assay | Assay
Assay | Assay
Assay | Assay
Assay | Assay
Assay | Assay
Assay | Assay
Assay as is (Without Drying) | Assay as is (Without Drying)
Assay as is Final (Without Drying) | Assay as is Final (Without Drying)
Assay Final | Assay Final
Assay Final | Assay Final
Assay Final | Assay Final
Assay Final | Assay Final
Assay Final | Assay Final
Assay Final | Assay Final
Assay Final | Assay Final
Assay Final | Assay Final
Assay Final | Assay Final
Assay Final | Assay Final
Assay Final | Assay Final
Assay Final | Assay Final
Assay Final | Assay Final
Assay Final | Assay Final
Assay Final | Assay Final
Assay Final | Assay Final
Assay on Dried Basis | Assay on Dried Basis
Assay on Dried Basis | Assay on Dried Basis
Assay on Dried Basis Final | Assay on Dried Basis Final
Assay on Dried Basis Final | Assay on Dried Basis Final
B.cepacia | B.cepacia
B.cepacia | B.cepacia
Back Flip | Back Flip
Back Flip | Back Flip
Back Height | Back Height
Bacteria | Bacteria
Bacteria | Bacteria
Bacteria Count (Obeserved) | Bacteria Count (Obeserved)
Bacteria Count (Obeserved) | Bacteria Count (Obeserved)
Bacteria Count (Reported) | Bacteria Count (Reported)
Bacteria Count (Reported) | Bacteria Count (Reported)
Barcode Legibility/Quality | Barcode Legibility/Quality
Barcode Legibility/Quality | Barcode Legibility/Quality
Barcode Legibility/Quality | Barcode Legibility/Quality
Barcode Legibility/Quality | Barcode Legibility/Quality
Barcode Legibility/Quality | Barcode Legibility/Quality
Basis Weight | Basis Weight
Basis Weight | Basis Weight
Basis Weight | Basis Weight
Basis Weight | Basis Weight
Basis Weight | Basis Weight
Basis Weight | Basis Weight
Basis Weight (GSM) (TM 101) | Basis Weight
Basis Weight (GSM) (TM 101) | Basis Weight
Basis Weight (GSM) (TM 101) | Basis Weight
Basis Weight (GSM) (TM 101) | Basis Weight
Basis Weight (GSM) (TM 101) | Basis Weight
Basis Weight (GSM) (TM 101) | Basis Weight
Basis Weight (GSM) (TM 101) | Basis Weight
Basis Weight (GSM) (TM 101) | Basis Weight
Basis Weight (GSM) (TM 101) | Basis Weight
Basis Weight (GSM) (TM 101) | Basis Weight
Basis Weight (GSM) (TM 101) | Basis Weight
Basis Weight (GSM) (TM 101) | Basis Weight
Basis Weight (GSM) (TM 101) | Basis Weight
Basis Weight (GSM) (TM 101) | Basis Weight
Basis Weight (GSM) (TM 101) | Basis Weight
Basis Weight (GSM) (TM 101) | Basis Weight
Basis Weight (GSM) (TM 101) | Basis Weight
Basis Weight (lbs/ream) (TM 101) | Basis Weight
Basis_Weight | Basis Weight
Basis_Weight | Basis Weight
Basis_Weight_(#/3000_sq_ft) | Basis Weight
Basis_Weight_(#/3000_sq_ft) | Basis Weight
Basis_Weight_(#/Ream-Printed) | Basis Weight
Basis_Weight_(#/Ream-Printed) | Basis Weight
Basis_Weight_(#/Ream-Printed) | Basis Weight
Basis_Weight_(#/Ream-Printed/Varnished) | Basis Weight
Basis_Weight_(#/Ream-Printed/Varnished) | Basis Weight
Basis_Weight_(#/Ream-Printed/Varnished) | Basis Weight
Basis_Weight_(#/Ream-Printed/Varnished) | Basis Weight
Basis_Weight_(#/Ream-Printed/Varnished) | Basis Weight
Basis_Weight_(#/Ream-Printed/Varnished) | Basis Weight
Basis_Weight_(#/Ream-Printed/Varnished) | Basis Weight
Basis_Weight_(#/Ream-Printed/Varnished) | Basis Weight
Basis_Weight_(#/Ream-Unprinted) | Basis Weight
Basis_Weight_(#/Ream) | Basis Weight
Basis_Weight_(#/Ream) | Basis Weight
Basis_Weight_(#/Ream) | Basis Weight
Basis_Weight_(#/Ream) | Basis Weight
Basis_Weight_(#/Ream) | Basis Weight
Basis_Weight_(#/Ream) | Basis Weight
Basis_Weight_(#/Ream) | Basis Weight
Basis_Weight_(#/Ream) | Basis Weight
Basis_Weight_(#/Ream) | Basis Weight
Basis_Weight_(#/Ream) | Basis Weight
Basis_Weight_(#/Ream) | Basis Weight
Basis_Weight_(#/Ream) | Basis Weight
Basis_Weight_(#/Ream) | Basis Weight
Basis_Weight_(#/Ream) | Basis Weight
Basis_Weight_(#/Ream) | Basis Weight
Basis_Weight_(#/Ream) | Basis Weight
Basis_Weight_(#/Ream) | Basis Weight
Basis_Weight_(#/Ream) | Basis Weight
Basis_Weight_(#/Ream) | Basis Weight
Basis_Weight_(#/Ream) | Basis Weight
Basis_Weight_(#/Ream) | Basis Weight
Basis_Weight_(#/Ream) | Basis Weight
Basis_Weight_(#/Ream) | Basis Weight
Basis_Weight_(#/Ream) | Basis Weight
Basis_Weight_(#/Ream) | Basis Weight
Basis_Weight_(#/Ream) | Basis Weight
Basis_Weight_(#/Ream) | Basis Weight
Basis_Weight_(#/Ream) | Basis Weight
Basis_Weight_(#/Ream) | Basis Weight
Basis_Weight_(#/Ream) | Basis Weight
Basis_Weight_(#/Ream) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Basis_Weight_(GSM) | Basis Weight
Bead Diameter | Bead Diameter
Bead Diameter | Bead Diameter
Bead Diameter | Bead Diameter
Bead Diameter | Bead Diameter
Bead Diameter | Bead Diameter
Bead Diameter | Bead Diameter
Bead Diameter | Bead Diameter
Bead Diameter | Bead Diameter
Bead Diameter | Bead Diameter
Bead Diameter | Bead Diameter
Benzoic Acid | Benzoic Acid
Biggy or SDA | Biggy or SDA
Biggy or SDA | Biggy or SDA
Biggy or SDA | Biggy or SDA
Biggy or SDA | Biggy or SDA
Biggy or SDA | Biggy or SDA
Biggy or SDA | Biggy or SDA
Blood Agar | Blood Agar
Blood Agar | Blood Agar
Bottle Internal Diameter at Flange | Bottle Internal Diameter at Flange
Bottle Internal Diameter at Flange | Bottle Internal Diameter at Flange
Bottle Internal Diameter at Flange | Bottle Internal Diameter at Flange
C12 + C14 Content of Total | C12 + C14 Content of Total
C12 Content of Total | C12 Content of Total
C14 Content of Total | C14 Content of Total
Calcium | Calcium
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
Caliper | Caliper
caliper (COA) | Caliper
caliper (COA) | Caliper
Caliper 1-ply (mm) (TM 102) | Caliper
Caliper 1-ply (mm) (TM 102) | Caliper
Caliper 1-ply (mm) (TM 102) | Caliper
Caliper after Slitter (inches) (TM 102) | Caliper
Caliper after Slitter (inches) (TM 102) | Caliper
Caliper after Slitter (mm) (CofA) | Caliper
Caliper after Slitter (mm) (CofA) | Caliper
Caliper after Slitter (mm) (CofA) | Caliper
Caliper after Slitter (mm) (CofA) | Caliper
Caliper after Slitter (mm) (TM 102) | Caliper
Caliper after Slitter (mm) (TM 102) | Caliper
Caliper after Slitter (mm) (TM 102) | Caliper
Caliper after Slitter (mm) (TM 102) | Caliper
Caliper after Slitter (mm) (TM 102) | Caliper
Caliper after Slitter (mm) (TM 102) | Caliper
Caliper after Slitter (mm) (TM 102) | Caliper
Caliper at point of manufacture (mils) (TM 102) | Caliper
Caliper_After_Slitter_(inch) | Caliper
Caliper_After_Slitter_(inch) | Caliper
Caliper_After_Slitter_(mm) | Caliper
Caliper_After_Slitter_(mm) | Caliper
Caliper_After_Slitter_(mm) | Caliper
Caliper_After_Slitter_(mm) | Caliper
Caliper_After_Slitter_(mm) | Caliper
Caliper_After_Slitter_(mm) | Caliper
Caliper_After_Slitter_(mm) | Caliper
Caliper_After_Slitter_(mm) | Caliper
Caliper_After_Slitter_(mm) | Caliper
Caliper_After_Slitter_(mm) | Caliper
Caliper_After_Slitter_(mm) | Caliper
Caliper_After_Slitter_(mm) | Caliper
Caliper_After_Slitter_(mm) | Caliper
Caliper_After_Slitter_CoA_(mm) | Caliper
Caliper_After_Slitter_CoA_(mm) | Caliper
Caliper_After_Slitter_CoA_(mm) | Caliper
Caliper_After_Slitter_CoA_(mm) | Caliper
Caliper_After_Slitter_CoA_(mm) | Caliper
Caliper_After_Slitter_CoA_(mm) | Caliper
Caliper_After_Slitter_CoA_(mm) | Caliper
Caliper_After_Slitter_CoA_(mm) | Caliper
Caliper_at_Pt_of_Mfr_(micron) | Caliper
Caliper_at_Pt_of_Mfr_(micron) | Caliper
Caliper_CoA_(uM) | Caliper
Caliper_CoA_(uM) | Caliper
Caliper_CoA_(uM) | Caliper
Caliper_CoA_(uM) | Caliper
Caliper_CoA_(uM) | Caliper
Caliper-1_Ply_(mm) | Caliper
Caliper-1_Ply_(mm) | Caliper
Caliper-1_Ply_(mm) | Caliper
Cap Thickness | Cap Thickness
Cap Thickness | Cap Thickness
Cap Thickness | Cap Thickness
Cap Thickness | Cap Thickness
Cap Thickness | Cap Thickness
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Material | Carton Material
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
Carton Style | Carton Style
cavity Inserts | cavity Inserts
cavity Inserts | cavity Inserts
cavity Inserts | cavity Inserts
cavity Inserts | cavity Inserts
cavity Inserts | cavity Inserts
cavity Inserts | cavity Inserts
Celsis_ Result | Celsis Result
Celsis_ Result | Celsis Result
Center to Center Distance | Center to Center Distance
Center to Center Distance | Center to Center Distance
Center to Center Distance | Center to Center Distance
Center to Center Distance | Center to Center Distance
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Analysis | Certificate of Analysis
Certificate of Compliance | Certificate of Compliance
Certificate of Compliance | Certificate of Compliance
Certificate of Compliance | Certificate of Compliance
Certificate of Compliance | Certificate of Compliance
Certificate of Compliance | Certificate of Compliance
Certificate of Compliance | Certificate of Compliance
Certificate of Compliance | Certificate of Compliance
Certificate of Compliance | Certificate of Compliance
Certificate of Compliance | Certificate of Compliance
Certificate of Compliance | Certificate of Compliance
Certificate of Compliance | Certificate of Compliance
Certificate of Compliance | Certificate of Compliance
Certificate of Compliance | Certificate of Compliance
Certificate of Compliance | Certificate of Compliance
Certificate of Compliance | Certificate of Compliance
Certificate of Compliance | Certificate of Compliance
Certificate of Compliance | Certificate of Compliance
Certificate of Compliance | Certificate of Compliance
Certificate of Compliance | Certificate of Compliance
Certificate of Compliance | Certificate of Compliance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certificate of Conformance | Certificate of Conformance
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Certification Stamp (ECT) | Certification Stamp (ECT)
Cetrimide | Cetrimide
Cetrimide | Cetrimide
Cetrimide | Cetrimide
Cetrimide | Cetrimide
Cetrimide | Cetrimide
Cetrimide | Cetrimide
Ch!orhexidine oxazinone analog | Ch!orhexidine oxazinone analog
Chlorhexidine | Chlorhexidine
Chlorhexidine urea | Chlorhexidine urea
Chlorhexidine amine | Chlorhexidine amine
Chlorhexidine dimer | Chlorhexidine dimer
Chlorhexidine glucityl triazine | Chlorhexidine glucityl triazine
Chlorhexidine guanidine | Chlorhexidine guanidine
Chlorhexidine nitrile | Chlorhexidine nitrile
Chloride | Chloride
Chloride | Chloride
Cloudpoint | Cloudpoint
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
Color | Color
color (Gardner) | color (Gardner)
color (Gardner) | color (Gardner)
Color (J&J TM-000839) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 208) | Color
Color (TM 230) | Color
Color (TM 230) | Color
Color (TM 230) | Color
Color (USP) | Color (USP)
Color / Appearance (TM 208) | Color / Appearance
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Color/Copy | Color/Copy
Colour (J&J TM7377) | Colour
Colour (J&J TM7377) | Colour
Colour (J&J TMP-552) | Colour
Contamination | Contamination
Contamination | Contamination
Contamination | Contamination
Contamination | Contamination
Contamination | Contamination
Contamination | Contamination
Contamination | Contamination
Contamination | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Contamination (TM 208) | Contamination
Core ID (Inches) | Core ID (Inches)
Core ID (Inches) | Core ID (Inches)
Dehydroacetic Acid | Dehydroacetic Acid
Density | Density
Density | Density
Density | Density
Density | Density
Density | Density
Description | Description
Description | Description
Description (TM 208) | Description
Description (TM 208) | Description
Description (TM 208) | Description
Destruct Bond | Destruct Bond
Destruct Bond | Destruct Bond
Destruct Bond | Destruct Bond
Diameter | Diameter
Diameter | Diameter
Diameter | Diameter
Diameter | Diameter
Diethylene Glycol | Diethylene Glycol
Diethylene Glycol | Diethylene Glycol
Dimensions - Bead OD | Dimensions - Bead OD
Dimensions - Bead OD | Dimensions - Bead OD
Dimensions - Bead OD | Dimensions - Bead OD
Dimensions - Bead OD | Dimensions - Bead OD
Dimensions - Bead OD | Dimensions - Bead OD
Dimensions - Bead OD | Dimensions - Bead OD
Dimensions - Bead OD | Dimensions - Bead OD
Dimensions - Canister Height OD | Dimensions - Canister Height OD
Dimensions - Canister OD | Dimensions - Canister OD
Dimensions - Canister OD | Dimensions - Canister OD
Dimensions - Canister OD | Dimensions - Canister OD
Dimensions - Canister OD | Dimensions - Canister OD
Dimensions - Canister OD | Dimensions - Canister OD
Dimensions - Canister OD | Dimensions - Canister OD
Dimensions - Label Length | Dimensions - Label Length
Dimensions - Label Length | Dimensions - Label Length
Dimensions - Label Length | Dimensions - Label Length
Dimensions - Label Length | Dimensions - Label Length
Dimensions - Label Length | Dimensions - Label Length
Dimensions - Label Length | Dimensions - Label Length
Dimensions - Label Length | Dimensions - Label Length
Dimensions - Label Length | Dimensions - Label Length
Dimensions - Label Length | Dimensions - Label Length
Dimensions - Label Length | Dimensions - Label Length
Dimensions - Label Length | Dimensions - Label Length
Dimensions - Label Length | Dimensions - Label Length
Dimensions - Label Length | Dimensions - Label Length
Dimensions - Label Length | Dimensions - Label Length
Dimensions - Label Length | Dimensions - Label Length
Dimensions - Label Length | Dimensions - Label Length
Dimensions - Label Length | Dimensions - Label Length
Dimensions - Label Length | Dimensions - Label Length
Dimensions - Label Length | Dimensions - Label Length
Dimensions - Label Width | Dimensions - Label Width
Dimensions - Label Width | Dimensions - Label Width
Dimensions - Label Width | Dimensions - Label Width
Dimensions - Label Width | Dimensions - Label Width
Dimensions - Label Width | Dimensions - Label Width
Dimensions - Label Width | Dimensions - Label Width
Dimensions - Label Width | Dimensions - Label Width
Dimensions - Label Width | Dimensions - Label Width
Dimensions - Label Width | Dimensions - Label Width
Dimensions - Label Width | Dimensions - Label Width
Dimensions - Label Width | Dimensions - Label Width
Dimensions - Label Width | Dimensions - Label Width
Dimensions - Label Width | Dimensions - Label Width
Dimensions - Label Width | Dimensions - Label Width
Dimensions - Label Width | Dimensions - Label Width
Dimensions - Label Width | Dimensions - Label Width
Dimensions - Label Width | Dimensions - Label Width
Dimensions - Label Width | Dimensions - Label Width
Dimensions - Label Width | Dimensions - Label Width
Dimensions - Lip (width) | Dimensions - Lip (width)
Dimensions - Lip (width) | Dimensions - Lip (width)
Dimensions - Lip (width) | Dimensions - Lip (width)
Dimensions - Neck OD | Dimensions - Neck OD
Dimensions - Neck OD | Dimensions - Neck OD
Dimensions - Neck OD | Dimensions - Neck OD
Dimensions - Neck OD | Dimensions - Neck OD
Dimensions - Neck OD | Dimensions - Neck OD
Dimensions - Total Height OD | Dimensions - Total Height OD
Dimensions - Total Height OD | Dimensions - Total Height OD
Dimensions - Total Height OD | Dimensions - Total Height OD
Dimensions - Total Height OD | Dimensions - Total Height OD
Dimensions - Total Height OD | Dimensions - Total Height OD
Dimensions - Total Height OD | Dimensions - Total Height OD
Dry CD (g/in) (TM103) | Dry CD (g/in)
Dry CD (g/in) (TM103) | Dry CD (g/in)
Dry CD (g/in) (TM103) | Dry CD (g/in)
Dry CD (g/in) (TM103) | Dry CD (g/in)
Dry CD (g/in) (TM103) | Dry CD (g/in)
Dry CD (g/in) (TM103) | Dry CD (g/in)
Dry MD (g/in) (TM 103) | Dry MD (g/in)
Dry MD (g/in) (TM 103) | Dry MD (g/in)
Dry MD (g/in) (TM 103) | Dry MD (g/in)
Dry MD (g/in) (TM 103) | Dry MD (g/in)
Dry MD (g/in) (TM 103) | Dry MD (g/in)
Dry MD (g/in) (TM 103) | Dry MD (g/in)
Dry_CD_(g/in) | Dry CD (g/in)
Dry_CD_(g/in) | Dry CD (g/in)
Dry_CD_(g/in) | Dry CD (g/in)
Dry_CD_(g/in) | Dry CD (g/in)
Dry_CD_(g/in) | Dry CD (g/in)
Dry_CD_(g/in) | Dry CD (g/in)
Dry_CD_(g/in) | Dry CD (g/in)
Dry_CD_(g/in) | Dry CD (g/in)
Dry_CD_(g/in) | Dry CD (g/in)
Dry_CD_(g/in) | Dry CD (g/in)
Dry_CD_(g/in) | Dry CD (g/in)
Dry_CD_CoA_(g/in) | Dry CD CoA (g/in)
Dry_MD_(g/in) | Dry MD (g/in)
Dry_MD_(g/in) | Dry MD (g/in)
Dry_MD_(g/in) | Dry MD (g/in)
Dry_MD_(g/in) | Dry MD (g/in)
Dry_MD_(g/in) | Dry MD (g/in)
Dry_MD_(g/in) | Dry MD (g/in)
Dry_MD_(g/in) | Dry MD (g/in)
Dry_MD_(g/in) | Dry MD (g/in)
Dry_MD_(g/in) | Dry MD (g/in)
Dry_MD_(g/in) | Dry MD (g/in)
Dry_MD_(g/in) | Dry MD (g/in)
Dry_MD_CoA_(g/in) | Dry MD CoA (g/in)
Dyne Level (outside film) | Dyne Level (outside film)
Dyne Level (outside film) | Dyne Level (outside film)
Dyne Level (outside film) | Dyne Level (outside film)
Dyne Level (outside film) | Dyne Level (outside film)
Dyne Level (outside film) | Dyne Level (outside film)
Dyne Level (outside film) | Dyne Level (outside film)
E Diameter | E Diameter
E Diameter | E Diameter
E Diameter | E Diameter
E.coli | E.coli
E.coli | E.coli
Elongation % Wet CD | Elongation % Wet CD
Elongation % Wet CD | Elongation % Wet CD
Elongation % Wet CD | Elongation % Wet CD
Elongation % Wet MD | Elongation % Wet MD
Elongation % Wet MD | Elongation % Wet MD
Elongation % Wet MD | Elongation % Wet MD
End Circle Diameter | End Circle Diameter
End Circle Diameter | End Circle Diameter
End Circle Diameter | End Circle Diameter
End Circle Diameter | End Circle Diameter
End Weight (g) | End Weight (g)
End Weight (g) | End Weight (g)
End Weight (g) | End Weight (g)
End Weight (g) | End Weight (g)
Ethanol | Ethanol
Ethanol Final | Ethanol Final
Ethylene Glycol | Ethylene Glycol
Ethylene Glycol | Ethylene Glycol
Eye spot repeat over 7 impressions | Eye spot repeat over 7 impressions
Flange - Sides | Flange - Sides
Flange - Top & Bottom | Flange - Top & Bottom
Flip Lid Force | Flip Lid Force
Flip Lid Force | Flip Lid Force
Flip Lid Force | Flip Lid Force
Flip Lid Force | Flip Lid Force
Flip Lid Force | Flip Lid Force
Flip Lid Force | Flip Lid Force
Flip Lid Force | Flip Lid Force
Flip Lid Force | Flip Lid Force
Flip Lid Force | Flip Lid Force
Foil diameter to fit Bottle(Item#2OX510B)opening | Foil diameter to fit Bottle(Item#2OX510B)opening
Front Height | Front Height
Functionality | Functionality
Functionality | Functionality
Functionality | Functionality
Functionality | Functionality
Functionality | Functionality
Functionality | Functionality
Functionality | Functionality
Functionality | Functionality
Functionality | Functionality
Functionality | Functionality
Functionality | Functionality
Functionality | Functionality
Functionality | Functionality
Functionality | Functionality
Functionality | Functionality
Functionality | Functionality
Functionality | Functionality
Functionality | Functionality
Functionality | Functionality
Functionality | Functionality
Functionality | Functionality
Functionality - Cap Closure | Functionality - Cap Closure
Functionality - Cap Closure | Functionality - Cap Closure
Functionality - Cap Closure | Functionality - Cap Closure
Functionality - Cap Closure | Functionality - Cap Closure
Functionality - Lid Fit | Functionality - Lid Fit
Functionality - Lid Fit | Functionality - Lid Fit
Functionality - Lid Fit | Functionality - Lid Fit
Functionality - Lid Fit | Functionality - Lid Fit
Functionality - Lid Fit | Functionality - Lid Fit
Functionality - Lid Fit | Functionality - Lid Fit
Functionality - Lid Fit | Functionality - Lid Fit
Functionality - Lid Fit | Functionality - Lid Fit
Functionality - Lid Fit | Functionality - Lid Fit
Functionality - Lid Fit | Functionality - Lid Fit
Functionality - Lid Fit | Functionality - Lid Fit
Functionality - Lid Fit | Functionality - Lid Fit
Functionality - Lid Fit | Functionality - Lid Fit
Functionality - Lid Fit | Functionality - Lid Fit
Functionality - Lid Fit | Functionality - Lid Fit
Functionality - Lid Fit | Functionality - Lid Fit
Fungal | Fungal
Fungal | Fungal
Fungal | Fungal
Fungal | Fungal
Fungal | Fungal
Fungal | Fungal
Fungal | Fungal
Fungal | Fungal
Fungal | Fungal
Fungal | Fungal
GC Profile | GC Profile
GC Profile | GC Profile
GC Profile | GC Profile
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
General Inspection | General Inspection
Gluconic_Acid (TM 3012) | Gluconic Acid
Gluconic_Acid_Final (TM 3012) | Gluconic Acid Final
Gram Stain_J&JCandida | Gram Stain J&JCandida
GramsStain | GramsStain
GramsStain | GramsStain
GramsStain | GramsStain
GramsStain | GramsStain
GramsStain | GramsStain
GramsStain | GramsStain
GramsStain | GramsStain
GramsStain | GramsStain
GramsStain | GramsStain
GramsStain | GramsStain
GramsStain | GramsStain
GramsStain | GramsStain
GramsStain | GramsStain
GramsStain | GramsStain
GramsStain | GramsStain
GramsStain | GramsStain
GramsStain | GramsStain
GramsStain | GramsStain
GramsStain | GramsStain
GramStain_Bioburden_TMYC | GramStain Bioburden TMYC
GramStain_Bioburden_TMYC | GramStain Bioburden TMYC
GramStain_Bioburden_TMYC | GramStain Bioburden TMYC
GramStain_Bioburden_TMYC | GramStain Bioburden TMYC
GramStain_Bioburden_TMYC | GramStain Bioburden TMYC
GramStain_Bioburden_TMYC | GramStain Bioburden TMYC
GramStain_Bioburden_TMYC | GramStain Bioburden TMYC
GramStain_Bioburden_TMYC | GramStain Bioburden TMYC
GramStain_Bioburden_TMYC | GramStain Bioburden TMYC
GramStain_Bioburden_TMYC | GramStain Bioburden TMYC
GramStain_Bioburden_TMYC | GramStain Bioburden TMYC
GramStain_BloodAgar | GramStain BloodAgar
GramStain_J&J_TSA | GramStain J&J TSA
Gussett | Gussett
Gussett | Gussett
H Dimension | H Dimension
H Dimension | H Dimension
Handle Internal Diameter | Handle Internal Diameter
Handle Internal Diameter | Handle Internal Diameter
Handle Internal Diameter | Handle Internal Diameter
HDPE Resin Identification Marker | HDPE Resin Identification Marker
HDPE Resin Identification Marker | HDPE Resin Identification Marker
HDPE Resin Identification Marker | HDPE Resin Identification Marker
HDPE Resin Identification Marker | HDPE Resin Identification Marker
HDPE Resin Identification Marker | HDPE Resin Identification Marker
HDPE Resin Identification Marker | HDPE Resin Identification Marker
HDPE Resin Identification Marker | HDPE Resin Identification Marker
HDPE Resin Identification Marker | HDPE Resin Identification Marker
HDPE Resin Identification Marker | HDPE Resin Identification Marker
HDPE Resin Identification Marker | HDPE Resin Identification Marker
HDPE Resin Identification Marker | HDPE Resin Identification Marker
Heavy Metals | Heavy Metals
Heavy Metals | Heavy Metals
Heavy Metals | Heavy Metals
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height | Height
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Height I.D. | Height I.D.
Hydroxyl Value | Hydroxyl Value
Hydroxyl Value Final | Hydroxyl Value Final
IdCoag_BloodAgar | IdCoag BloodAgar
Identification | Identification
Identification | Identification
Identification (FTIR) | Identification (FTIR)
Identification (TM 400) | Identification
Identification A | Identification A
Identification A | Identification A
Identification A | Identification A
Identification A | Identification A
Identification A | Identification A
Identification A - Chloride | Identification A - Chloride
Identification A - Sodium | Identification A - Sodium
Identification A (USP <197k>) | Identification A (USP <197k>)
Identification B | Identification B
Identification B | Identification B
Identification B | Identification B
Identification B | Identification B
Identification B | Identification B
Identification B | Identification B
Identification B | Identification B
Identification B - Benzalkonium | Identification B - Benzalkonium
Identification B - Citrate | Identification B - Citrate
Identification B (USP <201>) | Identification B (USP <201>)
Identification C | Identification C
Identification C | Identification C
Identification C | Identification C
Identification D | Identification D
Identification for Ammonium (USP <191>) | Identification for Ammonium (USP <191>)
Identification for Chloride (USP <191>) | Identification for Chloride (USP <191>)
Identification for Chloride (USP <191>) | Identification for Chloride (USP <191>)
Identification for Potassium (USP <191>) | Identification for Potassium (USP <191>)
Identification for Potassium (USP <191>) | Identification for Potassium (USP <191>)
Identification for Sodium (USP <191>) | Identification for Sodium (USP <191>)
Identification for Sodium (USP <191>) | Identification for Sodium (USP <191>)
Impurity | Impurity
Individual Impurity | Individual Impurity
Individual Impurity | Individual Impurity
Insert Fin Gap (per specification) | Insert Fin Gap (per specification)
Insert Fin Gap (per specification) | Insert Fin Gap (per specification)
Insert Fin Gap (per specification) | Insert Fin Gap (per specification)
Insert Fin Gap (per specification) | Insert Fin Gap (per specification)
Insert Fin Gap (per specification) | Insert Fin Gap (per specification)
Insert Fin Gap (per specification) | Insert Fin Gap (per specification)
Insert Fin Gap (per specification) | Insert Fin Gap (per specification)
Insert Fin Gap (per specification) | Insert Fin Gap (per specification)
Iodine Value | Iodine Value
Iodine Value | Iodine Value
Iodine Value Final | Iodine Value Final
Iodine Value Final | Iodine Value Final
IPBC | IPBC
IPBC | IPBC
IPBC Final | IPBC Final
IPBC Final | IPBC Final
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification | IR Identification
IR Identification (TM 411) | IR Identification
IR Identification (TM 411) | IR Identification
IR Identification (TM 411) | IR Identification
IR Identification (TM 411) | IR Identification
IR Identification (TM 411) | IR Identification
IR Identification (TM 411) | IR Identification
IR Identification (TM 411) | IR Identification
IR Identification (TM 411) | IR Identification
IR Identification (TM 411) | IR Identification
IR Identification (TM 411) | IR Identification
IR Identification (TM 411) | IR Identification
IR Identification (TM 411) | IR Identification
IR Identification (TM 411) | IR Identification
IR Identification (TM 411) | IR Identification
IR Identification (TM 411) | IR Identification
IR Identification (TM 411) | IR Identification
IR Identification (TM 411) | IR Identification
IR Identification (TM 411) | IR Identification
IR Identification (TM 411) | IR Identification
IR Identification (TM 411) | IR Identification
IR Identification (TM 411) | IR Identification
Iso-Menthone | Iso-Menthone
Iso-Pulegol | Iso-Pulegol
JJ Fungal_RM | JJ Fungal RM
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length | Length
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Length I.D. | Length I.D.
Lid Inside Diameter (as per specification) | Lid Inside Diameter (as per specification)
Lid Inside Diameter (as per specification) | Lid Inside Diameter (as per specification)
Lid Inside Diameter (as per specification) | Lid Inside Diameter (as per specification)
Lid Inside Diameter (as per specification) | Lid Inside Diameter (as per specification)
Lid Inside Diameter (as per specification) | Lid Inside Diameter (as per specification)
Lid Inside Diameter (as per specification) | Lid Inside Diameter (as per specification)
Lid Inside Diameter (as per specification) | Lid Inside Diameter (as per specification)
Lid Inside Diameter (as per specification) | Lid Inside Diameter (as per specification)
Lid Inside Diameter (as per specification) | Lid Inside Diameter (as per specification)
Lid Pull-Off Force | Lid Pull-Off Force
Lid Pull-Off Force | Lid Pull-Off Force
Lid Pull-Off Force | Lid Pull-Off Force
Lid Pull-Off Force | Lid Pull-Off Force
Lid Pull-Off Force | Lid Pull-Off Force
Lid Wall Thickness (per speciffication) | Lid Wall Thickness (per speciffication)
Lid Wall Thickness (per speciffication) | Lid Wall Thickness (per speciffication)
Lid Wall Thickness (per speciffication) | Lid Wall Thickness (per speciffication)
Lid Wall Thickness (per speciffication) | Lid Wall Thickness (per speciffication)
Lid Wall Thickness (per speciffication) | Lid Wall Thickness (per speciffication)
Lid Wall Thickness (per speciffication) | Lid Wall Thickness (per speciffication)
Lid Wall Thickness (per speciffication) | Lid Wall Thickness (per speciffication)
Lid Wall Thickness (per speciffication) | Lid Wall Thickness (per speciffication)
Lid Wall Thickness (per speciffication) | Lid Wall Thickness (per speciffication)
Lid Wall Thickness (per speciffication) | Lid Wall Thickness (per speciffication)
Lid Wall Thickness (per speciffication) | Lid Wall Thickness (per speciffication)
Lid Wall Thickness (per speciffication) | Lid Wall Thickness (per speciffication)
Lid Wall Thickness (per speciffication) | Lid Wall Thickness (per speciffication)
Lid Wall Thickness (per speciffication) | Lid Wall Thickness (per speciffication)
Lid Wall Thickness (per speciffication) | Lid Wall Thickness (per speciffication)
Lid Wall Thickness (per speciffication) | Lid Wall Thickness (per speciffication)
Lid Wall Thickness (per speciffication) | Lid Wall Thickness (per speciffication)
Lid Wall Thickness (per speciffication) | Lid Wall Thickness (per speciffication)
Lid Wall Thickness (per speciffication) | Lid Wall Thickness (per speciffication)
Lid Wall Thickness (per speciffication) | Lid Wall Thickness (per speciffication)
Lid Wall Thickness (per speciffication) | Lid Wall Thickness (per speciffication)
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Liner Board Color | Liner Board Color
Lip Width (for reference only) | Lip Width (for reference only)
Lip Width (for reference only) | Lip Width (for reference only)
Lip Width (for reference only) | Lip Width (for reference only)
Loss on Drying | Loss on Drying
Loss on Drying | Loss on Drying
Loss on Drying | Loss on Drying
Loss on Drying | Loss on Drying
Loss on Drying | Loss on Drying
Loss on Drying | Loss on Drying
Loss on Drying (80C Vacuum 1 hour) | Loss on Drying (80C Vacuum 1 hour)
Loss on Drying (80C Vacuum 1 hour) Final | Loss on Drying (80C Vacuum 1 hour) Final
Loss on Drying Final | Loss on Drying Final
Loss on Drying Final | Loss on Drying Final
Loss on Drying Final | Loss on Drying Final
Loss on Drying Final | Loss on Drying Final
Loss on Drying Final | Loss on Drying Final
MacConkey | MacConkey
MacConkey | MacConkey
MacConkey | MacConkey
MacConkey | MacConkey
MacConkey | MacConkey
MacConkey | MacConkey
MacConkey Broth | MacConkey Broth
MacConkey Broth | MacConkey Broth
Melting Range - End | Melting Range - End
Melting Range - End | Melting Range - End
Melting Range - Start | Melting Range - Start
Melting Range - Start | Melting Range - Start
Menthone | Menthone
Mitt_Length_(Inches) | Mitt Length (Inches)
Mitt_Length_(Inches) | Mitt Length (Inches)
Mitt_Length_(Inches) | Mitt Length (Inches)
Mitt_Thickness_(Inches) | Mitt Thickness (Inches)
Mitt_Thickness_(Inches) | Mitt Thickness (Inches)
Mitt_Width_(Inches) | Mitt Width (Inches)
Mitt_Width_(Inches) | Mitt Width (Inches)
Mitt_Width_(Inches) | Mitt Width (Inches)
Moisture Content | Moisture Content
Moisture Content | Moisture Content
Moisture Content (10 grams of sample) | Moisture Content (10 grams of sample)
Mold | Mold
Mold | Mold
Mold Count (Observed) | Mold Count (Observed)
Mold Count (Observed) | Mold Count (Observed)
Mold Count (Reported) | Mold Count (Reported)
Mold Count (Reported) | Mold Count (Reported)
MSA or VJ | MSA or VJ
MSA or VJ | MSA or VJ
MSA or VJ | MSA or VJ
MSA or VJ | MSA or VJ
MSA or VJ | MSA or VJ
MSA or VJ | MSA or VJ
Negative Control | Negative Control
Negative Control | Negative Control
Negative Control | Negative Control
Negative Control | Negative Control
Negative Control | Negative Control
Negative Control | Negative Control
Negative Control | Negative Control
Negative Control | Negative Control
Negative Control | Negative Control
Negative Control | Negative Control
Negative Control | Negative Control
NegativeControl | NegativeControl
NegativeControl | NegativeControl
NegativeControl | NegativeControl
NegativeControl | NegativeControl
NegativeControl | NegativeControl
NegativeControl | NegativeControl
NegativeControl | NegativeControl
NegativeControl | NegativeControl
NegativeControl | NegativeControl
NegativeControl | NegativeControl
NegativeControl | NegativeControl
Neo-Menthol | Neo-Menthol
Nitrilotriacetic Acid | Nitrilotriacetic Acid
Nonvolatile Residue | Nonvolatile Residue
Nonvolatile Residue (calc) | Nonvolatile Residue (calc)
Nonvolatile Residue (calc) | Nonvolatile Residue (calc)
Nonvolatile Residue (calc) | Nonvolatile Residue (calc)
Nonvolatile Residue Final (calc) | Nonvolatile Residue Final (calc)
Nonvolatile Residue Final (calc) | Nonvolatile Residue Final (calc)
Nonvolatile Residue Final (calc) | Nonvolatile Residue Final (calc)
Norphenylephrine | Norphenylephrine
NP Structure | NP Structure
o-Chlorhexidine and specified unidentified impurit | o-Chlorhexidine and specified unidentified impurit
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor | Odor
Odor (CoA) | Odor (CoA)
Odor (J&J TM-003188) | Odor
Odor (J&J TM-003188) | Odor
Odor (J&J TM-003188) | Odor
Odor (J&J TM-003188) | Odor
Odor (J&J TM-003188) | Odor
Odor (J&J TM-003188) | Odor
Odor (J&J TM-003188) | Odor
Odor (J&J TM-003188) | Odor
Odor (J&J TM-003188) | Odor
Odor (J&J TM-003188) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Odor (TM 208) | Odor
Open Lid Overall Width | Open Lid Overall Width
Open Lid Overall Width | Open Lid Overall Width
Open Lid Overall Width | Open Lid Overall Width
Open Lid Overall Width | Open Lid Overall Width
Open Lid Overall Width | Open Lid Overall Width
Open Lid Overall Width | Open Lid Overall Width
Open Lid Overall Width | Open Lid Overall Width
Open Lid Overall Width | Open Lid Overall Width
Optical Rotation | Optical Rotation
Orifice Diameter | Orifice Diameter
Orifice Diameter | Orifice Diameter
Orifice Diameter | Orifice Diameter
Orifice Flap Thickness (per specification) | Orifice Flap Thickness (per specification)
Orifice Flap Thickness (per specification) | Orifice Flap Thickness (per specification)
Orifice Flap Thickness (per specification) | Orifice Flap Thickness (per specification)
Orifice Flap Thickness (per specification) | Orifice Flap Thickness (per specification)
Orifice Flap Thickness (per specification) | Orifice Flap Thickness (per specification)
Orifice Flap Thickness (per specification) | Orifice Flap Thickness (per specification)
Orifice Flap Thickness (per specification) | Orifice Flap Thickness (per specification)
Orifice Flap Thickness (per specification) | Orifice Flap Thickness (per specification)
Orifice Flap Thickness (per specification) | Orifice Flap Thickness (per specification)
Outside Flap | Outside Flap
Overall closed lid height | Overall closed lid height
Overall closed lid height | Overall closed lid height
Overall closed lid height | Overall closed lid height
Overall closed lid height | Overall closed lid height
Overall closed lid height | Overall closed lid height
Overall closed lid height | Overall closed lid height
Overall closed lid height | Overall closed lid height
Overall closed lid height | Overall closed lid height
Overall Length | Overall Length
Overall Length | Overall Length
Overall Length | Overall Length
Overall Length | Overall Length
Overall Length | Overall Length
Overall Lid Height (per specification) | Overall Lid Height (per specification)
Overall Lid Height (per specification) | Overall Lid Height (per specification)
Overall Lid Height (per specification) | Overall Lid Height (per specification)
Overall Lid Height (per specification) | Overall Lid Height (per specification)
Overall Lid Height (per specification) | Overall Lid Height (per specification)
Overall Lid Height (per specification) | Overall Lid Height (per specification)
Overall Lid Height (per specification) | Overall Lid Height (per specification)
Overall Lid Height (per specification) | Overall Lid Height (per specification)
Overall Lid Height (per specification) | Overall Lid Height (per specification)
Overall Lid Height (per specification) | Overall Lid Height (per specification)
Overall Lid Height (per specification) | Overall Lid Height (per specification)
Overall Lid Height (per specification) | Overall Lid Height (per specification)
Overall Lid Height (per specification) | Overall Lid Height (per specification)
Overall_Swab_Length_(Inches) | Overall Swab Length (Inches)
Overall_Swab_Length_(Inches) | Overall Swab Length (Inches)
Overall_Swab_Length_(Inches) | Overall Swab Length (Inches)
Overall_Swab_Length_(Inches) | Overall Swab Length (Inches)
Oxochlorhexidine | Oxochlorhexidine
p-Chloroaniline | p-Chloroaniline
p-Chloroaniline Final | p-Chloroaniline Final
p-Chlorophenvl urea | p-Chlorophenvl urea
P. aeruginosa | P. aeruginosa
P. aeruginosa | P. aeruginosa
Packaging Damage or Contamination | Packaging Damage or Contamination
Packaging Damage or Contamination | Packaging Damage or Contamination
Packaging Damage or Contamination | Packaging Damage or Contamination
Packaging Damage or Contamination | Packaging Damage or Contamination
Packaging Damage or Contamination | Packaging Damage or Contamination
Packaging Damage or Contamination | Packaging Damage or Contamination
Packaging Damage or Contamination | Packaging Damage or Contamination
Packaging Damage or Contamination | Packaging Damage or Contamination
Packaging Damage or Contamination | Packaging Damage or Contamination
Packaging Damage or Contamination | Packaging Damage or Contamination
Packaging Damage or Contamination | Packaging Damage or Contamination
Packaging Damage or Contamination | Packaging Damage or Contamination
Packaging Damage or Contamination | Packaging Damage or Contamination
Packaging Damage or Contamination | Packaging Damage or Contamination
Packaging Damage or Contamination | Packaging Damage or Contamination
Packaging Damage or Contamination | Packaging Damage or Contamination
Packaging Damage or Contamination | Packaging Damage or Contamination
Packaging Damage or Contamination | Packaging Damage or Contamination
Packaging Damage or Contamination | Packaging Damage or Contamination
Packaging Damage or Contamination | Packaging Damage or Contamination
Packaging Damage or Contamination | Packaging Damage or Contamination
Packaging Damage or Contamination | Packaging Damage or Contamination
Packaging Damage or Contamination | Packaging Damage or Contamination
Packaging Damage or Contamination | Packaging Damage or Contamination
Packaging Damage or Contamination | Packaging Damage or Contamination
Packaging Damage or Contamination | Packaging Damage or Contamination
Packaging Damage or Contamination | Packaging Damage or Contamination
Packaging Damage or Contamination | Packaging Damage or Contamination
Packaging Damage or Contamination | Packaging Damage or Contamination
Packaging Damage or Contamination | Packaging Damage or Contamination
Packaging Damage or Contamination | Packaging Damage or Contamination
Pad Diameter | Pad Diameter
Peroxide Value | Peroxide Value
pH | pH
pH | pH
pH | pH
pH | pH
pH | pH
pH | pH
pH | pH
pH | pH
pH | pH
pH | pH
pH | pH
pH | pH
pH | pH
pH | pH
pH | pH
pH | pH
pH | pH
pH | pH
pH | pH
pH | pH
pH | pH
pH | pH
pH | pH
pH | pH
pH | pH
pH (1 in 20 solution) | pH (1 in 20 solution)
pH (1 in 20 solution) | pH (1 in 20 solution)
pH (1 in 20 solution) | pH (1 in 20 solution)
pH (1% Solution) | pH (1% Solution)
pH (1% Solution) | pH (1% Solution)
pH (1% w/v aqueous solution) | pH (1% w/v aqueous solution)
pH (1% w/v aqueous solution) | pH (1% w/v aqueous solution)
pH (1% w/v aqueous solution) | pH (1% w/v aqueous solution)
pH (1%) in 10:6 IPA: H2O | pH (1%) in 10:6 IPA: H2O
pH (10% in water) | pH (10% in water)
pH (10% in water) | pH (10% in water)
pH (10% in water) | pH (10% in water)
pH (10% in water) | pH (10% in water)
pH (10% Solution) | pH (10% Solution)
pH (10% Solution) | pH (10% Solution)
pH (10% Solution) | pH (10% Solution)
pH (10% w/v aqueous solution) | pH (10% w/v aqueous solution)
pH (5% aqueous @25 degree C) | pH (5% aqueous @25 degree C)
pH (5% Solution in DI Water:Methanol 1:1) | pH (5% Solution in DI Water:Methanol 1:1)
pH (5% Solution) | pH (5% Solution)
pH (Diluted 1 in 20) | pH (Diluted 1 in 20)
pH (TM 200) | pH
pH (TM 200) | pH
pH (TM 200) | pH
pH (TM 200) | pH
pH (TM 200) | pH
pH (TM 200) | pH
Phenol | Phenol
Phenol | Phenol
Phenol | Phenol
Phenoxyethanol | Phenoxyethanol
Phenoxyethanol | Phenoxyethanol
Phenoxyethanol | Phenoxyethanol
Phenoxyethanol | Phenoxyethanol
Phenoxyethanol Assay | Phenoxyethanol Assay
Phenoxyethanol Final | Phenoxyethanol Final
Phenoxyethanol Final | Phenoxyethanol Final
Phenoxyethanol_Final | Phenoxyethanol Final
Phenylephrine Related Compound C | Phenylephrine Related Compound C
Phenylephrine Related Compound D | Phenylephrine Related Compound D
Phenylephrine Related Compound E | Phenylephrine Related Compound E
Physical Properties | Physical Properties
Physical Properties | Physical Properties
Physical Properties | Physical Properties
Pop-Up Hole Outer Diameter (per specification) | Pop-Up Hole Outer Diameter (per specification)
Pop-Up Hole Outer Diameter (per specification) | Pop-Up Hole Outer Diameter (per specification)
Pop-Up Hole Outer Diameter (per specification) | Pop-Up Hole Outer Diameter (per specification)
Pop-Up Hole Outer Diameter (per specification) | Pop-Up Hole Outer Diameter (per specification)
Pop-Up Hole Outer Diameter (per specification) | Pop-Up Hole Outer Diameter (per specification)
Pop-Up Hole Outer Diameter (per specification) | Pop-Up Hole Outer Diameter (per specification)
Pop-Up Hole Outer Diameter (per specification) | Pop-Up Hole Outer Diameter (per specification)
Pop-Up Hole Outer Diameter (per specification) | Pop-Up Hole Outer Diameter (per specification)
Pop-Up Hole Thicknes | Pop-Up Hole Thicknes
Pop-Up Hole Thicknes | Pop-Up Hole Thicknes
Pop-Up Hole Thicknes | Pop-Up Hole Thicknes
Pop-Up Hole Thicknes | Pop-Up Hole Thicknes
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Print | Print
Pulegone | Pulegone
Purity | Purity
Quat | Quat
Quat | Quat
Quat | Quat
Quat | Quat
Quat | Quat
Quat | Quat
Quat_Final | Quat Final
Quat_Final | Quat Final
Quat_Final | Quat Final
Quat_Final | Quat Final
Quat_Final | Quat Final
Quat_Final | Quat Final
Ratio: Wet MD / Wet CD (TM 103) | Ratio: Wet MD / Wet CD
Ratio: Wet MD / Wet CD (TM 103) | Ratio: Wet MD / Wet CD
Reducing Sugars | Reducing Sugars
Refractive Index | Refractive Index
Refractive Index | Refractive Index
Refractive Index | Refractive Index
Refractive Index | Refractive Index
Refractive Index | Refractive Index
Refractive Index | Refractive Index
Refractive Index | Refractive Index
Refractive Index | Refractive Index
Refractive Index | Refractive Index
Refractive Index | Refractive Index
Refractive Index | Refractive Index
Refractive Index | Refractive Index
Refractive Index | Refractive Index
Refractive Index | Refractive Index
Refractive Index | Refractive Index
Refractive Index | Refractive Index
Refractive Index | Refractive Index
Refractive Index at 20°C | Refractive Index at 20°C
Refractive Index at 20°C | Refractive Index at 20°C
Refractive Index at 20°C | Refractive Index at 20°C
Refractive Index at 20°C | Refractive Index at 20°C
Refractive Index at 20°C | Refractive Index at 20°C
Refractive Index at 20°C | Refractive Index at 20°C
Refractive Index at 20°C | Refractive Index at 20°C
Refractive Index at 20°C | Refractive Index at 20°C
Refractive Index at 20°C | Refractive Index at 20°C
Refractive Index at 20°C | Refractive Index at 20°C
Refractive Index at 20°C | Refractive Index at 20°C
Refractive Index at 20°C | Refractive Index at 20°C
Refractive Index at 20°C | Refractive Index at 20°C
Refractive Index at 20°C | Refractive Index at 20°C
Refractive Index at 20°C | Refractive Index at 20°C
Refractive Index at 20°C | Refractive Index at 20°C
Refractive Index at 20°C | Refractive Index at 20°C
Refractive Index at 20°C | Refractive Index at 20°C
Related Compound A | Related Compound A
Related Compound B | Related Compound B
Report result from COA | Report result from COA
Residual Solvents | Residual Solvents
Residue on Ignition | Residue on Ignition
Residue on Ignition | Residue on Ignition
Residue on Ignition | Residue on Ignition
Roll Diameter (COA) | Roll Diameter (COA)
Roll Diameter (COA) | Roll Diameter (COA)
Roll Diameter (COA) | Roll Diameter (COA)
Roll Diameter (COA) | Roll Diameter (COA)
Roll Diameter (COA) | Roll Diameter (COA)
Roll Diameter (COA) | Roll Diameter (COA)
Roll Diameter (COA) | Roll Diameter (COA)
Roll Diameter (Inches) | Roll Diameter (Inches)
Roll Diameter (Inches) | Roll Diameter (Inches)
Roll Diameter (Inches) | Roll Diameter (Inches)
Roll Diameter (Inches) | Roll Diameter (Inches)
Roll Diameter (Inches) | Roll Diameter (Inches)
Roll Diameter (Inches) | Roll Diameter (Inches)
Roll Diameter (Inches) | Roll Diameter (Inches)
Roll Diameter (Inches) | Roll Diameter (Inches)
Roll Diameter (Inches) | Roll Diameter (Inches)
Roll Diameter (Inches) | Roll Diameter (Inches)
Roll Diameter (Inches) | Roll Diameter (Inches)
Roll Diameter (Inches) | Roll Diameter (Inches)
Roll Diameter (Inches) | Roll Diameter (Inches)
Roll Diameter (Inches) | Roll Diameter (Inches)
Roll Diameter (Inches) | Roll Diameter (Inches)
Roll Diameter (Inches) | Roll Diameter (Inches)
Roll Diameter (Inches) | Roll Diameter (Inches)
Roll Diameter (Inches) | Roll Diameter (Inches)
Roll Diameter (Inches) | Roll Diameter (Inches)
Roll Diameter (Inches) | Roll Diameter (Inches)
Roll Diameter (Inches) | Roll Diameter (Inches)
Roll Diameter (Inches) (Measurement) | Roll Diameter (Inches) (Measurement)
Roll Diameter (Inches) (Measurement) | Roll Diameter (Inches) (Measurement)
Roll Diameter (Inches) (Measurement) | Roll Diameter (Inches) (Measurement)
Roll Diameter (Inches) (Measurement) | Roll Diameter (Inches) (Measurement)
Roll Diameter (Inches) (Measurement) | Roll Diameter (Inches) (Measurement)
Roll Diameter (Inches) (Measurement) | Roll Diameter (Inches) (Measurement)
Roll Diameter (Inches) (Measurement) | Roll Diameter (Inches) (Measurement)
Roll Diameter (Inches) (Measurement) | Roll Diameter (Inches) (Measurement)
Roll Diameter (Inches) (Measurement) | Roll Diameter (Inches) (Measurement)
Roll Diameter (Inches) (Measurement) | Roll Diameter (Inches) (Measurement)
Roll Diameter (Inches) (Measurement) | Roll Diameter (Inches) (Measurement)
Roll Diameter (Inches) (Measurement) | Roll Diameter (Inches) (Measurement)
Roll Diameter (Inches) (Measurement) | Roll Diameter (Inches) (Measurement)
Roll OD | Roll OD
Roll OD | Roll OD
Roll width (COA) | Roll width (COA)
Roll width (COA) | Roll width (COA)
Roll width (COA) | Roll width (COA)
Roll width (COA) | Roll width (COA)
Roll width (COA) | Roll width (COA)
Roll width (COA) | Roll width (COA)
Roll Width (Inches) | Roll Width (Inches)
Roll Width (Inches) | Roll Width (Inches)
Roll Width (Inches) | Roll Width (Inches)
Roll Width (Inches) | Roll Width (Inches)
Roll Width (Inches) | Roll Width (Inches)
Roll Width (Inches) | Roll Width (Inches)
Roll Width (Inches) | Roll Width (Inches)
Roll Width (Inches) | Roll Width (Inches)
Roll Width (Inches) | Roll Width (Inches)
Roll Width (Inches) | Roll Width (Inches)
Roll Width (Inches) | Roll Width (Inches)
Roll Width (Inches) | Roll Width (Inches)
Roll Width (Inches) | Roll Width (Inches)
Roll Width (Inches) | Roll Width (Inches)
Roll Width (Inches) | Roll Width (Inches)
Roll Width (Inches) | Roll Width (Inches)
Roll Width (Inches) | Roll Width (Inches)
Roll Width (Inches) | Roll Width (Inches)
Roll Width (Inches) | Roll Width (Inches)
Roll Width (Inches) | Roll Width (Inches)
Roll Width (Inches) | Roll Width (Inches)
Roll Width (Inches) | Roll Width (Inches)
Roll Width (Inches) | Roll Width (Inches)
Roll Width (Inches) | Roll Width (Inches)
Roll Width (Inches) | Roll Width (Inches)
Roll Width (Inches) | Roll Width (Inches)
Roll Width (Inches) | Roll Width (Inches)
Roll Width (Inches) | Roll Width (Inches)
Roll Width (Inches) (Measurement) | Roll Width (Inches) (Measurement)
Roll Width (Inches) (Measurement) | Roll Width (Inches) (Measurement)
Roll Width (Inches) (Measurement) | Roll Width (Inches) (Measurement)
Roll Width (Inches) (Measurement) | Roll Width (Inches) (Measurement)
Roll Width (Inches) (Measurement) | Roll Width (Inches) (Measurement)
Roll Width (Inches) (Measurement) | Roll Width (Inches) (Measurement)
Roll Width (Inches) (Measurement) | Roll Width (Inches) (Measurement)
Roll Width (Inches) (Measurement) | Roll Width (Inches) (Measurement)
Roll Width (Inches) (Measurement) | Roll Width (Inches) (Measurement)
Roll Width (Inches) (Measurement) | Roll Width (Inches) (Measurement)
Roll Width (Inches) (Measurement) | Roll Width (Inches) (Measurement)
Roll Width (mm) | Roll Width (mm)
Roll Width (mm) (Measurement) | Roll Width (mm) (Measurement)
Roll Width (mm) (Measurement) | Roll Width (mm) (Measurement)
RRT of Any individual unspecified impuritv | RRT of Any individual unspecified impuritv
RRT of Ch!orhexidine oxazinone analog | RRT of Ch!orhexidine oxazinone analog
RRT of Chlorhexidine | RRT of Chlorhexidine
RRT of Chlorhexidine urea | RRT of Chlorhexidine urea
RRT of Chlorhexidine amine | RRT of Chlorhexidine amine
RRT of Chlorhexidine dimer | RRT of Chlorhexidine dimer
RRT of Chlorhexidine glucityl triazine | RRT of Chlorhexidine glucityl triazine
RRT of Chlorhexidine guanidine | RRT of Chlorhexidine guanidine
RRT of Chlorhexidine nitrile | RRT of Chlorhexidine nitrile
RRT of o-Chlorhexidine and specified unidentified | RRT of o-Chlorhexidine and specified unidentified
RRT of Oxochlorhexidine | RRT of Oxochlorhexidine
RRT of p-Chlorophenvl urea | RRT of p-Chlorophenvl urea
RRT of Specified unidentified impuritv 1 | RRT of Specified unidentified impuritv 1
S Dimension | S Dimension
S Dimension | S Dimension
S.aureus | S.aureus
S.aureus | S.aureus
Salicylic Acid | Salicylic Acid
Salicylic Acid Final | Salicylic Acid Final
Salicylic_Acid | Salicylic Acid
Salicylic_Acid_Final | Salicylic Acid Final
Sample Weight (g) | Sample Weight (g)
Saponification Value | Saponification Value
Saponification Value | Saponification Value
Saponification Value | Saponification Value
Saponification Value Final | Saponification Value Final
Saponification Value Final | Saponification Value Final
Saponification Value Final | Saponification Value Final
Scrub_Test | Scrub Test
Scrub_Test | Scrub Test
Scrub_Test | Scrub Test
Scrub_Test | Scrub Test
Seal Flange Diameter | Seal Flange Diameter
Seal Flange Diameter | Seal Flange Diameter
Seal Flange Diameter | Seal Flange Diameter
Seal Strength (@300DegF/40psi/0.5 sec dewll) | Seal Strength (@300DegF/40psi/0.5 sec dewll)
Seal Strength (g/inch@130DegC/0.5 sec/1 bar) | Seal Strength (g/inch@130DegC/0.5 sec/1 bar)
Seal Strength (g/inch@130DegC/0.5 sec/1 bar) | Seal Strength (g/inch@130DegC/0.5 sec/1 bar)
Seal Strength (g/inch@130DegC/14psi/1.0 sec dwell) | Seal Strength (g/inch@130DegC/14psi/1.0 sec dwell)
Seal Strength (g/inch@130DegC/14psi/1.0 sec dwell) | Seal Strength (g/inch@130DegC/14psi/1.0 sec dwell)
Seal Strength (g/inch@130DegC/14psi/1.0 sec dwell) | Seal Strength (g/inch@130DegC/14psi/1.0 sec dwell)
Seal Strength (g/inch@140C/14psi/1 s dwellTopJaw) | Seal Strength (g/inch@140C/14psi/1 s dwellTopJaw)
Seal Strength (g/inch@250DegF/30psi/0.5 sec dwell) | Seal Strength (g/inch@250DegF/30psi/0.5 sec dwell)
Seal Strength (g/inch@250DegF/30psi/0.5 sec dwell) | Seal Strength (g/inch@250DegF/30psi/0.5 sec dwell)
Seal Strength (g/inch@290DegF/40psi/0.5 sec dwell) | Seal Strength (g/inch@290DegF/40psi/0.5 sec dwell)
Seal Strength (g/inch@290DegF/40psi/0.5 sec dwell) | Seal Strength (g/inch@290DegF/40psi/0.5 sec dwell)
Seal Strength (g/inch@320DegF/40psi/0.5 sec dwell) | Seal Strength (g/inch@320DegF/40psi/0.5 sec dwell)
Seal Strength (g/inch@320DegF/40psi/0.5 sec dwell) | Seal Strength (g/inch@320DegF/40psi/0.5 sec dwell)
Seal Strength (g/inch@320DegF/40psi/0.5 sec dwell) | Seal Strength (g/inch@320DegF/40psi/0.5 sec dwell)
Seal Strength (g/mm@130DegC/1bar/0.5sec dwell) | Seal Strength (g/mm@130DegC/1bar/0.5sec dwell)
Seal Strength (g/mm@130DegC/1bar/0.5sec dwell) | Seal Strength (g/mm@130DegC/1bar/0.5sec dwell)
Seal strength 2500g/in@140DegC/1sec/14psi | Seal strength 2500g/in@140DegC/1sec/14psi
Seal strength 2500g/in@140DegC/1sec/14psi | Seal strength 2500g/in@140DegC/1sec/14psi
Seal Strength(g/in@350DegF/40psi/0.5sec dwell) | Seal Strength(g/in@350DegF/40psi/0.5sec dwell)
Seal Strength(g/in@350F/30psi/0.5s dwell Top Jaw) | Seal Strength(g/in@350F/30psi/0.5s dwell Top Jaw)
Seal Strength(g/in@350F/30psi/0.5s dwell Top Jaw) | Seal Strength(g/in@350F/30psi/0.5s dwell Top Jaw)
Seal Strength(g/in@350F/30psi/0.5s dwell Top Jaw) | Seal Strength(g/in@350F/30psi/0.5s dwell Top Jaw)
Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300DegF/30psi/1.5 sec dwell)
Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300DegF/30psi/1.5 sec dwell)
Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300DegF/30psi/1.5 sec dwell)
Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300DegF/30psi/1.5 sec dwell)
Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300DegF/30psi/1.5 sec dwell)
Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300DegF/30psi/1.5 sec dwell)
Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300DegF/30psi/1.5 sec dwell)
Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300DegF/30psi/1.5 sec dwell)
Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300DegF/30psi/1.5 sec dwell)
Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300DegF/30psi/1.5 sec dwell)
Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300DegF/30psi/1.5 sec dwell)
Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300DegF/30psi/1.5 sec dwell)
Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300DegF/30psi/1.5 sec dwell)
Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300DegF/30psi/1.5 sec dwell)
Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300DegF/30psi/1.5 sec dwell)
Seal_Strength_(@300DegF/40psi/0.5 sec dewll) | Seal Strength (@300DegF/40psi/0.5 sec dewll)
Seal_Strength_(@300DegF/40psi/0.5 sec dewll) | Seal Strength (@300DegF/40psi/0.5 sec dewll)
Seal_Strength_(@300DegF/40psi/0.5 sec dewll) | Seal Strength (@300DegF/40psi/0.5 sec dewll)
Seal_Strength_(@300DegF/40psi/0.5 sec dewll) | Seal Strength (@300DegF/40psi/0.5 sec dewll)
Seal_Strength_(@300DegF/40psi/0.5 sec dewll) | Seal Strength (@300DegF/40psi/0.5 sec dewll)
Seal_Strength_(@300DegF/40psi/0.5 sec dewll) | Seal Strength (@300DegF/40psi/0.5 sec dewll)
Seal_Strength_(@325DegF/40PSI/0.5Sec_Dwell) | Seal Strength (@325DegF/40PSI/0.5Sec Dwell)
Seal_Strength_(@325DegF/40PSI/0.5Sec_Dwell) | Seal Strength (@325DegF/40PSI/0.5Sec Dwell)
Seal_Strength_(@325DegF/40PSI/0.5Sec_Dwell) | Seal Strength (@325DegF/40PSI/0.5Sec Dwell)
Seal_Strength_(@325DegF/40PSI/0.5Sec_Dwell) | Seal Strength (@325DegF/40PSI/0.5Sec Dwell)
Seal_Strength_(@325DegF/40PSI/0.5Sec_Dwell) | Seal Strength (@325DegF/40PSI/0.5Sec Dwell)
Seal_Strength_(@325DegF/40PSI/0.5Sec_Dwell) | Seal Strength (@325DegF/40PSI/0.5Sec Dwell)
Seal_Strength_(@350DegF/40psi/1.0 sec dwell) | Seal Strength (@350DegF/40psi/1.0 sec dwell)
Seal_Strength_(@350DegF/40psi/1.0 sec dwell) | Seal Strength (@350DegF/40psi/1.0 sec dwell)
Seal_Strength_(@350DegF/40psi/1.0 sec dwell) | Seal Strength (@350DegF/40psi/1.0 sec dwell)
Seal_Strength_(@350DegF/40psi/1.0 sec dwell) | Seal Strength (@350DegF/40psi/1.0 sec dwell)
Seal_Strength_(@350DegF/40psi/1.0 sec dwell) | Seal Strength (@350DegF/40psi/1.0 sec dwell)
Seal_Strength_(@350DegF/40psi/1.0 sec dwell) | Seal Strength (@350DegF/40psi/1.0 sec dwell)
Seal_Strength_(@350DegF/40psi/1.0 sec dwell) | Seal Strength (@350DegF/40psi/1.0 sec dwell)
Seal_Strength_(@350DegF/40psi/1.0 sec dwell) | Seal Strength (@350DegF/40psi/1.0 sec dwell)
Seal_Strength_(@350DegF/50psi/1.0 sec dwell) | Seal Strength (@350DegF/50psi/1.0 sec dwell)
Seal_Strength_(@350DegF/50psi/1.0 sec dwell) | Seal Strength (@350DegF/50psi/1.0 sec dwell)
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
SG | SG
Shipment Temperatutre | Shipment Temperatutre
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Shipper Style | Shipper Style
Sodium Benzoate | Sodium Benzoate
Sodium Benzoate | Sodium Benzoate
Sodium Benzoate | Sodium Benzoate
Sodium Benzoate | Sodium Benzoate
Sodium Benzoate | Sodium Benzoate
Sodium Benzoate Final | Sodium Benzoate Final
Sodium Benzoate Final | Sodium Benzoate Final
Sodium Benzoate Final | Sodium Benzoate Final
Sodium Benzoate Final | Sodium Benzoate Final
Sodium Benzoate Final | Sodium Benzoate Final
Sodium_Benzoate | Sodium Benzoate
Sodium_Benzoate (TM 395) | Sodium Benzoate
Sodium_Benzoate_Final | Sodium Benzoate Final
Sodium_Benzoate_Final (TM 395) | Sodium Benzoate Final
Solubility | Solubility
Solubility (1% w/v aqueous solution) | Solubility (1% w/v aqueous solution)
Specific Gravity | Specific Gravity
Specific Gravity | Specific Gravity
Specific Gravity @ 15.56°C | Specific Gravity @ 15.56°C
Specific Rotation | Specific Rotation
Specified unidentified impuritv 1 | Specified unidentified impuritv 1
Spectrophotometric | Spectrophotometric
Spectrophotometric | Spectrophotometric
Stack Count | Stack Count
Start Weight (g) | Start Weight (g)
Start Weight (g) | Start Weight (g)
Start Weight (g) | Start Weight (g)
Start Weight (g) | Start Weight (g)
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials | Sterility Controls Raw Materials
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA
Stick Diameter | Stick Diameter
Sulfate | Sulfate
Sulfate | Sulfate
Sulfate Content | Sulfate Content
T Diameter | T Diameter
T Diameter | T Diameter
T Diameter | T Diameter
TAMC (Average per 10mL) | TAMC (Average per 10mL)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Device) | TAMC (Average per Device)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator)
TAMC (Average per Gram) | TAMC (Average per Gram)
TAMC (Average per Gram) | TAMC (Average per Gram)
TAMC (Average per Gram) | TAMC (Average per Gram)
TAMC (Average per Gram) | TAMC (Average per Gram)
TAMC (Average per Gram) | TAMC (Average per Gram)
TAMC (Average per Gram) | TAMC (Average per Gram)
TAMC (Average per mL - Undiluted) | TAMC (Average per mL - Undiluted)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
TAMC (Individual) | TAMC (Individual)
Temptale Reading | Temptale Reading
Thickness | Thickness
Thickness | Thickness
Thickness | Thickness
Thickness | Thickness
Thickness | Thickness
Thickness | Thickness
Thickness | Thickness
Titrant End Volume (mL) | Titrant End Volume (mL)
Titrant End Volume (mL) | Titrant End Volume (mL)
Titrant End Volume (mL) | Titrant End Volume (mL)
Titrant End Volume (mL) | Titrant End Volume (mL)
Titrant Start Volume (mL) | Titrant Start Volume (mL)
Titrant Start Volume (mL) | Titrant Start Volume (mL)
Titrant Start Volume (mL) | Titrant Start Volume (mL)
Titrant Start Volume (mL) | Titrant Start Volume (mL)
TMYC (Average per 10mL) | TMYC (Average per 10mL)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Device) | TMYC (Average per Device)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator)
TMYC (Average per Gram) | TMYC (Average per Gram)
TMYC (Average per Gram) | TMYC (Average per Gram)
TMYC (Average per Gram) | TMYC (Average per Gram)
TMYC (Average per Gram) | TMYC (Average per Gram)
TMYC (Average per Gram) | TMYC (Average per Gram)
TMYC (Average per mL - Undiluted) | TMYC (Average per mL - Undiluted)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
TMYC (Individual) | TMYC (Individual)
Top load Compression Strength (0.5"/mSR,0.450"TD) | Top load Compression Strength (0.5"/mSR,0.450"TD)
Top load Compression Strength (0.5"/mSR,0.450"TD) | Top load Compression Strength (0.5"/mSR,0.450"TD)
Top load Compression Strength (0.5"/mSR,0.450"TD) | Top load Compression Strength (0.5"/mSR,0.450"TD)
Top load Compression Strength (0.5"/mSR,0.450"TD) | Top load Compression Strength (0.5"/mSR,0.450"TD)
Top load Compression Strength (0.5"/mSR,0.450"TD) | Top load Compression Strength (0.5"/mSR,0.450"TD)
Top load Compression Strength (0.5"/mSR,0.450"TD) | Top load Compression Strength (0.5"/mSR,0.450"TD)
Total Height | Total Height
Total Impurities | Total Impurities
Total Impurities | Total Impurities
Total Impurities | Total Impurities
Total Impurities | Total Impurities
Total Impurities | Total Impurities
Total Impurities | Total Impurities
Total non-volatile reside (calc) | Total non-volatile reside (calc)
Total Solids (2 g for 3 hr @ 105°C) | Total Solids (2 g for 3 hr @ 105°C)
Total Solids Final (2 g for 3 hr @ 105°C) | Total Solids Final (2 g for 3 hr @ 105°C)
Total Sugar | Total Sugar
Total Volume of Titrant (mL) | Total Volume of Titrant (mL)
TSA | TSA
TSA | TSA
TSA | TSA
TSA | TSA
TSA | TSA
TSA | TSA
TSA JJ | TSA JJ
TSA JJ | TSA JJ
TSA JJ | TSA JJ
TSA JJ | TSA JJ
TSA JJ | TSA JJ
TSA JJ | TSA JJ
TSA JJ | TSA JJ
TSA JJ | TSA JJ
TSA JJ | TSA JJ
TSA JJ | TSA JJ
TSA JJ RM | TSA JJ RM
TSB | TSB
TSB | TSB
TSB | TSB
TSB | TSB
TSB | TSB
TSB | TSB
TSB | TSB
TSB | TSB
Tuft_Adhesion_(Lb.) | Tuft Adhesion (Lb.)
Tuft_Adhesion_(Lb.) | Tuft Adhesion (Lb.)
Tuft_Adhesion_(Lb.) | Tuft Adhesion (Lb.)
Tuft_Diameter_(Inches) | Tuft Diameter (Inches)
Unknown impurities | Unknown impurities
UPC Verification | UPC Verification
Use_Test-Solvents | Use Test-Solvents
Use_Test-Solvents | Use Test-Solvents
Use_Test-Solvents | Use Test-Solvents
Verify KC Part Number | Verify KC Part Number
Viscosity | Viscosity
Viscosity (Spindle 3 @ 60RPM in 300 ml beaker) | Viscosity (Spindle 3 @ 60RPM in 300 ml beaker)
Viscosity (Spindle 7@20RPM@30 sec 0.5% solution) | Viscosity (Spindle 7@20RPM@30 sec 0.5% solution)
Wall Thickness (Bottom base area) | Wall Thickness (Bottom base area)
Wall Thickness (Bottom base area) | Wall Thickness (Bottom base area)
Wall Thickness (Bottom base area) | Wall Thickness (Bottom base area)
Wall Thickness (Bottom base area) | Wall Thickness (Bottom base area)
Wall Thickness (Top area) | Wall Thickness (Top area)
Wall Thickness (Top area) | Wall Thickness (Top area)
Wall Thickness (Top area) | Wall Thickness (Top area)
Wall Thickness (Top area) | Wall Thickness (Top area)
Water | Water
Water Content | Water Content
Water Content | Water Content
Water Content | Water Content
Water Content | Water Content
Water Content | Water Content
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (Inches) | Web Width (Inches)
Web Width (mm) | Web Width (mm)
Web Width (mm) | Web Width (mm)
Web Width (mm) | Web Width (mm)
Web Width (mm) | Web Width (mm)
Web Width (mm) | Web Width (mm)
Web_Thickness_(mils) | Web Thickness (mils)
Web_Thickness_(mils) | Web Thickness (mils)
Web_Thickness_(mils) | Web Thickness (mils)
Web_Thickness_(mils) | Web Thickness (mils)
Web_Thickness_(mils) | Web Thickness (mils)
Web_Thickness_(mils) | Web Thickness (mils)
Web_Thickness_(mils) | Web Thickness (mils)
Web_Thickness_(mils) | Web Thickness (mils)
Web_Thickness_(mils) | Web Thickness (mils)
Web_Thickness_(mils) | Web Thickness (mils)
Web_Thickness_(mils) | Web Thickness (mils)
Web_Thickness_(mils) | Web Thickness (mils)
Web_Thickness_(mils) | Web Thickness (mils)
Web_Thickness_(mils) | Web Thickness (mils)
Web_Thickness_(mils) | Web Thickness (mils)
Web_Thickness_(mils) | Web Thickness (mils)
Web_Thickness_(mils) | Web Thickness (mils)
Web_Thickness_(mils) | Web Thickness (mils)
Web_Thickness_(mils) | Web Thickness (mils)
Web_Thickness_(mils) | Web Thickness (mils)
Web_Thickness_(mils) | Web Thickness (mils)
Web_Thickness_(mils) | Web Thickness (mils)
Web_Thickness_(mils) | Web Thickness (mils)
Web_Thickness_(mils) | Web Thickness (mils)
Web_Thickness_(mils) | Web Thickness (mils)
Weight | Weight
Weight | Weight
Weight | Weight
Weight | Weight
Weight | Weight
Weight | Weight
Weight | Weight
Weight | Weight
Weight | Weight
Weight | Weight
Weight | Weight
Weight | Weight
Weight | Weight
Weight | Weight
Weight | Weight
Weight | Weight
Weight | Weight
Weight | Weight
Weight | Weight
Weight | Weight
Weight | Weight
Weight (per specification) | Weight (per specification)
Weight (per specification) | Weight (per specification)
Weight (per specification) | Weight (per specification)
Weight (per specification) | Weight (per specification)
Weight (per specification) | Weight (per specification)
Weight (per specification) | Weight (per specification)
Weight (per specification) | Weight (per specification)
Weight (per specification) | Weight (per specification)
Weight (per specification) | Weight (per specification)
Weight (per specification) | Weight (per specification)
Weight (per specification) | Weight (per specification)
Weight (per specification) | Weight (per specification)
Weight (per specification) | Weight (per specification)
Weight (per specification) | Weight (per specification)
Weight of Residue (g) | Weight of Residue (g)
Weight of Residue (g) | Weight of Residue (g)
Weight of Residue (g) | Weight of Residue (g)
Wet CD (g/in) (TM 103) | Wet CD (g/in)
Wet CD (g/in) (TM 103) | Wet CD (g/in)
Wet CD (g/in) (TM 103) | Wet CD (g/in)
Wet CD (g/in) (TM 103) | Wet CD (g/in)
Wet CD (g/in) (TM 103) | Wet CD (g/in)
Wet CD (g/in) (TM 103) | Wet CD (g/in)
Wet CD (g/in) (TM 103) | Wet CD (g/in)
Wet CD (g/in) (TM 103) | Wet CD (g/in)
Wet CD (g/in) (TM 103) | Wet CD (g/in)
Wet CD (g/in) (TM 103) | Wet CD (g/in)
Wet CD (g/in) (TM 103) | Wet CD (g/in)
Wet MD (g/in) (TM 103) | Wet MD (g/in)
Wet MD (g/in) (TM 103) | Wet MD (g/in)
Wet MD (g/in) (TM 103) | Wet MD (g/in)
Wet MD (g/in) (TM 103) | Wet MD (g/in)
Wet MD (g/in) (TM 103) | Wet MD (g/in)
Wet MD (g/in) (TM 103) | Wet MD (g/in)
Wet MD (g/in) (TM 103) | Wet MD (g/in)
Wet MD (g/in) (TM 103) | Wet MD (g/in)
Wet MD (g/in) (TM 103) | Wet MD (g/in)
Wet MD (g/in) (TM 103) | Wet MD (g/in)
Wet MD (g/in) (TM 103) | Wet MD (g/in)
Wet_CD_(g/in) | Wet CD (g/in)
Wet_CD_(g/in) | Wet CD (g/in)
Wet_CD_(g/in) | Wet CD (g/in)
Wet_CD_(g/in) | Wet CD (g/in)
Wet_CD_(g/in) | Wet CD (g/in)
Wet_CD_(g/in) | Wet CD (g/in)
Wet_CD_(g/in) | Wet CD (g/in)
Wet_CD_(g/in) | Wet CD (g/in)
Wet_CD_(g/in) | Wet CD (g/in)
Wet_CD_(g/in) | Wet CD (g/in)
Wet_CD_(g/in) | Wet CD (g/in)
Wet_CD_(g/in) | Wet CD (g/in)
Wet_CD_(g/in) | Wet CD (g/in)
Wet_CD_(g/in) | Wet CD (g/in)
Wet_CD_(g/in) | Wet CD (g/in)
Wet_CD_CoA_(g/25mm) | Wet CD CoA (g/25mm)
Wet_CD_CoA_(g/in) | Wet CD CoA (g/in)
Wet_CD_CoA_(g/in) | Wet CD CoA (g/in)
Wet_CD_CoA_(g/in) | Wet CD CoA (g/in)
Wet_CD_CoA_(g/in) | Wet CD CoA (g/in)
Wet_CD_CoA_(g/in) | Wet CD CoA (g/in)
Wet_CD_CoA_(g/in) | Wet CD CoA (g/in)
Wet_MD_(g/in) | Wet MD (g/in)
Wet_MD_(g/in) | Wet MD (g/in)
Wet_MD_(g/in) | Wet MD (g/in)
Wet_MD_(g/in) | Wet MD (g/in)
Wet_MD_(g/in) | Wet MD (g/in)
Wet_MD_(g/in) | Wet MD (g/in)
Wet_MD_(g/in) | Wet MD (g/in)
Wet_MD_(g/in) | Wet MD (g/in)
Wet_MD_(g/in) | Wet MD (g/in)
Wet_MD_(g/in) | Wet MD (g/in)
Wet_MD_(g/in) | Wet MD (g/in)
Wet_MD_(g/in) | Wet MD (g/in)
Wet_MD_(g/in) | Wet MD (g/in)
Wet_MD_(g/in) | Wet MD (g/in)
Wet_MD_(g/in) | Wet MD (g/in)
Wet_MD_CoA_(g/25mm) | Wet MD CoA (g/25mm)
Wet_MD_CoA_(g/in) | Wet MD CoA (g/in)
Wet_MD_CoA_(g/in) | Wet MD CoA (g/in)
Wet_MD_CoA_(g/in) | Wet MD CoA (g/in)
Wet_MD_CoA_(g/in) | Wet MD CoA (g/in)
Wet_MD_CoA_(g/in) | Wet MD CoA (g/in)
Wet_MD_CoA_(g/in) | Wet MD CoA (g/in)
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width | Width
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Width I.D. | Width I.D.
Wind Direction | Wind Direction

## mainActive
plan.a | Mat_MatCode | Mat_MatDesc | ITEM_DESCRIPTION | SP_TestPlan | Test_Lab | Res_TestCode | Res_TestName | Analyte.C | Numeric_Result | Res_Units | Rounding_Rule | Low_Release_Limit | High_Release_Limit | METHOD | METHOD.C | Facility Name | FACILITY_DESCRIPTION | SAMPLE_PLAN | _merge
2AA284SLI/Lid_01B | 2AA284SLI | Preparation H Solo | Prep H Solo | Lid_01B | NY-Analytical | 765 | Canister & Lid Inspection | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AA284SLI/Lid_01B | 2AA284SLI | Preparation H Solo | Prep H Solo | Lid_01B | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AA284SLI/Lid_01B | 2AA284SLI | Preparation H Solo | Prep H Solo | Lid_01B | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AA284SLI/Lid_01B | 2AA284SLI | Preparation H Solo | Prep H Solo | Lid_01B | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AA284SLI/Lid_01B | 2AA284SLI | Preparation H Solo | Prep H Solo | Lid_01B | NY-Analytical | 765 | Canister & Lid Inspection | Label Length | 3.6875 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AA284SLI/Lid_01B | 2AA284SLI | Preparation H Solo | Prep H Solo | Lid_01B | NY-Analytical | 765 | Canister & Lid Inspection | Label Width | 1.875 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AA284SLI/Lid_01B | 2AA284SLI | Preparation H Solo | Prep H Solo | Lid_01B | NY-Analytical | 765 | Canister & Lid Inspection | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AA347/Web_NP8550_10.25 | 2AA347 | Summer's Eve Fresh Cycle Cleans. Cloths | Summer's Eve Fresh Cycle Cleans. Cloths | Web_NP8550_10.25 | NY-Analytical | 885 | Web Inspection | AQL Sample Size | 315 |  | FDA,0 | 0 |  | TM101/TM107/TM132/Measure | TM101_TM107_TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2AA347/Web_NP8550_10.25 | 2AA347 | Summer's Eve Fresh Cycle Cleans. Cloths | Summer's Eve Fresh Cycle Cleans. Cloths | Web_NP8550_10.25 | NY-Analytical | 885 | Web Inspection | Artwork / Copy | Sample(s) all meet specifications (Done) |  |  |  |  | TM101/TM107/TM132/Measure | TM101_TM107_TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2AA347/Web_NP8550_10.25 | 2AA347 | Summer's Eve Fresh Cycle Cleans. Cloths | Summer's Eve Fresh Cycle Cleans. Cloths | Web_NP8550_10.25 | NY-Analytical | 885 | Web Inspection | Basis Weight | 41.3 | #/Ream | FDA,1 | 37.8 | 48.2 | TM101/TM107/TM132/Measure | TM101_TM107_TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2AA347/Web_NP8550_10.25 | 2AA347 | Summer's Eve Fresh Cycle Cleans. Cloths | Summer's Eve Fresh Cycle Cleans. Cloths | Web_NP8550_10.25 | NY-Analytical | 885 | Web Inspection | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | TM101/TM107/TM132/Measure | TM101_TM107_TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2AA347/Web_NP8550_10.25 | 2AA347 | Summer's Eve Fresh Cycle Cleans. Cloths | Summer's Eve Fresh Cycle Cleans. Cloths | Web_NP8550_10.25 | NY-Analytical | 885 | Web Inspection | Seal Strength (@300°F/30psi/1.5 sec dwell) | 5762.95 | g/in | FDA,0 | 3500 |  | TM101/TM107/TM132/Measure | TM101_TM107_TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2AA347/Web_NP8550_10.25 | 2AA347 | Summer's Eve Fresh Cycle Cleans. Cloths | Summer's Eve Fresh Cycle Cleans. Cloths | Web_NP8550_10.25 | NY-Analytical | 885 | Web Inspection | Web Width | 10.25 | in | FDA,4 | 10.1875 | 10.3125 | TM101/TM107/TM132/Measure | TM101_TM107_TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2AA592/Web_NP8504_11.0 | 2AA592 | Prep H Medicated Wipes 48ct. | Prep H Medicated Wipes 48Ct. | Web_NP8504_11.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Dyne Level | Dyne Level | NY | RAW MATERIAL - NY | RM_WEB | both
2AA592/Web_NP8504_11.0 | 2AA592 | Prep H Medicated Wipes 48ct. | Prep H Medicated Wipes 48Ct. | Web_NP8504_11.0 | NY-Analytical | 836 | Dyne Level | Dyne Level (outside film) | 36 |  | FDA,0 | 36 | 38 | Dyne Test Kit | Dyne_Test_Kit | NY | RAW MATERIAL - NY | RM_WEB | both
2AA592/Web_NP8504_11.0 | 2AA592 | Prep H Medicated Wipes 48ct. | Prep H Medicated Wipes 48Ct. | Web_NP8504_11.0 | NY-Analytical | 761 | Web Measurement | Web Width | 10.968 | in | FDA,4 | 10.9375 | 11.0625 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2AA592/Web_NP8504_11.0 | 2AA592 | Prep H Medicated Wipes 48ct. | Prep H Medicated Wipes 48Ct. | Web_NP8504_11.0 | NY-Analytical | 756 | Basis Weight | Basis Weight | 57.54 | #/Ream | FDA,1 | 53.3 | 65.1 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2AA592/Web_NP8504_11.0 | 2AA592 | Prep H Medicated Wipes 48ct. | Prep H Medicated Wipes 48Ct. | Web_NP8504_11.0 | NY-Analytical | 762 | Seal Strength | Seal Strength (@350°F/30psi/0.5s dwell Top Jaw) | 5339.63 | g/in | FDA,0 | 2700 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2AA592/Web_NP8504_11.0 | 2AA592 | Prep H Medicated Wipes 48ct. | Prep H Medicated Wipes 48Ct. | Web_NP8504_11.0 | NY-Analytical | 763 | Web Thickness | Web Thickness | 3.65 | mil | FDA,1 | 3.2 | 3.9 | TM 132 | TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Preparation H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Lid_NY_NP7600_22R1 | NY-Analytical | 765 | Canister & Lid Inspection | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Preparation H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Lid_NY_NP7600_22R1 | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Preparation H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Lid_NY_NP7600_22R1 | NY-Analytical | 818 | Solo Lid Inspection | Color | Meets the specification (Done) |  |  |  |  | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Preparation H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Lid_NY_NP7600_22R1 | NY-Analytical | 818 | Solo Lid Inspection | Functionality | Meet specification (Done) |  |  |  |  | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Preparation H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Lid_NY_NP7600_22R1 | NY-Analytical | 818 | Solo Lid Inspection | Lid Wall Thickness | 0.028 | in | FDA,3 | 0.025 | 0.035 | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Preparation H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Lid_NY_NP7600_22R1 | NY-Analytical | 818 | Solo Lid Inspection | Open Lid Overall Width | 4.36 | in | FDA,3 | 4.351 | 4.401 | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Preparation H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Lid_NY_NP7600_22R1 | NY-Analytical | 818 | Solo Lid Inspection | Overall closed lid height | 0.37 | in | FDA,3 | 0.355 | 0.385 | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Preparation H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Lid_NY_NP7600_22R1 | NY-Analytical | 818 | Solo Lid Inspection | Weight | 11.241 | g | FDA,1 | 10.5 | 12.5 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_LID | both
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Preparation H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Lid_NY_NP7600_22R1 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Preparation H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Lid_NY_NP7600_22R1 | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Preparation H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Lid_NY_NP7600_22R1 | NY-Analytical | 765 | Canister & Lid Inspection | Label Length | 3.7025 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Preparation H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Lid_NY_NP7600_22R1 | NY-Analytical | 765 | Canister & Lid Inspection | Label Width | 1.8705 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AA592SLI/Lid_NY_NP7600_22R1 | 2AA592SLI | Preparation H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Lid_NY_NP7600_22R1 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AA632/Web_NP8504_11.75 | 2AA632 | Prep H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Web_NP8504_11.75 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Dyne Level | Dyne Level | NY | RAW MATERIAL - NY | RM_WEB | both
2AA632/Web_NP8504_11.75 | 2AA632 | Prep H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Web_NP8504_11.75 | NY-Analytical | 836 | Dyne Level | Dyne Level (outside film) | 38 |  | FDA,0 | 36 | 38 | Dyne Test Kit | Dyne_Test_Kit | NY | RAW MATERIAL - NY | RM_WEB | both
2AA632/Web_NP8504_11.75 | 2AA632 | Prep H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Web_NP8504_11.75 | NY-Analytical | 761 | Web Measurement | Web Width | 11.75 | in | FDA,4 | 11.6875 | 11.8125 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2AA632/Web_NP8504_11.75 | 2AA632 | Prep H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Web_NP8504_11.75 | NY-Analytical | 756 | Basis Weight | Basis Weight | 59.05 | #/Ream | FDA,1 | 53.3 | 65.1 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2AA632/Web_NP8504_11.75 | 2AA632 | Prep H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Web_NP8504_11.75 | NY-Analytical | 762 | Seal Strength | Seal Strength (@350°F/30psi/0.5s dwell Top Jaw) | 5664 | g/in | FDA,0 | 2700 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2AA632/Web_NP8504_11.75 | 2AA632 | Prep H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Web_NP8504_11.75 | NY-Analytical | 763 | Web Thickness | Web Thickness | 3.7 | mil | FDA,1 | 3.2 | 3.9 | TM 132 | TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2AA714SLI/Lid_01V | 2AA714SLI | Dove Baby Wipes Solo | Dove Baby Wipes Solo | Lid_01V | NY-Analytical | 765 | Canister & Lid Inspection | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AA714SLI/Lid_01V | 2AA714SLI | Dove Baby Wipes Solo | Dove Baby Wipes Solo | Lid_01V | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AA714SLI/Lid_01V | 2AA714SLI | Dove Baby Wipes Solo | Dove Baby Wipes Solo | Lid_01V | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AA714SLI/Lid_01V | 2AA714SLI | Dove Baby Wipes Solo | Dove Baby Wipes Solo | Lid_01V | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AA714SLI/Lid_01V | 2AA714SLI | Dove Baby Wipes Solo | Dove Baby Wipes Solo | Lid_01V | NY-Analytical | 765 | Canister & Lid Inspection | Label Length | 3.687 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AA714SLI/Lid_01V | 2AA714SLI | Dove Baby Wipes Solo | Dove Baby Wipes Solo | Lid_01V | NY-Analytical | 765 | Canister & Lid Inspection | Label Width | 1.8125 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AA714SLI/Lid_01V | 2AA714SLI | Dove Baby Wipes Solo | Dove Baby Wipes Solo | Lid_01V | NY-Analytical | 765 | Canister & Lid Inspection | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AA922SLI/Lid_01V | 2AA922SLI | Dove Baby Wipes Solo | Dove Baby Wipes Solo | Lid_01V | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis | COA | NY | RAW MATERIAL - NY | RM_LID | both
2AA922SLI/Lid_NY_Skip lot testing | 2AA922SLI | Dove Baby Wipes Solo | Dove Baby Wipes Solo | Lid_NY_Skip lot testing | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Compliance | COA | NY | RAW MATERIAL - NY | RM_LID | both
2AA922SLI/Lid_NY_Skip lot testing | 2AA922SLI | Dove Baby Wipes Solo | Dove Baby Wipes Solo | Lid_NY_Skip lot testing | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AJ123/Web_NY_Skip lot testing | 2AJ123 | Cetaphil Gentle Skin Cleansing Cloths 25ct. | Cetaphil Gentle Skin Cleansing Cloths 25Ct. | Web_NY_Skip lot testing | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2AJ143SLI/Lid_NY_NP7600_01C | 2AJ143SLI | Hygea Premium Multi Purpose - NY | Hygea Premium Multi Purpose - Ny | Lid_NY_NP7600_01C | IN-Analytical | 765 | Canister & Lid Inspection | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AJ143SLI/Lid_NY_NP7600_01C | 2AJ143SLI | Hygea Premium Multi Purpose - NY | Hygea Premium Multi Purpose - Ny | Lid_NY_NP7600_01C | IN-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AJ143SLI/Lid_NY_NP7600_01C | 2AJ143SLI | Hygea Premium Multi Purpose - NY | Hygea Premium Multi Purpose - Ny | Lid_NY_NP7600_01C | IN-Analytical | 765 | Canister & Lid Inspection | AQL Sample Size | 200 |  |  | 0 |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AJ143SLI/Lid_NY_NP7600_01C | 2AJ143SLI | Hygea Premium Multi Purpose - NY | Hygea Premium Multi Purpose - Ny | Lid_NY_NP7600_01C | IN-Analytical | 765 | Canister & Lid Inspection | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AJ143SLI/Lid_NY_NP7600_01C | 2AJ143SLI | Hygea Premium Multi Purpose - NY | Hygea Premium Multi Purpose - Ny | Lid_NY_NP7600_01C | IN-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AJ143SLI/Lid_NY_NP7600_01C | 2AJ143SLI | Hygea Premium Multi Purpose - NY | Hygea Premium Multi Purpose - Ny | Lid_NY_NP7600_01C | IN-Analytical | 765 | Canister & Lid Inspection | Label Length | 3.7088 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AJ143SLI/Lid_NY_NP7600_01C | 2AJ143SLI | Hygea Premium Multi Purpose - NY | Hygea Premium Multi Purpose - Ny | Lid_NY_NP7600_01C | IN-Analytical | 765 | Canister & Lid Inspection | Label Width | 1.8438 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AJ143SLI/Lid_NY_NP7600_01C | 2AJ143SLI | Hygea Premium Multi Purpose - NY | Hygea Premium Multi Purpose - Ny | Lid_NY_NP7600_01C | IN-Analytical | 765 | Canister & Lid Inspection | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AJ170/Web_NP8530_11.0 | 2AJ170 | Cetaphil Gentle Skin Cleansing wipes 25ct | Cetaphil Gentle Skin Cleansing Wipes 25Ct. | Web_NP8530_11.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2AJ170/Web_NP8530_11.0 | 2AJ170 | Cetaphil Gentle Skin Cleansing wipes 25ct | Cetaphil Gentle Skin Cleansing Wipes 25Ct. | Web_NP8530_11.0 | NY-Analytical | 761 | Web Measurement | Web Width | 11 | in | FDA,4 | 10.9375 | 11.0625 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2AJ170/Web_NP8530_11.0 | 2AJ170 | Cetaphil Gentle Skin Cleansing wipes 25ct | Cetaphil Gentle Skin Cleansing Wipes 25Ct. | Web_NP8530_11.0 | NY-Analytical | 756 | Basis Weight | Basis Weight | 34.39 | #/Ream | FDA,2 | 29.4 | 37.42 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2AJ170/Web_NP8530_11.0 | 2AJ170 | Cetaphil Gentle Skin Cleansing wipes 25ct | Cetaphil Gentle Skin Cleansing Wipes 25Ct. | Web_NP8530_11.0 | NY-Analytical | 762 | Seal Strength | Seal Strength (@300°F/30psi/1.5 sec dwell) | 5442 | g/in | FDA,0 | 3500 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2AJ170/Web_NP8530_11.0 | 2AJ170 | Cetaphil Gentle Skin Cleansing wipes 25ct | Cetaphil Gentle Skin Cleansing Wipes 25Ct. | Web_NP8530_11.0 | NY-Analytical | 763 | Web Thickness | Web Thickness | 1.9 | mil | FDA,1 | 1.8 | 2.2 | TM 132 | TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22FA | NY-Analytical | 765 | Canister & Lid Inspection | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22FA | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22FA | NY-Analytical | 818 | Solo Lid Inspection | Color | Meets the specification (Done) |  |  |  |  | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22FA | NY-Analytical | 818 | Solo Lid Inspection | Functionality | Meet specification (Done) |  |  |  |  | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22FA | NY-Analytical | 818 | Solo Lid Inspection | Lid Wall Thickness | 0.028 | in | FDA,3 | 0.025 | 0.035 | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22FA | NY-Analytical | 818 | Solo Lid Inspection | Open Lid Overall Width | 4.36 | in | FDA,3 | 4.351 | 4.401 | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22FA | NY-Analytical | 818 | Solo Lid Inspection | Overall closed lid height | 0.37 | in | FDA,3 | 0.355 | 0.385 | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22FA | NY-Analytical | 818 | Solo Lid Inspection | Weight | 11.25 | g | FDA,1 | 10.5 | 12.5 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_LID | both
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22FA | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22FA | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22FA | NY-Analytical | 765 | Canister & Lid Inspection | Label Length | 3.67 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22FA | NY-Analytical | 765 | Canister & Lid Inspection | Label Width | 1.84 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AJ170SLI/Lid_NY_NP7600_22FA | 2AJ170SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22FA | NY-Analytical | 765 | Canister & Lid Inspection | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22V | NY-Analytical | 765 | Canister & Lid Inspection | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22V | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22V | NY-Analytical | 818 | Solo Lid Inspection | Appearance | Meet specification (Done) |  |  |  |  | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22V | NY-Analytical | 818 | Solo Lid Inspection | Color | Meets the specification (Done) |  |  |  |  | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22V | NY-Analytical | 818 | Solo Lid Inspection | Functionality | Meet specification (Done) |  |  |  |  | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22V | NY-Analytical | 818 | Solo Lid Inspection | Lid Wall Thickness | 0.0295 | in | FDA,3 | 0.025 | 0.035 | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22V | NY-Analytical | 818 | Solo Lid Inspection | Odor | None detectable (Done) |  |  |  |  | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22V | NY-Analytical | 818 | Solo Lid Inspection | Open Lid Overall Width | 4.374 | in | FDA,3 | 4.351 | 4.401 | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22V | NY-Analytical | 818 | Solo Lid Inspection | Overall closed lid height | 0.373 | in | FDA,3 | 0.355 | 0.385 | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22V | NY-Analytical | 818 | Solo Lid Inspection | Weight | 11.19 | g | FDA,1 | 10.5 | 12.5 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_LID | both
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22V | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22V | NY-Analytical | 765 | Canister & Lid Inspection | Label Length | 3.7188 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AJ190SLI/Lid_NY_NP7600_22V | 2AJ190SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22V | NY-Analytical | 765 | Canister & Lid Inspection | Label Width | 1.875 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22X | NY-Analytical | 765 | Canister & Lid Inspection | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22X | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22X | NY-Analytical | 818 | Solo Lid Inspection | Color | Meets the specification (Done) |  |  |  |  | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22X | NY-Analytical | 818 | Solo Lid Inspection | Functionality | Meet specification (Done) |  |  |  |  | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22X | NY-Analytical | 818 | Solo Lid Inspection | Lid Wall Thickness | 0.0297 | in | FDA,3 | 0.025 | 0.035 | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22X | NY-Analytical | 818 | Solo Lid Inspection | Open Lid Overall Width | 4.371 | in | FDA,3 | 4.351 | 4.401 | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22X | NY-Analytical | 818 | Solo Lid Inspection | Overall closed lid height | 0.3705 | in | FDA,3 | 0.355 | 0.385 | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22X | NY-Analytical | 818 | Solo Lid Inspection | Weight | 11.385 | g | FDA,1 | 10.5 | 12.5 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_LID | both
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22X | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22X | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22X | NY-Analytical | 765 | Canister & Lid Inspection | Label Length | 3.7088 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22X | NY-Analytical | 765 | Canister & Lid Inspection | Label Width | 1.875 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AJ324SLI/Lid_NY_NP7600_22X | 2AJ324SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22X | NY-Analytical | 765 | Canister & Lid Inspection | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AP075/Web_NP8542_12.205 | 2AP075 | L' Oreal paris Revitalift MUR Cleans. Twlts. | L'Oreal Paris Revitalift Mur Cleans. Twlts. | Web_NP8542_12.205 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2AP075/Web_NP8542_12.205 | 2AP075 | L' Oreal paris Revitalift MUR Cleans. Twlts. | L'Oreal Paris Revitalift Mur Cleans. Twlts. | Web_NP8542_12.205 | NY-Analytical | 761 | Web Measurement | Web Width | 12.25 | in | FDA,4 | 12.1425 | 12.2675 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2AP075/Web_NP8542_12.205 | 2AP075 | L' Oreal paris Revitalift MUR Cleans. Twlts. | L'Oreal Paris Revitalift Mur Cleans. Twlts. | Web_NP8542_12.205 | NY-Analytical | 756 | Basis Weight | Basis Weight | 60.85 | #/Ream | FDA,2 | 52.53 | 64.21 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2AP075/Web_NP8542_12.205 | 2AP075 | L' Oreal paris Revitalift MUR Cleans. Twlts. | L'Oreal Paris Revitalift Mur Cleans. Twlts. | Web_NP8542_12.205 | NY-Analytical | 762 | Seal Strength | Seal Strength (@130°C/1bar/0.5sec dwell) | 7718 | g/mm | FDA,0 | 25 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2AP087/Web_NP8550_12.205 | 2AP087 | Cetaphil Gentle Makeup Removing Wipes | Cetaphil Gentle Makeup Removing Wipes | Web_NP8550_12.205 | NY-Analytical | 885 | Web Inspection | AQL Sample Size | 200 |  | FDA,0 | 0 |  | TM101/TM107/TM132/Measure | TM101_TM107_TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2AP087/Web_NP8550_12.205 | 2AP087 | Cetaphil Gentle Makeup Removing Wipes | Cetaphil Gentle Makeup Removing Wipes | Web_NP8550_12.205 | NY-Analytical | 885 | Web Inspection | Artwork / Copy | Sample(s) all meet specifications (Done) |  |  |  |  | TM101/TM107/TM132/Measure | TM101_TM107_TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2AP087/Web_NP8550_12.205 | 2AP087 | Cetaphil Gentle Makeup Removing Wipes | Cetaphil Gentle Makeup Removing Wipes | Web_NP8550_12.205 | NY-Analytical | 885 | Web Inspection | Basis Weight | 43.3 | #/Ream | FDA,1 | 37.8 | 48.2 | TM101/TM107/TM132/Measure | TM101_TM107_TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2AP087/Web_NP8550_12.205 | 2AP087 | Cetaphil Gentle Makeup Removing Wipes | Cetaphil Gentle Makeup Removing Wipes | Web_NP8550_12.205 | NY-Analytical | 885 | Web Inspection | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | TM101/TM107/TM132/Measure | TM101_TM107_TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2AP087/Web_NP8550_12.205 | 2AP087 | Cetaphil Gentle Makeup Removing Wipes | Cetaphil Gentle Makeup Removing Wipes | Web_NP8550_12.205 | NY-Analytical | 885 | Web Inspection | Seal Strength (@300°F/30psi/1.5 sec dwell) | 5039.27 | g/in | FDA,0 | 3500 |  | TM101/TM107/TM132/Measure | TM101_TM107_TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2AP087/Web_NP8550_12.205 | 2AP087 | Cetaphil Gentle Makeup Removing Wipes | Cetaphil Gentle Makeup Removing Wipes | Web_NP8550_12.205 | NY-Analytical | 885 | Web Inspection | UPC Verification | UPC Scan is acceptable (Scan into comment section below - Done) |  |  |  |  | TM101/TM107/TM132/Measure | TM101_TM107_TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2AP087/Web_NP8550_12.205 | 2AP087 | Cetaphil Gentle Makeup Removing Wipes | Cetaphil Gentle Makeup Removing Wipes | Web_NP8550_12.205 | NY-Analytical | 885 | Web Inspection | Web Width | 12.1425 | in | FDA,4 | 12.1425 | 12.2675 | TM101/TM107/TM132/Measure | TM101_TM107_TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2AP087/Web_NP8550_12.205 | 2AP087 | Cetaphil Gentle Makeup Removing Wipes | Cetaphil Gentle Makeup Removing Wipes | Web_NP8550_12.205 | NY-Analytical | 885 | Web Inspection | Wind Direction | Meets specification (Indicate numeric value - Done) |  |  |  |  | TM101/TM107/TM132/Measure | TM101_TM107_TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2AP113/Web_NP8542_11.0 | 2AP113 | Garnier Skin Active Clean Refreshing Remover Cleans Twlts | Garnier Skin ACt. ive Clean Refreshing Remover Cleans Twlts | Web_NP8542_11.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2AP113/Web_NP8542_11.0 | 2AP113 | Garnier Skin Active Clean Refreshing Remover Cleans Twlts | Garnier Skin ACt. ive Clean Refreshing Remover Cleans Twlts | Web_NP8542_11.0 | NY-Analytical | 761 | Web Measurement | Web Width | 11 | in | FDA,4 | 10.9375 | 11.0625 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2AP113/Web_NP8542_11.0 | 2AP113 | Garnier Skin Active Clean Refreshing Remover Cleans Twlts | Garnier Skin ACt. ive Clean Refreshing Remover Cleans Twlts | Web_NP8542_11.0 | NY-Analytical | 756 | Basis Weight | Basis Weight | 61.11 | #/Ream | FDA,2 | 52.53 | 64.21 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2AP113/Web_NP8542_11.0 | 2AP113 | Garnier Skin Active Clean Refreshing Remover Cleans Twlts | Garnier Skin ACt. ive Clean Refreshing Remover Cleans Twlts | Web_NP8542_11.0 | NY-Analytical | 762 | Seal Strength | Seal Strength (@130°C/1bar/0.5sec dwell) | 180.6 | g/mm | FDA,0 | 25 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2AP260/Web_NP8535_12.75 | 2AP260 | Avon ANEW Clean MUR Wipes 24ct | Avon Anew Clean Mur Wipes 24Ct. | Web_NP8535_12.75 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2AP260/Web_NP8535_12.75 | 2AP260 | Avon ANEW Clean MUR Wipes 24ct | Avon Anew Clean Mur Wipes 24Ct. | Web_NP8535_12.75 | NY-Analytical | 761 | Web Measurement | Web Width | 12.75 | in | FDA,4 | 12.6875 | 12.8125 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2AP260/Web_NP8535_12.75 | 2AP260 | Avon ANEW Clean MUR Wipes 24ct | Avon Anew Clean Mur Wipes 24Ct. | Web_NP8535_12.75 | NY-Analytical | 756 | Basis Weight | Basis Weight | 41.14 | #/Ream | FDA,2 | 35.82 | 45.6 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2AP260/Web_NP8535_12.75 | 2AP260 | Avon ANEW Clean MUR Wipes 24ct | Avon Anew Clean Mur Wipes 24Ct. | Web_NP8535_12.75 | NY-Analytical | 762 | Seal Strength | Seal Strength (@300°F/30psi/1.5 sec dwell) | 6659.74 | g/in | FDA,0 | 3500 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2AP260/Web_NP8535_12.75 | 2AP260 | Avon ANEW Clean MUR Wipes 24ct | Avon Anew Clean Mur Wipes 24Ct. | Web_NP8535_12.75 | NY-Analytical | 763 | Web Thickness | Web Thickness | 2.4 | mil | FDA,1 | 2.2 | 2.8 | TM 132 | TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Medium Deep Well Lid | Medium Deep Well Lid | Lid_04B_560_1 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Medium Deep Well Lid | Medium Deep Well Lid | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Medium Deep Well Lid | Medium Deep Well Lid | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Cap Thickness | 0.154 | in | FDA,3 | 0.147 | 0.157 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Medium Deep Well Lid | Medium Deep Well Lid | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Blue (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Medium Deep Well Lid | Medium Deep Well Lid | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Pass (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Medium Deep Well Lid | Medium Deep Well Lid | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter | 3.7695 | in | FDA,3 | 3.76 | 3.77 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Medium Deep Well Lid | Medium Deep Well Lid | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness | 0.05 | in | FDA,3 | 0.04 | 0.05 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Medium Deep Well Lid | Medium Deep Well Lid | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Medium Deep Well Lid | Medium Deep Well Lid | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height | 1.5845 | in | FDA,3 | 1.58 | 1.6 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Medium Deep Well Lid | Medium Deep Well Lid | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Outer Diameter | 0.565 | in | FDA,3 | 0.555 | 0.565 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Medium Deep Well Lid | Medium Deep Well Lid | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Thicknes | 0.02 | in | FDA,3 | 0.015 | 0.025 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AQ471LI/Lid_04B_560_1 | 2AQ471LI | Medium Deep Well Lid | Medium Deep Well Lid | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Weight | 22.773 | g | FDA,1 | 20 | 24 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2AU191/Web_NP5115_7.5 | 2AU191 | Cetaphil Gentle Skin Cleans Cloths 1ct. | Cetaphil Gentle Skin Cleans Cloths 1Ct. | Web_NP5115_7.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2AU191/Web_NP5115_7.5 | 2AU191 | Cetaphil Gentle Skin Cleans Cloths 1ct. | Cetaphil Gentle Skin Cleans Cloths 1Ct. | Web_NP5115_7.5 | NY-Analytical | 761 | Web Measurement | Web Width | 7.4375 | in | FDA,4 | 7.4375 | 7.5625 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2AU191/Web_NP5115_7.5 | 2AU191 | Cetaphil Gentle Skin Cleans Cloths 1ct. | Cetaphil Gentle Skin Cleans Cloths 1Ct. | Web_NP5115_7.5 | NY-Analytical | 756 | Basis Weight | Basis Weight | 60.5 | #/Ream | FDA,1 | 55.6 | 68 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2AU191/Web_NP5115_7.5 | 2AU191 | Cetaphil Gentle Skin Cleans Cloths 1ct. | Cetaphil Gentle Skin Cleans Cloths 1Ct. | Web_NP5115_7.5 | NY-Analytical | 762 | Seal Strength | Seal Strength (@325°F/40PSI/0.5Sec Dwell) | 1823.34 | g/in | FDA,0 | 1500 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2AU252/Web_NP3459_8.0 | 2AU252 | L' Oreal paris Sublime Bronze Self-Tanning Twlts. | L'Oreal Paris Sublime Bronze Self-Tanning Twlts. | Web_NP3459_8.0 |  | 1044 | Basis Weight (lbs/ream) | Basis Weight | 63.4 | Lb/Ream | FDA,2 | 57.56 | 70.4 | ASTM D 646 | ASTMD_646 | NY | RAW MATERIAL - NY | RM_WEB | both
2AU252/Web_NP3459_8.0 | 2AU252 | L' Oreal paris Sublime Bronze Self-Tanning Twlts. | L'Oreal Paris Sublime Bronze Self-Tanning Twlts. | Web_NP3459_8.0 |  | 757 | Caliper | Caliper | 3.8 | mil | FDA,3 | 3.515 | 3.515 | Caliper | TM102 | NY | RAW MATERIAL - NY | RM_WEB | both
2AU252/Web_3459_8 | 2AU252 | L' Oreal paris Sublime Bronze Self-Tanning Twlts. | L'Oreal Paris Sublime Bronze Self-Tanning Twlts. | Web_3459_8 | NY-Analytical | 1044 | Basis Weight (lbs/ream) | Basis Weight |  | Lb/Ream | FDA,2 | 57.56 | 70.4 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2AU252/Web_NP3459_8.0 | 2AU252 | L' Oreal paris Sublime Bronze Self-Tanning Twlts. | L'Oreal Paris Sublime Bronze Self-Tanning Twlts. | Web_NP3459_8.0 | NY-Analytical | 762 | Seal Strength | Seal Strength (@290°F/40psi/0.5 sec dwell) |  | g/in | FDA,0 | 7500 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2AU252/Web_3459_8 | 2AU252 | L' Oreal paris Sublime Bronze Self-Tanning Twlts. | L'Oreal Paris Sublime Bronze Self-Tanning Twlts. | Web_3459_8 | NY-Analytical | 761 | Web Measurement | Core ID |  | in | FDA,3 | 2.9375 | 3.0625 | TM 143 | TM143 | NY | RAW MATERIAL - NY | RM_WEB | both
2AU252/Web_3459_8 | 2AU252 | L' Oreal paris Sublime Bronze Self-Tanning Twlts. | L'Oreal Paris Sublime Bronze Self-Tanning Twlts. | Web_3459_8 | NY-Analytical | 761 | Web Measurement | Roll OD | 15 | in | FDA,3 | 14.9375 | 15.0625 | TM 143 | TM143 | NY | RAW MATERIAL - NY | RM_WEB | both
2AU252/Web_NP3459_8.0 | 2AU252 | L' Oreal paris Sublime Bronze Self-Tanning Twlts. | L'Oreal Paris Sublime Bronze Self-Tanning Twlts. | Web_NP3459_8.0 |  | 761 | Web Measurement | Web Width | 8 | in | FDA,4 | 7.9375 | 8.0625 | TM 143 | TM143 | NY | RAW MATERIAL - NY | RM_WEB | both
2AU252/Web_3459_8 | 2AU252 | L' Oreal paris Sublime Bronze Self-Tanning Twlts. | L'Oreal Paris Sublime Bronze Self-Tanning Twlts. | Web_3459_8 | NY-Analytical | 761 | Web Measurement | Web Width | 8 | in | FDA,4 | 7.9375 | 8.0625 | TM 143 | TM143 | NY | RAW MATERIAL - NY | RM_WEB | both
2AU252/Web_NP3459_8.0 | 2AU252 | L' Oreal paris Sublime Bronze Self-Tanning Twlts. | L'Oreal Paris Sublime Bronze Self-Tanning Twlts. | Web_NP3459_8.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | [5 empty] | Visual | Visual | NY | RAW MATERIAL - NY | RM_WEB | both
2AU252/Web_3459_8 | 2AU252 | L' Oreal paris Sublime Bronze Self-Tanning Twlts. | L'Oreal Paris Sublime Bronze Self-Tanning Twlts. | Web_3459_8 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | [5 empty] | Visual | Visual | NY | RAW MATERIAL - NY | RM_WEB | both
2AU252/Web_NP3459_8.0 | 2AU252 | L' Oreal paris Sublime Bronze Self-Tanning Twlts. | L'Oreal Paris Sublime Bronze Self-Tanning Twlts. | Web_NP3459_8.0 |  | 748 | Visual Inspection | Appearance | Meets specifications (Done) |  |  |  |  | Visual Observation | Visual | NY | RAW MATERIAL - NY | RM_WEB | both
2AU252/Web_NP3459_8.0 | 2AU252 | L' Oreal paris Sublime Bronze Self-Tanning Twlts. | L'Oreal Paris Sublime Bronze Self-Tanning Twlts. | Web_NP3459_8.0 |  | 748 | Visual Inspection | Color/Copy | Meets specifications (Done) |  |  |  |  | Visual Observation | Visual | NY | RAW MATERIAL - NY | RM_WEB | both
2AU555/Web_NP3460_7.5 | 2AU555 | Prep H Rapid Relief with Lidocaine Totables | Prep H Rapid Relief With Lidocaine Totables | Web_NP3460_7.5 | NY-Analytical | 1044 | Basis Weight (lbs/ream) | Basis Weight | 72.89 | Lb/Ream | FDA,1 | 66.6 | 81.4 | ASTM D 646 | ASTMD_646 | NY | RAW MATERIAL - NY | RM_WEB | both
2AU555/Web_NP3460_7.5 | 2AU555 | Prep H Rapid Relief with Lidocaine Totables | Prep H Rapid Relief With Lidocaine Totables | Web_NP3460_7.5 | NY-Analytical | 757 | Caliper | Caliper | 4.22 | mil | FDA,4 | 3.995 | 3.995 | Caliper | TM102 | NY | RAW MATERIAL - NY | RM_WEB | both
2AU555/Web_NP3460_7.5 | 2AU555 | Prep H Rapid Relief with Lidocaine Totables | Prep H Rapid Relief With Lidocaine Totables | Web_NP3460_7.5 | NY-Analytical | 836 | Dyne Level | Dyne Level (outside film) | 36 |  | FDA,0 |  |  | Dyne Test Kit | Dyne_Test_Kit | NY | RAW MATERIAL - NY | RM_WEB | both
2AU555/Web_NP3460_7.5 | 2AU555 | Prep H Rapid Relief with Lidocaine Totables | Prep H Rapid Relief With Lidocaine Totables | Web_NP3460_7.5 | NY-Analytical | 762 | Seal Strength | Seal Strength (@290°F/40psi/0.5 sec dwell) | 8720.24 | g/in | FDA,0 | 5850 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2AU555/Web_NP3460_7.5 | 2AU555 | Prep H Rapid Relief with Lidocaine Totables | Prep H Rapid Relief With Lidocaine Totables | Web_NP3460_7.5 | NY-Analytical | 761 | Web Measurement | Core ID | 3 | in | FDA,4 | 2.9375 | 3.0625 | TM 143 | TM143 | NY | RAW MATERIAL - NY | RM_WEB | both
2AU555/Web_NP3460_7.5 | 2AU555 | Prep H Rapid Relief with Lidocaine Totables | Prep H Rapid Relief With Lidocaine Totables | Web_NP3460_7.5 | NY-Analytical | 761 | Web Measurement | Roll OD | 15 | in | FDA,4 | 14.9375 | 15.0625 | TM 143 | TM143 | NY | RAW MATERIAL - NY | RM_WEB | both
2AU555/Web_NP3460_7.5 | 2AU555 | Prep H Rapid Relief with Lidocaine Totables | Prep H Rapid Relief With Lidocaine Totables | Web_NP3460_7.5 | NY-Analytical | 761 | Web Measurement | Web Width | 7.5 | in | FDA,4 | 7.4375 | 7.5625 | TM 143 | TM143 | NY | RAW MATERIAL - NY | RM_WEB | both
2AU555/Web_NP3460_7.5 | 2AU555 | Prep H Rapid Relief with Lidocaine Totables | Prep H Rapid Relief With Lidocaine Totables | Web_NP3460_7.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Visual | Visual | NY | RAW MATERIAL - NY | RM_WEB | both
2AU555/Web_NP3460_7.5 | 2AU555 | Prep H Rapid Relief with Lidocaine Totables | Prep H Rapid Relief With Lidocaine Totables | Web_NP3460_7.5 | NY-Analytical | 748 | Visual Inspection | Appearance | Meets specifications (Done) |  |  |  |  | Visual Observation | Visual | NY | RAW MATERIAL - NY | RM_WEB | both
2AU555/Web_NP3460_7.5 | 2AU555 | Prep H Rapid Relief with Lidocaine Totables | Prep H Rapid Relief With Lidocaine Totables | Web_NP3460_7.5 | NY-Analytical | 748 | Visual Inspection | Color/Copy | Meets specifications (Done) |  |  |  |  | Visual Observation | Visual | NY | RAW MATERIAL - NY | RM_WEB | both
2BA972/Web_NP8540_13.5 | 2BA972 | Sani Pro. - Table Turners Web 75ct | Sani Professional - Table Turners Web 75Ct. | Web_NP8540_13.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | [5 empty] | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2BA972/Web_NP8540_13.5 | 2BA972 | Sani Pro. - Table Turners Web 75ct | Sani Professional - Table Turners Web 75Ct. | Web_NP8540_13.5 | NY-Analytical | 761 | Web Measurement | Web Width | 13.5 | in | FDA,4 | 13.4375 | 13.5625 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2BA972/Web_NP8540_13.5 | 2BA972 | Sani Pro. - Table Turners Web 75ct | Sani Professional - Table Turners Web 75Ct. | Web_NP8540_13.5 | NY-Analytical | 756 | Basis Weight | Basis Weight | 53.97 | #/Ream | FDA,1 | 45.4 | 57.8 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2BA972/Web_NP8540_13.5 | 2BA972 | Sani Pro. - Table Turners Web 75ct | Sani Professional - Table Turners Web 75Ct. | Web_NP8540_13.5 | NY-Analytical | 762 | Seal Strength | Seal Strength (@300°F/30psi/1.5 sec dwell) |  | g/in | FDA,0 | 3500 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2BA972/Web_NP8540_13.5 | 2BA972 | Sani Pro. - Table Turners Web 75ct | Sani Professional - Table Turners Web 75Ct. | Web_NP8540_13.5 | NY-Analytical | 763 | Web Thickness | Web Thickness | 3.1 | mil | FDA,2 | 2.86 | 3.64 | TM 132 | TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2BC801/Web_NP3870_6.0 | 2BC801 | Philips Sali- Wipes Sterile Saline Wipes | Philips Sali- Wipes Sterile Saline Wipes | Web_NP3870_6.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2BC801/Web_NP3870_6.0 | 2BC801 | Philips Sali- Wipes Sterile Saline Wipes | Philips Sali- Wipes Sterile Saline Wipes | Web_NP3870_6.0 | NY-Analytical | 761 | Web Measurement | Web Width | 6 | in | FDA,4 | 5.9375 | 6.0625 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2BC801/Web_NP3870_6.0 | 2BC801 | Philips Sali- Wipes Sterile Saline Wipes | Philips Sali- Wipes Sterile Saline Wipes | Web_NP3870_6.0 | NY-Analytical | 756 | Basis Weight | Basis Weight | 92.46 | #/Ream | FDA,1 | 81 | 99 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2BC801/Web_NP3870_6.0 | 2BC801 | Philips Sali- Wipes Sterile Saline Wipes | Philips Sali- Wipes Sterile Saline Wipes | Web_NP3870_6.0 | NY-Analytical | 762 | Seal Strength | Seal Strength (@350°F/40psi/1.0 sec dwell) | 10097 | g/in | FDA,0 | 5000 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2BH592/Web_NP3895_6.0 | 2BH592 | Sani-Cloth AF3 | Sani-Cloth AF3 | Web_NP3895_6.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2BH592/Web_NP3895_6.0 | 2BH592 | Sani-Cloth AF3 | Sani-Cloth AF3 | Web_NP3895_6.0 | NY-Analytical | 761 | Web Measurement | Web Width | 6 | in | FDA,4 | 5.9375 | 6.0625 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2BH592/Web_NP3895_6.0 | 2BH592 | Sani-Cloth AF3 | Sani-Cloth AF3 | Web_NP3895_6.0 | NY-Analytical | 756 | Basis Weight | Basis Weight | 69.71 | #/3000sqft | FDA,1 | 65.3 | 79.8 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2BH592/Web_NP3895_6.0 | 2BH592 | Sani-Cloth AF3 | Sani-Cloth AF3 | Web_NP3895_6.0 | NY-Analytical | 762 | Seal Strength | Seal Strength (@250°F/30psi/0.5 sec dwell) | 5740.1 | g/in | FDA,0 | 2270 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2BH592/Web_NP3895_6.0 | 2BH592 | Sani-Cloth AF3 | Sani-Cloth AF3 | Web_NP3895_6.0 | NY-Analytical | 763 | Web Thickness | Web Thickness | 4.5 | mil | FDA,1 | 3.9 | 4.7 | TM 132 | TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2BJ123/Web_NP8550_11 | 2BJ123 | Cetaphil Gentle Skin Cleansing Cloths 25ct | Cetaphil Gentle Skin Cleansing Cloths 25Ct. | Web_NP8550_11 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2BJ123/Web_NP8550_11 | 2BJ123 | Cetaphil Gentle Skin Cleansing Cloths 25ct | Cetaphil Gentle Skin Cleansing Cloths 25Ct. | Web_NP8550_11 | NY-Analytical | 761 | Web Measurement | Web Width | 11 | in | FDA,4 | 10.9375 | 11.0625 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2BJ123/Web_NP8550_11 | 2BJ123 | Cetaphil Gentle Skin Cleansing Cloths 25ct | Cetaphil Gentle Skin Cleansing Cloths 25Ct. | Web_NP8550_11 | NY-Analytical | 756 | Basis Weight | Basis Weight | 43.12 | #/Ream | FDA,1 | 37.8 | 48.2 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2BJ123/Web_NP8550_11 | 2BJ123 | Cetaphil Gentle Skin Cleansing Cloths 25ct | Cetaphil Gentle Skin Cleansing Cloths 25Ct. | Web_NP8550_11 | NY-Analytical | 762 | Seal Strength | Seal Strength (@300°F/30psi/1.5 sec dwell) | 6855 | g/in | FDA,0 | 3500 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2BJ143/Web_NP8530_16.25 | 2BJ143 | Hygea Premium Multi Purpose Washcloths | Hygea Premium Multi Purpose Washcloths | Web_NP8530_16.25 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2BJ143/Web_NP8530_16.25 | 2BJ143 | Hygea Premium Multi Purpose Washcloths | Hygea Premium Multi Purpose Washcloths | Web_NP8530_16.25 | NY-Analytical | 761 | Web Measurement | Web Width | 16.25 | in | FDA,4 | 16.1875 | 16.3125 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2BJ143/Web_NP8530_16.25 | 2BJ143 | Hygea Premium Multi Purpose Washcloths | Hygea Premium Multi Purpose Washcloths | Web_NP8530_16.25 | NY-Analytical | 756 | Basis Weight | Basis Weight | 34.26 | #/Ream | FDA,2 | 29.4 | 37.42 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2BJ143/Web_NP8530_16.25 | 2BJ143 | Hygea Premium Multi Purpose Washcloths | Hygea Premium Multi Purpose Washcloths | Web_NP8530_16.25 | NY-Analytical | 762 | Seal Strength | Seal Strength (@300°F/30psi/1.5 sec dwell) | 7619 | g/in | FDA,0 | 3500 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2BJ143/Web_NP8530_16.25 | 2BJ143 | Hygea Premium Multi Purpose Washcloths | Hygea Premium Multi Purpose Washcloths | Web_NP8530_16.25 | NY-Analytical | 763 | Web Thickness | Web Thickness | 2.1 | mil | FDA,1 | 1.8 | 2.2 | TM 132 | TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2BJ213/Web_NP8530_18 | 2BJ213 | Hygea Multi-Purpose Washcloths | Hygea Multi-Purpose Washcloths | Web_NP8530_18 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2BJ213/Web_NP8530_18 | 2BJ213 | Hygea Multi-Purpose Washcloths | Hygea Multi-Purpose Washcloths | Web_NP8530_18 | NY-Analytical | 761 | Web Measurement | Web Width | 18 | in | FDA,4 | 17.9375 | 18.0625 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2BJ213/Web_NP8530_18 | 2BJ213 | Hygea Multi-Purpose Washcloths | Hygea Multi-Purpose Washcloths | Web_NP8530_18 | NY-Analytical | 756 | Basis Weight | Basis Weight | 33.33 | #/Ream | FDA,2 | 29.4 | 37.42 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2BJ213/Web_NP8530_18 | 2BJ213 | Hygea Multi-Purpose Washcloths | Hygea Multi-Purpose Washcloths | Web_NP8530_18 | NY-Analytical | 762 | Seal Strength | Seal Strength (@300°F/30psi/1.5 sec dwell) | 6908 | g/in | FDA,0 | 3500 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2BJ213/Web_NP8530_18 | 2BJ213 | Hygea Multi-Purpose Washcloths | Hygea Multi-Purpose Washcloths | Web_NP8530_18 | NY-Analytical | 763 | Web Thickness | Web Thickness | 1.85 | mil | FDA,1 | 1.8 | 2.2 | TM 132 | TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2BJ227/Web_NP8530_15.75 | 2BJ227 | Hygea Multi Purpose Washcloths | Hygea Multi Purpose Washcloths | Web_NP8530_15.75 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2BJ227/Web_NP8530_15.75 | 2BJ227 | Hygea Multi Purpose Washcloths | Hygea Multi Purpose Washcloths | Web_NP8530_15.75 | NY-Analytical | 761 | Web Measurement | Web Width | 15.75 | in | FDA,4 | 15.6875 | 15.8125 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2BJ227/Web_NP8530_15.75 | 2BJ227 | Hygea Multi Purpose Washcloths | Hygea Multi Purpose Washcloths | Web_NP8530_15.75 | NY-Analytical | 756 | Basis Weight | Basis Weight | 33.9 | #/Ream | FDA,2 | 29.4 | 37.42 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2BJ227/Web_NP8530_15.75 | 2BJ227 | Hygea Multi Purpose Washcloths | Hygea Multi Purpose Washcloths | Web_NP8530_15.75 | NY-Analytical | 762 | Seal Strength | Seal Strength (@300°F/30psi/1.5 sec dwell) | 5068.96 | g/in | FDA,0 | 3500 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2BJ227/Web_NP8530_15.75 | 2BJ227 | Hygea Multi Purpose Washcloths | Hygea Multi Purpose Washcloths | Web_NP8530_15.75 | NY-Analytical | 763 | Web Thickness | Web Thickness | 2.18 | mil | FDA,1 | 1.8 | 2.2 | TM 132 | TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2BQ084LI/Lid_11A | 2BQ084LI | Large Deep Well Lid 560 Orifice | Large Deep Well 560 Orifice | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Lid fits with a tight seal on |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ084LI/Lid_11A | 2BQ084LI | Large Deep Well Lid 560 Orifice | Large Deep Well 560 Orifice | Lid_11A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Flip Lid Force |  | NY | RAW MATERIAL - NY | RM_LID | both
2BQ084LI/Lid_11A | 2BQ084LI | Large Deep Well Lid 560 Orifice | Large Deep Well 560 Orifice | Lid_11A | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | 2.32 | lb | FDA,2 | 0.75 |  | TM 143 | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ084LI/Lid_11A | 2BQ084LI | Large Deep Well Lid 560 Orifice | Large Deep Well 560 Orifice | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Meets Specifications (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ084LI/Lid_11A | 2BQ084LI | Large Deep Well Lid 560 Orifice | Large Deep Well 560 Orifice | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | 4.85 | in | FDA,3 | 4.847 | 4.867 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ084LI/Lid_11A | 2BQ084LI | Large Deep Well Lid 560 Orifice | Large Deep Well 560 Orifice | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Color | Translucent Teal (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ084LI/Lid_11A | 2BQ084LI | Large Deep Well Lid 560 Orifice | Large Deep Well 560 Orifice | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Insert Fin Gap | 0.071 | in | FDA,3 | 0.065 | 0.075 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ084LI/Lid_11A | 2BQ084LI | Large Deep Well Lid 560 Orifice | Large Deep Well 560 Orifice | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter | 4.92 | in | FDA,3 | 4.917 | 4.937 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ084LI/Lid_11A | 2BQ084LI | Large Deep Well Lid 560 Orifice | Large Deep Well 560 Orifice | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness | 0.044 | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ084LI/Lid_11A | 2BQ084LI | Large Deep Well Lid 560 Orifice | Large Deep Well 560 Orifice | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ084LI/Lid_11A | 2BQ084LI | Large Deep Well Lid 560 Orifice | Large Deep Well 560 Orifice | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness | 0.023 | in | FDA,3 | 0.018 | 0.028 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ084LI/Lid_11A | 2BQ084LI | Large Deep Well Lid 560 Orifice | Large Deep Well 560 Orifice | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height | 1.48 | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ084LI/Lid_11A | 2BQ084LI | Large Deep Well Lid 560 Orifice | Large Deep Well 560 Orifice | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Outer Diameter | 0.166 | in | FDA,3 | 0.16 | 0.17 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ084LI/Lid_11A | 2BQ084LI | Large Deep Well Lid 560 Orifice | Large Deep Well 560 Orifice | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Weight | 34.11 | g | FDA,1 | 32 | 42 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ471LI/Lid_04F_560 | 2BQ471LI | Medium Deep Well Lid with 560 Orifice Trans Blue | Medium Deep Well With 560 Orifice Trans Blue | Lid_04F_560 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ471LI/Lid_04F_560 | 2BQ471LI | Medium Deep Well Lid with 560 Orifice Trans Blue | Medium Deep Well With 560 Orifice Trans Blue | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ471LI/Lid_04F_560 | 2BQ471LI | Medium Deep Well Lid with 560 Orifice Trans Blue | Medium Deep Well With 560 Orifice Trans Blue | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Cap Thickness | 0.151 | in | FDA,3 | 0.147 | 0.157 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ471LI/Lid_04F_560 | 2BQ471LI | Medium Deep Well Lid with 560 Orifice Trans Blue | Medium Deep Well With 560 Orifice Trans Blue | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Translucent Blue (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ471LI/Lid_04F_560 | 2BQ471LI | Medium Deep Well Lid with 560 Orifice Trans Blue | Medium Deep Well With 560 Orifice Trans Blue | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Pass (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ471LI/Lid_04F_560 | 2BQ471LI | Medium Deep Well Lid with 560 Orifice Trans Blue | Medium Deep Well With 560 Orifice Trans Blue | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter | 3.7695 | in | FDA,3 | 3.76 | 3.77 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ471LI/Lid_04F_560 | 2BQ471LI | Medium Deep Well Lid with 560 Orifice Trans Blue | Medium Deep Well With 560 Orifice Trans Blue | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness | 0.046 | in | FDA,3 | 0.04 | 0.05 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ471LI/Lid_04F_560 | 2BQ471LI | Medium Deep Well Lid with 560 Orifice Trans Blue | Medium Deep Well With 560 Orifice Trans Blue | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ471LI/Lid_04F_560 | 2BQ471LI | Medium Deep Well Lid with 560 Orifice Trans Blue | Medium Deep Well With 560 Orifice Trans Blue | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height | 1.486 | in | FDA,3 | 1.473 | 1.493 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ471LI/Lid_04F_560 | 2BQ471LI | Medium Deep Well Lid with 560 Orifice Trans Blue | Medium Deep Well With 560 Orifice Trans Blue | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Outer Diameter | 0.555 | in | FDA,3 | 0.555 | 0.565 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ471LI/Lid_04F_560 | 2BQ471LI | Medium Deep Well Lid with 560 Orifice Trans Blue | Medium Deep Well With 560 Orifice Trans Blue | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Thicknes | 0.015 | in | FDA,3 | 0.015 | 0.025 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ471LI/Lid_04F_560 | 2BQ471LI | Medium Deep Well Lid with 560 Orifice Trans Blue | Medium Deep Well With 560 Orifice Trans Blue | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Weight | 22.425 | g | FDA,1 | 20 | 24 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ551LI/Skip Lot Release_NY_2OQ850 | 2BQ551LI | Large Deep Well Lid PDI Orifice | Large Deep Well PDI Orifice | Skip Lot Release_NY_2OQ850 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA | COA | NY | RAW MATERIAL - NY | RM_LID | both
2BQ551LI/Applicator_NY_Skip lot testing | 2BQ551LI | Large Deep Well Lid PDI Orifice | Large Deep Well PDI Orifice | Applicator_NY_Skip lot testing | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | [5 empty] | Certificate of Analysis | COA | NY | RAW MATERIAL - NY | RM_LID | both
2BQ551LI/Lid_13A | 2BQ551LI | Large Deep Well Lid PDI Orifice | Large Deep Well PDI Orifice | Lid_13A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Flip Lid Force |  | NY | RAW MATERIAL - NY | RM_LID | both
2BQ551LI/Lid_13A | 2BQ551LI | Large Deep Well Lid PDI Orifice | Large Deep Well PDI Orifice | Lid_13A | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | 1.61 | lb | FDA,2 | 0.75 |  | TM 143 | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ551LI/Lid_13A | 2BQ551LI | Large Deep Well Lid PDI Orifice | Large Deep Well PDI Orifice | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Meets Specifications (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ551LI/Lid_13A | 2BQ551LI | Large Deep Well Lid PDI Orifice | Large Deep Well PDI Orifice | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | 4.8545 | in | FDA,3 | 4.847 | 4.867 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ551LI/Lid_13A | 2BQ551LI | Large Deep Well Lid PDI Orifice | Large Deep Well PDI Orifice | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Center to Center Distance | 0.45 | in | FDA,3 | 0.445 | 0.455 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ551LI/Lid_13A | 2BQ551LI | Large Deep Well Lid PDI Orifice | Large Deep Well PDI Orifice | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Color | Translucent Purple (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ551LI/Lid_13A | 2BQ551LI | Large Deep Well Lid PDI Orifice | Large Deep Well PDI Orifice | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | End Circle Diameter | 0.123 | in | FDA,3 | 0.12 | 0.13 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ551LI/Lid_13A | 2BQ551LI | Large Deep Well Lid PDI Orifice | Large Deep Well PDI Orifice | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Pass (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ551LI/Lid_13A | 2BQ551LI | Large Deep Well Lid PDI Orifice | Large Deep Well PDI Orifice | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Insert Fin Gap | 0.076 | in | FDA,3 | 0.07 | 0.08 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ551LI/Lid_13A | 2BQ551LI | Large Deep Well Lid PDI Orifice | Large Deep Well PDI Orifice | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter | 4.923 | in | FDA,3 | 4.917 | 4.937 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ551LI/Lid_13A | 2BQ551LI | Large Deep Well Lid PDI Orifice | Large Deep Well PDI Orifice | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness | 0.04 | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ551LI/Lid_13A | 2BQ551LI | Large Deep Well Lid PDI Orifice | Large Deep Well PDI Orifice | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ551LI/Lid_13A | 2BQ551LI | Large Deep Well Lid PDI Orifice | Large Deep Well PDI Orifice | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness | 0.04 | in | FDA,3 | 0.03 | 0.04 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ551LI/Lid_13A | 2BQ551LI | Large Deep Well Lid PDI Orifice | Large Deep Well PDI Orifice | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height | 1.483 | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BQ551LI/Lid_13A | 2BQ551LI | Large Deep Well Lid PDI Orifice | Large Deep Well PDI Orifice | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Weight | 34.73 | g | FDA,1 | 32 | 42 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2BS420/Web_NP3870_11.5 | 2BS420 | PDI Povidone-Iodine Swab 1's | Povidone-Iodine Swab 1'S | Web_NP3870_11.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2BS420/Web_NP3870_11.5 | 2BS420 | PDI Povidone-Iodine Swab 1's | Povidone-Iodine Swab 1'S | Web_NP3870_11.5 | NY-Analytical | 761 | Web Measurement | Web Width | 11.5 | in | FDA,4 | 11.4375 | 11.5625 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2BS420/Web_NP3870_11.5 | 2BS420 | PDI Povidone-Iodine Swab 1's | Povidone-Iodine Swab 1'S | Web_NP3870_11.5 | NY-Analytical | 756 | Basis Weight | Basis Weight | 91.8 | #/Ream | FDA,1 | 81 | 99 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2BS420/Web_NP3870_11.5 | 2BS420 | PDI Povidone-Iodine Swab 1's | Povidone-Iodine Swab 1'S | Web_NP3870_11.5 | NY-Analytical | 762 | Seal Strength | Seal Strength (@350°F/40psi/1.0 sec dwell) | 10192.62 | g/in | FDA,0 | 5000 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2BU275/Web_NP3895_8.5 | 2BU275 | PDI Sani-Cloth AF3 | Sani-Cloth AF3 | Web_NP3895_8.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2BU275/Web_NP3895_8.5 | 2BU275 | PDI Sani-Cloth AF3 | Sani-Cloth AF3 | Web_NP3895_8.5 | NY-Analytical | 761 | Web Measurement | Web Width | 8.5 | in | FDA,4 | 8.4375 | 8.5625 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2BU275/Web_NP3895_8.5 | 2BU275 | PDI Sani-Cloth AF3 | Sani-Cloth AF3 | Web_NP3895_8.5 | NY-Analytical | 756 | Basis Weight | Basis Weight | 71.69 | #/3000sqft | FDA,1 | 65.3 | 79.8 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2BU275/Web_NP3895_8.5 | 2BU275 | PDI Sani-Cloth AF3 | Sani-Cloth AF3 | Web_NP3895_8.5 | NY-Analytical | 762 | Seal Strength | Seal Strength (@250°F/30psi/0.5 sec dwell) | 5589 | g/in | FDA,0 | 2270 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2BU275/Web_NP3895_8.5 | 2BU275 | PDI Sani-Cloth AF3 | Sani-Cloth AF3 | Web_NP3895_8.5 | NY-Analytical | 763 | Web Thickness | Web Thickness | 4.05 | mil | FDA,1 | 3.9 | 4.7 | TM 132 | TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2CA580SLI/Lid_01R | 2CA580SLI | Table Turners Cleaning Multi-Surface Wipes | Table Turners Cleaning Multi-Surface Wipes | Lid_01R | NY-Analytical | 765 | Canister & Lid Inspection | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2CA580SLI/Lid_01R | 2CA580SLI | Table Turners Cleaning Multi-Surface Wipes | Table Turners Cleaning Multi-Surface Wipes | Lid_01R | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2CA580SLI/Lid_01R | 2CA580SLI | Table Turners Cleaning Multi-Surface Wipes | Table Turners Cleaning Multi-Surface Wipes | Lid_01R | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2CA580SLI/Lid_01R | 2CA580SLI | Table Turners Cleaning Multi-Surface Wipes | Table Turners Cleaning Multi-Surface Wipes | Lid_01R | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2CA580SLI/Lid_01R | 2CA580SLI | Table Turners Cleaning Multi-Surface Wipes | Table Turners Cleaning Multi-Surface Wipes | Lid_01R | NY-Analytical | 765 | Canister & Lid Inspection | Label Length | 3.73 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2CA580SLI/Lid_01R | 2CA580SLI | Table Turners Cleaning Multi-Surface Wipes | Table Turners Cleaning Multi-Surface Wipes | Lid_01R | NY-Analytical | 765 | Canister & Lid Inspection | Label Width | 1.875 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2CA580SLI/Lid_01R | 2CA580SLI | Table Turners Cleaning Multi-Surface Wipes | Table Turners Cleaning Multi-Surface Wipes | Lid_01R | NY-Analytical | 765 | Canister & Lid Inspection | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2CA735/Web_NP8810_12.125 | 2CA735 | Dove Baby Sensitive Moisture Wipes | Dove Baby Sensitive Moisture Wipes | Web_NP8810_12.125 | NY-Analytical | 1019 | Moisture Vapor Transmission rate | Report result from COA | 5 | GSM | FDA,0 |  | 10 | Report result from COA | 0 | NY | RAW MATERIAL - NY | RM_WEB | both
2CA735/Web_NP8810_12.125 | 2CA735 | Dove Baby Sensitive Moisture Wipes | Dove Baby Sensitive Moisture Wipes | Web_NP8810_12.125 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Dyne Level | Dyne Level | NY | RAW MATERIAL - NY | RM_WEB | both
2CA735/Web_NP8810_12.125 | 2CA735 | Dove Baby Sensitive Moisture Wipes | Dove Baby Sensitive Moisture Wipes | Web_NP8810_12.125 | NY-Analytical | 836 | Dyne Level | Dyne Level (outside film) | 36 |  | FDA,0 | 36 | 38 | Dyne Test Kit | Dyne_Test_Kit | NY | RAW MATERIAL - NY | RM_WEB | both
2CA735/Web_NP8810_12.125 | 2CA735 | Dove Baby Sensitive Moisture Wipes | Dove Baby Sensitive Moisture Wipes | Web_NP8810_12.125 | NY-Analytical | 761 | Web Measurement | Web Width | 12.125 | in | FDA,4 | 12.0625 | 12.1875 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2CA735/Web_NP8810_12.125 | 2CA735 | Dove Baby Sensitive Moisture Wipes | Dove Baby Sensitive Moisture Wipes | Web_NP8810_12.125 | NY-Analytical | 756 | Basis Weight | Basis Weight | 42.33 | #/Ream | FDA,1 | 39.4 | 48.2 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2CA735/Web_NP8810_12.125 | 2CA735 | Dove Baby Sensitive Moisture Wipes | Dove Baby Sensitive Moisture Wipes | Web_NP8810_12.125 | NY-Analytical | 762 | Seal Strength | Seal Strength (@350°F/30psi/0.5s dwell Top Jaw) | 5303 | gms/f | FDA,0 | 500 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2CA735/Web_NP8810_12.125 | 2CA735 | Dove Baby Sensitive Moisture Wipes | Dove Baby Sensitive Moisture Wipes | Web_NP8810_12.125 | NY-Analytical | 763 | Web Thickness | Web Thickness | 2.7 | mil | FDA,2 | 2.48 | 3.03 | TM 132 | TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2CA735/Web_NP8810_12.125 | 2CA735 | Dove Baby Sensitive Moisture Wipes | Dove Baby Sensitive Moisture Wipes | Web_NP8810_12.125 | NY-Analytical | 397 | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_WEB | both
2CB196/Web_NP5114_41.0 | 2CB196 | PDI Prevantics Device Swab | Prevantics Device Swab | Web_NP5114_41.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2CB196/Web_NP5114_41.0 | 2CB196 | PDI Prevantics Device Swab | Prevantics Device Swab | Web_NP5114_41.0 | NY-Analytical | 761 | Web Measurement | Web Width | 40.9375 | in | FDA,4 | 40.9375 | 41.0625 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2CB196/Web_NP5114_41.0 | 2CB196 | PDI Prevantics Device Swab | Prevantics Device Swab | Web_NP5114_41.0 | NY-Analytical | 756 | Basis Weight | Basis Weight | 59.79 | #/Ream | FDA,1 | 53.1 | 64.9 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2CB196/Web_NP5114_41.0 | 2CB196 | PDI Prevantics Device Swab | Prevantics Device Swab | Web_NP5114_41.0 | NY-Analytical | 762 | Seal Strength | Seal Strength (@300°F/40psi/0.5 sec dewll) | 2891.81 | g/in | FDA,0 | 1300 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2CB196/Web_NP5114_41.0 | 2CB196 | PDI Prevantics Device Swab | Prevantics Device Swab | Web_NP5114_41.0 | NY-Analytical | 763 | Web Thickness | Web Thickness | 3.2 | mil | FDA,2 | 3.15 | 3.85 | TM 132 | TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2CB512/Web_NP3870_4.5 | 2CB512 | PDI Povidone-Iodine Prep Pad | Povidone-Iodine Prep Pad | Web_NP3870_4.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2CB512/Web_NP3870_4.5 | 2CB512 | PDI Povidone-Iodine Prep Pad | Povidone-Iodine Prep Pad | Web_NP3870_4.5 | NY-Analytical | 761 | Web Measurement | Web Width | 4.5 | in | FDA,4 | 4.4375 | 4.5625 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2CB512/Web_NP3870_4.5 | 2CB512 | PDI Povidone-Iodine Prep Pad | Povidone-Iodine Prep Pad | Web_NP3870_4.5 | NY-Analytical | 756 | Basis Weight | Basis Weight | 92.6 | #/Ream | FDA,1 | 81 | 99 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2CB512/Web_NP3870_4.5 | 2CB512 | PDI Povidone-Iodine Prep Pad | Povidone-Iodine Prep Pad | Web_NP3870_4.5 | NY-Analytical | 762 | Seal Strength | Seal Strength (@350°F/40psi/1.0 sec dwell) | 9864.87 | g/in | FDA,0 | 5000 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2DH211/Web_NP5116_6.0 | 2DH211 | CVS Total Home Antibacterial Moist Wipes | CVS Total Home AntibaCt. erial Moist Wipes | Web_NP5116_6.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2DH211/Web_NP5116_6.0 | 2DH211 | CVS Total Home Antibacterial Moist Wipes | CVS Total Home AntibaCt. erial Moist Wipes | Web_NP5116_6.0 | NY-Analytical | 761 | Web Measurement | Web Width | 6 | in | FDA,4 | 5.9375 | 6.0625 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2DH211/Web_NP5116_6.0 | 2DH211 | CVS Total Home Antibacterial Moist Wipes | CVS Total Home AntibaCt. erial Moist Wipes | Web_NP5116_6.0 | NY-Analytical | 756 | Basis Weight | Basis Weight | 57.67 | #/Ream | FDA,1 | 55.8 | 68.2 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2DH211/Web_NP5116_6.0 | 2DH211 | CVS Total Home Antibacterial Moist Wipes | CVS Total Home AntibaCt. erial Moist Wipes | Web_NP5116_6.0 | NY-Analytical | 762 | Seal Strength | Seal Strength (@300°F/40psi/0.5 sec dewll) | 1800.28 | g/in | FDA,0 | 1500 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2ED110/Web_NP5410_4.5 | 2ED110 | WEB ORIGINAL W/N 4.5"/5410 | Web Original W/N 4.5"/5410 | Web_NP5410_4.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2ED110/Web_NP5410_4.5 | 2ED110 | WEB ORIGINAL W/N 4.5"/5410 | Web Original W/N 4.5"/5410 | Web_NP5410_4.5 | NY-Analytical | 761 | Web Measurement | Web Width | 4.5 | in | FDA,4 | 4.4375 | 4.5625 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2ED110/Web_NP5410_4.5 | 2ED110 | WEB ORIGINAL W/N 4.5"/5410 | Web Original W/N 4.5"/5410 | Web_NP5410_4.5 | NY-Analytical | 756 | Basis Weight | Basis Weight | 54.5 | #/Ream | FDA,1 | 49.8 | 60.8 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2ED110/Web_NP5410_4.5 | 2ED110 | WEB ORIGINAL W/N 4.5"/5410 | Web Original W/N 4.5"/5410 | Web_NP5410_4.5 | NY-Analytical | 762 | Seal Strength | Destruct Bond | Break occurs off the seal (Done) | g/in |  |  |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2ED110/Web_NP5410_4.5 | 2ED110 | WEB ORIGINAL W/N 4.5"/5410 | Web Original W/N 4.5"/5410 | Web_NP5410_4.5 | NY-Analytical | 762 | Seal Strength | Seal Strength (@300°F/40psi/0.5 sec dewll) | 1814.1 | g/in | FDA,0 | 1500 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2EU351/Web_NP5211_7.0 | 2EU351 | Qbrexza Web 1ct | Qbrexza Web 1Ct. | Web_NP5211_7.0 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Eyespot Repeat |  | NY | RAW MATERIAL - NY | RM_WEB | both
2EU351/Web_NP5211_7.0 | 2EU351 | Qbrexza Web 1ct | Qbrexza Web 1Ct. | Web_NP5211_7.0 | NY-Analytical | 999 | Eyespot Repeat | Eye spot repeat over 7 impressions | 488 | mm | FDA,0 | 488 | 490 | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_WEB | both
2EU351/Web_NP5211_7.0 | 2EU351 | Qbrexza Web 1ct | Qbrexza Web 1Ct. | Web_NP5211_7.0 | NY-Analytical | 761 | Web Measurement | Web Width | 7 | in | FDA,4 | 6.9375 | 7.0625 | TM 143 | TM143 | NY | RAW MATERIAL - NY | RM_WEB | both
2EU351/Web_NP5211_7.0 | 2EU351 | Qbrexza Web 1ct | Qbrexza Web 1Ct. | Web_NP5211_7.0 | NY-Analytical | 1031 | 2D Matrix (Barcode) | 2D Matrix (Barcode) | Meet Specifications (Done) |  |  |  |  | Visual | Visual | NY | RAW MATERIAL - NY | RM_WEB | both
2EU351/Web_NP5211_7.0 | 2EU351 | Qbrexza Web 1ct | Qbrexza Web 1Ct. | Web_NP5211_7.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Visual | Visual | NY | RAW MATERIAL - NY | RM_WEB | both
2EU351/Web_NP5211_7.0 | 2EU351 | Qbrexza Web 1ct | Qbrexza Web 1Ct. | Web_NP5211_7.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | COA Meets Specification (Done) |  |  |  |  | Visual | Visual | NY | RAW MATERIAL - NY | RM_WEB | both
2EU351/Web_NP5211_7.0 | 2EU351 | Qbrexza Web 1ct | Qbrexza Web 1Ct. | Web_NP5211_7.0 | NY-Analytical | 748 | Visual Inspection | Appearance | Meets specifications (Done) |  |  |  |  | Visual Observation | Visual | NY | RAW MATERIAL - NY | RM_WEB | both
2EU351/Web_NP5211_7.0 | 2EU351 | Qbrexza Web 1ct | Qbrexza Web 1Ct. | Web_NP5211_7.0 | NY-Analytical | 748 | Visual Inspection | Color/Copy | Meets specifications (Done) |  |  |  |  | Visual Observation | Visual | NY | RAW MATERIAL - NY | RM_WEB | both
2FB6038/Web_NP5410_8.0 | 2FB6038 | PDI Alcohol Prep Pad | PDI Alcohol Prep Pad | Web_NP5410_8.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2FB6038/Web_NP5410_8.0 | 2FB6038 | PDI Alcohol Prep Pad | PDI Alcohol Prep Pad | Web_NP5410_8.0 | NY-Analytical | 761 | Web Measurement | Web Width |  | in | FDA,4 | 7.9375 | 8.0625 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2FB6038/Web_NP5410_8.0 | 2FB6038 | PDI Alcohol Prep Pad | PDI Alcohol Prep Pad | Web_NP5410_8.0 | NY-Analytical | 756 | Basis Weight | Basis Weight | 52.91 | #/Ream | FDA,1 | 49.8 | 60.8 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2FB6038/Web_NP5410_8.0 | 2FB6038 | PDI Alcohol Prep Pad | PDI Alcohol Prep Pad | Web_NP5410_8.0 | NY-Analytical | 762 | Seal Strength | Destruct Bond |  | g/in |  |  |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2FB6038/Web_NP5410_8.0 | 2FB6038 | PDI Alcohol Prep Pad | PDI Alcohol Prep Pad | Web_NP5410_8.0 | NY-Analytical | 762 | Seal Strength | Seal Strength (@300°F/40psi/0.5 sec dewll) |  | g/in | FDA,0 | 1500 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2FH070/Web_NP5115_6.0 | 2FH070 | EZ Mammo Wipe | EZ Mammo Wipe | Web_NP5115_6.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2FH070/Web_NP5115_6.0 | 2FH070 | EZ Mammo Wipe | EZ Mammo Wipe | Web_NP5115_6.0 | NY-Analytical | 761 | Web Measurement | Web Width | 6 | in | FDA,4 | 5.9375 | 6.0625 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2FH070/Web_NP5115_6.0 | 2FH070 | EZ Mammo Wipe | EZ Mammo Wipe | Web_NP5115_6.0 | NY-Analytical | 756 | Basis Weight | Basis Weight | 59.66 | #/Ream | FDA,1 | 55.6 | 68 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2FH070/Web_NP5115_6.0 | 2FH070 | EZ Mammo Wipe | EZ Mammo Wipe | Web_NP5115_6.0 | NY-Analytical | 762 | Seal Strength | Seal Strength (@325°F/40PSI/0.5Sec Dwell) | 2304.16 | g/in | FDA,0 | 1500 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2FJ141R/Web_NP8535_16 | 2FJ141R | Hygea Adult Flow Wrap | Hygea Adult Flow Wrap | Web_NP8535_16 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2FJ141R/Web_NP8535_16 | 2FJ141R | Hygea Adult Flow Wrap | Hygea Adult Flow Wrap | Web_NP8535_16 | NY-Analytical | 761 | Web Measurement | Web Width | 16 | in | FDA,4 | 15.9375 | 16.0625 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2FJ141R/Web_NP8535_16 | 2FJ141R | Hygea Adult Flow Wrap | Hygea Adult Flow Wrap | Web_NP8535_16 | NY-Analytical | 756 | Basis Weight | Basis Weight | 42.06 | #/Ream | FDA,2 | 35.82 | 45.6 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2FJ141R/Web_NP8535_16 | 2FJ141R | Hygea Adult Flow Wrap | Hygea Adult Flow Wrap | Web_NP8535_16 | NY-Analytical | 762 | Seal Strength | Seal Strength (@300°F/30psi/1.5 sec dwell) | 7567 | g/in | FDA,0 | 3500 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2FJ141R/Web_NP8535_16 | 2FJ141R | Hygea Adult Flow Wrap | Hygea Adult Flow Wrap | Web_NP8535_16 | NY-Analytical | 763 | Web Thickness | Web Thickness | 2.45 | mil | FDA,1 | 2.2 | 2.8 | TM 132 | TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2HA580/Web_NP8540_17.0 | 2HA580 | Sani Table Turners Web | Sani-Table Turners Web | Web_NP8540_17.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2HA580/Web_NP8540_17.0 | 2HA580 | Sani Table Turners Web | Sani-Table Turners Web | Web_NP8540_17.0 | NY-Analytical | 761 | Web Measurement | Web Width | 17 | in | FDA,4 | 16.9375 | 17.0625 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2HA580/Web_NP8540_17.0 | 2HA580 | Sani Table Turners Web | Sani-Table Turners Web | Web_NP8540_17.0 | NY-Analytical | 756 | Basis Weight | Basis Weight | 52.91 | #/Ream | FDA,1 | 45.4 | 57.8 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2HA580/Web_NP8540_17.0 | 2HA580 | Sani Table Turners Web | Sani-Table Turners Web | Web_NP8540_17.0 | NY-Analytical | 762 | Seal Strength | Seal Strength (@300°F/30psi/1.5 sec dwell) | 7254 | g/in | FDA,0 | 3500 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2HA580/Web_NP8540_17.0 | 2HA580 | Sani Table Turners Web | Sani-Table Turners Web | Web_NP8540_17.0 | NY-Analytical | 763 | Web Thickness | Web Thickness | 3 | mil | FDA,2 | 2.86 | 3.64 | TM 132 | TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2HU120/Web_NP5115_8.5 | 2HU120 | PDI Hygea Personal Cleansing Washcloths | PDI Hygea Personal Cleansing Washcloths | Web_NP5115_8.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2HU120/Web_NP5115_8.5 | 2HU120 | PDI Hygea Personal Cleansing Washcloths | PDI Hygea Personal Cleansing Washcloths | Web_NP5115_8.5 | NY-Analytical | 761 | Web Measurement | Web Width | 8.5 | in | FDA,4 | 8.4375 | 8.5625 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2HU120/Web_NP5115_8.5 | 2HU120 | PDI Hygea Personal Cleansing Washcloths | PDI Hygea Personal Cleansing Washcloths | Web_NP5115_8.5 | NY-Analytical | 756 | Basis Weight | Basis Weight | 58.47 | #/Ream | FDA,1 | 55.6 | 68 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2HU120/Web_NP5115_8.5 | 2HU120 | PDI Hygea Personal Cleansing Washcloths | PDI Hygea Personal Cleansing Washcloths | Web_NP5115_8.5 | NY-Analytical | 762 | Seal Strength | Seal Strength (@325°F/40PSI/0.5Sec Dwell) | 2199 | g/in | FDA,0 | 1500 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2JC699/Web_NP5410_6.0 | 2JC699 | PDI Alcohol Prep Pad | PDI Alcohol Prep Pad | Web_NP5410_6.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2JC699/Web_NP5410_6.0 | 2JC699 | PDI Alcohol Prep Pad | PDI Alcohol Prep Pad | Web_NP5410_6.0 | NY-Analytical | 761 | Web Measurement | Web Width | 6 | in | FDA,4 | 5.9375 | 6.0625 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2JC699/Web_NP5410_6.0 | 2JC699 | PDI Alcohol Prep Pad | PDI Alcohol Prep Pad | Web_NP5410_6.0 | NY-Analytical | 756 | Basis Weight | Basis Weight | 55.6 | #/Ream | FDA,1 | 49.8 | 60.8 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2JC699/Web_NP5410_6.0 | 2JC699 | PDI Alcohol Prep Pad | PDI Alcohol Prep Pad | Web_NP5410_6.0 | NY-Analytical | 762 | Seal Strength | Destruct Bond | Break occurs off the seal (Done) | g/in |  |  |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2JC699/Web_NP5410_6.0 | 2JC699 | PDI Alcohol Prep Pad | PDI Alcohol Prep Pad | Web_NP5410_6.0 | NY-Analytical | 762 | Seal Strength | Seal Strength (@300°F/40psi/0.5 sec dewll) | 1686.57 | g/in | FDA,0 | 1500 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2LB406/Web_NP3870_10.0 | 2LB406 | PDI PVP Prep Pad Med | PVP Prep Pad Med | Web_NP3870_10.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2LB406/Web_NP3870_10.0 | 2LB406 | PDI PVP Prep Pad Med | PVP Prep Pad Med | Web_NP3870_10.0 | NY-Analytical | 761 | Web Measurement | Web Width | 10 | in | FDA,4 | 9.9375 | 10.0625 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2LB406/Web_NP3870_10.0 | 2LB406 | PDI PVP Prep Pad Med | PVP Prep Pad Med | Web_NP3870_10.0 | NY-Analytical | 756 | Basis Weight | Basis Weight | 93.53 | #/Ream | FDA,1 | 83.7 | 96.3 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2LB406/Web_NP3870_10.0 | 2LB406 | PDI PVP Prep Pad Med | PVP Prep Pad Med | Web_NP3870_10.0 | NY-Analytical | 762 | Seal Strength | Seal Strength (@350°F/40psi/1.0 sec dwell) | 10196.54 | g/in | FDA,0 | 5000 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Summer's Eve Fresh Cycle Cleans Cloths | Summer's Eve Fresh Cycle Cleans Cloths | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 765 | Canister & Lid Inspection | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Summer's Eve Fresh Cycle Cleans Cloths | Summer's Eve Fresh Cycle Cleans Cloths | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Summer's Eve Fresh Cycle Cleans Cloths | Summer's Eve Fresh Cycle Cleans Cloths | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 818 | Solo Lid Inspection | Appearance | Meet specification (Done) |  |  |  |  | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Summer's Eve Fresh Cycle Cleans Cloths | Summer's Eve Fresh Cycle Cleans Cloths | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 818 | Solo Lid Inspection | Color | Meets the specification (Done) |  |  |  |  | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Summer's Eve Fresh Cycle Cleans Cloths | Summer's Eve Fresh Cycle Cleans Cloths | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 818 | Solo Lid Inspection | Functionality | Meet specification (Done) |  |  |  |  | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Summer's Eve Fresh Cycle Cleans Cloths | Summer's Eve Fresh Cycle Cleans Cloths | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 818 | Solo Lid Inspection | Lid Wall Thickness | 0.0285 | in | FDA,3 | 0.025 | 0.035 | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Summer's Eve Fresh Cycle Cleans Cloths | Summer's Eve Fresh Cycle Cleans Cloths | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 818 | Solo Lid Inspection | Odor | None detectable (Done) |  |  |  |  | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Summer's Eve Fresh Cycle Cleans Cloths | Summer's Eve Fresh Cycle Cleans Cloths | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 818 | Solo Lid Inspection | Open Lid Overall Width | 4.367 | in | FDA,3 | 4.351 | 4.401 | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Summer's Eve Fresh Cycle Cleans Cloths | Summer's Eve Fresh Cycle Cleans Cloths | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 818 | Solo Lid Inspection | Overall closed lid height | 0.3688 | in | FDA,3 | 0.355 | 0.385 | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Summer's Eve Fresh Cycle Cleans Cloths | Summer's Eve Fresh Cycle Cleans Cloths | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 818 | Solo Lid Inspection | Weight | 11.33 | g | FDA,1 | 10.5 | 12.5 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_LID | both
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Summer's Eve Fresh Cycle Cleans Cloths | Summer's Eve Fresh Cycle Cleans Cloths | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Summer's Eve Fresh Cycle Cleans Cloths | Summer's Eve Fresh Cycle Cleans Cloths | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 765 | Canister & Lid Inspection | Label Length | 3.7187 | in | FDA,4 | 3.6875 | 3.8125 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | 2OA347SLI | Summer's Eve Fresh Cycle Cleans Cloths | Summer's Eve Fresh Cycle Cleans Cloths | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 765 | Canister & Lid Inspection | Label Width | 1.875 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Prep-H Medicated Wipes Solo | Prep-H Medicated Wipes Solo | Lid_NY_NP7600_22S | NY-Analytical | 765 | Canister & Lid Inspection | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection | TM143 | IN | RAW MATERIAL - INDIANA | RM_LID | both
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Prep-H Medicated Wipes Solo | Prep-H Medicated Wipes Solo | Lid_NY_NP7600_22S | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Solo Lid Inspection | SOLO | IN | RAW MATERIAL - INDIANA | RM_LID | both
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Prep-H Medicated Wipes Solo | Prep-H Medicated Wipes Solo | Lid_NY_NP7600_22S | NY-Analytical | 818 | Solo Lid Inspection | Appearance | Meet specification (Done) |  |  |  |  | Solo Lid Inspection | SOLO | IN | RAW MATERIAL - INDIANA | RM_LID | both
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Prep-H Medicated Wipes Solo | Prep-H Medicated Wipes Solo | Lid_NY_NP7600_22S | NY-Analytical | 818 | Solo Lid Inspection | Color | Teal Pantone 3145C |  |  |  |  | Solo Lid Inspection | SOLO | IN | RAW MATERIAL - INDIANA | RM_LID | both
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Prep-H Medicated Wipes Solo | Prep-H Medicated Wipes Solo | Lid_NY_NP7600_22S | NY-Analytical | 818 | Solo Lid Inspection | Functionality | Meet specification (Done) |  |  |  |  | Solo Lid Inspection | SOLO | IN | RAW MATERIAL - INDIANA | RM_LID | both
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Prep-H Medicated Wipes Solo | Prep-H Medicated Wipes Solo | Lid_NY_NP7600_22S | NY-Analytical | 818 | Solo Lid Inspection | Lid Wall Thickness | 0.028 | in | FDA,3 | 0.025 | 0.035 | Solo Lid Inspection | SOLO | IN | RAW MATERIAL - INDIANA | RM_LID | both
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Prep-H Medicated Wipes Solo | Prep-H Medicated Wipes Solo | Lid_NY_NP7600_22S | NY-Analytical | 818 | Solo Lid Inspection | Odor | None detectable (Done) |  |  |  |  | Solo Lid Inspection | SOLO | IN | RAW MATERIAL - INDIANA | RM_LID | both
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Prep-H Medicated Wipes Solo | Prep-H Medicated Wipes Solo | Lid_NY_NP7600_22S | NY-Analytical | 818 | Solo Lid Inspection | Open Lid Overall Width | 4.361 | in | FDA,3 | 4.351 | 4.401 | Solo Lid Inspection | SOLO | IN | RAW MATERIAL - INDIANA | RM_LID | both
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Prep-H Medicated Wipes Solo | Prep-H Medicated Wipes Solo | Lid_NY_NP7600_22S | NY-Analytical | 818 | Solo Lid Inspection | Overall closed lid height | 0.3815 | in | FDA,3 | 0.355 | 0.385 | Solo Lid Inspection | SOLO | IN | RAW MATERIAL - INDIANA | RM_LID | both
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Prep-H Medicated Wipes Solo | Prep-H Medicated Wipes Solo | Lid_NY_NP7600_22S | NY-Analytical | 818 | Solo Lid Inspection | Weight | 11.26 | g | FDA,1 | 10.5 | 12.5 | Measurement | Measurement | IN | RAW MATERIAL - INDIANA | RM_LID | both
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Prep-H Medicated Wipes Solo | Prep-H Medicated Wipes Solo | Lid_NY_NP7600_22S | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual | TM143 | IN | RAW MATERIAL - INDIANA | RM_LID | both
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Prep-H Medicated Wipes Solo | Prep-H Medicated Wipes Solo | Lid_NY_NP7600_22S | NY-Analytical | 765 | Canister & Lid Inspection | Label Length | 3.75 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | TM143 | IN | RAW MATERIAL - INDIANA | RM_LID | both
2OA592SLI/Lid_NY_NP7600_22S | 2OA592SLI | Prep-H Medicated Wipes Solo | Prep-H Medicated Wipes Solo | Lid_NY_NP7600_22S | NY-Analytical | 765 | Canister & Lid Inspection | Label Width | 1.875 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | TM143 | IN | RAW MATERIAL - INDIANA | RM_LID | both
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Preparation H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 765 | Canister & Lid Inspection | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Preparation H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Preparation H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 818 | Solo Lid Inspection | Color | Meets the specification (Done) |  |  |  |  | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Preparation H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 818 | Solo Lid Inspection | Functionality | Meet specification (Done) |  |  |  |  | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Preparation H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 818 | Solo Lid Inspection | Lid Wall Thickness | 0.027 | in | FDA,3 | 0.025 | 0.035 | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Preparation H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 818 | Solo Lid Inspection | Open Lid Overall Width | 4.358 | in | FDA,3 | 4.351 | 4.401 | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Preparation H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 818 | Solo Lid Inspection | Overall closed lid height | 0.3785 | in | FDA,3 | 0.355 | 0.385 | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Preparation H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 818 | Solo Lid Inspection | Weight | 11.235 | g | FDA,1 | 10.5 | 12.5 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_LID | both
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Preparation H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Preparation H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Preparation H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 765 | Canister & Lid Inspection | Label Length | 3.69 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Preparation H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 765 | Canister & Lid Inspection | Label Width | 1.83 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | 2OA632SLI | Preparation H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 765 | Canister & Lid Inspection | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OJ124SLI/Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | 2OJ124SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | NY-Analytical | 765 | Canister & Lid Inspection | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OJ124SLI/Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | 2OJ124SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OJ124SLI/Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | 2OJ124SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OJ124SLI/Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | 2OJ124SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OJ124SLI/Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | 2OJ124SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | NY-Analytical | 765 | Canister & Lid Inspection | Label Length | 3.7187 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OJ124SLI/Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | 2OJ124SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | NY-Analytical | 765 | Canister & Lid Inspection | Label Width | 1.875 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OJ124SLI/Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | 2OJ124SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | NY-Analytical | 765 | Canister & Lid Inspection | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22Y | NY-Analytical | 765 | Canister & Lid Inspection | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22Y | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22Y | NY-Analytical | 818 | Solo Lid Inspection | Color | Meets the specification (Done) |  |  |  |  | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22Y | NY-Analytical | 818 | Solo Lid Inspection | Functionality | Meet specification (Done) |  |  |  |  | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22Y | NY-Analytical | 818 | Solo Lid Inspection | Lid Wall Thickness | 0.0285 | in | FDA,3 | 0.025 | 0.035 | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22Y | NY-Analytical | 818 | Solo Lid Inspection | Open Lid Overall Width | 4.373 | in | FDA,3 | 4.351 | 4.401 | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22Y | NY-Analytical | 818 | Solo Lid Inspection | Overall closed lid height | 0.3715 | in | FDA,3 | 0.355 | 0.385 | Solo Lid Inspection | SOLO | NY | RAW MATERIAL - NY | RM_LID | both
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22Y | NY-Analytical | 818 | Solo Lid Inspection | Weight | 11.4 | g | FDA,1 | 10.5 | 12.5 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_LID | both
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22Y | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22Y | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22Y | NY-Analytical | 765 | Canister & Lid Inspection | Label Length | 3.7187 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22Y | NY-Analytical | 765 | Canister & Lid Inspection | Label Width | 1.875 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OJ130SLI/Lid_NY_NP7600_22Y | 2OJ130SLI | Cetaphil Gentle Skin Cleansing Wipes Solo | Cetaphil Gentle Skin Cleansing Wipes Solo | Lid_NY_NP7600_22Y | NY-Analytical | 765 | Canister & Lid Inspection | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OJ213SLI/Lid_01A | 2OJ213SLI | LID SOLO HYGEA MULTIPURPOSE | Solo Hygea Multipurpose | Lid_01A | NY-Analytical | 765 | Canister & Lid Inspection | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OJ213SLI/Lid_01A | 2OJ213SLI | LID SOLO HYGEA MULTIPURPOSE | Solo Hygea Multipurpose | Lid_01A | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OJ213SLI/Lid_01A | 2OJ213SLI | LID SOLO HYGEA MULTIPURPOSE | Solo Hygea Multipurpose | Lid_01A | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OJ213SLI/Lid_01A | 2OJ213SLI | LID SOLO HYGEA MULTIPURPOSE | Solo Hygea Multipurpose | Lid_01A | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OJ213SLI/Lid_01A | 2OJ213SLI | LID SOLO HYGEA MULTIPURPOSE | Solo Hygea Multipurpose | Lid_01A | NY-Analytical | 765 | Canister & Lid Inspection | Label Length | 1.875 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OJ213SLI/Lid_01A | 2OJ213SLI | LID SOLO HYGEA MULTIPURPOSE | Solo Hygea Multipurpose | Lid_01A | NY-Analytical | 765 | Canister & Lid Inspection | Label Width | 3.6875 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OJ213SLI/Lid_01A | 2OJ213SLI | LID SOLO HYGEA MULTIPURPOSE | Solo Hygea Multipurpose | Lid_01A | NY-Analytical | 765 | Canister & Lid Inspection | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OM304/Web_NP8540_16.0 | 2OM304 | Sani Pro - Table Turners Web 72ct. | Sani Professional - Table Turners Web 72Ct. | Web_NP8540_16.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2OM304/Web_NP8540_16.0 | 2OM304 | Sani Pro - Table Turners Web 72ct. | Sani Professional - Table Turners Web 72Ct. | Web_NP8540_16.0 | NY-Analytical | 761 | Web Measurement | Web Width | 16 | in | FDA,4 | 15.9375 | 16.0625 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2OM304/Web_NP8540_16.0 | 2OM304 | Sani Pro - Table Turners Web 72ct. | Sani Professional - Table Turners Web 72Ct. | Web_NP8540_16.0 | NY-Analytical | 756 | Basis Weight | Basis Weight | 53.176 | #/Ream | FDA,1 | 45.4 | 57.8 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2OM304/Web_NP8540_16.0 | 2OM304 | Sani Pro - Table Turners Web 72ct. | Sani Professional - Table Turners Web 72Ct. | Web_NP8540_16.0 | NY-Analytical | 762 | Seal Strength | Seal Strength (@300°F/30psi/1.5 sec dwell) | 7038.28 | g/in | FDA,0 | 3500 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2OM304/Web_NP8540_16.0 | 2OM304 | Sani Pro - Table Turners Web 72ct. | Sani Professional - Table Turners Web 72Ct. | Web_NP8540_16.0 | NY-Analytical | 763 | Web Thickness | Web Thickness | 3.3 | mil | FDA,2 | 2.86 | 3.64 | TM 132 | TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2OM304SLI/Lid_No-Rinse sanitizing Solo | 2OM304SLI | Sani Professional No-Rinse Sanitizing Solo | Sani Professional No-Rinse Sanitizing Solo | Lid_No-Rinse sanitizing Solo | NY-Analytical | 765 | Canister & Lid Inspection | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OM304SLI/Lid_No-Rinse sanitizing Solo | 2OM304SLI | Sani Professional No-Rinse Sanitizing Solo | Sani Professional No-Rinse Sanitizing Solo | Lid_No-Rinse sanitizing Solo | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OM304SLI/Lid_No-Rinse sanitizing Solo | 2OM304SLI | Sani Professional No-Rinse Sanitizing Solo | Sani Professional No-Rinse Sanitizing Solo | Lid_No-Rinse sanitizing Solo | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OM304SLI/Lid_No-Rinse sanitizing Solo | 2OM304SLI | Sani Professional No-Rinse Sanitizing Solo | Sani Professional No-Rinse Sanitizing Solo | Lid_No-Rinse sanitizing Solo | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OM304SLI/Lid_No-Rinse sanitizing Solo | 2OM304SLI | Sani Professional No-Rinse Sanitizing Solo | Sani Professional No-Rinse Sanitizing Solo | Lid_No-Rinse sanitizing Solo | NY-Analytical | 765 | Canister & Lid Inspection | Label Length | 3.7187 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OM304SLI/Lid_No-Rinse sanitizing Solo | 2OM304SLI | Sani Professional No-Rinse Sanitizing Solo | Sani Professional No-Rinse Sanitizing Solo | Lid_No-Rinse sanitizing Solo | NY-Analytical | 765 | Canister & Lid Inspection | Label Width | 1.875 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OM304SLI/Lid_No-Rinse sanitizing Solo | 2OM304SLI | Sani Professional No-Rinse Sanitizing Solo | Sani Professional No-Rinse Sanitizing Solo | Lid_No-Rinse sanitizing Solo | NY-Analytical | 765 | Canister & Lid Inspection | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | 2OM701SLI | Sani Professional No-Rinse Sanitizing Solo | Sani Professional No-Rinse Sanitizing Solo | Lid_No-Rinse sanitizing Solo_01 | NY-Analytical | 765 | Canister & Lid Inspection | Print | No smears, streaks, |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | 2OM701SLI | Sani Professional No-Rinse Sanitizing Solo | Sani Professional No-Rinse Sanitizing Solo | Lid_No-Rinse sanitizing Solo_01 | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | 2OM701SLI | Sani Professional No-Rinse Sanitizing Solo | Sani Professional No-Rinse Sanitizing Solo | Lid_No-Rinse sanitizing Solo_01 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | 2OM701SLI | Sani Professional No-Rinse Sanitizing Solo | Sani Professional No-Rinse Sanitizing Solo | Lid_No-Rinse sanitizing Solo_01 | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Copy/color is in accordance with the item # print specification (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | 2OM701SLI | Sani Professional No-Rinse Sanitizing Solo | Sani Professional No-Rinse Sanitizing Solo | Lid_No-Rinse sanitizing Solo_01 | NY-Analytical | 765 | Canister & Lid Inspection | Label Length | 3.687 | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | 2OM701SLI | Sani Professional No-Rinse Sanitizing Solo | Sani Professional No-Rinse Sanitizing Solo | Lid_No-Rinse sanitizing Solo_01 | NY-Analytical | 765 | Canister & Lid Inspection | Label Width | 1.8125 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | 2OM701SLI | Sani Professional No-Rinse Sanitizing Solo | Sani Professional No-Rinse Sanitizing Solo | Lid_No-Rinse sanitizing Solo_01 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | No solvent odor is present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP037/Web_NP8511_245 | 2OP037 | Simple Kind To Skin Facial Cleansing Wipes 7ct US | Simple Kind To Skin Facial Cleansing Wipes 7Ct. US | Web_NP8511_245 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2OP037/Web_NP8511_245 | 2OP037 | Simple Kind To Skin Facial Cleansing Wipes 7ct US | Simple Kind To Skin Facial Cleansing Wipes 7Ct. US | Web_NP8511_245 | NY-Analytical | 761 | Web Measurement | Web Width | 245 | mm | FDA,0 | 243 | 247 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2OP037/Web_NP8511_245 | 2OP037 | Simple Kind To Skin Facial Cleansing Wipes 7ct US | Simple Kind To Skin Facial Cleansing Wipes 7Ct. US | Web_NP8511_245 | NY-Analytical | 756 | Basis Weight | Basis Weight | 47.88 | Lbs/Ream | FDA,1 | 43.1 | 52.7 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2OP037/Web_NP8511_245 | 2OP037 | Simple Kind To Skin Facial Cleansing Wipes 7ct US | Simple Kind To Skin Facial Cleansing Wipes 7Ct. US | Web_NP8511_245 | NY-Analytical | 1048 | Seal strength (NP8511) | Seal Strength (@140°C/1sec/14psi | 6296.05 | g/in | FDA,0 | 2500 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2OP037/Web_NP8511_245 | 2OP037 | Simple Kind To Skin Facial Cleansing Wipes 7ct US | Simple Kind To Skin Facial Cleansing Wipes 7Ct. US | Web_NP8511_245 | NY-Analytical | 763 | Web Thickness | Thickness | 71 | micron | FDA,1 | 64.8 | 79.2 | TM 132 | TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2OP037/Web_NP8511_245 | 2OP037 | Simple Kind To Skin Facial Cleansing Wipes 7ct US | Simple Kind To Skin Facial Cleansing Wipes 7Ct. US | Web_NP8511_245 | NY-Analytical | 885 | Web Inspection | Artwork / Copy | Sample(s) all meet specifications (Done) |  |  |  |  | Visual Observation | Visual | NY | RAW MATERIAL - NY | RM_WEB | both
2OP063/Web_NP8511_320 | 2OP063 | Simple Kind To Skin Facial Cleansing Wipes 25ct Canada | Simple Kind To Skin Facial Cleansing Wipes 25Ct. Canada | Web_NP8511_320 | NY-Analytical | 836 | Dyne Level | Dyne Level (outside film) | 36 |  | FDA,0 |  |  | Dyne Level | Dyne Level | NY | RAW MATERIAL - NY | RM_WEB | both
2OP063/Web_NP8511_320 | 2OP063 | Simple Kind To Skin Facial Cleansing Wipes 25ct Canada | Simple Kind To Skin Facial Cleansing Wipes 25Ct. Canada | Web_NP8511_320 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2OP063/Web_NP8511_320 | 2OP063 | Simple Kind To Skin Facial Cleansing Wipes 25ct Canada | Simple Kind To Skin Facial Cleansing Wipes 25Ct. Canada | Web_NP8511_320 | NY-Analytical | 761 | Web Measurement | Web Width | 320 | mm | FDA,0 | 318 | 322 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2OP063/Web_NP8511_320 | 2OP063 | Simple Kind To Skin Facial Cleansing Wipes 25ct Canada | Simple Kind To Skin Facial Cleansing Wipes 25Ct. Canada | Web_NP8511_320 | NY-Analytical | 756 | Basis Weight | Basis Weight | 47.8 | Lbs/Ream | FDA,1 | 43.1 | 52.7 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2OP063/Web_NP8511_320 | 2OP063 | Simple Kind To Skin Facial Cleansing Wipes 25ct Canada | Simple Kind To Skin Facial Cleansing Wipes 25Ct. Canada | Web_NP8511_320 | NY-Analytical | 762 | Seal Strength | Seal Strength (@140°C/14psi/1 s dwellTopJaw) | 7207.05 | g/in | FDA,0 | 2500 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2OP063/Web_NP8511_320 | 2OP063 | Simple Kind To Skin Facial Cleansing Wipes 25ct Canada | Simple Kind To Skin Facial Cleansing Wipes 25Ct. Canada | Web_NP8511_320 | NY-Analytical | 1048 | Seal strength (NP8511) | Seal Strength (@140°C/1sec/14psi | 5948 | g/in | FDA,0 | 2500 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2OP063/Web_NP8511_320 | 2OP063 | Simple Kind To Skin Facial Cleansing Wipes 25ct Canada | Simple Kind To Skin Facial Cleansing Wipes 25Ct. Canada | Web_NP8511_320 | NY-Analytical | 763 | Web Thickness | Thickness | 74 | micron | FDA,1 | 64.8 | 79.2 | TM 132 | TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2OP065/Web_NP8536_11.875 | 2OP065 | Ponds Evening Soothe Twlts. | Pond's Evening Soothe Twlts. | Web_NP8536_11.875 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2OP065/Web_NP8536_11.875 | 2OP065 | Ponds Evening Soothe Twlts. | Pond's Evening Soothe Twlts. | Web_NP8536_11.875 | NY-Analytical | 761 | Web Measurement | Web Width | 11.875 | in | FDA,4 | 11.8125 | 11.9375 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2OP065/Web_NP8536_11.875 | 2OP065 | Ponds Evening Soothe Twlts. | Pond's Evening Soothe Twlts. | Web_NP8536_11.875 | NY-Analytical | 1044 | Basis Weight (lbs/ream) | Basis Weight | 41.4 | Lb/Ream | FDA,2 | 40.1 | 49.02 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2OP065/Web_NP8536_11.875 | 2OP065 | Ponds Evening Soothe Twlts. | Pond's Evening Soothe Twlts. | Web_NP8536_11.875 | NY-Analytical | 1049 | Seal strength (NP8536) | Seal Strength (@320°F/40psi/0.5 sec dwell) | 8624 | g/in | FDA,0 | 2000 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2OP065/Web_NP8536_11.875 | 2OP065 | Ponds Evening Soothe Twlts. | Pond's Evening Soothe Twlts. | Web_NP8536_11.875 | NY-Analytical | 763 | Web Thickness | Thickness | 2.4 | mil | FDA,3 | 2.367 | 2.893 | TM 132 | TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2OP065/Web_NP8536_11.875 | 2OP065 | Ponds Evening Soothe Twlts. | Pond's Evening Soothe Twlts. | Web_NP8536_11.875 | NY-Analytical | 748 | Visual Inspection | Appearance | Meets specifications (Done) |  |  |  |  | Visual | Visual | NY | RAW MATERIAL - NY | RM_WEB | both
2OP065/Web_NP8536_11.875 | 2OP065 | Ponds Evening Soothe Twlts. | Pond's Evening Soothe Twlts. | Web_NP8536_11.875 | NY-Analytical | 748 | Visual Inspection | Barcode Legibility/Quality | Barcode should scan and code should be matched under the barcode (Done) |  |  |  |  | Visual | Visual | NY | RAW MATERIAL - NY | RM_WEB | both
2OP065/Web_NP8536_11.875 | 2OP065 | Ponds Evening Soothe Twlts. | Pond's Evening Soothe Twlts. | Web_NP8536_11.875 | NY-Analytical | 748 | Visual Inspection | Color/Copy | Meets specifications (Done) |  |  |  |  | Visual | Visual | NY | RAW MATERIAL - NY | RM_WEB | both
2OP072/Web_NP8511_245 | 2OP072 | Simple Kind To Skin Facial Cleansing Wipes 7ct Canada | Simple Kind To Skin Facial Cleansing Wipes 7Ct. Canada | Web_NP8511_245 | NY-Analytical | 836 | Dyne Level | Dyne Level (outside film) | 34 |  | FDA,0 |  |  | Dyne Level | Dyne Level | NY | RAW MATERIAL - NY | RM_WEB | both
2OP073/Web_NP8536_11.875 | 2OP073 | Ponds Luminous Clean Twlts. | Pond's Luminous Clean Twlts. | Web_NP8536_11.875 | NY-Analytical | 756 | Basis Weight | Basis Weight |  | lbs/rm | FDA,2 |  |  | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2OP073/Web_NP8536_11.875 | 2OP073 | Ponds Luminous Clean Twlts. | Pond's Luminous Clean Twlts. | Web_NP8536_11.875 | NY-Analytical | 762 | Seal Strength | Seal Strength (@350°F/40psi/0.5sec dwell) |  | g/in | FDA,0 | 2500 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2OP075/Web_NP8541_12.205 | 2OP075 | L'Oreal Paris Revitalift MUR Cleans. Twlt | L'Oreal Paris Revitalift Mur Cleans. Twlt | Web_NP8541_12.205 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2OP075/Web_NP8541_12.205 | 2OP075 | L'Oreal Paris Revitalift MUR Cleans. Twlt | L'Oreal Paris Revitalift Mur Cleans. Twlt | Web_NP8541_12.205 | NY-Analytical | 761 | Web Measurement | Web Width | 12.187 | in | FDA,4 | 12.1425 | 12.2675 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2OP075/Web_NP8541_12.205 | 2OP075 | L'Oreal Paris Revitalift MUR Cleans. Twlt | L'Oreal Paris Revitalift Mur Cleans. Twlt | Web_NP8541_12.205 | NY-Analytical | 756 | Basis Weight | Basis Weight | 86.11 | GSM | FDA,1 | 74.3 | 90.9 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2OP075/Web_NP8541_12.205 | 2OP075 | L'Oreal Paris Revitalift MUR Cleans. Twlt | L'Oreal Paris Revitalift Mur Cleans. Twlt | Web_NP8541_12.205 | NY-Analytical | 762 | Seal Strength | Seal Strength (@130°C/0.5 sec/1 bar) | 8916 | g/in | FDA,0 | 1000 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2OP075/Web_NP8541_12.205 | 2OP075 | L'Oreal Paris Revitalift MUR Cleans. Twlt | L'Oreal Paris Revitalift Mur Cleans. Twlt | Web_NP8541_12.205 | NY-Analytical | 763 | Web Thickness | Thickness | 0.084 | mm | FDA,3 | 0.074 | 0.094 | TM 132 | TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2OP075/Web_NP8541_12.205 | 2OP075 | L'Oreal Paris Revitalift MUR Cleans. Twlt | L'Oreal Paris Revitalift Mur Cleans. Twlt | Web_NP8541_12.205 | NY-Analytical | 748 | Visual Inspection | Appearance | Meets specifications (Done) |  |  |  |  | Visual | Visual | NY | RAW MATERIAL - NY | RM_WEB | both
2OP075/Web_NP8541_12.205 | 2OP075 | L'Oreal Paris Revitalift MUR Cleans. Twlt | L'Oreal Paris Revitalift Mur Cleans. Twlt | Web_NP8541_12.205 | NY-Analytical | 748 | Visual Inspection | Barcode Legibility/Quality | Barcode should scan and code should be matched under the barcode (Done) |  |  |  |  | Visual | Visual | NY | RAW MATERIAL - NY | RM_WEB | both
2OP075/Web_NP8541_12.205 | 2OP075 | L'Oreal Paris Revitalift MUR Cleans. Twlt | L'Oreal Paris Revitalift Mur Cleans. Twlt | Web_NP8541_12.205 | NY-Analytical | 748 | Visual Inspection | Color/Copy | Meets specifications (Done) |  |  |  |  | Visual | Visual | NY | RAW MATERIAL - NY | RM_WEB | both
2OP092/Web_NP8536_10 | 2OP092 | Ponds Original Fresh 15ct Twlts. | Pond's Original Fresh 15Ct. Twlts. | Web_NP8536_10 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2OP092/Web_NP8536_10 | 2OP092 | Ponds Original Fresh 15ct Twlts. | Pond's Original Fresh 15Ct. Twlts. | Web_NP8536_10 | NY-Analytical | 761 | Web Measurement | Web Width | 10 | in | FDA,4 | 9.9375 | 10.0625 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2OP092/Web_NP8536_10 | 2OP092 | Ponds Original Fresh 15ct Twlts. | Pond's Original Fresh 15Ct. Twlts. | Web_NP8536_10 | NY-Analytical | 1044 | Basis Weight (lbs/ream) | Basis Weight | 42.3 | Lb/Ream | FDA,2 | 40.1 | 49.02 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2OP092/Web_NP8536_10 | 2OP092 | Ponds Original Fresh 15ct Twlts. | Pond's Original Fresh 15Ct. Twlts. | Web_NP8536_10 | NY-Analytical | 1049 | Seal strength (NP8536) | Seal Strength (@320°F/40psi/0.5 sec dwell) | 8116.37 | g/in | FDA,0 | 2000 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2OP092/Web_NP8536_10 | 2OP092 | Ponds Original Fresh 15ct Twlts. | Pond's Original Fresh 15Ct. Twlts. | Web_NP8536_10 | NY-Analytical | 763 | Web Thickness | Thickness | 2.8 | mil | FDA,3 | 2.367 | 2.893 | TM 132 | TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2OP092/Web_NP8536_10 | 2OP092 | Ponds Original Fresh 15ct Twlts. | Pond's Original Fresh 15Ct. Twlts. | Web_NP8536_10 | NY-Analytical | 748 | Visual Inspection | Appearance | Meets specifications (Done) |  |  |  |  | Visual | Visual | NY | RAW MATERIAL - NY | RM_WEB | both
2OP092/Web_NP8536_10 | 2OP092 | Ponds Original Fresh 15ct Twlts. | Pond's Original Fresh 15Ct. Twlts. | Web_NP8536_10 | NY-Analytical | 748 | Visual Inspection | Barcode Legibility/Quality | Barcode should scan and code should be matched under the barcode (Done) |  |  |  |  | Visual | Visual | NY | RAW MATERIAL - NY | RM_WEB | both
2OP092/Web_NP8536_10 | 2OP092 | Ponds Original Fresh 15ct Twlts. | Pond's Original Fresh 15Ct. Twlts. | Web_NP8536_10 | NY-Analytical | 748 | Visual Inspection | Color/Copy | Meets specifications (Done) |  |  |  |  | Visual | Visual | NY | RAW MATERIAL - NY | RM_WEB | both
2OP093/Skip Lot Release_NY_2OP093 | 2OP093 | 10oz. Blue HDPE Bottle Grade Jar | Blue Hdpe Bottle Grade Jar 10Oz. | Skip Lot Release_NY_2OP093 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA | COA | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | 10oz. Blue HDPE Bottle Grade Jar | Blue Hdpe Bottle Grade Jar 10Oz. | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis | COA | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP093/Skip Lot Release_NY_2OP093 | 2OP093 | 10oz. Blue HDPE Bottle Grade Jar | Blue Hdpe Bottle Grade Jar 10Oz. | Skip Lot Release_NY_2OP093 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | 10oz. Blue HDPE Bottle Grade Jar | Blue Hdpe Bottle Grade Jar 10Oz. | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | 10oz. Blue HDPE Bottle Grade Jar | Blue Hdpe Bottle Grade Jar 10Oz. | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | Color | Matches Specification (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | 10oz. Blue HDPE Bottle Grade Jar | Blue Hdpe Bottle Grade Jar 10Oz. | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | Diameter | 2.746 | in | FDA,3 | 2.709 | 2.803 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | 10oz. Blue HDPE Bottle Grade Jar | Blue Hdpe Bottle Grade Jar 10Oz. | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | E Diameter | 2.51 | in | FDA,3 | 2.509 | 2.539 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | 10oz. Blue HDPE Bottle Grade Jar | Blue Hdpe Bottle Grade Jar 10Oz. | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Pass (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | 10oz. Blue HDPE Bottle Grade Jar | Blue Hdpe Bottle Grade Jar 10Oz. | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | H Dimension | 0.458 | in | FDA,3 | 0.457 | 0.487 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | 10oz. Blue HDPE Bottle Grade Jar | Blue Hdpe Bottle Grade Jar 10Oz. | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | Height | 3.52 | in | FDA,3 | 3.45 | 3.544 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | 10oz. Blue HDPE Bottle Grade Jar | Blue Hdpe Bottle Grade Jar 10Oz. | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | Odor | None detectable (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | 10oz. Blue HDPE Bottle Grade Jar | Blue Hdpe Bottle Grade Jar 10Oz. | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | S Dimension | 0.053 | in | FDA,3 | 0.035 | 0.065 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | 10oz. Blue HDPE Bottle Grade Jar | Blue Hdpe Bottle Grade Jar 10Oz. | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | T Diameter | 2.6 | in | FDA,3 | 2.593 | 2.623 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP093/ Jar _10 oz Blue HDPE Grade Jar | 2OP093 | 10oz. Blue HDPE Bottle Grade Jar | Blue Hdpe Bottle Grade Jar 10Oz. | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | Weight | 26.19 | g | FDA,2 | 24.5 | 27.5 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP093LI/Lid_NY_Threaded Blue Cap | 2OP093LI | PP Threaded Blue cap | PP Threaded Blue Cap | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Lid fits with a tight seal on |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP093LI/Lid_NY_Threaded Blue Cap | 2OP093LI | PP Threaded Blue cap | PP Threaded Blue Cap | Lid_NY_Threaded Blue Cap | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP093LI/Lid_NY_Threaded Blue Cap | 2OP093LI | PP Threaded Blue cap | PP Threaded Blue Cap | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Meets Specifications (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP093LI/Lid_NY_Threaded Blue Cap | 2OP093LI | PP Threaded Blue cap | PP Threaded Blue Cap | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | Color | Matches Specification (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP093LI/Lid_NY_Threaded Blue Cap | 2OP093LI | PP Threaded Blue cap | PP Threaded Blue Cap | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | Diameter | 2.7515 | in | FDA,3 | 2.748 | 2.764 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP093LI/Lid_NY_Threaded Blue Cap | 2OP093LI | PP Threaded Blue cap | PP Threaded Blue Cap | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | E Diameter | 2.5495 | in | FDA,3 | 2.548 | 2.562 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP093LI/Lid_NY_Threaded Blue Cap | 2OP093LI | PP Threaded Blue cap | PP Threaded Blue Cap | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | Height | 0.5625 | in | FDA,3 | 0.56 | 0.58 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP093LI/Lid_NY_Threaded Blue Cap | 2OP093LI | PP Threaded Blue cap | PP Threaded Blue Cap | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP093LI/Lid_NY_Threaded Blue Cap | 2OP093LI | PP Threaded Blue cap | PP Threaded Blue Cap | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | T Diameter | 2.619 | in | FDA,3 | 2.619 | 2.639 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP093LI/Lid_NY_Threaded Blue Cap | 2OP093LI | PP Threaded Blue cap | PP Threaded Blue Cap | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | Weight | 10.465 | g | FDA,2 |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Canister - Medium, Tapered White 4.1" | Canister - Medium, Tapered White 4.1" | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Cap Closure | Cap of lid |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Canister - Medium, Tapered White 4.1" | Canister - Medium, Tapered White 4.1" | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Lid fits with a tight seal on |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP134/Skip Lot Release_NY_2OP134 | 2OP134 | Canister - Medium, Tapered White 4.1" | Canister - Medium, Tapered White 4.1" | Skip Lot Release_NY_2OP134 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA | COA | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Canister - Medium, Tapered White 4.1" | Canister - Medium, Tapered White 4.1" | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA | COA | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Canister - Medium, Tapered White 4.1" | Canister - Medium, Tapered White 4.1" | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Canister - Medium, Tapered White 4.1" | Canister - Medium, Tapered White 4.1" | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Canister - Medium, Tapered White 4.1" | Canister - Medium, Tapered White 4.1" | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Color | White (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Canister - Medium, Tapered White 4.1" | Canister - Medium, Tapered White 4.1" | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Bead OD | 3.77 | in | FDA,3 | 3.768 | 3.792 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Canister - Medium, Tapered White 4.1" | Canister - Medium, Tapered White 4.1" | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Canister Height OD | 6.66 | in | FDA,3 | 6.608 | 6.732 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Canister - Medium, Tapered White 4.1" | Canister - Medium, Tapered White 4.1" | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Lip Width | 0.138 | in | FDA,3 | 0.127 | 0.143 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Canister - Medium, Tapered White 4.1" | Canister - Medium, Tapered White 4.1" | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Neck OD | 3.68 | in | FDA,3 | 3.673 | 3.697 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Canister - Medium, Tapered White 4.1" | Canister - Medium, Tapered White 4.1" | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | Meets the Specification (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Canister - Medium, Tapered White 4.1" | Canister - Medium, Tapered White 4.1" | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | None detectable (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Canister - Medium, Tapered White 4.1" | Canister - Medium, Tapered White 4.1" | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Weight | 53.94 | g | FDA,1 | 52 | 56 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP134/Reduced Testing_NY_Can 2OP134 | 2OP134 | Canister - Medium, Tapered White 4.1" | Canister - Medium, Tapered White 4.1" | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 766 | Lid Pull Force | Lid Pull-Off Force | 2997 | g | FDA,0 | 2000 | 3500 | TM 148 | TM148 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP138/Skip Lot Relrase_NY_2OP138C | 2OP138 | Large Ribbed Straight Walled Canister | Large Ribbed Straight Walled Canister | Skip Lot Relrase_NY_2OP138C | NY-Analytical | 1014 | Canister Release | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA | COA | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP138/Skip Lot Relrase_NY_2OP638 | 2OP138 | Large Ribbed Straight Walled Canister | Large Ribbed Straight Walled Canister | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA | COA | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Large Ribbed Straight Walled Canister | Large Ribbed Straight Walled Canister | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA | COA | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Large Ribbed Straight Walled Canister | Large Ribbed Straight Walled Canister | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Large Ribbed Straight Walled Canister | Large Ribbed Straight Walled Canister | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | [5 empty] | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Large Ribbed Straight Walled Canister | Large Ribbed Straight Walled Canister | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Color | [5 empty] | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Large Ribbed Straight Walled Canister | Large Ribbed Straight Walled Canister | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Bead OD |  | in | FDA,3 | 4.922 | 4.958 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Large Ribbed Straight Walled Canister | Large Ribbed Straight Walled Canister | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Canister OD |  | in | FDA,3 | 4.956 | 5.076 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Large Ribbed Straight Walled Canister | Large Ribbed Straight Walled Canister | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Neck OD |  | in | FDA,3 | 4.824 | 4.86 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Large Ribbed Straight Walled Canister | Large Ribbed Straight Walled Canister | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Total Height OD |  | in | FDA,3 | 6.745 | 6.865 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Large Ribbed Straight Walled Canister | Large Ribbed Straight Walled Canister | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | [5 empty] | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Large Ribbed Straight Walled Canister | Large Ribbed Straight Walled Canister | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | [5 empty] | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Large Ribbed Straight Walled Canister | Large Ribbed Straight Walled Canister | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Lip Width (for reference only) |  | in | FDA,3 | 0.14 |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Large Ribbed Straight Walled Canister | Large Ribbed Straight Walled Canister | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | [5 empty] | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Large Ribbed Straight Walled Canister | Large Ribbed Straight Walled Canister | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Wall Thickness (Bottom base area) |  | in | FDA,3 | 0.01 |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Large Ribbed Straight Walled Canister | Large Ribbed Straight Walled Canister | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Wall Thickness (Top area) |  | in | FDA,3 | 0.02 |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Large Ribbed Straight Walled Canister | Large Ribbed Straight Walled Canister | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Weight |  | g | FDA,1 | 66.5 | 73.5 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Large Ribbed Straight Walled Canister | Large Ribbed Straight Walled Canister | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 1052 | Top Load Compression Strength | Top load Compression Strength (0.5"/mSR,0.450"TD) | 150.3 | lb | FDA,0 | 84 |  | NPTM 1019 | TM1019 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP138/Reduced Testing_NY_Can 2OP138 | 2OP138 | Large Ribbed Straight Walled Canister | Large Ribbed Straight Walled Canister | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 1053 | Cavity | cavity Inserts | [5 empty] | Visual | Visual | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Large Ribbed Straight Walled Canister | Large Ribbed Straight Walled Canister | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Large Ribbed Straight Walled Canister | Large Ribbed Straight Walled Canister | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | [5 empty] | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Large Ribbed Straight Walled Canister | Large Ribbed Straight Walled Canister | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Color | [5 empty] | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Large Ribbed Straight Walled Canister | Large Ribbed Straight Walled Canister | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Bead OD |  | in | FDA,3 | 4.922 | 4.958 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Large Ribbed Straight Walled Canister | Large Ribbed Straight Walled Canister | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Canister OD |  | in | FDA,3 | 4.956 | 5.076 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Large Ribbed Straight Walled Canister | Large Ribbed Straight Walled Canister | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Neck OD |  | in | FDA,3 | 4.824 | 4.86 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Large Ribbed Straight Walled Canister | Large Ribbed Straight Walled Canister | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Total Height OD | 6.785 | in | FDA,3 | 6.745 | 6.865 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Large Ribbed Straight Walled Canister | Large Ribbed Straight Walled Canister | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | [5 empty] | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Large Ribbed Straight Walled Canister | Large Ribbed Straight Walled Canister | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | [5 empty] | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Large Ribbed Straight Walled Canister | Large Ribbed Straight Walled Canister | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Odor | [5 empty] | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Large Ribbed Straight Walled Canister | Large Ribbed Straight Walled Canister | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Wall Thickness (Bottom base area) |  | in | FDA,3 | 0.01 |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Large Ribbed Straight Walled Canister | Large Ribbed Straight Walled Canister | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Wall Thickness (Top area) |  | in | FDA,3 | 0.02 |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Large Ribbed Straight Walled Canister | Large Ribbed Straight Walled Canister | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Weight |  | g | FDA,1 | 66.5 | 73.5 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Large Ribbed Straight Walled Canister | Large Ribbed Straight Walled Canister | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 1052 | Top Load Compression Strength | Top load Compression Strength (0.5"/mSR,0.450"TD) |  | lb | FDA,0 | 84 |  | NPTM 1019 | TM1019 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP138C/Reduced Testing_NY_Can 2OP138C | 2OP138C | Large Ribbed Straight Walled Canister | Large Ribbed Straight Walled Canister | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 1053 | Cavity | cavity Inserts | [5 empty] | Visual | Visual | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP159LI/Lid_03B | 2OP159LI | Large Deep Well Lid .560 Orifice Blue | Large Deep Well .560 Orifice Blue | Lid_03B | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP159LI/Lid_03B | 2OP159LI | Large Deep Well Lid .560 Orifice Blue | Large Deep Well .560 Orifice Blue | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP159LI/Lid_03B | 2OP159LI | Large Deep Well Lid .560 Orifice Blue | Large Deep Well .560 Orifice Blue | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | 4.856 | in | FDA,3 | 4.847 | 4.867 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP159LI/Lid_03B | 2OP159LI | Large Deep Well Lid .560 Orifice Blue | Large Deep Well .560 Orifice Blue | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Cap Thickness | 0.049 | in | FDA,3 | 0.045 | 0.055 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP159LI/Lid_03B | 2OP159LI | Large Deep Well Lid .560 Orifice Blue | Large Deep Well .560 Orifice Blue | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Color | Blue (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP159LI/Lid_03B | 2OP159LI | Large Deep Well Lid .560 Orifice Blue | Large Deep Well .560 Orifice Blue | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Pass (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP159LI/Lid_03B | 2OP159LI | Large Deep Well Lid .560 Orifice Blue | Large Deep Well .560 Orifice Blue | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Insert Fin Gap | 0.07 | in | FDA,3 | 0.065 | 0.075 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP159LI/Lid_03B | 2OP159LI | Large Deep Well Lid .560 Orifice Blue | Large Deep Well .560 Orifice Blue | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter | 4.929 | in | FDA,3 | 4.922 | 4.932 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP159LI/Lid_03B | 2OP159LI | Large Deep Well Lid .560 Orifice Blue | Large Deep Well .560 Orifice Blue | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness | 0.042 | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP159LI/Lid_03B | 2OP159LI | Large Deep Well Lid .560 Orifice Blue | Large Deep Well .560 Orifice Blue | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP159LI/Lid_03B | 2OP159LI | Large Deep Well Lid .560 Orifice Blue | Large Deep Well .560 Orifice Blue | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height | 1.486 | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP159LI/Lid_03B | 2OP159LI | Large Deep Well Lid .560 Orifice Blue | Large Deep Well .560 Orifice Blue | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Outer Diameter | 0.163 | in | FDA,3 | 0.16 | 0.17 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP159LI/Lid_03B | 2OP159LI | Large Deep Well Lid .560 Orifice Blue | Large Deep Well .560 Orifice Blue | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Thicknes | 0.02 | in | FDA,3 | 0.012 | 0.022 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP159LI/Lid_03B | 2OP159LI | Large Deep Well Lid .560 Orifice Blue | Large Deep Well .560 Orifice Blue | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Weight | 34.335 | g | FDA,1 | 32 | 42 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP213/Web_NP8541_11 | 2OP213 | Garnier Skin Active Clean Micellar Makeup Remover Cleans. Twlts. - Canada | Garnier Skin ACt. ive Clean Micellar Makeup Remover Cleans. Twlts. - Canada | Web_NP8541_11 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2OP213/Web_NP8541_11 | 2OP213 | Garnier Skin Active Clean Micellar Makeup Remover Cleans. Twlts. - Canada | Garnier Skin ACt. ive Clean Micellar Makeup Remover Cleans. Twlts. - Canada | Web_NP8541_11 | NY-Analytical | 761 | Web Measurement | Web Width | 11 | in | FDA,4 | 10.9375 | 11.0625 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2OP213/Web_NP8541_11 | 2OP213 | Garnier Skin Active Clean Micellar Makeup Remover Cleans. Twlts. - Canada | Garnier Skin ACt. ive Clean Micellar Makeup Remover Cleans. Twlts. - Canada | Web_NP8541_11 | NY-Analytical | 756 | Basis Weight | Basis Weight | 86.76 | GSM | FDA,1 | 74.3 | 90.9 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2OP213/Web_NP8541_11 | 2OP213 | Garnier Skin Active Clean Micellar Makeup Remover Cleans. Twlts. - Canada | Garnier Skin ACt. ive Clean Micellar Makeup Remover Cleans. Twlts. - Canada | Web_NP8541_11 | NY-Analytical | 762 | Seal Strength | Seal Strength (@130°C/0.5 sec/1 bar) | 8561 | g/in | FDA,0 | 1000 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2OP213/Web_NP8541_11 | 2OP213 | Garnier Skin Active Clean Micellar Makeup Remover Cleans. Twlts. - Canada | Garnier Skin ACt. ive Clean Micellar Makeup Remover Cleans. Twlts. - Canada | Web_NP8541_11 | NY-Analytical | 763 | Web Thickness | Thickness | 0.088 | mm | FDA,3 | 0.074 | 0.094 | TM 132 | TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2OP213/Web_NP8541_11 | 2OP213 | Garnier Skin Active Clean Micellar Makeup Remover Cleans. Twlts. - Canada | Garnier Skin ACt. ive Clean Micellar Makeup Remover Cleans. Twlts. - Canada | Web_NP8541_11 | NY-Analytical | 748 | Visual Inspection | Appearance | Meets specifications (Done) |  |  |  |  | Visual | Visual | NY | RAW MATERIAL - NY | RM_WEB | both
2OP213/Web_NP8541_11 | 2OP213 | Garnier Skin Active Clean Micellar Makeup Remover Cleans. Twlts. - Canada | Garnier Skin ACt. ive Clean Micellar Makeup Remover Cleans. Twlts. - Canada | Web_NP8541_11 | NY-Analytical | 748 | Visual Inspection | Barcode Legibility/Quality | Barcode should scan and code should be matched under the barcode (Done) |  |  |  |  | Visual | Visual | NY | RAW MATERIAL - NY | RM_WEB | both
2OP213/Web_NP8541_11 | 2OP213 | Garnier Skin Active Clean Micellar Makeup Remover Cleans. Twlts. - Canada | Garnier Skin ACt. ive Clean Micellar Makeup Remover Cleans. Twlts. - Canada | Web_NP8541_11 | NY-Analytical | 748 | Visual Inspection | Color/Copy | Meets specifications (Done) |  |  |  |  | Visual | Visual | NY | RAW MATERIAL - NY | RM_WEB | both
2OP223/Web_NP8511_320 | 2OP223 | Simple Kind To Skin Facial Cleansing Wipes 25ct US | Simple Kind To Skin Facial Cleansing Wipes 25Ct. US | Web_NP8511_320 | NY-Analytical | 885 | Web Inspection | Artwork / Copy | Sample(s) all meet specifications (Done) |  |  |  |  | Visual Observation | Visual | NY | RAW MATERIAL - NY | RM_WEB | both
2OP242LI/Lid_17C | 2OP242LI | Large Deep Well Lid-Sani Cloth Prime Lobe Orifice | Large Deep Well Lid-Sani-Cloth Prime Lobe Orifice | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Cap Closure | Cap of lid |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP242LI/Lid_17C | 2OP242LI | Large Deep Well Lid-Sani Cloth Prime Lobe Orifice | Large Deep Well Lid-Sani-Cloth Prime Lobe Orifice | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Lid fits with a tight seal on |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP242LI/Lid_17C | 2OP242LI | Large Deep Well Lid-Sani Cloth Prime Lobe Orifice | Large Deep Well Lid-Sani-Cloth Prime Lobe Orifice | Lid_17C | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Flip Lid Force |  | NY | RAW MATERIAL - NY | RM_LID | both
2OP242LI/Lid_17C | 2OP242LI | Large Deep Well Lid-Sani Cloth Prime Lobe Orifice | Large Deep Well Lid-Sani-Cloth Prime Lobe Orifice | Lid_17C | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | 1.11 | lb | FDA,2 | 0.75 |  | TM 143 | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP242LI/Lid_17C | 2OP242LI | Large Deep Well Lid-Sani Cloth Prime Lobe Orifice | Large Deep Well Lid-Sani-Cloth Prime Lobe Orifice | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Meets Specifications (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP242LI/Lid_17C | 2OP242LI | Large Deep Well Lid-Sani Cloth Prime Lobe Orifice | Large Deep Well Lid-Sani-Cloth Prime Lobe Orifice | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | 4.835 | in | FDA,3 | 4.831 | 4.851 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP242LI/Lid_17C | 2OP242LI | Large Deep Well Lid-Sani Cloth Prime Lobe Orifice | Large Deep Well Lid-Sani-Cloth Prime Lobe Orifice | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Color | Translucent Plum (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP242LI/Lid_17C | 2OP242LI | Large Deep Well Lid-Sani Cloth Prime Lobe Orifice | Large Deep Well Lid-Sani-Cloth Prime Lobe Orifice | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness | 0.045 | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP242LI/Lid_17C | 2OP242LI | Large Deep Well Lid-Sani Cloth Prime Lobe Orifice | Large Deep Well Lid-Sani-Cloth Prime Lobe Orifice | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP242LI/Lid_17C | 2OP242LI | Large Deep Well Lid-Sani Cloth Prime Lobe Orifice | Large Deep Well Lid-Sani-Cloth Prime Lobe Orifice | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Diameter | 0.62 | in | FDA,3 | 0.62 | 0.63 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP242LI/Lid_17C | 2OP242LI | Large Deep Well Lid-Sani Cloth Prime Lobe Orifice | Large Deep Well Lid-Sani-Cloth Prime Lobe Orifice | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness | 0.021 | in | FDA,3 | 0.017 | 0.023 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP242LI/Lid_17C | 2OP242LI | Large Deep Well Lid-Sani Cloth Prime Lobe Orifice | Large Deep Well Lid-Sani-Cloth Prime Lobe Orifice | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height | 1.483 | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP242LI/Lid_17C | 2OP242LI | Large Deep Well Lid-Sani Cloth Prime Lobe Orifice | Large Deep Well Lid-Sani-Cloth Prime Lobe Orifice | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Weight | 38.41 | g | FDA,1 | 32.5 | 41.5 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP280LI/Lid_17A | 2OP280LI | Large Deep Well Lid- Bleach Orifice | Large Deep Well Lid- Bleach Orifice | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Cap Closure | Cap of lid |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP280LI/Lid_17A | 2OP280LI | Large Deep Well Lid- Bleach Orifice | Large Deep Well Lid- Bleach Orifice | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Lid fits with a tight seal on |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP280LI/Lid_17A | 2OP280LI | Large Deep Well Lid- Bleach Orifice | Large Deep Well Lid- Bleach Orifice | Lid_17A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Flip Lid Force |  | NY | RAW MATERIAL - NY | RM_LID | both
2OP280LI/Lid_17A | 2OP280LI | Large Deep Well Lid- Bleach Orifice | Large Deep Well Lid- Bleach Orifice | Lid_17A | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | 2.88 | lb | FDA,2 | 0.75 |  | TM 143 | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP280LI/Lid_17A | 2OP280LI | Large Deep Well Lid- Bleach Orifice | Large Deep Well Lid- Bleach Orifice | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Meets Specifications (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP280LI/Lid_17A | 2OP280LI | Large Deep Well Lid- Bleach Orifice | Large Deep Well Lid- Bleach Orifice | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | 4.851 | in | FDA,3 | 4.831 | 4.851 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP280LI/Lid_17A | 2OP280LI | Large Deep Well Lid- Bleach Orifice | Large Deep Well Lid- Bleach Orifice | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Color | Translucent Orange (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP280LI/Lid_17A | 2OP280LI | Large Deep Well Lid- Bleach Orifice | Large Deep Well Lid- Bleach Orifice | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness | 0.035 | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP280LI/Lid_17A | 2OP280LI | Large Deep Well Lid- Bleach Orifice | Large Deep Well Lid- Bleach Orifice | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP280LI/Lid_17A | 2OP280LI | Large Deep Well Lid- Bleach Orifice | Large Deep Well Lid- Bleach Orifice | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Diameter | 1.468 | in | FDA,3 | 1.468 | 1.478 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP280LI/Lid_17A | 2OP280LI | Large Deep Well Lid- Bleach Orifice | Large Deep Well Lid- Bleach Orifice | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness | 0.035 | in | FDA,3 | 0.03 | 0.04 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP280LI/Lid_17A | 2OP280LI | Large Deep Well Lid- Bleach Orifice | Large Deep Well Lid- Bleach Orifice | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height | 1.484 | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP280LI/Lid_17A | 2OP280LI | Large Deep Well Lid- Bleach Orifice | Large Deep Well Lid- Bleach Orifice | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Weight | 37.555 | g | FDA,1 | 36 | 42 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP300LI/Lid_13E | 2OP300LI | Large DeepWell Lid Genesis Orifice TR Blue | Large Deepwell Genesis Orifice Tr Blue | Lid_13E | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP300LI/Lid_13E | 2OP300LI | Large DeepWell Lid Genesis Orifice TR Blue | Large Deepwell Genesis Orifice Tr Blue | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP300LI/Lid_13E | 2OP300LI | Large DeepWell Lid Genesis Orifice TR Blue | Large Deepwell Genesis Orifice Tr Blue | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | 4.852 | in | FDA,3 | 4.847 | 4.867 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP300LI/Lid_13E | 2OP300LI | Large DeepWell Lid Genesis Orifice TR Blue | Large Deepwell Genesis Orifice Tr Blue | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Center to Center Distance | 0.455 | in | FDA,3 | 0.445 | 0.455 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP300LI/Lid_13E | 2OP300LI | Large DeepWell Lid Genesis Orifice TR Blue | Large Deepwell Genesis Orifice Tr Blue | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Color | Translucent Blue (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP300LI/Lid_13E | 2OP300LI | Large DeepWell Lid Genesis Orifice TR Blue | Large Deepwell Genesis Orifice Tr Blue | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | End Circle Diameter | 0.128 | in | FDA,3 | 0.12 | 0.13 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP300LI/Lid_13E | 2OP300LI | Large DeepWell Lid Genesis Orifice TR Blue | Large Deepwell Genesis Orifice Tr Blue | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Pass (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP300LI/Lid_13E | 2OP300LI | Large DeepWell Lid Genesis Orifice TR Blue | Large Deepwell Genesis Orifice Tr Blue | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Insert Fin Gap | 0.08 | in | FDA,3 | 0.07 | 0.08 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP300LI/Lid_13E | 2OP300LI | Large DeepWell Lid Genesis Orifice TR Blue | Large Deepwell Genesis Orifice Tr Blue | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness | 0.04 | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP300LI/Lid_13E | 2OP300LI | Large DeepWell Lid Genesis Orifice TR Blue | Large Deepwell Genesis Orifice Tr Blue | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP300LI/Lid_13E | 2OP300LI | Large DeepWell Lid Genesis Orifice TR Blue | Large Deepwell Genesis Orifice Tr Blue | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness | 0.031 | in | FDA,3 | 0.03 | 0.04 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP300LI/Lid_13E | 2OP300LI | Large DeepWell Lid Genesis Orifice TR Blue | Large Deepwell Genesis Orifice Tr Blue | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height | 1.4885 | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP300LI/Lid_13E | 2OP300LI | Large DeepWell Lid Genesis Orifice TR Blue | Large Deepwell Genesis Orifice Tr Blue | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Weight | 35.535 | g | FDA,1 | 32 | 42 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP300LI/Lid_13E | 2OP300LI | Large DeepWell Lid Genesis Orifice TR Blue | Large Deepwell Genesis Orifice Tr Blue | Lid_13E | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | 1.05 | lb | FDA,2 | 0.75 |  | TM 148 | TM148 | NY | RAW MATERIAL - NY | RM_LID | both
2OP435LI/Lid_04C_560 | 2OP435LI | LID MED DW LID TRANSLUCE PURPLE | Med Dw Transluce Purple | Lid_04C_560 | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP435LI/Lid_04C_560 | 2OP435LI | LID MED DW LID TRANSLUCE PURPLE | Med Dw Transluce Purple | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP435LI/Lid_04C_560 | 2OP435LI | LID MED DW LID TRANSLUCE PURPLE | Med Dw Transluce Purple | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Cap Thickness | 0.157 | in | FDA,3 | 0.147 | 0.157 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP435LI/Lid_04C_560 | 2OP435LI | LID MED DW LID TRANSLUCE PURPLE | Med Dw Transluce Purple | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Translucent Purple (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP435LI/Lid_04C_560 | 2OP435LI | LID MED DW LID TRANSLUCE PURPLE | Med Dw Transluce Purple | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Pass (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP435LI/Lid_04C_560 | 2OP435LI | LID MED DW LID TRANSLUCE PURPLE | Med Dw Transluce Purple | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter | 3.765 | in | FDA,3 | 3.76 | 3.77 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP435LI/Lid_04C_560 | 2OP435LI | LID MED DW LID TRANSLUCE PURPLE | Med Dw Transluce Purple | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness | 0.04 | in | FDA,3 | 0.04 | 0.05 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP435LI/Lid_04C_560 | 2OP435LI | LID MED DW LID TRANSLUCE PURPLE | Med Dw Transluce Purple | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP435LI/Lid_04C_560 | 2OP435LI | LID MED DW LID TRANSLUCE PURPLE | Med Dw Transluce Purple | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height | 1.492 | in | FDA,3 | 1.473 | 1.493 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP435LI/Lid_04C_560 | 2OP435LI | LID MED DW LID TRANSLUCE PURPLE | Med Dw Transluce Purple | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Outer Diameter | 0.565 | in | FDA,3 | 0.555 | 0.565 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP435LI/Lid_04C_560 | 2OP435LI | LID MED DW LID TRANSLUCE PURPLE | Med Dw Transluce Purple | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Thicknes | 0.02 | in | FDA,3 | 0.015 | 0.025 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP435LI/Lid_04C_560 | 2OP435LI | LID MED DW LID TRANSLUCE PURPLE | Med Dw Transluce Purple | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Weight | 21.1 | g | FDA,1 | 20 | 24 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Extra-Large Straight Walled Canister | Extra-Large Straight Walled Canister | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA | COA | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP638/Skip Lot Relrase_NY_2OP638 | 2OP638 | Extra-Large Straight Walled Canister | Extra-Large Straight Walled Canister | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis | COA | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Extra-Large Straight Walled Canister | Extra-Large Straight Walled Canister | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Extra-Large Straight Walled Canister | Extra-Large Straight Walled Canister | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Bead OD | 4.9545 | in | FDA,3 | 4.922 | 4.958 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Extra-Large Straight Walled Canister | Extra-Large Straight Walled Canister | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Canister OD | 5.031 | in | FDA,3 | 4.956 | 5.076 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Extra-Large Straight Walled Canister | Extra-Large Straight Walled Canister | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Neck OD | 4.8425 | in | FDA,3 | 4.824 | 4.86 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Extra-Large Straight Walled Canister | Extra-Large Straight Walled Canister | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Total Height OD | 8.75 | in | FDA,3 | 8.688 | 8.808 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Extra-Large Straight Walled Canister | Extra-Large Straight Walled Canister | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | [5 empty] | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Extra-Large Straight Walled Canister | Extra-Large Straight Walled Canister | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | [5 empty] | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Extra-Large Straight Walled Canister | Extra-Large Straight Walled Canister | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Lip Width (for reference only) | 0.15 | in | FDA,3 | 0.14 |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Extra-Large Straight Walled Canister | Extra-Large Straight Walled Canister | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Wall Thickness (Bottom base area) |  | in | FDA,3 | 0.01 |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Extra-Large Straight Walled Canister | Extra-Large Straight Walled Canister | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Wall Thickness (Top area) |  | in | FDA,3 | 0.02 |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Extra-Large Straight Walled Canister | Extra-Large Straight Walled Canister | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Weight | 90.2 | g | FDA,1 | 85.5 | 94.5 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Extra-Large Straight Walled Canister | Extra-Large Straight Walled Canister | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 1052 | Top Load Compression Strength | Top load Compression Strength (0.5"/mSR,0.450"TD) |  | lb | FDA,0 | 97 |  | NPTM 1019 | TM1019 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP638/Bulk_IN_4ABBY12 - 20190218 | 2OP638 | Extra-Large Straight Walled Canister | Extra-Large Straight Walled Canister | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 902 | Bulk Liquid Inspection | Appearance | [5 empty] | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP638/Bulk_IN_4ABBY12 - 20190218 | 2OP638 | Extra-Large Straight Walled Canister | Extra-Large Straight Walled Canister | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 902 | Bulk Liquid Inspection | Color | [5 empty] | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP638/Bulk_IN_4ABBY12 - 20190218 | 2OP638 | Extra-Large Straight Walled Canister | Extra-Large Straight Walled Canister | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 902 | Bulk Liquid Inspection | Contamination | [5 empty] | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP638/Bulk_IN_4ABBY12 - 20190218 | 2OP638 | Extra-Large Straight Walled Canister | Extra-Large Straight Walled Canister | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 902 | Bulk Liquid Inspection | Odor | [5 empty] | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP638/Bulk_IN_4ABBY12 - 20190218 | 2OP638 | Extra-Large Straight Walled Canister | Extra-Large Straight Walled Canister | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 902 | Bulk Liquid Inspection | pH |  |  | FDA,2 | 4.5 | 4.9 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Extra-Large Straight Walled Canister | Extra-Large Straight Walled Canister | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Meets Specifications (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Extra-Large Straight Walled Canister | Extra-Large Straight Walled Canister | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Color | White (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Extra-Large Straight Walled Canister | Extra-Large Straight Walled Canister | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP638/Bulk_IN_4ABBY12 - 20190218 | 2OP638 | Extra-Large Straight Walled Canister | Extra-Large Straight Walled Canister | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 943 | %Phenoxyethanol & Sodium Benzoate | Phenoxyethanol |  | % | FDA,3 | 0.63 | 0.77 | TM 395 | TM395 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP638/Bulk_IN_4ABBY12 - 20190218 | 2OP638 | Extra-Large Straight Walled Canister | Extra-Large Straight Walled Canister | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 943 | %Phenoxyethanol & Sodium Benzoate | Phenoxyethanol Final |  | % | FDA,3 | 0.63 | 0.77 | TM 395 | TM395 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP638/Bulk_IN_4ABBY12 - 20190218 | 2OP638 | Extra-Large Straight Walled Canister | Extra-Large Straight Walled Canister | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 943 | %Phenoxyethanol & Sodium Benzoate | Sodium Benzoate |  | % | FDA,3 | 0.423 | 0.517 | TM 395 | TM395 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP638/Bulk_IN_4ABBY12 - 20190218 | 2OP638 | Extra-Large Straight Walled Canister | Extra-Large Straight Walled Canister | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 943 | %Phenoxyethanol & Sodium Benzoate | Sodium Benzoate Final |  | % | FDA,3 | 0.423 | 0.517 | TM 395 | TM395 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP638/Reduced Testing_NY_Can 2OP638 | 2OP638 | Extra-Large Straight Walled Canister | Extra-Large Straight Walled Canister | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 1053 | Cavity | cavity Inserts | [5 empty] | Visual | Visual | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OP638LI/Lid_13D | 2OP638LI | Large Deep Well Lid PDI Orifice Translucent Gray | Large Deep Well PDI Orifice Translucent Gray | Lid_13D | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP638LI/Lid_13D | 2OP638LI | Large Deep Well Lid PDI Orifice Translucent Gray | Large Deep Well PDI Orifice Translucent Gray | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP638LI/Lid_13D | 2OP638LI | Large Deep Well Lid PDI Orifice Translucent Gray | Large Deep Well PDI Orifice Translucent Gray | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | 4.8615 | in | FDA,3 | 4.847 | 4.867 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP638LI/Lid_13D | 2OP638LI | Large Deep Well Lid PDI Orifice Translucent Gray | Large Deep Well PDI Orifice Translucent Gray | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Center to Center Distance | 0.4535 | in | FDA,3 | 0.445 | 0.455 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP638LI/Lid_13D | 2OP638LI | Large Deep Well Lid PDI Orifice Translucent Gray | Large Deep Well PDI Orifice Translucent Gray | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Color | Translucent Gray (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP638LI/Lid_13D | 2OP638LI | Large Deep Well Lid PDI Orifice Translucent Gray | Large Deep Well PDI Orifice Translucent Gray | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | End Circle Diameter | 0.122 | in | FDA,3 | 0.12 | 0.13 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP638LI/Lid_13D | 2OP638LI | Large Deep Well Lid PDI Orifice Translucent Gray | Large Deep Well PDI Orifice Translucent Gray | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Pass (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP638LI/Lid_13D | 2OP638LI | Large Deep Well Lid PDI Orifice Translucent Gray | Large Deep Well PDI Orifice Translucent Gray | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Insert Fin Gap | 0.0745 | in | FDA,3 | 0.07 | 0.08 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP638LI/Lid_13D | 2OP638LI | Large Deep Well Lid PDI Orifice Translucent Gray | Large Deep Well PDI Orifice Translucent Gray | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter | 4.9218 | in | FDA,3 | 4.917 | 4.937 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP638LI/Lid_13D | 2OP638LI | Large Deep Well Lid PDI Orifice Translucent Gray | Large Deep Well PDI Orifice Translucent Gray | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness | 0.041 | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP638LI/Lid_13D | 2OP638LI | Large Deep Well Lid PDI Orifice Translucent Gray | Large Deep Well PDI Orifice Translucent Gray | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP638LI/Lid_13D | 2OP638LI | Large Deep Well Lid PDI Orifice Translucent Gray | Large Deep Well PDI Orifice Translucent Gray | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness | 0.0345 | in | FDA,3 | 0.03 | 0.04 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP638LI/Lid_13D | 2OP638LI | Large Deep Well Lid PDI Orifice Translucent Gray | Large Deep Well PDI Orifice Translucent Gray | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height | 1.487 | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP638LI/Lid_13D | 2OP638LI | Large Deep Well Lid PDI Orifice Translucent Gray | Large Deep Well PDI Orifice Translucent Gray | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Weight | 33.92 | g | FDA,1 | 32 | 42 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OP638LI/Lid_13D | 2OP638LI | Large Deep Well Lid PDI Orifice Translucent Gray | Large Deep Well PDI Orifice Translucent Gray | Lid_13D | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | 1.79 | lb | FDA,2 | 0.75 |  | TM 148 | TM148 | NY | RAW MATERIAL - NY | RM_LID | both
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | CAN LARGE CANISTER COPOLYMER | Can Large Canister Copolymer | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Lid has a tight friction fit |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ753/Skip Lot Release_NY_2OQ753 | 2OQ753 | CAN LARGE CANISTER COPOLYMER | Can Large Canister Copolymer | Skip Lot Release_NY_2OQ753 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA | COA | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | CAN LARGE CANISTER COPOLYMER | Can Large Canister Copolymer | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA | COA | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | CAN LARGE CANISTER COPOLYMER | Can Large Canister Copolymer | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis | COA | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ753/Skip Lot Release_NY_2OQ753 | 2OQ753 | CAN LARGE CANISTER COPOLYMER | Can Large Canister Copolymer | Skip Lot Release_NY_2OQ753 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | CAN LARGE CANISTER COPOLYMER | Can Large Canister Copolymer | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | CAN LARGE CANISTER COPOLYMER | Can Large Canister Copolymer | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Color | White (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | CAN LARGE CANISTER COPOLYMER | Can Large Canister Copolymer | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Bead OD | 4.9355 | in | FDA,3 | 4.908 | 4.944 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | CAN LARGE CANISTER COPOLYMER | Can Large Canister Copolymer | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Canister OD | 5.0145 | in | FDA,3 | 4.998 | 5.034 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | CAN LARGE CANISTER COPOLYMER | Can Large Canister Copolymer | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Lip Width | 0.17 | in | FDA,3 | 0.142 | 0.178 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | CAN LARGE CANISTER COPOLYMER | Can Large Canister Copolymer | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Total Height OD | 6.748 | in | FDA,3 | 6.706 | 6.826 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | CAN LARGE CANISTER COPOLYMER | Can Large Canister Copolymer | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | Meets the Specification (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | CAN LARGE CANISTER COPOLYMER | Can Large Canister Copolymer | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | None detectable (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | CAN LARGE CANISTER COPOLYMER | Can Large Canister Copolymer | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Weight | 69.21 | g | FDA,1 | 67.5 | 72.5 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ753/Reduced Testing_NY_Can 2OQ753 | 2OQ753 | CAN LARGE CANISTER COPOLYMER | Can Large Canister Copolymer | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 766 | Lid Pull Force | Lid Pull-Off Force | 2859.39 | g | FDA,0 | 2000 | 3500 | TM 148 | TM148 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 2OQ753X | Canister - Large, Ribbed, White | Canister - Large, Ribbed, White | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Lid has a tight friction fit |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ753X/Skip Lot Relrase_NY_2OQ753X | 2OQ753X | Canister - Large, Ribbed, White | Canister - Large, Ribbed, White | Skip Lot Relrase_NY_2OQ753X | NY-Analytical | 1014 | Canister Release | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA | COA | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 2OQ753X | Canister - Large, Ribbed, White | Canister - Large, Ribbed, White | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 1014 | Canister Release | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA | COA | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ753X/Skip Lot Release_NY_2OQ850X | 2OQ753X | Canister - Large, Ribbed, White | Canister - Large, Ribbed, White | Skip Lot Release_NY_2OQ850X | NY-Analytical | 1014 | Canister Release | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA | COA | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 2OQ753X | Canister - Large, Ribbed, White | Canister - Large, Ribbed, White | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | [5 empty] | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 2OQ753X | Canister - Large, Ribbed, White | Canister - Large, Ribbed, White | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 2OQ753X | Canister - Large, Ribbed, White | Canister - Large, Ribbed, White | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 765 | Canister & Lid Inspection | Color | White (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 2OQ753X | Canister - Large, Ribbed, White | Canister - Large, Ribbed, White | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | Meets the Specification (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 2OQ753X | Canister - Large, Ribbed, White | Canister - Large, Ribbed, White | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 765 | Canister & Lid Inspection | Odor | None detectable (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 2OQ753X | Canister - Large, Ribbed, White | Canister - Large, Ribbed, White | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 765 | Canister & Lid Inspection | Weight | 70.655 | g | FDA,1 | 66.5 | 73.5 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 2OQ753X | Canister - Large, Ribbed, White | Canister - Large, Ribbed, White | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 1052 | Top Load Compression Strength | Top load Compression Strength (0.5"/mSR,0.450"TD) |  | lb | FDA,0 | 84 |  | NPTM 1019 | TM1019 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ753X/Reduced Testing_NY_Can 2OQ753 X | 2OQ753X | Canister - Large, Ribbed, White | Canister - Large, Ribbed, White | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 1053 | Cavity | cavity Inserts | Meets specification (Done) |  |  |  |  | Visual | Visual | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ850/Reduced Testing_NY_Can 2OQ850X | 2OQ850 | Canister - Extra Large, Ribbed, White | Canister - Extra Large, Ribbed, White | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Lid has a tight friction fit |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ850/Reduced Testing_NY_Can 2OQ850 | 2OQ850 | Canister - Extra Large, Ribbed, White | Canister - Extra Large, Ribbed, White | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Lid has a tight friction fit |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ850/Reduced Testing_NY_Can 2OQ850 | 2OQ850 | Canister - Extra Large, Ribbed, White | Canister - Extra Large, Ribbed, White | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis | C of A Meet Specifications (Done) |  |  |  |  | COA | COA | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ850/Skip Lot Release_NY_2OQ850 | 2OQ850 | Canister - Extra Large, Ribbed, White | Canister - Extra Large, Ribbed, White | Skip Lot Release_NY_2OQ850 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis | COA | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ850/Reduced Testing_NY_Can 2OQ850 | 2OQ850 | Canister - Extra Large, Ribbed, White | Canister - Extra Large, Ribbed, White | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis | COA | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ850/Reduced Testing_NY_Can 2OQ850X | 2OQ850 | Canister - Extra Large, Ribbed, White | Canister - Extra Large, Ribbed, White | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ850/Reduced Testing_NY_Can 2OQ850X | 2OQ850 | Canister - Extra Large, Ribbed, White | Canister - Extra Large, Ribbed, White | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ850/Reduced Testing_NY_Can 2OQ850 | 2OQ850 | Canister - Extra Large, Ribbed, White | Canister - Extra Large, Ribbed, White | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ850/Reduced Testing_NY_Can 2OQ850X | 2OQ850 | Canister - Extra Large, Ribbed, White | Canister - Extra Large, Ribbed, White | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 765 | Canister & Lid Inspection | Color | White (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ850/Reduced Testing_NY_Can 2OQ850 | 2OQ850 | Canister - Extra Large, Ribbed, White | Canister - Extra Large, Ribbed, White | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Color | White (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ850/Reduced Testing_NY_Can 2OQ850X | 2OQ850 | Canister - Extra Large, Ribbed, White | Canister - Extra Large, Ribbed, White | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | Meets the Specification (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ850/Reduced Testing_NY_Can 2OQ850 | 2OQ850 | Canister - Extra Large, Ribbed, White | Canister - Extra Large, Ribbed, White | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | Meets the Specification (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ850/Reduced Testing_NY_Can 2OQ850X | 2OQ850 | Canister - Extra Large, Ribbed, White | Canister - Extra Large, Ribbed, White | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 765 | Canister & Lid Inspection | Odor | None detectable (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ850/Reduced Testing_NY_Can 2OQ850 | 2OQ850 | Canister - Extra Large, Ribbed, White | Canister - Extra Large, Ribbed, White | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | None detectable (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ850/Reduced Testing_NY_Can 2OQ850X | 2OQ850 | Canister - Extra Large, Ribbed, White | Canister - Extra Large, Ribbed, White | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 765 | Canister & Lid Inspection | Weight | 88.41 | g | FDA,1 | 85.5 | 94.5 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ850/Reduced Testing_NY_Can 2OQ850 | 2OQ850 | Canister - Extra Large, Ribbed, White | Canister - Extra Large, Ribbed, White | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Weight | 88.41 | g | FDA,1 | 87 | 93 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ850/Reduced Testing_NY_Can 2OQ850X | 2OQ850 | Canister - Extra Large, Ribbed, White | Canister - Extra Large, Ribbed, White | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 1052 | Top Load Compression Strength | Top load Compression Strength (0.5"/mSR,0.450"TD) |  | lb | FDA,0 | 97 |  | NPTM 1019 | TM1019 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ850/Reduced Testing_NY_Can 2OQ850 | 2OQ850 | Canister - Extra Large, Ribbed, White | Canister - Extra Large, Ribbed, White | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 766 | Lid Pull Force | Lid Pull-Off Force |  | g | FDA,0 | 2000 | 3500 | TM 148 | TM148 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ850/Reduced Testing_NY_Can 2OQ850X | 2OQ850 | Canister - Extra Large, Ribbed, White | Canister - Extra Large, Ribbed, White | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 1053 | Cavity | cavity Inserts | [5 empty] | Visual | Visual | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ850LI/Lid_11B | 2OQ850LI | Large Deep Well Lid .560 Orifice | Large Deep Well .560 Orifice | Lid_11B | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Flip Lid Force |  | NY | RAW MATERIAL - NY | RM_LID | both
2OQ850LI/Lid_11B | 2OQ850LI | Large Deep Well Lid .560 Orifice | Large Deep Well .560 Orifice | Lid_11B | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | 2.41 | lb | FDA,2 | 0.75 |  | TM 143 | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OQ850LI/Lid_11B | 2OQ850LI | Large Deep Well Lid .560 Orifice | Large Deep Well .560 Orifice | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OQ850LI/Lid_11B | 2OQ850LI | Large Deep Well Lid .560 Orifice | Large Deep Well .560 Orifice | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | 4.861 | in | FDA,3 | 4.847 | 4.867 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OQ850LI/Lid_11B | 2OQ850LI | Large Deep Well Lid .560 Orifice | Large Deep Well .560 Orifice | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Color | Translucent Red (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OQ850LI/Lid_11B | 2OQ850LI | Large Deep Well Lid .560 Orifice | Large Deep Well .560 Orifice | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Pass (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OQ850LI/Lid_11B | 2OQ850LI | Large Deep Well Lid .560 Orifice | Large Deep Well .560 Orifice | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Insert Fin Gap | 0.073 | in | FDA,3 | 0.065 | 0.075 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OQ850LI/Lid_11B | 2OQ850LI | Large Deep Well Lid .560 Orifice | Large Deep Well .560 Orifice | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter | 4.919 | in | FDA,3 | 4.917 | 4.937 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OQ850LI/Lid_11B | 2OQ850LI | Large Deep Well Lid .560 Orifice | Large Deep Well .560 Orifice | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness | 0.0425 | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OQ850LI/Lid_11B | 2OQ850LI | Large Deep Well Lid .560 Orifice | Large Deep Well .560 Orifice | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OQ850LI/Lid_11B | 2OQ850LI | Large Deep Well Lid .560 Orifice | Large Deep Well .560 Orifice | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness | 0.024 | in | FDA,3 | 0.018 | 0.028 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OQ850LI/Lid_11B | 2OQ850LI | Large Deep Well Lid .560 Orifice | Large Deep Well .560 Orifice | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height | 1.488 | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OQ850LI/Lid_11B | 2OQ850LI | Large Deep Well Lid .560 Orifice | Large Deep Well .560 Orifice | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Outer Diameter | 0.165 | in | FDA,3 | 0.16 | 0.17 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OQ850LI/Lid_11B | 2OQ850LI | Large Deep Well Lid .560 Orifice | Large Deep Well .560 Orifice | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Weight | 34.04 | g | FDA,1 | 33 | 42 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | 2OQ850X | Canister - Extra Large, Ribbed, White | Canister - Extra Large, Ribbed, White | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size |  |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | 2OQ850X | Canister - Extra Large, Ribbed, White | Canister - Extra Large, Ribbed, White | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | [5 empty] | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | 2OQ850X | Canister - Extra Large, Ribbed, White | Canister - Extra Large, Ribbed, White | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | 2OQ850X | Canister - Extra Large, Ribbed, White | Canister - Extra Large, Ribbed, White | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | [5 empty] | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | 2OQ850X | Canister - Extra Large, Ribbed, White | Canister - Extra Large, Ribbed, White | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | Height |  | in | FDA,4 | 4.0625 | 4.3125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | 2OQ850X | Canister - Extra Large, Ribbed, White | Canister - Extra Large, Ribbed, White | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | Length |  | in | FDA,4 | 12.3125 | 12.5625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | 2OQ850X | Canister - Extra Large, Ribbed, White | Canister - Extra Large, Ribbed, White | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | [5 empty] | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | 2OQ850X | Canister - Extra Large, Ribbed, White | Canister - Extra Large, Ribbed, White | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | [5 empty] | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | 2OQ850X | Canister - Extra Large, Ribbed, White | Canister - Extra Large, Ribbed, White | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | Width |  | in | FDA,4 | 5.3125 | 5.5625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OQ882LI/Lid_12A | 2OQ882LI | LID LG DW DP GREEN 560 ORIFICE | Lg Dw Dp Green 560 Orifice | Lid_12A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Flip Lid Force |  | NY | RAW MATERIAL - NY | RM_LID | both
2OQ882LI/Lid_12A | 2OQ882LI | LID LG DW DP GREEN 560 ORIFICE | Lg Dw Dp Green 560 Orifice | Lid_12A | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | 1.61 | lb | FDA,2 | 0.75 |  | TM 143 | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OQ882LI/Lid_12A | 2OQ882LI | LID LG DW DP GREEN 560 ORIFICE | Lg Dw Dp Green 560 Orifice | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance Material is clean and free of particulate matter. (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OQ882LI/Lid_12A | 2OQ882LI | LID LG DW DP GREEN 560 ORIFICE | Lg Dw Dp Green 560 Orifice | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | 4.854 | in | FDA,3 | 4.852 | 4.862 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OQ882LI/Lid_12A | 2OQ882LI | LID LG DW DP GREEN 560 ORIFICE | Lg Dw Dp Green 560 Orifice | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Cap Thickness | 0.048 | in | FDA,3 | 0.045 | 0.055 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OQ882LI/Lid_12A | 2OQ882LI | LID LG DW DP GREEN 560 ORIFICE | Lg Dw Dp Green 560 Orifice | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Color | Dark Green (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OQ882LI/Lid_12A | 2OQ882LI | LID LG DW DP GREEN 560 ORIFICE | Lg Dw Dp Green 560 Orifice | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Pass (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OQ882LI/Lid_12A | 2OQ882LI | LID LG DW DP GREEN 560 ORIFICE | Lg Dw Dp Green 560 Orifice | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Insert Fin Gap | 0.074 | in | FDA,3 | 0.065 | 0.075 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OQ882LI/Lid_12A | 2OQ882LI | LID LG DW DP GREEN 560 ORIFICE | Lg Dw Dp Green 560 Orifice | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness | 0.045 | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OQ882LI/Lid_12A | 2OQ882LI | LID LG DW DP GREEN 560 ORIFICE | Lg Dw Dp Green 560 Orifice | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Odor | None detectable (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OQ882LI/Lid_12A | 2OQ882LI | LID LG DW DP GREEN 560 ORIFICE | Lg Dw Dp Green 560 Orifice | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness | 0.021 | in | FDA,3 | 0.018 | 0.028 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OQ882LI/Lid_12A | 2OQ882LI | LID LG DW DP GREEN 560 ORIFICE | Lg Dw Dp Green 560 Orifice | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height | 1.488 | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OQ882LI/Lid_12A | 2OQ882LI | LID LG DW DP GREEN 560 ORIFICE | Lg Dw Dp Green 560 Orifice | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Outer Diameter | 0.168 | in | FDA,3 | 0.16 | 0.17 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OQ882LI/Lid_12A | 2OQ882LI | LID LG DW DP GREEN 560 ORIFICE | Lg Dw Dp Green 560 Orifice | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Weight | 35.675 | g | FDA,1 | 32 | 42 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LID | both
2OU130/Web_NP5115_9.0 | 2OU130 | Grime Boss Heavy Duty Cleaning Wipes | Grime Boss Heavy Duty Cleaning Wipes | Web_NP5115_9.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2OU130/Web_NP5115_9.0 | 2OU130 | Grime Boss Heavy Duty Cleaning Wipes | Grime Boss Heavy Duty Cleaning Wipes | Web_NP5115_9.0 | NY-Analytical | 761 | Web Measurement | Web Width | 8.9375 | in | FDA,4 | 8.9375 | 9.0625 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2OU130/Web_NP5115_9.0 | 2OU130 | Grime Boss Heavy Duty Cleaning Wipes | Grime Boss Heavy Duty Cleaning Wipes | Web_NP5115_9.0 | NY-Analytical | 756 | Basis Weight | Basis Weight | 60.8 | #/Ream | FDA,1 | 55.6 | 68 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2OU130/Web_NP5115_9.0 | 2OU130 | Grime Boss Heavy Duty Cleaning Wipes | Grime Boss Heavy Duty Cleaning Wipes | Web_NP5115_9.0 | NY-Analytical | 762 | Seal Strength | Seal Strength (@325°F/40PSI/0.5Sec Dwell) | 2198.05 | g/in | FDA,0 | 1500 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | Bottle, 26Ml | Bottle, 26ml | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | Bottle, 26Ml | Bottle, 26ml | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | Bottle, 26Ml | Bottle, 26ml | NY-Analytical | 765 | Canister & Lid Inspection | Bottle Internal Diameter at Flange | 0.722 | in | FDA,3 | 0.709 | 0.729 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | Bottle, 26Ml | Bottle, 26ml | NY-Analytical | 765 | Canister & Lid Inspection | Color | Natural (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | Bottle, 26Ml | Bottle, 26ml | NY-Analytical | 765 | Canister & Lid Inspection | Overall Length | 4.855 | in | FDA,3 | 4.825 | 4.875 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | Bottle, 26Ml | Bottle, 26ml | NY-Analytical | 765 | Canister & Lid Inspection | Seal Flange Diameter | 0.8735 | in | FDA,3 | 0.864 | 0.874 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | Bottle, 26Ml | Bottle, 26ml | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | 2 | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | Bottle, 26Ml | Bottle, 26ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | Bottle, 26Ml | Bottle, 26ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | Bottle, 26Ml | Bottle, 26ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | Bottle, 26Ml | Bottle, 26ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | Bottle, 26Ml | Bottle, 26ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | Bottle, 26Ml | Bottle, 26ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 1 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | Bottle, 26Ml | Bottle, 26ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | 0.0028852 | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | Bottle, 26Ml | Bottle, 26ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | Bottle, 26Ml | Bottle, 26ml | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | Gram Negative Rods (OOS) |  |  |  |  | WI QAS 076 | WI_QAS_076 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX465B/Bottle, 26ml | 2OX465B | Bottle, 26ml | Bottle, 26Ml | Bottle, 26ml | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Mold (OOS) |  |  |  |  | WI QAS 076 | WI_QAS_076 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | Bottle, 5.1Ml | Bottle, 5.1ml | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Certificate of Conformance | COA | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Bottle, 5.1ml | Bottle, 5.1Ml | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Bottle, 5.1ml | Bottle, 5.1Ml | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | Bottle, 5.1Ml | Bottle, 5.1ml | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Bottle, 5.1ml | Bottle, 5.1Ml | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 765 | Canister & Lid Inspection | Bottle Internal Diameter at Flange | 0.4425 | in | FDA,3 | 0.442 | 0.452 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | Bottle, 5.1Ml | Bottle, 5.1ml | NY-Analytical | 765 | Canister & Lid Inspection | Bottle Internal Diameter at Flange | 0.443 | in | FDA,3 | 0.442 | 0.452 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Bottle, 5.1ml | Bottle, 5.1Ml | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 765 | Canister & Lid Inspection | Color | Natural (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | Bottle, 5.1Ml | Bottle, 5.1ml | NY-Analytical | 765 | Canister & Lid Inspection | Color | Natural (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Bottle, 5.1ml | Bottle, 5.1Ml | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 765 | Canister & Lid Inspection | Overall Length | 2.8805 | in | FDA,3 | 2.848 | 2.888 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | Bottle, 5.1Ml | Bottle, 5.1ml | NY-Analytical | 765 | Canister & Lid Inspection | Overall Length | 2.8737 | in | FDA,3 | 2.848 | 2.888 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Bottle, 5.1ml | Bottle, 5.1Ml | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 765 | Canister & Lid Inspection | Seal Flange Diameter | 0.55 | in | FDA,3 | 0.544 | 0.554 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | Bottle, 5.1Ml | Bottle, 5.1ml | NY-Analytical | 765 | Canister & Lid Inspection | Seal Flange Diameter | 0.55 | in | FDA,3 | 0.544 | 0.554 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Bottle, 5.1ml | Bottle, 5.1Ml | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight |  | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | Bottle, 5.1Ml | Bottle, 5.1ml | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | 4.5 | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Bottle, 5.1ml | Bottle, 5.1Ml | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | Bottle, 5.1Ml | Bottle, 5.1ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Bottle, 5.1ml | Bottle, 5.1Ml | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | Bottle, 5.1Ml | Bottle, 5.1ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Bottle, 5.1ml | Bottle, 5.1Ml | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | Bottle, 5.1Ml | Bottle, 5.1ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | 1.1 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Bottle, 5.1ml | Bottle, 5.1Ml | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | Bottle, 5.1Ml | Bottle, 5.1ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | 0.24444 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Bottle, 5.1ml | Bottle, 5.1Ml | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | Bottle, 5.1Ml | Bottle, 5.1ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Bottle, 5.1ml | Bottle, 5.1Ml | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 1.1 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | Bottle, 5.1Ml | Bottle, 5.1ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 1 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Bottle, 5.1ml | Bottle, 5.1Ml | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | 0.015027 | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | Bottle, 5.1Ml | Bottle, 5.1ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | 0.022222 | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Bottle, 5.1ml | Bottle, 5.1Ml | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | Bottle, 5.1Ml | Bottle, 5.1ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 1 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Bottle, 5.1ml | Bottle, 5.1Ml | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | Gram Positive Rods (Done) |  |  |  |  | WI QAS 076 | WI_QAS_076 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | Bottle, 5.1Ml | Bottle, 5.1ml | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | Gram Positive Rods (Done) |  |  |  |  | WI QAS 076 | WI_QAS_076 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510B/Microbiologu - Bioburden_Applicator_NY | 2OX510B | Bottle, 5.1ml | Bottle, 5.1Ml | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Mold (OOS) |  |  |  |  | WI QAS 076 | WI_QAS_076 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510B/Bottle, 5.1ml | 2OX510B | Bottle, 5.1ml | Bottle, 5.1Ml | Bottle, 5.1ml | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Mold (OOS) |  |  |  |  | WI QAS 076 | WI_QAS_076 | NY | RAW MATERIAL - NY | RM_CANNISTER | both
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil Seal 5.1 ml | Foil Seal 5.1 Ml | Foil seal 5.1 ml | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA | COA | NY | RAW MATERIAL - NY | RM_WEB | both
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil Seal 5.1 ml | Foil Seal 5.1 Ml | Foil seal 5.1 ml | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | 1.78 | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_WEB | both
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil Seal 5.1 ml | Foil Seal 5.1 Ml | Foil seal 5.1 ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_WEB | both
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil Seal 5.1 ml | Foil Seal 5.1 Ml | Foil seal 5.1 ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_WEB | both
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil Seal 5.1 ml | Foil Seal 5.1 Ml | Foil seal 5.1 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | 11.6 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_WEB | both
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil Seal 5.1 ml | Foil Seal 5.1 Ml | Foil seal 5.1 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | 6.5169 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_WEB | both
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil Seal 5.1 ml | Foil Seal 5.1 Ml | Foil seal 5.1 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 11 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_WEB | both
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil Seal 5.1 ml | Foil Seal 5.1 Ml | Foil seal 5.1 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_WEB | both
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil Seal 5.1 ml | Foil Seal 5.1 Ml | Foil seal 5.1 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_WEB | both
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil Seal 5.1 ml | Foil Seal 5.1 Ml | Foil seal 5.1 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_WEB | both
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil Seal 5.1 ml | Foil Seal 5.1 Ml | Foil seal 5.1 ml | NY-Analytical | 1056 | Functionality | Foil diameter to fit Bottle(Item#2OX510B)opening | Pass |  |  |  |  | Visual | Visual | NY | RAW MATERIAL - NY | RM_WEB | both
2OX510FS/Foil seal 5.1 ml | 2OX510FS | Foil Seal 5.1 ml | Foil Seal 5.1 Ml | Foil seal 5.1 ml | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | Gram Negative Rods (OOS) |  |  |  |  | WI QAS 076 | WI_QAS_076 | NY | RAW MATERIAL - NY | RM_WEB | both
2PD351/Web_NP5116_4.5 | 2PD351 | PDI Hygea BZK Antiseptic Towelette | PDI Hygea BZK Antiseptic Towelette | Web_NP5116_4.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2PD351/Web_NP5116_4.5 | 2PD351 | PDI Hygea BZK Antiseptic Towelette | PDI Hygea BZK Antiseptic Towelette | Web_NP5116_4.5 | NY-Analytical | 761 | Web Measurement | Web Width | 4.5 | in | FDA,4 | 4.4375 | 4.5625 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2PD351/Web_NP5116_4.5 | 2PD351 | PDI Hygea BZK Antiseptic Towelette | PDI Hygea BZK Antiseptic Towelette | Web_NP5116_4.5 | NY-Analytical | 756 | Basis Weight | Basis Weight | 59.82 | #/Ream | FDA,1 | 55.8 | 68.2 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2PD351/Web_NP5116_4.5 | 2PD351 | PDI Hygea BZK Antiseptic Towelette | PDI Hygea BZK Antiseptic Towelette | Web_NP5116_4.5 | NY-Analytical | 762 | Seal Strength | Seal Strength (@300°F/40psi/0.5 sec dewll) | 2163 | g/in | FDA,0 | 1500 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2QB114X/Web_NP3860_5.0 | 2QB114X | PDI Prevantics Skin Swab Packette | Prevantics Skin Swab Packette | Web_NP3860_5.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2QB114X/Web_NP3860_5.0 | 2QB114X | PDI Prevantics Skin Swab Packette | Prevantics Skin Swab Packette | Web_NP3860_5.0 | NY-Analytical | 761 | Web Measurement | Web Width | 5 | in | FDA,4 | 4.9375 | 5.0625 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2QB114X/Web_NP3860_5.0 | 2QB114X | PDI Prevantics Skin Swab Packette | Prevantics Skin Swab Packette | Web_NP3860_5.0 | NY-Analytical | 756 | Basis Weight | Basis Weight | 89.29 | #/Ream | FDA,1 | 78.9 | 96.5 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2QB114X/Web_NP3860_5.0 | 2QB114X | PDI Prevantics Skin Swab Packette | Prevantics Skin Swab Packette | Web_NP3860_5.0 | NY-Analytical | 762 | Seal Strength | Seal Strength (@350°F/40psi/1.0 sec dwell) | 9928 | g/in | FDA,0 | 7250 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2QB114X/Web_NP3860_5.0 | 2QB114X | PDI Prevantics Skin Swab Packette | Prevantics Skin Swab Packette | Web_NP3860_5.0 | NY-Analytical | 763 | Web Thickness | Web Thickness | 5 | mil | FDA,1 | 4.7 | 5.7 | TM 132 | TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2RQ890/Web_NP9508_18.125 | 2RQ890 | PIL WH CL-250XWH 18.125 NP9508 | Pil Wh Cl-250Xwh 18.125 Np9508 | Web_NP9508_18.125 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2RQ890/Web_NP9508_18.125 | 2RQ890 | PIL WH CL-250XWH 18.125 NP9508 | Pil Wh Cl-250Xwh 18.125 Np9508 | Web_NP9508_18.125 | NY-Analytical | 761 | Web Measurement | Web Width | 18.125 | in | FDA,4 | 18.0625 | 18.1875 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2RQ890/Web_NP9508_18.125 | 2RQ890 | PIL WH CL-250XWH 18.125 NP9508 | Pil Wh Cl-250Xwh 18.125 Np9508 | Web_NP9508_18.125 | NY-Analytical | 756 | Basis Weight | Basis Weight |  | #/Ream | FDA,1 | 37.8 | 46.2 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2RQ890/Web_NP9508_18.125 | 2RQ890 | PIL WH CL-250XWH 18.125 NP9508 | Pil Wh Cl-250Xwh 18.125 Np9508 | Web_NP9508_18.125 | NY-Analytical | 762 | Seal Strength | Seal Strength (@300°F/40psi/0.5 sec dewll) | Sample peeled apart and showed stress whitening when sealed (Done) |  |  |  |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2RQ890/Web_NP9508_18.125 | 2RQ890 | PIL WH CL-250XWH 18.125 NP9508 | Pil Wh Cl-250Xwh 18.125 Np9508 | Web_NP9508_18.125 | NY-Analytical | 763 | Web Thickness | Web Thickness | 2.6 | mil | FDA,1 | 2.3 | 2.8 | TM 132 | TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2RS324/Web_NP3860_12.5 | 2RS324 | PDI Prevantics Swabstick | Prevantics Swabstick | Web_NP3860_12.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2RS324/Web_NP3860_12.5 | 2RS324 | PDI Prevantics Swabstick | Prevantics Swabstick | Web_NP3860_12.5 | NY-Analytical | 761 | Web Measurement | Web Width | 12.5 | in | FDA,4 | 12.4375 | 12.5625 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2RS324/Web_NP3860_12.5 | 2RS324 | PDI Prevantics Swabstick | Prevantics Swabstick | Web_NP3860_12.5 | NY-Analytical | 756 | Basis Weight | Basis Weight | 89 | #/Ream | FDA,1 | 78.9 | 96.5 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2RS324/Web_NP3860_12.5 | 2RS324 | PDI Prevantics Swabstick | Prevantics Swabstick | Web_NP3860_12.5 | NY-Analytical | 762 | Seal Strength | Seal Strength (@350°F/40psi/1.0 sec dwell) | 9622 | g/in | FDA,0 | 7250 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2RS324/Web_NP3860_12.5 | 2RS324 | PDI Prevantics Swabstick | Prevantics Swabstick | Web_NP3860_12.5 | NY-Analytical | 763 | Web Thickness | Web Thickness | 4.91 | mil | FDA,1 | 4.7 | 5.7 | TM 132 | TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2RS407/Web_NP3885_13.0 | 2RS407 | PDI Prevantics Skin Swabsticks | Prevantics Skin Swabsticks | Web_NP3885_13.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2RS407/Web_NP3885_13.0 | 2RS407 | PDI Prevantics Skin Swabsticks | Prevantics Skin Swabsticks | Web_NP3885_13.0 | NY-Analytical | 761 | Web Measurement | Web Width | 12.9375 | in | FDA,4 | 12.9375 | 13.0625 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2RS407/Web_NP3885_13.0 | 2RS407 | PDI Prevantics Skin Swabsticks | Prevantics Skin Swabsticks | Web_NP3885_13.0 | NY-Analytical | 756 | Basis Weight | Basis Weight | 80.69 | #/Ream | FDA,0 | 74 | 90 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2RS407/Web_NP3885_13.0 | 2RS407 | PDI Prevantics Skin Swabsticks | Prevantics Skin Swabsticks | Web_NP3885_13.0 | NY-Analytical | 762 | Seal Strength | Seal Strength (@350°F/50psi/1.0 sec dwell) | 1941 | g/in | FDA,0 | 700 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2RS419/Web_NP3885_16 | 2RS419 | PDI Prevantics Maxi Swabstick | Prevantics Maxi Swabstick | Web_NP3885_16 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2RS419/Web_NP3885_16 | 2RS419 | PDI Prevantics Maxi Swabstick | Prevantics Maxi Swabstick | Web_NP3885_16 | NY-Analytical | 761 | Web Measurement | Web Width | 16 | in | FDA,4 | 15.9375 | 16.0625 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2RS419/Web_NP3885_16 | 2RS419 | PDI Prevantics Maxi Swabstick | Prevantics Maxi Swabstick | Web_NP3885_16 | NY-Analytical | 756 | Basis Weight | Basis Weight | 81.85 | #/Ream | FDA,0 | 74 | 90 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2RS419/Web_NP3885_16 | 2RS419 | PDI Prevantics Maxi Swabstick | Prevantics Maxi Swabstick | Web_NP3885_16 | NY-Analytical | 762 | Seal Strength | Seal Strength (@350°F/50psi/1.0 sec dwell) | 2030.5 | g/in | FDA,0 | 700 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2SS273/Web_NP3860_15.5 | 2SS273 | PDI Prevantics Maxi Swabstick | Prevantics Maxi Swabstick | Web_NP3860_15.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Web Measurement |  | NY | RAW MATERIAL - NY | RM_WEB | both
2SS273/Web_NP3860_15.5 | 2SS273 | PDI Prevantics Maxi Swabstick | Prevantics Maxi Swabstick | Web_NP3860_15.5 | NY-Analytical | 761 | Web Measurement | Web Width | 15.5 | in | FDA,4 | 15.4375 | 15.5625 | Measurement | Measurement | NY | RAW MATERIAL - NY | RM_WEB | both
2SS273/Web_NP3860_15.5 | 2SS273 | PDI Prevantics Maxi Swabstick | Prevantics Maxi Swabstick | Web_NP3860_15.5 | NY-Analytical | 756 | Basis Weight | Basis Weight | 87.7 | #/Ream | FDA,1 | 78.9 | 96.5 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2SS273/Web_NP3860_15.5 | 2SS273 | PDI Prevantics Maxi Swabstick | Prevantics Maxi Swabstick | Web_NP3860_15.5 | NY-Analytical | 762 | Seal Strength | Seal Strength (@350°F/40psi/1.0 sec dwell) | 10111.34 | g/in | FDA,0 | 7250 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2SS273/Web_NP3860_15.5 | 2SS273 | PDI Prevantics Maxi Swabstick | Prevantics Maxi Swabstick | Web_NP3860_15.5 | NY-Analytical | 763 | Web Thickness | Web Thickness | 5.2 | mil | FDA,1 | 4.7 | 5.7 | TM 132 | TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
2SS428/Web_NP3860_8.0 | 2SS428 | PDI Prevantics Skin Swabstick | Prevantics Skin Swabstick | Web_NP3860_8.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Basis Weight |  | NY | RAW MATERIAL - NY | RM_WEB | both
2SS428/Web_NP3860_8.0 | 2SS428 | PDI Prevantics Skin Swabstick | Prevantics Skin Swabstick | Web_NP3860_8.0 | NY-Analytical | 756 | Basis Weight | Basis Weight | 87.36 | #/Ream | FDA,1 | 78.9 | 96.5 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_WEB | both
2SS428/Web_NP3860_8.0 | 2SS428 | PDI Prevantics Skin Swabstick | Prevantics Skin Swabstick | Web_NP3860_8.0 | NY-Analytical | 762 | Seal Strength | Seal Strength (@350°F/40psi/1.0 sec dwell) | 9981.29 | g/in | FDA,0 | 7250 |  | TM 107 | TM107 | NY | RAW MATERIAL - NY | RM_WEB | both
2SS428/Web_NP3860_8.0 | 2SS428 | PDI Prevantics Skin Swabstick | Prevantics Skin Swabstick | Web_NP3860_8.0 | NY-Analytical | 763 | Web Thickness | Web Thickness | 5.2 | mil | FDA,1 | 4.7 | 5.7 | TM 132 | TM132 | NY | RAW MATERIAL - NY | RM_WEB | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Spunlaced Nonwoven 55gsm 37 1/2" | Spunlaced Nonwoven 55 GSM 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Spunlaced Nonwoven 55gsm 37 1/2" | Spunlaced Nonwoven 55 GSM 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 759 | Applicator Measurements | Roll Diameter | 44 | in | FDA,0 | 44 | 46 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Spunlaced Nonwoven 55gsm 37 1/2" | Spunlaced Nonwoven 55 GSM 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 759 | Applicator Measurements | Roll Width | 37.5 | in | FDA,1 | 37.3 | 37.7 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Spunlaced Nonwoven 55gsm 37 1/2" | Spunlaced Nonwoven 55 GSM 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 756 | Basis Weight | Basis Weight | 55.54 | GSM | FDA,1 | 49.5 | 60.5 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Spunlaced Nonwoven 55gsm 37 1/2" | Spunlaced Nonwoven 55 GSM 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 758 | Tensile Strength | Wet CD | 1147.63 | g/in | FDA,0 | 692 | 1438 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Spunlaced Nonwoven 55gsm 37 1/2" | Spunlaced Nonwoven 55 GSM 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 758 | Tensile Strength | Wet MD | 4131.08 | g/in | FDA,0 | 2785 | 5785 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Spunlaced Nonwoven 55gsm 37 1/2" | Spunlaced Nonwoven 55 GSM 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 748 | Visual Inspection | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Spunlaced Nonwoven 55gsm 37 1/2" | Spunlaced Nonwoven 55 GSM 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 748 | Visual Inspection | Color | White to Off-White (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Spunlaced Nonwoven 55gsm 37 1/2" | Spunlaced Nonwoven 55 GSM 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 748 | Visual Inspection | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Spunlaced Nonwoven 55gsm 37 1/2" | Spunlaced Nonwoven 55 GSM 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 509 | J&J_Pathogens | Biggy or SDA | No Growth_SDA |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Spunlaced Nonwoven 55gsm 37 1/2" | Spunlaced Nonwoven 55 GSM 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 509 | J&J_Pathogens | Cetrimide | No Growth |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Spunlaced Nonwoven 55gsm 37 1/2" | Spunlaced Nonwoven 55 GSM 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 524 | J&J Fungal | JJ Fungal RM | 0 | CFU/g |  |  | 200 | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Spunlaced Nonwoven 55gsm 37 1/2" | Spunlaced Nonwoven 55 GSM 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 509 | J&J_Pathogens | MacConkey | No Growth |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Spunlaced Nonwoven 55gsm 37 1/2" | Spunlaced Nonwoven 55 GSM 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 509 | J&J_Pathogens | MSA or VJ | No Growth_MSA |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Spunlaced Nonwoven 55gsm 37 1/2" | Spunlaced Nonwoven 55 GSM 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 700 | J&J SDA Controls | Negative Control | No Growth SDA |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Spunlaced Nonwoven 55gsm 37 1/2" | Spunlaced Nonwoven 55 GSM 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | No Growth TSB |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Spunlaced Nonwoven 55gsm 37 1/2" | Spunlaced Nonwoven 55 GSM 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 509 | J&J_Pathogens | TSA | No Growth |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Spunlaced Nonwoven 55gsm 37 1/2" | Spunlaced Nonwoven 55 GSM 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ RM | 0 | CFU/g |  |  | 200 | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Spunlaced Nonwoven 55gsm 37 1/2" | Spunlaced Nonwoven 55 GSM 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 517 | J&J_Enrichment | TSB | Not Turbid |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Spunlaced Nonwoven 55gsm 37 1/2" | Spunlaced Nonwoven 55 GSM 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 573 | GramStain_Candida | Gram Stain J&JCandida | Non-Subculturable Growth (Done) |  |  |  |  | WI QAS 076 | WI_QAS_076 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | 3AP211 | Spunlaced Nonwoven 55gsm 37 1/2" | Spunlaced Nonwoven 55 GSM 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 545 | GramStain_J&J_TSA | GramStain J&J TSA | Gram Positive Rods (Done) |  |  |  |  | WI QAS 076 | WI_QAS_076 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 3AQ470-A | Rolls Spunlaced Nonwoven – 37.75” – 50gsm | Rolls Spunlaced Nonwoven - 37.75" - 50gsm | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Color | White with no off colors or shading |  |  |  |  | Color (TM 208) | TM208 | IN | RAW MATERIAL - INDIANA | RM_APPLICATOR | both
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 3AQ470-A | Rolls Spunlaced Nonwoven – 37.75” – 50gsm | Rolls Spunlaced Nonwoven - 37.75" - 50gsm | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Appearance | Free of hard wrinkles. Formation uniform (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | IN | RAW MATERIAL - INDIANA | RM_APPLICATOR | both
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 3AQ470-A | Rolls Spunlaced Nonwoven – 37.75” – 50gsm | Rolls Spunlaced Nonwoven - 37.75" - 50gsm | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | 80 |  | FDA,0 |  |  | See Analyte for Method # | See_Analyte_for_Method | IN | RAW MATERIAL - INDIANA | RM_APPLICATOR | both
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 3AQ470-A | Rolls Spunlaced Nonwoven – 37.75” – 50gsm | Rolls Spunlaced Nonwoven - 37.75" - 50gsm | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Basis Weight | 49.51 | GSM | FDA,1 | 45 | 55 | See Analyte for Method # | See_Analyte_for_Method | IN | RAW MATERIAL - INDIANA | RM_APPLICATOR | both
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 3AQ470-A | Rolls Spunlaced Nonwoven – 37.75” – 50gsm | Rolls Spunlaced Nonwoven - 37.75" - 50gsm | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Caliper | 0.59 | mm | FDA,2 | 0.51 | 0.71 | See Analyte for Method # | See_Analyte_for_Method | IN | RAW MATERIAL - INDIANA | RM_APPLICATOR | both
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 3AQ470-A | Rolls Spunlaced Nonwoven – 37.75” – 50gsm | Rolls Spunlaced Nonwoven - 37.75" - 50gsm | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | IN | RAW MATERIAL - INDIANA | RM_APPLICATOR | both
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 3AQ470-A | Rolls Spunlaced Nonwoven – 37.75” – 50gsm | Rolls Spunlaced Nonwoven - 37.75" - 50gsm | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Odor | May have slight odor when wet (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | IN | RAW MATERIAL - INDIANA | RM_APPLICATOR | both
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 3AQ470-A | Rolls Spunlaced Nonwoven – 37.75” – 50gsm | Rolls Spunlaced Nonwoven - 37.75" - 50gsm | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter | 50 | in | FDA,0 | 45 | 60 | See Analyte for Method # | See_Analyte_for_Method | IN | RAW MATERIAL - INDIANA | RM_APPLICATOR | both
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 3AQ470-A | Rolls Spunlaced Nonwoven – 37.75” – 50gsm | Rolls Spunlaced Nonwoven - 37.75" - 50gsm | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Width | 37.75 | in | FDA,2 | 37.5 | 38 | See Analyte for Method # | See_Analyte_for_Method | IN | RAW MATERIAL - INDIANA | RM_APPLICATOR | both
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 3AQ470-A | Rolls Spunlaced Nonwoven – 37.75” – 50gsm | Rolls Spunlaced Nonwoven - 37.75" - 50gsm | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Wet CD | 1634 | g/in | FDA,0 | 1100 | 2300 | See Analyte for Method # | See_Analyte_for_Method | IN | RAW MATERIAL - INDIANA | RM_APPLICATOR | both
3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | 3AQ470-A | Rolls Spunlaced Nonwoven – 37.75” – 50gsm | Rolls Spunlaced Nonwoven - 37.75" - 50gsm | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Wet MD | 6147 | g/in | FDA,0 | 3000 | 7000 | See Analyte for Method # | See_Analyte_for_Method | IN | RAW MATERIAL - INDIANA | RM_APPLICATOR | both
3AU228/Applicator_06I | 3AU228 | AP1 SP SPUNLACE 55G 50/50 7.5 | Ap1 Sp Spunlace 55G 50/50 7.5 | Applicator_06I | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AU228/Applicator_06I | 3AU228 | AP1 SP SPUNLACE 55G 50/50 7.5 | Ap1 Sp Spunlace 55G 50/50 7.5 | Applicator_06I | NY-Analytical | 759 | Applicator Measurements | Roll Diameter | 30 | in | FDA,0 | 28 | 32 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AU228/Applicator_06I | 3AU228 | AP1 SP SPUNLACE 55G 50/50 7.5 | Ap1 Sp Spunlace 55G 50/50 7.5 | Applicator_06I | NY-Analytical | 759 | Applicator Measurements | Roll Width | 7.5 | in | FDA,3 | 7.375 | 7.625 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AU228/Applicator_06I | 3AU228 | AP1 SP SPUNLACE 55G 50/50 7.5 | Ap1 Sp Spunlace 55G 50/50 7.5 | Applicator_06I | NY-Analytical | 756 | Basis Weight | Basis Weight | 50.29 | GSM | FDA,1 | 49.5 | 60.5 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AU228/Applicator_06I | 3AU228 | AP1 SP SPUNLACE 55G 50/50 7.5 | Ap1 Sp Spunlace 55G 50/50 7.5 | Applicator_06I | NY-Analytical | 757 | Caliper | Caliper | 0.52 | mm | FDA,2 | 0.5 | 0.64 | TM 102 | TM102 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AU228/Applicator_06I | 3AU228 | AP1 SP SPUNLACE 55G 50/50 7.5 | Ap1 Sp Spunlace 55G 50/50 7.5 | Applicator_06I | NY-Analytical | 758 | Tensile Strength | Wet CD | 1925.71 | g/in | FDA,0 | 1443 | 2463 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AU228/Applicator_06I | 3AU228 | AP1 SP SPUNLACE 55G 50/50 7.5 | Ap1 Sp Spunlace 55G 50/50 7.5 | Applicator_06I | NY-Analytical | 758 | Tensile Strength | Wet MD | 6181.64 | g/in | FDA,0 | 4883 | 7967 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AU228/Applicator_06I | 3AU228 | AP1 SP SPUNLACE 55G 50/50 7.5 | Ap1 Sp Spunlace 55G 50/50 7.5 | Applicator_06I | NY-Analytical | 748 | Visual Inspection | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AU228/Applicator_06I | 3AU228 | AP1 SP SPUNLACE 55G 50/50 7.5 | Ap1 Sp Spunlace 55G 50/50 7.5 | Applicator_06I | NY-Analytical | 748 | Visual Inspection | Color | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AU228/Applicator_06I | 3AU228 | AP1 SP SPUNLACE 55G 50/50 7.5 | Ap1 Sp Spunlace 55G 50/50 7.5 | Applicator_06I | NY-Analytical | 748 | Visual Inspection | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AU228/Applicator_06I | 3AU228 | AP1 SP SPUNLACE 55G 50/50 7.5 | Ap1 Sp Spunlace 55G 50/50 7.5 | Applicator_06I | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | 1.03 | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AU228/Applicator_06I | 3AU228 | AP1 SP SPUNLACE 55G 50/50 7.5 | Ap1 Sp Spunlace 55G 50/50 7.5 | Applicator_06I | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AU228/Applicator_06I | 3AU228 | AP1 SP SPUNLACE 55G 50/50 7.5 | Ap1 Sp Spunlace 55G 50/50 7.5 | Applicator_06I | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AU228/Applicator_06I | 3AU228 | AP1 SP SPUNLACE 55G 50/50 7.5 | Ap1 Sp Spunlace 55G 50/50 7.5 | Applicator_06I | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AU228/Applicator_06I | 3AU228 | AP1 SP SPUNLACE 55G 50/50 7.5 | Ap1 Sp Spunlace 55G 50/50 7.5 | Applicator_06I | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AU228/Applicator_06I | 3AU228 | AP1 SP SPUNLACE 55G 50/50 7.5 | Ap1 Sp Spunlace 55G 50/50 7.5 | Applicator_06I | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AU228/Applicator_06I | 3AU228 | AP1 SP SPUNLACE 55G 50/50 7.5 | Ap1 Sp Spunlace 55G 50/50 7.5 | Applicator_06I | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AU228/Applicator_06I | 3AU228 | AP1 SP SPUNLACE 55G 50/50 7.5 | Ap1 Sp Spunlace 55G 50/50 7.5 | Applicator_06I | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3AU228/Applicator_06I | 3AU228 | AP1 SP SPUNLACE 55G 50/50 7.5 | Ap1 Sp Spunlace 55G 50/50 7.5 | Applicator_06I | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | Applicator Textured Thermal Bond – 52.6 gsm | Textured Thermal Bond – 52.6 GSM | REDUCED TESTING_NY_3BB406 | NY-Analytical | 884 | Applicator_Inspection | Color | White with no off colors or shading |  |  |  |  | Color (TM 208) | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | Applicator Textured Thermal Bond – 52.6 gsm | Textured Thermal Bond – 52.6 GSM | REDUCED TESTING_NY_3BB406 | NY-Analytical | 884 | Applicator_Inspection | Odor | May have residual sulfur odor when wet |  |  |  |  | Odor (TM 208) | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | Applicator Textured Thermal Bond – 52.6 gsm | Textured Thermal Bond – 52.6 GSM | REDUCED TESTING_NY_3BB406 | NY-Analytical | 884 | Applicator_Inspection | Appearance | Meets specifications (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | Applicator Textured Thermal Bond – 52.6 gsm | Textured Thermal Bond – 52.6 GSM | REDUCED TESTING_NY_3BB406 | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | 200 |  | FDA,0 |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | Applicator Textured Thermal Bond – 52.6 gsm | Textured Thermal Bond – 52.6 GSM | REDUCED TESTING_NY_3BB406 | NY-Analytical | 884 | Applicator_Inspection | Basis Weight | 52.74 | GSM | FDA,1 | 48.9 | 56.3 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | Applicator Textured Thermal Bond – 52.6 gsm | Textured Thermal Bond – 52.6 GSM | REDUCED TESTING_NY_3BB406 | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | Applicator Textured Thermal Bond – 52.6 gsm | Textured Thermal Bond – 52.6 GSM | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight |  | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | Applicator Textured Thermal Bond – 52.6 gsm | Textured Thermal Bond – 52.6 GSM | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | Applicator Textured Thermal Bond – 52.6 gsm | Textured Thermal Bond – 52.6 GSM | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | Applicator Textured Thermal Bond – 52.6 gsm | Textured Thermal Bond – 52.6 GSM | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | Applicator Textured Thermal Bond – 52.6 gsm | Textured Thermal Bond – 52.6 GSM | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | Applicator Textured Thermal Bond – 52.6 gsm | Textured Thermal Bond – 52.6 GSM | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | Applicator Textured Thermal Bond – 52.6 gsm | Textured Thermal Bond – 52.6 GSM | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 1 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | Applicator Textured Thermal Bond – 52.6 gsm | Textured Thermal Bond – 52.6 GSM | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | Applicator Textured Thermal Bond – 52.6 gsm | Textured Thermal Bond – 52.6 GSM | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) |  | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3BB406/REDUCED TESTING_NY_3BB406 | 3BB406 | Applicator Textured Thermal Bond – 52.6 gsm | Textured Thermal Bond – 52.6 GSM | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Mold (OOS) |  |  |  |  | WI QAS 076 | WI_QAS_076 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3BJ141/Applicator_07K | 3BJ141 | AP1 VJ 43G 40/60 BGDOT 11.625 | Ap1 Vj 43G 40/60 Bgdot 11.625 | Applicator_07K | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3BJ141/Applicator_07K | 3BJ141 | AP1 VJ 43G 40/60 BGDOT 11.625 | Ap1 Vj 43G 40/60 Bgdot 11.625 | Applicator_07K | NY-Analytical | 759 | Applicator Measurements | Roll Width | 11.625 | in | FDA,2 | 11.5 | 11.75 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3BJ141/Applicator_07K | 3BJ141 | AP1 VJ 43G 40/60 BGDOT 11.625 | Ap1 Vj 43G 40/60 Bgdot 11.625 | Applicator_07K | NY-Analytical | 756 | Basis Weight | Basis Weight | 41.33 | GSM | FDA,1 | 38.7 | 47.3 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3BJ141/Applicator_07K | 3BJ141 | AP1 VJ 43G 40/60 BGDOT 11.625 | Ap1 Vj 43G 40/60 Bgdot 11.625 | Applicator_07K | NY-Analytical | 757 | Caliper | Caliper | 0.0225 | in | FDA,3 | 0.022 | 0.03 | TM 102 | TM102 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3BJ141/Applicator_07K | 3BJ141 | AP1 VJ 43G 40/60 BGDOT 11.625 | Ap1 Vj 43G 40/60 Bgdot 11.625 | Applicator_07K | NY-Analytical | 758 | Tensile Strength | Wet CD | 1064.23 | g/in | FDA,0 | 921 | 1535 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3BJ141/Applicator_07K | 3BJ141 | AP1 VJ 43G 40/60 BGDOT 11.625 | Ap1 Vj 43G 40/60 Bgdot 11.625 | Applicator_07K | NY-Analytical | 758 | Tensile Strength | Wet MD | 5060.98 | g/in | FDA,0 | 3623 | 5650 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3BJ141/Applicator_07K | 3BJ141 | AP1 VJ 43G 40/60 BGDOT 11.625 | Ap1 Vj 43G 40/60 Bgdot 11.625 | Applicator_07K | NY-Analytical | 748 | Visual Inspection | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3BJ141/Applicator_07K | 3BJ141 | AP1 VJ 43G 40/60 BGDOT 11.625 | Ap1 Vj 43G 40/60 Bgdot 11.625 | Applicator_07K | NY-Analytical | 748 | Visual Inspection | Color | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3BJ141/Applicator_07K | 3BJ141 | AP1 VJ 43G 40/60 BGDOT 11.625 | Ap1 Vj 43G 40/60 Bgdot 11.625 | Applicator_07K | NY-Analytical | 748 | Visual Inspection | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3DQ595/Applicator_10D | 3DQ595 | AP1 SG THERMAL SH-PP-40 30" | Ap1 Sg Thermal Sh-Pp-40 30" | Applicator_10D | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3DQ595/Applicator_10D | 3DQ595 | AP1 SG THERMAL SH-PP-40 30" | Ap1 Sg Thermal Sh-Pp-40 30" | Applicator_10D | NY-Analytical | 759 | Applicator Measurements | Roll Diameter | 46 | in | FDA,0 | 44 | 47 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3DQ595/Applicator_10D | 3DQ595 | AP1 SG THERMAL SH-PP-40 30" | Ap1 Sg Thermal Sh-Pp-40 30" | Applicator_10D | NY-Analytical | 759 | Applicator Measurements | Roll Width | 30 | in | FDA,3 | 29.875 | 30.125 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3DQ595/Applicator_10D | 3DQ595 | AP1 SG THERMAL SH-PP-40 30" | Ap1 Sg Thermal Sh-Pp-40 30" | Applicator_10D | NY-Analytical | 756 | Basis Weight | Basis Weight | 39.18 | GSM | FDA,1 | 38 | 42 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3DQ595/Applicator_10D | 3DQ595 | AP1 SG THERMAL SH-PP-40 30" | Ap1 Sg Thermal Sh-Pp-40 30" | Applicator_10D | NY-Analytical | 757 | Caliper | Caliper | 0.36 | mm | FDA,2 | 0.35 | 0.45 | TM 102 | TM102 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3DQ595/Applicator_10D | 3DQ595 | AP1 SG THERMAL SH-PP-40 30" | Ap1 Sg Thermal Sh-Pp-40 30" | Applicator_10D | NY-Analytical | 758 | Tensile Strength | Dry CD | 667.15 | g/in | FDA,0 | 513 | 1130 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3DQ595/Applicator_10D | 3DQ595 | AP1 SG THERMAL SH-PP-40 30" | Ap1 Sg Thermal Sh-Pp-40 30" | Applicator_10D | NY-Analytical | 758 | Tensile Strength | Dry CD | 2695.76 | g/in | FDA,0 | 2054 | 4108 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3DQ595/Applicator_10D | 3DQ595 | AP1 SG THERMAL SH-PP-40 30" | Ap1 Sg Thermal Sh-Pp-40 30" | Applicator_10D | NY-Analytical | 748 | Visual Inspection | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3DQ595/Applicator_10D | 3DQ595 | AP1 SG THERMAL SH-PP-40 30" | Ap1 Sg Thermal Sh-Pp-40 30" | Applicator_10D | NY-Analytical | 748 | Visual Inspection | Color | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3DQ595/Applicator_10D | 3DQ595 | AP1 SG THERMAL SH-PP-40 30" | Ap1 Sg Thermal Sh-Pp-40 30" | Applicator_10D | NY-Analytical | 748 | Visual Inspection | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3DQ595N/Applicator_10I | 3DQ595N | AP1 SG THERMAL SH-PP-40 30" | Ap1 Sg Thermal Sh-Pp-40 30" | Applicator_10I | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3DQ595N/Applicator_10I | 3DQ595N | AP1 SG THERMAL SH-PP-40 30" | Ap1 Sg Thermal Sh-Pp-40 30" | Applicator_10I | NY-Analytical | 759 | Applicator Measurements | Roll Diameter | 46 | in | FDA,0 | 44 | 47 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3DQ595N/Applicator_10I | 3DQ595N | AP1 SG THERMAL SH-PP-40 30" | Ap1 Sg Thermal Sh-Pp-40 30" | Applicator_10I | NY-Analytical | 759 | Applicator Measurements | Roll Width | 30 | in | FDA,3 | 29.875 | 30.125 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3DQ595N/Applicator_10I | 3DQ595N | AP1 SG THERMAL SH-PP-40 30" | Ap1 Sg Thermal Sh-Pp-40 30" | Applicator_10I | NY-Analytical | 756 | Basis Weight | Basis Weight | 39.83 | GSM | FDA,0 | 38 | 42 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3DQ595N/Applicator_10I | 3DQ595N | AP1 SG THERMAL SH-PP-40 30" | Ap1 Sg Thermal Sh-Pp-40 30" | Applicator_10I | NY-Analytical | 757 | Caliper | Caliper | 0.41 | mm | FDA,2 | 0.35 | 0.45 | TM 102 | TM102 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3DQ595N/Applicator_10I | 3DQ595N | AP1 SG THERMAL SH-PP-40 30" | Ap1 Sg Thermal Sh-Pp-40 30" | Applicator_10I | NY-Analytical | 758 | Tensile Strength | Dry CD | 573.11 | g/in | FDA,0 | 513 | 1130 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3DQ595N/Applicator_10I | 3DQ595N | AP1 SG THERMAL SH-PP-40 30" | Ap1 Sg Thermal Sh-Pp-40 30" | Applicator_10I | NY-Analytical | 758 | Tensile Strength | Dry CD | 3431.99 | g/in | FDA,0 | 2054 | 4108 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3DQ595N/Applicator_10I | 3DQ595N | AP1 SG THERMAL SH-PP-40 30" | Ap1 Sg Thermal Sh-Pp-40 30" | Applicator_10I | NY-Analytical | 748 | Visual Inspection | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3DQ595N/Applicator_10I | 3DQ595N | AP1 SG THERMAL SH-PP-40 30" | Ap1 Sg Thermal Sh-Pp-40 30" | Applicator_10I | NY-Analytical | 748 | Visual Inspection | Color | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3DQ595N/Applicator_10I | 3DQ595N | AP1 SG THERMAL SH-PP-40 30" | Ap1 Sg Thermal Sh-Pp-40 30" | Applicator_10I | NY-Analytical | 748 | Visual Inspection | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Color | White with no off colors or shading |  |  |  |  | Color (TM 208) | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Color | White with no off colors or shading |  |  |  |  | Color (TM 208) | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Odor | May have residual sulfur odor when wet |  |  |  |  | Odor (TM 208) | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Odor | May have residual sulfur odor when wet |  |  |  |  | Odor (TM 208) | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis | COA | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator_Inspection |  | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Appearance | Smooth free of hard wrinkles, formation uniform matches embossing pattern (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Appearance | Smooth free of hard wrinkles, formation uniform matches embossing pattern (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | 125 |  | FDA,0 |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | 80 |  | FDA,0 |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Basis Weight | 61.79 | GSM | FDA,1 | 60 | 70 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Basis Weight | 65.66 | GSM | FDA,1 | 60 | 70 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Caliper | 0.64 | mm | FDA,2 | 0.54 | 0.66 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Caliper | 0.66 | mm | FDA,2 | 0.54 | 0.66 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Ratio: Wet MD / Wet CD | 0 |  | FDA,1 |  | 4.5 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Ratio: Wet MD / Wet CD | 0 |  | FDA,1 |  | 4.5 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter | 45 | in | FDA,0 | 42 | 48 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter | 43 | in | FDA,0 | 42 | 48 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Width | 6.062 | in | FDA,2 | 6 | 6.125 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Roll Width | 6 | in | FDA,2 | 6 | 6.125 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Wet CD | 416 | g/in | FDA,0 | 375 | 575 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Wet CD | 426 | g/in | FDA,0 | 375 | 575 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Wet MD | 692 | g/in | FDA,0 | 485 | 915 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Wet MD | 769 | g/in | FDA,0 | 485 | 915 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 756 | Basis Weight | Basis Weight | 67.38 | GSM | FDA,1 | 60 | 70 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 756 | Basis Weight | Basis Weight |  | GSM | FDA,1 | 60 | 70 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 757 | Caliper | Caliper | 0.64 | mm | FDA,2 | 0.54 | 0.66 | TM 102 | TM102 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 757 | Caliper | Caliper | 0.63 | mm | FDA,2 | 0.54 | 0.66 | TM 102 | TM102 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 758 | Tensile Strength | Elongation % Wet CD | 43.05 | % | FDA,0 | 15 | 65 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 758 | Tensile Strength | Elongation % Wet CD | 25.08 | % | FDA,0 | 15 | 65 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 758 | Tensile Strength | Elongation % Wet MD | 30.3 | % | FDA,0 | 2 | 50 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 758 | Tensile Strength | Elongation % Wet MD | 17.02 | % | FDA,0 | 2 | 50 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 758 | Tensile Strength | Wet CD | 561.7 | g/in | FDA,0 | 375 | 575 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 758 | Tensile Strength | Wet CD | 495 | g/in | FDA,0 | 375 | 575 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 758 | Tensile Strength | Wet MD | 694.79 | g/in | FDA,0 | 485 | 915 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 758 | Tensile Strength | Wet MD | 648 | g/in | FDA,0 | 485 | 915 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 748 | Visual Inspection | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 748 | Visual Inspection | Color | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 748 | Visual Inspection | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Microbiologu - Bioburden_Applicator_NY | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 748 | Visual Inspection | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | 2.7 | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 1 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | 0.96774 | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | Gram Positive Rods (Done) |  |  |  |  | WI QAS 076 | WI_QAS_076 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | 3OA347 | Nonwoven Rolls Dispersible Unembossed - 6" width-65 gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width-65 GSM | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Mold (OOS) |  |  |  |  | WI QAS 076 | WI_QAS_076 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | 3OA426 | Rolls Spunlaced Nonwoven – 8” – 52 gsm | Rolls Spunlaced Nonwoven – 8” – 52 GSM | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Color | White with no off colors or shading |  |  |  |  | Color (TM 208) | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | 3OA426 | Rolls Spunlaced Nonwoven – 8” – 52 gsm | Rolls Spunlaced Nonwoven – 8” – 52 GSM | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Appearance | Free of hard wrinkles, formation uniform, large dot embossing pattern. (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | 3OA426 | Rolls Spunlaced Nonwoven – 8” – 52 gsm | Rolls Spunlaced Nonwoven – 8” – 52 GSM | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | 200 |  | FDA,0 |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | 3OA426 | Rolls Spunlaced Nonwoven – 8” – 52 gsm | Rolls Spunlaced Nonwoven – 8” – 52 GSM | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Basis Weight | 52.01 | GSM | FDA,0 | 47 | 57 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | 3OA426 | Rolls Spunlaced Nonwoven – 8” – 52 gsm | Rolls Spunlaced Nonwoven – 8” – 52 GSM | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Caliper | 0.79 | mm | FDA,2 | 0.74 | 0.83 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | 3OA426 | Rolls Spunlaced Nonwoven – 8” – 52 gsm | Rolls Spunlaced Nonwoven – 8” – 52 GSM | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | 3OA426 | Rolls Spunlaced Nonwoven – 8” – 52 gsm | Rolls Spunlaced Nonwoven – 8” – 52 GSM | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Odor | May have slight odor when wet (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | 3OA426 | Rolls Spunlaced Nonwoven – 8” – 52 gsm | Rolls Spunlaced Nonwoven – 8” – 52 GSM | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter | 45 | in | FDA,0 | 43 | 48 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | 3OA426 | Rolls Spunlaced Nonwoven – 8” – 52 gsm | Rolls Spunlaced Nonwoven – 8” – 52 GSM | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Width | 8 | in | FDA,3 | 7.875 | 8.125 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | 3OA426 | Rolls Spunlaced Nonwoven – 8” – 52 gsm | Rolls Spunlaced Nonwoven – 8” – 52 GSM | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Wet CD | 1337.35 | g/in | FDA,0 | 1200 | 2100 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | 3OA426 | Rolls Spunlaced Nonwoven – 8” – 52 gsm | Rolls Spunlaced Nonwoven – 8” – 52 GSM | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Wet MD | 5490.37 | g/in | FDA,0 | 3800 | 6270 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA580/Applicator_38A | 3OA580 | 45 gsm, Hydroentangled Nonwoven 11-1/2” | Hydroentangled Nonwoven 11-1/2” 45 GSM | Applicator_38A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA580/Applicator_38A | 3OA580 | 45 gsm, Hydroentangled Nonwoven 11-1/2” | Hydroentangled Nonwoven 11-1/2” 45 GSM | Applicator_38A | NY-Analytical | 759 | Applicator Measurements | Roll Diameter | 45 | in | FDA,0 | 43 | 47 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA580/Applicator_38A | 3OA580 | 45 gsm, Hydroentangled Nonwoven 11-1/2” | Hydroentangled Nonwoven 11-1/2” 45 GSM | Applicator_38A | NY-Analytical | 759 | Applicator Measurements | Roll Width | 11.5 | in | FDA,3 | 11.375 | 11.625 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA580/Applicator_38A | 3OA580 | 45 gsm, Hydroentangled Nonwoven 11-1/2” | Hydroentangled Nonwoven 11-1/2” 45 GSM | Applicator_38A | NY-Analytical | 756 | Basis Weight | Basis Weight | 44.35 | GSM | FDA,0 | 40 | 50 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA580/Applicator_38A | 3OA580 | 45 gsm, Hydroentangled Nonwoven 11-1/2” | Hydroentangled Nonwoven 11-1/2” 45 GSM | Applicator_38A | NY-Analytical | 757 | Caliper | Caliper | 500 | um | FDA,0 | 416 | 534 | TM 102 | TM102 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA580/Applicator_38A | 3OA580 | 45 gsm, Hydroentangled Nonwoven 11-1/2” | Hydroentangled Nonwoven 11-1/2” 45 GSM | Applicator_38A | NY-Analytical | 758 | Tensile Strength | Dry CD (CoA) | 1326 | g/in | FDA,0 | 610 |  | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA580/Applicator_38A | 3OA580 | 45 gsm, Hydroentangled Nonwoven 11-1/2” | Hydroentangled Nonwoven 11-1/2” 45 GSM | Applicator_38A | NY-Analytical | 758 | Tensile Strength | Dry MD (CoA) | 3086 | g/in | FDA,0 | 1825 |  | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA580/Applicator_38A | 3OA580 | 45 gsm, Hydroentangled Nonwoven 11-1/2” | Hydroentangled Nonwoven 11-1/2” 45 GSM | Applicator_38A | NY-Analytical | 748 | Visual Inspection | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA580/Applicator_38A | 3OA580 | 45 gsm, Hydroentangled Nonwoven 11-1/2” | Hydroentangled Nonwoven 11-1/2” 45 GSM | Applicator_38A | NY-Analytical | 748 | Visual Inspection | Color | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA580/Applicator_38A | 3OA580 | 45 gsm, Hydroentangled Nonwoven 11-1/2” | Hydroentangled Nonwoven 11-1/2” 45 GSM | Applicator_38A | NY-Analytical | 748 | Visual Inspection | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Nonwoven Rolls Dispersible Unembossed - 6" width - 70gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width - 70 GSM | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Nonwoven Rolls Dispersible Unembossed - 6" width - 70gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width - 70 GSM | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 759 | Applicator Measurements | Roll Diameter | 45 | in | FDA,0 | 45 | 48 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Nonwoven Rolls Dispersible Unembossed - 6" width - 70gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width - 70 GSM | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 759 | Applicator Measurements | Roll Width | 6 | in | FDA,3 | 6 | 6.125 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Nonwoven Rolls Dispersible Unembossed - 6" width - 70gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width - 70 GSM | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 756 | Basis Weight | Basis Weight | 69.3 | GSM | FDA,0 | 65 | 75 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Nonwoven Rolls Dispersible Unembossed - 6" width - 70gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width - 70 GSM | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 757 | Caliper | Caliper | 0.6 | mm | FDA,2 | 0.54 | 0.66 | TM 102 | TM102 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Nonwoven Rolls Dispersible Unembossed - 6" width - 70gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width - 70 GSM | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 758 | Tensile Strength | Wet CD | 387.69 | g/in | FDA,0 | 300 | 500 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Nonwoven Rolls Dispersible Unembossed - 6" width - 70gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width - 70 GSM | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 758 | Tensile Strength | Wet MD | 793.49 | g/in | FDA,0 | 610 | 1040 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Nonwoven Rolls Dispersible Unembossed - 6" width - 70gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width - 70 GSM | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 748 | Visual Inspection | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Nonwoven Rolls Dispersible Unembossed - 6" width - 70gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width - 70 GSM | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 748 | Visual Inspection | Color | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Nonwoven Rolls Dispersible Unembossed - 6" width - 70gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width - 70 GSM | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 748 | Visual Inspection | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Nonwoven Rolls Dispersible Unembossed - 6" width - 70gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width - 70 GSM | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | 2.76 | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Nonwoven Rolls Dispersible Unembossed - 6" width - 70gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width - 70 GSM | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Nonwoven Rolls Dispersible Unembossed - 6" width - 70gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width - 70 GSM | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Nonwoven Rolls Dispersible Unembossed - 6" width - 70gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width - 70 GSM | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Nonwoven Rolls Dispersible Unembossed - 6" width - 70gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width - 70 GSM | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Nonwoven Rolls Dispersible Unembossed - 6" width - 70gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width - 70 GSM | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Nonwoven Rolls Dispersible Unembossed - 6" width - 70gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width - 70 GSM | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Nonwoven Rolls Dispersible Unembossed - 6" width - 70gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width - 70 GSM | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | 3OA592 | Nonwoven Rolls Dispersible Unembossed - 6" width - 70gsm | Nonwoven Rolls Dispersible Unembossed - 6" Width - 70 GSM | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | 55gsm Big Dot Embossed Nonwoven | Big Dot Embossed Nonwoven 55 GSM | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | 55gsm Big Dot Embossed Nonwoven | Big Dot Embossed Nonwoven 55 GSM | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 759 | Applicator Measurements | Roll Width | 7.9 | in | FDA,3 | 7.9 | 8.025 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | 55gsm Big Dot Embossed Nonwoven | Big Dot Embossed Nonwoven 55 GSM | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 756 | Basis Weight | Basis Weight | 54.9 | GSM | FDA,0 | 50 | 60 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | 55gsm Big Dot Embossed Nonwoven | Big Dot Embossed Nonwoven 55 GSM | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 757 | Caliper | Caliper | 0.72 | mm | FDA,2 | 0.63 | 0.87 | TM 102 | TM102 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | 55gsm Big Dot Embossed Nonwoven | Big Dot Embossed Nonwoven 55 GSM | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 758 | Tensile Strength | Dry CD | 1382.58 | g/in | FDA,0 | 1036 |  | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | 55gsm Big Dot Embossed Nonwoven | Big Dot Embossed Nonwoven 55 GSM | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 758 | Tensile Strength | Dry CD | 6608.28 | g/in | FDA,0 | 4404 |  | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | 55gsm Big Dot Embossed Nonwoven | Big Dot Embossed Nonwoven 55 GSM | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 748 | Visual Inspection | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | 55gsm Big Dot Embossed Nonwoven | Big Dot Embossed Nonwoven 55 GSM | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 748 | Visual Inspection | Color | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | 55gsm Big Dot Embossed Nonwoven | Big Dot Embossed Nonwoven 55 GSM | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 748 | Visual Inspection | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | 55gsm Big Dot Embossed Nonwoven | Big Dot Embossed Nonwoven 55 GSM | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | 1.05 | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | 55gsm Big Dot Embossed Nonwoven | Big Dot Embossed Nonwoven 55 GSM | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | [5 empty] | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | 55gsm Big Dot Embossed Nonwoven | Big Dot Embossed Nonwoven 55 GSM | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | [5 empty] | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | 55gsm Big Dot Embossed Nonwoven | Big Dot Embossed Nonwoven 55 GSM | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | 55gsm Big Dot Embossed Nonwoven | Big Dot Embossed Nonwoven 55 GSM | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | 55gsm Big Dot Embossed Nonwoven | Big Dot Embossed Nonwoven 55 GSM | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | 55gsm Big Dot Embossed Nonwoven | Big Dot Embossed Nonwoven 55 GSM | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | 55gsm Big Dot Embossed Nonwoven | Big Dot Embossed Nonwoven 55 GSM | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | 3OA911 | 55gsm Big Dot Embossed Nonwoven | Big Dot Embossed Nonwoven 55 GSM | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | 3OA972 | 40 gsm, Hydroentangled Nonwoven 9" with Diamond Emboss | Hydroentangled Nonwoven 9" With Diamond Emboss 40 GSM | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | 3OA972 | 40 gsm, Hydroentangled Nonwoven 9" with Diamond Emboss | Hydroentangled Nonwoven 9" With Diamond Emboss 40 GSM | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 759 | Applicator Measurements | Roll Diameter | 44 | in | FDA,0 | 43 | 47 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | 3OA972 | 40 gsm, Hydroentangled Nonwoven 9" with Diamond Emboss | Hydroentangled Nonwoven 9" With Diamond Emboss 40 GSM | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 759 | Applicator Measurements | Roll Width | 9 | in | FDA,3 | 8.875 | 9.125 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | 3OA972 | 40 gsm, Hydroentangled Nonwoven 9" with Diamond Emboss | Hydroentangled Nonwoven 9" With Diamond Emboss 40 GSM | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 756 | Basis Weight | Basis Weight | 40.9 | GSM | FDA,1 | 35.2 | 44.8 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | 3OA972 | 40 gsm, Hydroentangled Nonwoven 9" with Diamond Emboss | Hydroentangled Nonwoven 9" With Diamond Emboss 40 GSM | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 757 | Caliper | Caliper | 437.8 | micron | FDA,0 | 340 | 540 | TM 102 | TM102 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | 3OA972 | 40 gsm, Hydroentangled Nonwoven 9" with Diamond Emboss | Hydroentangled Nonwoven 9" With Diamond Emboss 40 GSM | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 758 | Tensile Strength | Wet CD CoA | 1145 | g/in | FDA,0 | 400 | 1800 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | 3OA972 | 40 gsm, Hydroentangled Nonwoven 9" with Diamond Emboss | Hydroentangled Nonwoven 9" With Diamond Emboss 40 GSM | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 758 | Tensile Strength | Wet MD CoA | 2050 | g/in | FDA,0 | 1000 | 3400 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | 3OA972 | 40 gsm, Hydroentangled Nonwoven 9" with Diamond Emboss | Hydroentangled Nonwoven 9" With Diamond Emboss 40 GSM | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 748 | Visual Inspection | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | 3OA972 | 40 gsm, Hydroentangled Nonwoven 9" with Diamond Emboss | Hydroentangled Nonwoven 9" With Diamond Emboss 40 GSM | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 748 | Visual Inspection | Color | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | 3OA972 | 40 gsm, Hydroentangled Nonwoven 9" with Diamond Emboss | Hydroentangled Nonwoven 9" With Diamond Emboss 40 GSM | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 748 | Visual Inspection | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | Applicator Textured Thermal Bond - 55.8 gsm - 25 ½" | Textured Thermal Bond - 55.8 GSM - 25 ½" | REDUCED TESTING_NY_3OB108 | NY-Analytical | 884 | Applicator_Inspection | Color | White with no off colors or shading |  |  |  |  | Color (TM 208) | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | Applicator Textured Thermal Bond - 55.8 gsm - 25 ½" | Textured Thermal Bond - 55.8 GSM - 25 ½" | REDUCED TESTING_NY_3OB108 | NY-Analytical | 884 | Applicator_Inspection | Appearance | Meets specifications (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | Applicator Textured Thermal Bond - 55.8 gsm - 25 ½" | Textured Thermal Bond - 55.8 GSM - 25 ½" | REDUCED TESTING_NY_3OB108 | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | 125 |  | FDA,0 |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | Applicator Textured Thermal Bond - 55.8 gsm - 25 ½" | Textured Thermal Bond - 55.8 GSM - 25 ½" | REDUCED TESTING_NY_3OB108 | NY-Analytical | 884 | Applicator_Inspection | Basis Weight | 54.9 | GSM | FDA,1 | 52.3 | 59.6 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | Applicator Textured Thermal Bond - 55.8 gsm - 25 ½" | Textured Thermal Bond - 55.8 GSM - 25 ½" | REDUCED TESTING_NY_3OB108 | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | Applicator Textured Thermal Bond - 55.8 gsm - 25 ½" | Textured Thermal Bond - 55.8 GSM - 25 ½" | REDUCED TESTING_NY_3OB108 | NY-Analytical | 884 | Applicator_Inspection | Odor | Meets specifications (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | Applicator Textured Thermal Bond - 55.8 gsm - 25 ½" | Textured Thermal Bond - 55.8 GSM - 25 ½" | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | 1.25 | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | Applicator Textured Thermal Bond - 55.8 gsm - 25 ½" | Textured Thermal Bond - 55.8 GSM - 25 ½" | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | Applicator Textured Thermal Bond - 55.8 gsm - 25 ½" | Textured Thermal Bond - 55.8 GSM - 25 ½" | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | Applicator Textured Thermal Bond - 55.8 gsm - 25 ½" | Textured Thermal Bond - 55.8 GSM - 25 ½" | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | Applicator Textured Thermal Bond - 55.8 gsm - 25 ½" | Textured Thermal Bond - 55.8 GSM - 25 ½" | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | Applicator Textured Thermal Bond - 55.8 gsm - 25 ½" | Textured Thermal Bond - 55.8 GSM - 25 ½" | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | Applicator Textured Thermal Bond - 55.8 gsm - 25 ½" | Textured Thermal Bond - 55.8 GSM - 25 ½" | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | Applicator Textured Thermal Bond - 55.8 gsm - 25 ½" | Textured Thermal Bond - 55.8 GSM - 25 ½" | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OB108/REDUCED TESTING_NY_3OB108 | 3OB108 | Applicator Textured Thermal Bond - 55.8 gsm - 25 ½" | Textured Thermal Bond - 55.8 GSM - 25 ½" | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | 55 gsm unembossed spunlaced nonwoven- 6" | Unembossed Spunlaced Nonwoven- 6" 55 GSM | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Color | White with no off colors or shading |  |  |  |  | Color (TM 208) | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | 55 gsm unembossed spunlaced nonwoven- 6" | Unembossed Spunlaced Nonwoven- 6" 55 GSM | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Odor | May have residual sulfur odor when wet |  |  |  |  | Odor (TM 208) | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | 55 gsm unembossed spunlaced nonwoven- 6" | Unembossed Spunlaced Nonwoven- 6" 55 GSM | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Appearance | Free of hard wrinkles, and formation uniform, with micro crepe texture (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | 55 gsm unembossed spunlaced nonwoven- 6" | Unembossed Spunlaced Nonwoven- 6" 55 GSM | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | 80 |  | FDA,0 |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | 55 gsm unembossed spunlaced nonwoven- 6" | Unembossed Spunlaced Nonwoven- 6" 55 GSM | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Basis Weight | 54.04 | GSM | FDA,0 | 49.5 | 60.5 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | 55 gsm unembossed spunlaced nonwoven- 6" | Unembossed Spunlaced Nonwoven- 6" 55 GSM | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Caliper | 0.55 | mm | FDA,2 | 0.52 | 0.74 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | 55 gsm unembossed spunlaced nonwoven- 6" | Unembossed Spunlaced Nonwoven- 6" 55 GSM | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | 55 gsm unembossed spunlaced nonwoven- 6" | Unembossed Spunlaced Nonwoven- 6" 55 GSM | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter | 32 | in | FDA,0 | 28 | 36 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | 55 gsm unembossed spunlaced nonwoven- 6" | Unembossed Spunlaced Nonwoven- 6" 55 GSM | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Roll Width | 6 | in | FDA,0 | 5.875 | 6.125 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | 55 gsm unembossed spunlaced nonwoven- 6" | Unembossed Spunlaced Nonwoven- 6" 55 GSM | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Wet CD | 1004.84 | g/in | FDA,0 | 992 | 2208 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | 55 gsm unembossed spunlaced nonwoven- 6" | Unembossed Spunlaced Nonwoven- 6" 55 GSM | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Wet MD | 4435.12 | g/in | FDA,0 | 4392 | 7808 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | 55 gsm unembossed spunlaced nonwoven- 6" | Unembossed Spunlaced Nonwoven- 6" 55 GSM | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | 1.5 | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | 55 gsm unembossed spunlaced nonwoven- 6" | Unembossed Spunlaced Nonwoven- 6" 55 GSM | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | 55 gsm unembossed spunlaced nonwoven- 6" | Unembossed Spunlaced Nonwoven- 6" 55 GSM | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | 55 gsm unembossed spunlaced nonwoven- 6" | Unembossed Spunlaced Nonwoven- 6" 55 GSM | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | 55 gsm unembossed spunlaced nonwoven- 6" | Unembossed Spunlaced Nonwoven- 6" 55 GSM | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | 55 gsm unembossed spunlaced nonwoven- 6" | Unembossed Spunlaced Nonwoven- 6" 55 GSM | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | 55 gsm unembossed spunlaced nonwoven- 6" | Unembossed Spunlaced Nonwoven- 6" 55 GSM | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | 55 gsm unembossed spunlaced nonwoven- 6" | Unembossed Spunlaced Nonwoven- 6" 55 GSM | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | 3OC585 | 55 gsm unembossed spunlaced nonwoven- 6" | Unembossed Spunlaced Nonwoven- 6" 55 GSM | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Paper Crepe 23 lb., 7 3/4" | Paper Crepe 23 Lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Color | White with no off colors or shading |  |  |  |  | Color (TM 208) | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Paper Crepe 23 lb., 7 3/4" | Paper Crepe 23 Lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Odor | May have residual sulfur odor when wet |  |  |  |  | Odor (TM 208) | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Paper Crepe 23 lb., 7 3/4" | Paper Crepe 23 Lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Appearance | Meets specifications (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Paper Crepe 23 lb., 7 3/4" | Paper Crepe 23 Lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | 200 |  | FDA,0 |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Paper Crepe 23 lb., 7 3/4" | Paper Crepe 23 Lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Basis Weight | 23.81 | #/Ream | FDA,1 | 21.5 | 25 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Paper Crepe 23 lb., 7 3/4" | Paper Crepe 23 Lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Caliper | 4.9 | mil | FDA,1 | 4.2 | 5.6 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Paper Crepe 23 lb., 7 3/4" | Paper Crepe 23 Lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Paper Crepe 23 lb., 7 3/4" | Paper Crepe 23 Lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter | 32 | in | FDA,0 | 31 | 33 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Paper Crepe 23 lb., 7 3/4" | Paper Crepe 23 Lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Roll Width | 7.75 | in | FDA,4 | 7.6875 | 7.8125 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Paper Crepe 23 lb., 7 3/4" | Paper Crepe 23 Lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Wet CD | 1261.89 | g/in | FDA,0 | 1000 | 2500 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Paper Crepe 23 lb., 7 3/4" | Paper Crepe 23 Lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Wet MD | 2429.95 | g/in | FDA,0 | 1500 | 3850 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Paper Crepe 23 lb., 7 3/4" | Paper Crepe 23 Lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | 1.1 | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Paper Crepe 23 lb., 7 3/4" | Paper Crepe 23 Lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Paper Crepe 23 lb., 7 3/4" | Paper Crepe 23 Lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Paper Crepe 23 lb., 7 3/4" | Paper Crepe 23 Lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Paper Crepe 23 lb., 7 3/4" | Paper Crepe 23 Lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Paper Crepe 23 lb., 7 3/4" | Paper Crepe 23 Lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Paper Crepe 23 lb., 7 3/4" | Paper Crepe 23 Lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Paper Crepe 23 lb., 7 3/4" | Paper Crepe 23 Lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | 3OD186 | Paper Crepe 23 lb., 7 3/4" | Paper Crepe 23 Lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Color | White with no off colors or shading |  |  |  |  | Color (TM 208) | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | REDUCED TESTING_NY_3OD438 | NY-Analytical | 884 | Applicator_Inspection | Color | White with no off colors or shading |  |  |  |  | Color (TM 208) | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Odor | May have residual sulfur odor when wet |  |  |  |  | Odor (TM 208) | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | REDUCED TESTING_NY_3OD438 | NY-Analytical | 884 | Applicator_Inspection | Odor | May have residual sulfur odor when wet |  |  |  |  | Odor (TM 208) | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Appearance | Smooth free of hard wrinkles, formation uniform matches embossing pattern (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | REDUCED TESTING_NY_3OD438 | NY-Analytical | 884 | Applicator_Inspection | Appearance | Smooth free of hard wrinkles, formation uniform matches embossing pattern (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | 125 |  | FDA,0 |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | REDUCED TESTING_NY_3OD438 | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | 315 |  | FDA,0 |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Basis Weight | 34.66 | GSM | FDA,1 | 31.5 | 38.5 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | REDUCED TESTING_NY_3OD438 | NY-Analytical | 884 | Applicator_Inspection | Basis Weight | 35.09 | GSM | FDA,1 | 31.5 | 38.5 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Caliper | 0.0077 | in | FDA,4 | 0.006 | 0.0087 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | REDUCED TESTING_NY_3OD438 | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Dry CD | 1868.87 | g/in | FDA,0 | 900 | 2200 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Dry CD | 4655.69 | g/in | FDA,0 | 2000 | 5000 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter | 32 | in | FDA,0 | 30 | 34 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Roll Width | 7.875 | in | FDA,2 | 7.75 | 8 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Wet CD | 789.96 | g/in | FDA,0 | 250 | 1250 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Wet MD | 2141.89 | g/in | FDA,0 | 800 | 2600 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | 1.1 | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | 0.97 | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | 1 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | 0.18182 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OD438/REDUCED TESTING_NY_3OD438 | 3OD438 | 35 gsm Unembossed Nonwoven | Unembossed Nonwoven 35 GSM | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | Gram Positive Rods (Done) |  |  |  |  | WI QAS 076 | WI_QAS_076 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Embossed Spunlaced Nonwoven 45 gsm – 7” | Embossed Spunlaced Nonwoven 45 GSM – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Color | White with no off colors or shading |  |  |  |  | Color (TM 208) | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Embossed Spunlaced Nonwoven 45 gsm – 7” | Embossed Spunlaced Nonwoven 45 GSM – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Odor | May have residual sulfur odor when wet |  |  |  |  | Odor (TM 208) | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Embossed Spunlaced Nonwoven 45 gsm – 7” | Embossed Spunlaced Nonwoven 45 GSM – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Appearance | Free of hard wrinkles and formation uniform (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Embossed Spunlaced Nonwoven 45 gsm – 7” | Embossed Spunlaced Nonwoven 45 GSM – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | 315 |  | FDA,0 |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Embossed Spunlaced Nonwoven 45 gsm – 7” | Embossed Spunlaced Nonwoven 45 GSM – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Basis Weight | 46.72 | GSM | FDA,1 | 40.5 | 49.5 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Embossed Spunlaced Nonwoven 45 gsm – 7” | Embossed Spunlaced Nonwoven 45 GSM – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Caliper | 0.64 | mm | FDA,2 | 0.55 | 0.67 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Embossed Spunlaced Nonwoven 45 gsm – 7” | Embossed Spunlaced Nonwoven 45 GSM – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Embossed Spunlaced Nonwoven 45 gsm – 7” | Embossed Spunlaced Nonwoven 45 GSM – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter | 32 | in | FDA,0 | 31 | 33 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Embossed Spunlaced Nonwoven 45 gsm – 7” | Embossed Spunlaced Nonwoven 45 GSM – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Roll Width | 7 | in | FDA,0 | 6.875 | 7.125 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Embossed Spunlaced Nonwoven 45 gsm – 7” | Embossed Spunlaced Nonwoven 45 GSM – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Wet CD | 2598 | g/in | FDA,0 | 1131 | 1593 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Embossed Spunlaced Nonwoven 45 gsm – 7” | Embossed Spunlaced Nonwoven 45 GSM – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Wet MD | 3711 | g/in | FDA,0 | 3906 | 4934 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Embossed Spunlaced Nonwoven 45 gsm – 7” | Embossed Spunlaced Nonwoven 45 GSM – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | 1.75 | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Embossed Spunlaced Nonwoven 45 gsm – 7” | Embossed Spunlaced Nonwoven 45 GSM – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Embossed Spunlaced Nonwoven 45 gsm – 7” | Embossed Spunlaced Nonwoven 45 GSM – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Embossed Spunlaced Nonwoven 45 gsm – 7” | Embossed Spunlaced Nonwoven 45 GSM – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Embossed Spunlaced Nonwoven 45 gsm – 7” | Embossed Spunlaced Nonwoven 45 GSM – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Embossed Spunlaced Nonwoven 45 gsm – 7” | Embossed Spunlaced Nonwoven 45 GSM – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Embossed Spunlaced Nonwoven 45 gsm – 7” | Embossed Spunlaced Nonwoven 45 GSM – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Embossed Spunlaced Nonwoven 45 gsm – 7” | Embossed Spunlaced Nonwoven 45 GSM – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | 3OH143 | Embossed Spunlaced Nonwoven 45 gsm – 7” | Embossed Spunlaced Nonwoven 45 GSM – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Spunlaced Nonwoven 50 GSM 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Color | White with no off colors or shading |  |  |  |  | Color (TM 208) | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Spunlaced Nonwoven 50 GSM 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Appearance | Smooth free of hard wrinkles, formation uniform matches embossing pattern (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Spunlaced Nonwoven 50 GSM 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | 200 |  | FDA,0 |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Spunlaced Nonwoven 50 GSM 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Basis Weight | 52.74 | GSM | FDA,1 | 45 | 55 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Spunlaced Nonwoven 50 GSM 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Caliper | 0.55 | mm | FDA,2 | 0.51 | 0.63 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Spunlaced Nonwoven 50 GSM 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Spunlaced Nonwoven 50 GSM 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Dry CD | 2389.39 | g/in | FDA,0 | 1801 | 2821 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Spunlaced Nonwoven 50 GSM 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Dry CD | 8511.74 | g/in | FDA,0 | 6675 | 9447 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Spunlaced Nonwoven 50 GSM 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Odor | No objectionable odor detectable wet or dry (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Spunlaced Nonwoven 50 GSM 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter | 30 | in | FDA,0 | 30 | 32 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Spunlaced Nonwoven 50 GSM 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Roll Width | 7 | in | FDA,3 | 6.875 | 7.125 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Spunlaced Nonwoven 50 GSM 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | 1.07 | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Spunlaced Nonwoven 50 GSM 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Spunlaced Nonwoven 50 GSM 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Spunlaced Nonwoven 50 GSM 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Spunlaced Nonwoven 50 GSM 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Spunlaced Nonwoven 50 GSM 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Spunlaced Nonwoven 50 GSM 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Spunlaced Nonwoven 50 GSM 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | 3OH581 | Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Spunlaced Nonwoven 50 GSM 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH592/Applicator_36A | 3OH592 | 40 gsm, Hydroentangled Nonwoven 7 7/8” with Diamond Emboss | Hydroentangled Nonwoven 7 7/8” With Diamond Emboss 40 GSM | Applicator_36A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH592/Applicator_36A | 3OH592 | 40 gsm, Hydroentangled Nonwoven 7 7/8” with Diamond Emboss | Hydroentangled Nonwoven 7 7/8” With Diamond Emboss 40 GSM | Applicator_36A | NY-Analytical | 759 | Applicator Measurements | Roll Diameter | 32 | in | FDA,0 | 30 | 34 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH592/Applicator_36A | 3OH592 | 40 gsm, Hydroentangled Nonwoven 7 7/8” with Diamond Emboss | Hydroentangled Nonwoven 7 7/8” With Diamond Emboss 40 GSM | Applicator_36A | NY-Analytical | 759 | Applicator Measurements | Roll Width | 7.875 | in | FDA,3 | 7.75 | 8 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH592/Applicator_36A | 3OH592 | 40 gsm, Hydroentangled Nonwoven 7 7/8” with Diamond Emboss | Hydroentangled Nonwoven 7 7/8” With Diamond Emboss 40 GSM | Applicator_36A | NY-Analytical | 756 | Basis Weight | Basis Weight | 39.181 | GSM | FDA,0 | 35 | 45 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH592/Applicator_36A | 3OH592 | 40 gsm, Hydroentangled Nonwoven 7 7/8” with Diamond Emboss | Hydroentangled Nonwoven 7 7/8” With Diamond Emboss 40 GSM | Applicator_36A | NY-Analytical | 757 | Caliper | Caliper | 424 | um | FDA,0 | 340 | 540 | TM 102 | TM102 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH592/Applicator_36A | 3OH592 | 40 gsm, Hydroentangled Nonwoven 7 7/8” with Diamond Emboss | Hydroentangled Nonwoven 7 7/8” With Diamond Emboss 40 GSM | Applicator_36A | NY-Analytical | 758 | Tensile Strength | Wet CD CoA | 1235 | g/in | FDA,0 | 400 |  | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH592/Applicator_36A | 3OH592 | 40 gsm, Hydroentangled Nonwoven 7 7/8” with Diamond Emboss | Hydroentangled Nonwoven 7 7/8” With Diamond Emboss 40 GSM | Applicator_36A | NY-Analytical | 758 | Tensile Strength | Wet MD CoA | 2903 | g/in | FDA,0 | 1000 |  | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH592/Applicator_36A | 3OH592 | 40 gsm, Hydroentangled Nonwoven 7 7/8” with Diamond Emboss | Hydroentangled Nonwoven 7 7/8” With Diamond Emboss 40 GSM | Applicator_36A | NY-Analytical | 748 | Visual Inspection | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH592/Applicator_36A | 3OH592 | 40 gsm, Hydroentangled Nonwoven 7 7/8” with Diamond Emboss | Hydroentangled Nonwoven 7 7/8” With Diamond Emboss 40 GSM | Applicator_36A | NY-Analytical | 748 | Visual Inspection | Color | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OH592/Applicator_36A | 3OH592 | 40 gsm, Hydroentangled Nonwoven 7 7/8” with Diamond Emboss | Hydroentangled Nonwoven 7 7/8” With Diamond Emboss 40 GSM | Applicator_36A | NY-Analytical | 748 | Visual Inspection | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OJ213/Applicator_08A | 3OJ213 | AP1 VJ 40 GSM 50% VISC/50% PET 12" | Ap1 Vj 40 GSM 50% Visc/50% Pet 12" | Applicator_08A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | [5 empty] | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OJ213/Applicator_08A | 3OJ213 | AP1 VJ 40 GSM 50% VISC/50% PET 12" | Ap1 Vj 40 GSM 50% Visc/50% Pet 12" | Applicator_08A | NY-Analytical | 759 | Applicator Measurements | Roll Width |  | in | FDA,3 | 11.875 | 12.125 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OJ213/Applicator_08A | 3OJ213 | AP1 VJ 40 GSM 50% VISC/50% PET 12" | Ap1 Vj 40 GSM 50% Visc/50% Pet 12" | Applicator_08A | NY-Analytical | 756 | Basis Weight | Basis Weight | 41.76 | GSM | FDA,0 | 36 | 44 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OJ213/Applicator_08A | 3OJ213 | AP1 VJ 40 GSM 50% VISC/50% PET 12" | Ap1 Vj 40 GSM 50% Visc/50% Pet 12" | Applicator_08A | NY-Analytical | 757 | Caliper | Caliper | 0.54 | mm | FDA,2 | 0.48 | 0.61 | TM 102 | TM102 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OJ213/Applicator_08A | 3OJ213 | AP1 VJ 40 GSM 50% VISC/50% PET 12" | Ap1 Vj 40 GSM 50% Visc/50% Pet 12" | Applicator_08A | NY-Analytical | 758 | Tensile Strength | Wet CD | 1096.49 | g/in | FDA,0 | 1030 | 1546 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OJ213/Applicator_08A | 3OJ213 | AP1 VJ 40 GSM 50% VISC/50% PET 12" | Ap1 Vj 40 GSM 50% Visc/50% Pet 12" | Applicator_08A | NY-Analytical | 758 | Tensile Strength | Wet MD | 4699.94 | g/in | FDA,0 | 3343 | 4350 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OJ213/Applicator_08A | 3OJ213 | AP1 VJ 40 GSM 50% VISC/50% PET 12" | Ap1 Vj 40 GSM 50% Visc/50% Pet 12" | Applicator_08A | NY-Analytical | 748 | Visual Inspection | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OJ213/Applicator_08A | 3OJ213 | AP1 VJ 40 GSM 50% VISC/50% PET 12" | Ap1 Vj 40 GSM 50% Visc/50% Pet 12" | Applicator_08A | NY-Analytical | 748 | Visual Inspection | Color | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OJ213/Applicator_08A | 3OJ213 | AP1 VJ 40 GSM 50% VISC/50% PET 12" | Ap1 Vj 40 GSM 50% Visc/50% Pet 12" | Applicator_08A | NY-Analytical | 748 | Visual Inspection | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OJ213/Applicator_08A | 3OJ213 | AP1 VJ 40 GSM 50% VISC/50% PET 12" | Ap1 Vj 40 GSM 50% Visc/50% Pet 12" | Applicator_08A | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight |  | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OJ213/Applicator_08A | 3OJ213 | AP1 VJ 40 GSM 50% VISC/50% PET 12" | Ap1 Vj 40 GSM 50% Visc/50% Pet 12" | Applicator_08A | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | [5 empty] | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OJ213/Applicator_08A | 3OJ213 | AP1 VJ 40 GSM 50% VISC/50% PET 12" | Ap1 Vj 40 GSM 50% Visc/50% Pet 12" | Applicator_08A | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | [5 empty] | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OJ213/Applicator_08A | 3OJ213 | AP1 VJ 40 GSM 50% VISC/50% PET 12" | Ap1 Vj 40 GSM 50% Visc/50% Pet 12" | Applicator_08A | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OJ213/Applicator_08A | 3OJ213 | AP1 VJ 40 GSM 50% VISC/50% PET 12" | Ap1 Vj 40 GSM 50% Visc/50% Pet 12" | Applicator_08A | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OJ213/Applicator_08A | 3OJ213 | AP1 VJ 40 GSM 50% VISC/50% PET 12" | Ap1 Vj 40 GSM 50% Visc/50% Pet 12" | Applicator_08A | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) |  | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OJ213/Applicator_08A | 3OJ213 | AP1 VJ 40 GSM 50% VISC/50% PET 12" | Ap1 Vj 40 GSM 50% Visc/50% Pet 12" | Applicator_08A | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OJ213/Applicator_08A | 3OJ213 | AP1 VJ 40 GSM 50% VISC/50% PET 12" | Ap1 Vj 40 GSM 50% Visc/50% Pet 12" | Applicator_08A | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OJ213/Applicator_08A | 3OJ213 | AP1 VJ 40 GSM 50% VISC/50% PET 12" | Ap1 Vj 40 GSM 50% Visc/50% Pet 12" | Applicator_08A | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) |  | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Spunlaced Nonwoven 50 gsm 100% Polyester - 9" | Spunlaced Nonwoven 50 GSM 100% Polyester - 9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Appearance | Meets specifications (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Spunlaced Nonwoven 50 gsm 100% Polyester - 9" | Spunlaced Nonwoven 50 GSM 100% Polyester - 9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | 315 |  | FDA,0 |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Spunlaced Nonwoven 50 gsm 100% Polyester - 9" | Spunlaced Nonwoven 50 GSM 100% Polyester - 9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Basis Weight | 47.36 | GSM | FDA,0 | 45 | 55 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Spunlaced Nonwoven 50 gsm 100% Polyester - 9" | Spunlaced Nonwoven 50 GSM 100% Polyester - 9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Caliper | 0.46 | mm | FDA,2 | 0.51 | 0.63 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Spunlaced Nonwoven 50 gsm 100% Polyester - 9" | Spunlaced Nonwoven 50 GSM 100% Polyester - 9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Spunlaced Nonwoven 50 gsm 100% Polyester - 9" | Spunlaced Nonwoven 50 GSM 100% Polyester - 9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Color | Meets specifications (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Spunlaced Nonwoven 50 gsm 100% Polyester - 9" | Spunlaced Nonwoven 50 GSM 100% Polyester - 9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Dry CD | 2142.67 | g/in | FDA,0 | 1801 | 2821 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Spunlaced Nonwoven 50 gsm 100% Polyester - 9" | Spunlaced Nonwoven 50 GSM 100% Polyester - 9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Dry CD | 7892 | g/in | FDA,0 | 6675 | 9447 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Spunlaced Nonwoven 50 gsm 100% Polyester - 9" | Spunlaced Nonwoven 50 GSM 100% Polyester - 9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Odor | Meets specifications (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Spunlaced Nonwoven 50 gsm 100% Polyester - 9" | Spunlaced Nonwoven 50 GSM 100% Polyester - 9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (COA) | 43 |  | FDA,0 | 41 | 45 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Spunlaced Nonwoven 50 gsm 100% Polyester - 9" | Spunlaced Nonwoven 50 GSM 100% Polyester - 9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Roll width (COA) | 9 | in | FDA,2 | 8.75 | 9.25 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Spunlaced Nonwoven 50 gsm 100% Polyester - 9" | Spunlaced Nonwoven 50 GSM 100% Polyester - 9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | 0.89 | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Spunlaced Nonwoven 50 gsm 100% Polyester - 9" | Spunlaced Nonwoven 50 GSM 100% Polyester - 9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Spunlaced Nonwoven 50 gsm 100% Polyester - 9" | Spunlaced Nonwoven 50 GSM 100% Polyester - 9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Spunlaced Nonwoven 50 gsm 100% Polyester - 9" | Spunlaced Nonwoven 50 GSM 100% Polyester - 9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Spunlaced Nonwoven 50 gsm 100% Polyester - 9" | Spunlaced Nonwoven 50 GSM 100% Polyester - 9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Spunlaced Nonwoven 50 gsm 100% Polyester - 9" | Spunlaced Nonwoven 50 GSM 100% Polyester - 9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Spunlaced Nonwoven 50 gsm 100% Polyester - 9" | Spunlaced Nonwoven 50 GSM 100% Polyester - 9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Spunlaced Nonwoven 50 gsm 100% Polyester - 9" | Spunlaced Nonwoven 50 GSM 100% Polyester - 9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | 3OM304 | Spunlaced Nonwoven 50 gsm 100% Polyester - 9" | Spunlaced Nonwoven 50 GSM 100% Polyester - 9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | 3OP065 | Rolls Spunlaced Nonwoven Diamond - 60gsm | Rolls Spunlaced Nonwoven Diamond - 60 GSM | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Color | White with no off colors or shading |  |  |  |  | Color (TM 208) | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | 3OP065 | Rolls Spunlaced Nonwoven Diamond - 60gsm | Rolls Spunlaced Nonwoven Diamond - 60 GSM | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Appearance | Free of hard wrinkles. Formation uniform (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | 3OP065 | Rolls Spunlaced Nonwoven Diamond - 60gsm | Rolls Spunlaced Nonwoven Diamond - 60 GSM | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | 125 |  | FDA,0 |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | 3OP065 | Rolls Spunlaced Nonwoven Diamond - 60gsm | Rolls Spunlaced Nonwoven Diamond - 60 GSM | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Basis Weight | 57.26 | GSM | FDA,1 | 54 | 66 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | 3OP065 | Rolls Spunlaced Nonwoven Diamond - 60gsm | Rolls Spunlaced Nonwoven Diamond - 60 GSM | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Caliper | 0.8 | mm | FDA,2 | 0.67 | 0.81 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | 3OP065 | Rolls Spunlaced Nonwoven Diamond - 60gsm | Rolls Spunlaced Nonwoven Diamond - 60 GSM | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | 3OP065 | Rolls Spunlaced Nonwoven Diamond - 60gsm | Rolls Spunlaced Nonwoven Diamond - 60 GSM | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Odor | Meets specifications (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | 3OP065 | Rolls Spunlaced Nonwoven Diamond - 60gsm | Rolls Spunlaced Nonwoven Diamond - 60 GSM | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter | 46 | in | FDA,0 | 45 | 47 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | 3OP065 | Rolls Spunlaced Nonwoven Diamond - 60gsm | Rolls Spunlaced Nonwoven Diamond - 60 GSM | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Width | 30 | in | FDA,2 | 29.75 | 30.25 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | 3OP065 | Rolls Spunlaced Nonwoven Diamond - 60gsm | Rolls Spunlaced Nonwoven Diamond - 60 GSM | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Wet CD | 1584 | g/in | FDA,0 | 1300 | 2000 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | 3OP065 | Rolls Spunlaced Nonwoven Diamond - 60gsm | Rolls Spunlaced Nonwoven Diamond - 60 GSM | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Wet MD | 5195.31 | g/in | FDA,0 | 4000 | 5500 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Rolls Flat Spunlaced Nonwoven - 55 gsm | Rolls Flat Spunlaced Nonwoven - 55 GSM | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Rolls Flat Spunlaced Nonwoven - 55 gsm | Rolls Flat Spunlaced Nonwoven - 55 GSM | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 759 | Applicator Measurements | Roll Diameter |  | in | FDA,0 | 45 | 47 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Rolls Flat Spunlaced Nonwoven - 55 gsm | Rolls Flat Spunlaced Nonwoven - 55 GSM | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 759 | Applicator Measurements | Roll Width |  | in | FDA,2 | 29.75 | 30.25 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Rolls Flat Spunlaced Nonwoven - 55 gsm | Rolls Flat Spunlaced Nonwoven - 55 GSM | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 756 | Basis Weight | Basis Weight | 57.7 | GSM | FDA,1 | 49.5 | 60.5 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Rolls Flat Spunlaced Nonwoven - 55 gsm | Rolls Flat Spunlaced Nonwoven - 55 GSM | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 757 | Caliper | Caliper | 0.51 | mm | FDA,2 | 0.4 | 0.6 | TM 102 | TM102 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Rolls Flat Spunlaced Nonwoven - 55 gsm | Rolls Flat Spunlaced Nonwoven - 55 GSM | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 758 | Tensile Strength | Elongation % Wet CD | 103.07 | % | FDA,0 | 80 | 130 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Rolls Flat Spunlaced Nonwoven - 55 gsm | Rolls Flat Spunlaced Nonwoven - 55 GSM | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 758 | Tensile Strength | Elongation % Wet MD | 24.68 | % | FDA,0 | 20 | 35 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Rolls Flat Spunlaced Nonwoven - 55 gsm | Rolls Flat Spunlaced Nonwoven - 55 GSM | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 758 | Tensile Strength | Wet CD | 893.28 | g/in | FDA,0 | 650 | 1300 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Rolls Flat Spunlaced Nonwoven - 55 gsm | Rolls Flat Spunlaced Nonwoven - 55 GSM | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 758 | Tensile Strength | Wet MD | 2898.19 | g/in | FDA,0 | 2400 | 3400 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Rolls Flat Spunlaced Nonwoven - 55 gsm | Rolls Flat Spunlaced Nonwoven - 55 GSM | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 748 | Visual Inspection | Appearance | [5 empty] | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Rolls Flat Spunlaced Nonwoven - 55 gsm | Rolls Flat Spunlaced Nonwoven - 55 GSM | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 748 | Visual Inspection | Color | [5 empty] | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Rolls Flat Spunlaced Nonwoven - 55 gsm | Rolls Flat Spunlaced Nonwoven - 55 GSM | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 748 | Visual Inspection | Odor | [5 empty] | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Rolls Flat Spunlaced Nonwoven - 55 gsm | Rolls Flat Spunlaced Nonwoven - 55 GSM | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | 0.84 | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Rolls Flat Spunlaced Nonwoven - 55 gsm | Rolls Flat Spunlaced Nonwoven - 55 GSM | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Rolls Flat Spunlaced Nonwoven - 55 gsm | Rolls Flat Spunlaced Nonwoven - 55 GSM | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Rolls Flat Spunlaced Nonwoven - 55 gsm | Rolls Flat Spunlaced Nonwoven - 55 GSM | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Rolls Flat Spunlaced Nonwoven - 55 gsm | Rolls Flat Spunlaced Nonwoven - 55 GSM | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Rolls Flat Spunlaced Nonwoven - 55 gsm | Rolls Flat Spunlaced Nonwoven - 55 GSM | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Rolls Flat Spunlaced Nonwoven - 55 gsm | Rolls Flat Spunlaced Nonwoven - 55 GSM | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Rolls Flat Spunlaced Nonwoven - 55 gsm | Rolls Flat Spunlaced Nonwoven - 55 GSM | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | 3OP081 | Rolls Flat Spunlaced Nonwoven - 55 gsm | Rolls Flat Spunlaced Nonwoven - 55 GSM | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Die Cut Pads 2.25" 90 Count | Die Cut Pads 2.25" 90 Count | Applicator_NY_Die Cut Pads | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Die Cut Pads 2.25" 90 Count | Die Cut Pads 2.25" 90 Count | Applicator_NY_Die Cut Pads | NY-Analytical | 759 | Applicator Measurements | Pad Diameter | 2.2188 | in | FDA,2 | 2.19 | 2.31 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Die Cut Pads 2.25" 90 Count | Die Cut Pads 2.25" 90 Count | Applicator_NY_Die Cut Pads | NY-Analytical | 759 | Applicator Measurements | Stack Count | 90 | Units | FDA,0 | 90 | 90 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Die Cut Pads 2.25" 90 Count | Die Cut Pads 2.25" 90 Count | Applicator_NY_Die Cut Pads | NY-Analytical | 748 | Visual Inspection | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Die Cut Pads 2.25" 90 Count | Die Cut Pads 2.25" 90 Count | Applicator_NY_Die Cut Pads | NY-Analytical | 748 | Visual Inspection | Color | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Die Cut Pads 2.25" 90 Count | Die Cut Pads 2.25" 90 Count | Applicator_NY_Die Cut Pads | NY-Analytical | 748 | Visual Inspection | Odor | No objectionable odor Wet or Dry [Done] |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Die Cut Pads 2.25" 90 Count | Die Cut Pads 2.25" 90 Count | Applicator_NY_Die Cut Pads | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | 2.07 | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Die Cut Pads 2.25" 90 Count | Die Cut Pads 2.25" 90 Count | Applicator_NY_Die Cut Pads | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Die Cut Pads 2.25" 90 Count | Die Cut Pads 2.25" 90 Count | Applicator_NY_Die Cut Pads | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Die Cut Pads 2.25" 90 Count | Die Cut Pads 2.25" 90 Count | Applicator_NY_Die Cut Pads | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | 1 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Die Cut Pads 2.25" 90 Count | Die Cut Pads 2.25" 90 Count | Applicator_NY_Die Cut Pads | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | 0.042735 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Die Cut Pads 2.25" 90 Count | Die Cut Pads 2.25" 90 Count | Applicator_NY_Die Cut Pads | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Die Cut Pads 2.25" 90 Count | Die Cut Pads 2.25" 90 Count | Applicator_NY_Die Cut Pads | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Die Cut Pads 2.25" 90 Count | Die Cut Pads 2.25" 90 Count | Applicator_NY_Die Cut Pads | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Die Cut Pads 2.25" 90 Count | Die Cut Pads 2.25" 90 Count | Applicator_NY_Die Cut Pads | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Die Cut Pads 2.25" 90 Count | Die Cut Pads 2.25" 90 Count | Applicator_NY_Die Cut Pads | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | Gram Positive Cocci (OOS) |  |  |  |  | WI QAS 076 | WI_QAS_076 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP093PD/Applicator_NY_Die Cut Pads | 3OP093PD | Die Cut Pads 2.25" 90 Count | Die Cut Pads 2.25" 90 Count | Applicator_NY_Die Cut Pads | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Non-Subculturable Growth (Done) |  |  |  |  | WI QAS 076 | WI_QAS_076 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | 35gsm Stora Flat - 30" | Stora Flat - 30" 35 GSM | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | 35gsm Stora Flat - 30" | Stora Flat - 30" 35 GSM | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 759 | Applicator Measurements | Roll Diameter | 40 | in | FDA,0 | 39 | 40 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | 35gsm Stora Flat - 30" | Stora Flat - 30" 35 GSM | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 759 | Applicator Measurements | Roll Width | 30 | in | FDA,3 | 29.875 | 30.125 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | 35gsm Stora Flat - 30" | Stora Flat - 30" 35 GSM | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 756 | Basis Weight | Basis Weight | 34.01 | GSM | FDA,1 | 31.5 | 38.5 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | 35gsm Stora Flat - 30" | Stora Flat - 30" 35 GSM | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 757 | Caliper | Caliper | 0.008 | in | FDA,4 | 0.006 | 0.009 | TM 102 | TM102 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | 35gsm Stora Flat - 30" | Stora Flat - 30" 35 GSM | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 757 | Caliper | Caliper | 0.0079 | in |  |  |  | TM 102 | TM102 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | 35gsm Stora Flat - 30" | Stora Flat - 30" 35 GSM | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 758 | Tensile Strength | Wet CD | 812.99 | g/in | FDA,0 | 250 | 1250 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | 35gsm Stora Flat - 30" | Stora Flat - 30" 35 GSM | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 758 | Tensile Strength | Wet MD | 1821.41 | g/in | FDA,0 | 800 | 2600 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | 35gsm Stora Flat - 30" | Stora Flat - 30" 35 GSM | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 748 | Visual Inspection | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | 35gsm Stora Flat - 30" | Stora Flat - 30" 35 GSM | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 748 | Visual Inspection | Color | White with no off colors or shading [Done] |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | 35gsm Stora Flat - 30" | Stora Flat - 30" 35 GSM | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 748 | Visual Inspection | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | 35gsm Stora Flat - 30" | Stora Flat - 30" 35 GSM | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | 0.97 | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | 35gsm Stora Flat - 30" | Stora Flat - 30" 35 GSM | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | 35gsm Stora Flat - 30" | Stora Flat - 30" 35 GSM | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | 35gsm Stora Flat - 30" | Stora Flat - 30" 35 GSM | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | 35gsm Stora Flat - 30" | Stora Flat - 30" 35 GSM | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | 35gsm Stora Flat - 30" | Stora Flat - 30" 35 GSM | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | 35gsm Stora Flat - 30" | Stora Flat - 30" 35 GSM | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | 35gsm Stora Flat - 30" | Stora Flat - 30" 35 GSM | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | 35gsm Stora Flat - 30" | Stora Flat - 30" 35 GSM | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP134/Applicator_NY_35gsm Stora Flat - 30" | 3OP134 | 35gsm Stora Flat - 30" | Stora Flat - 30" 35 GSM | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | [5 empty] | WI QAS 076 | WI_QAS_076 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP138/Applicator_36D | 3OP138 | 40 gsm, Hydroentangled Nonwoven 30" with Diamond Emboss | Hydroentangled Nonwoven 30" With Diamond Emboss 40 GSM | Applicator_36D | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | [5 empty] | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP138/Applicator_36D | 3OP138 | 40 gsm, Hydroentangled Nonwoven 30" with Diamond Emboss | Hydroentangled Nonwoven 30" With Diamond Emboss 40 GSM | Applicator_36D | NY-Analytical | 759 | Applicator Measurements | Roll Diameter |  | in | FDA,0 | 43 | 47 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP138/Applicator_36D | 3OP138 | 40 gsm, Hydroentangled Nonwoven 30" with Diamond Emboss | Hydroentangled Nonwoven 30" With Diamond Emboss 40 GSM | Applicator_36D | NY-Analytical | 759 | Applicator Measurements | Roll Width | 30 | in | FDA,3 | 29.875 | 30.125 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP138/Applicator_36D | 3OP138 | 40 gsm, Hydroentangled Nonwoven 30" with Diamond Emboss | Hydroentangled Nonwoven 30" With Diamond Emboss 40 GSM | Applicator_36D | NY-Analytical | 756 | Basis Weight | Basis Weight | 40.26 | GSM | FDA,0 | 35 | 45 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP138/Applicator_36D | 3OP138 | 40 gsm, Hydroentangled Nonwoven 30" with Diamond Emboss | Hydroentangled Nonwoven 30" With Diamond Emboss 40 GSM | Applicator_36D | NY-Analytical | 757 | Caliper | Caliper |  | um | FDA,0 | 340 | 540 | TM 102 | TM102 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP138/Applicator_36D | 3OP138 | 40 gsm, Hydroentangled Nonwoven 30" with Diamond Emboss | Hydroentangled Nonwoven 30" With Diamond Emboss 40 GSM | Applicator_36D | NY-Analytical | 758 | Tensile Strength | Wet CD CoA | 1228 | g/in | FDA,0 | 400 |  | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP138/Applicator_36D | 3OP138 | 40 gsm, Hydroentangled Nonwoven 30" with Diamond Emboss | Hydroentangled Nonwoven 30" With Diamond Emboss 40 GSM | Applicator_36D | NY-Analytical | 758 | Tensile Strength | Wet MD CoA | 2545 | g/in | FDA,0 | 1000 |  | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP138/Applicator_36D | 3OP138 | 40 gsm, Hydroentangled Nonwoven 30" with Diamond Emboss | Hydroentangled Nonwoven 30" With Diamond Emboss 40 GSM | Applicator_36D | NY-Analytical | 748 | Visual Inspection | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP138/Applicator_36D | 3OP138 | 40 gsm, Hydroentangled Nonwoven 30" with Diamond Emboss | Hydroentangled Nonwoven 30" With Diamond Emboss 40 GSM | Applicator_36D | NY-Analytical | 748 | Visual Inspection | Color | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP138/Applicator_36D | 3OP138 | 40 gsm, Hydroentangled Nonwoven 30" with Diamond Emboss | Hydroentangled Nonwoven 30" With Diamond Emboss 40 GSM | Applicator_36D | NY-Analytical | 748 | Visual Inspection | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP197A/Applicator_NY_Skip lot testing | 3OP197A | Spunlaced Embossed Nonwoven - 35" | Spunlaced Embossed Nonwoven - 35" | Applicator_NY_Skip lot testing | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Bioburden_TMYC |  | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP197A-SU/Label_NY_NP7200_6.25-4.5 | 3OP197A-SU | Spunlaced Embossed Nonwoven - 900mm | Spunlaced Embossed Nonwoven - 900Mm | Label_NY_NP7200_6.25-4.5 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP197A-SU/Label_NY_NP7200_6.25-4.5 | 3OP197A-SU | Spunlaced Embossed Nonwoven - 900mm | Spunlaced Embossed Nonwoven - 900Mm | Label_NY_NP7200_6.25-4.5 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP197A-SU/Label_NY_NP7200_6.25-4.5 | 3OP197A-SU | Spunlaced Embossed Nonwoven - 900mm | Spunlaced Embossed Nonwoven - 900Mm | Label_NY_NP7200_6.25-4.5 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP197A-SU/Label_NY_NP7200_6.25-4.5 | 3OP197A-SU | Spunlaced Embossed Nonwoven - 900mm | Spunlaced Embossed Nonwoven - 900Mm | Label_NY_NP7200_6.25-4.5 | NY-Analytical | 770 | Label Inspection | Length | 6.25 | in | FDA,4 | 6.1875 | 6.3125 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP197A-SU/Label_NY_NP7200_6.25-4.5 | 3OP197A-SU | Spunlaced Embossed Nonwoven - 900mm | Spunlaced Embossed Nonwoven - 900Mm | Label_NY_NP7200_6.25-4.5 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP197A-SU/Label_NY_NP7200_6.25-4.5 | 3OP197A-SU | Spunlaced Embossed Nonwoven - 900mm | Spunlaced Embossed Nonwoven - 900Mm | Label_NY_NP7200_6.25-4.5 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP197A-SU/Label_NY_NP7200_6.25-4.5 | 3OP197A-SU | Spunlaced Embossed Nonwoven - 900mm | Spunlaced Embossed Nonwoven - 900Mm | Label_NY_NP7200_6.25-4.5 | NY-Analytical | 770 | Label Inspection | Width | 4.5 | in | FDA,4 | 4.4375 | 4.5625 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
| 3OP242 | 35gsm Flat Thermoembonded Nonwoven 37.5" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 759 | Applicator Measurements | Roll Width | 37.5 | in | FDA,3 | 37.375 | 37.625 | Applicator Measurements | TM142 |  |  |  | left_only
| 3OP242 | 35gsm Flat Thermoembonded Nonwoven 37.5" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Basis Weight |  |  |  |  | left_only
| 3OP242 | 35gsm Flat Thermoembonded Nonwoven 37.5" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 756 | Basis Weight | Basis Weight | 35.95 | GSM | FDA,1 | 31.5 | 38.5 | TM 101 | TM101 |  |  |  | left_only
| 3OP242 | 35gsm Flat Thermoembonded Nonwoven 37.5" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 757 | Caliper | Caliper | 0.34 | mm | FDA,2 | 0.27 | 0.41 | TM 102 | TM102 |  |  |  | left_only
| 3OP242 | 35gsm Flat Thermoembonded Nonwoven 37.5" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 758 | Tensile Strength | Wet CD CoA | 520.88 | g/in | FDA,0 | 300 |  | TM 103 | TM103 |  |  |  | left_only
| 3OP242 | 35gsm Flat Thermoembonded Nonwoven 37.5" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 758 | Tensile Strength | Wet MD CoA | 3556.57 | g/in | FDA,0 | 2000 | 4400 | TM 103 | TM103 |  |  |  | left_only
| 3OP242 | 35gsm Flat Thermoembonded Nonwoven 37.5" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 748 | Visual Inspection | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 |  |  |  | left_only
| 3OP242 | 35gsm Flat Thermoembonded Nonwoven 37.5" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 748 | Visual Inspection | Color | White with no off colors or shading [Done] |  |  |  |  | TM 208 | TM208 |  |  |  | left_only
| 3OP242 | 35gsm Flat Thermoembonded Nonwoven 37.5" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 748 | Visual Inspection | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 |  |  |  | left_only
| 3OP242 | 35gsm Flat Thermoembonded Nonwoven 37.5" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | 0.95 | g | FDA,2 |  |  | TM 610 | TM610 |  |  |  | left_only
| 3OP242 | 35gsm Flat Thermoembonded Nonwoven 37.5" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | TM610 |  |  |  | left_only
| 3OP242 | 35gsm Flat Thermoembonded Nonwoven 37.5" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610 | TM610 |  |  |  | left_only
| 3OP242 | 35gsm Flat Thermoembonded Nonwoven 37.5" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | 1.5 | cfu/device | FDA,1 |  |  | TM 610 | TM610 |  |  |  | left_only
| 3OP242 | 35gsm Flat Thermoembonded Nonwoven 37.5" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | 1.5789 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 |  |  |  | left_only
| 3OP242 | 35gsm Flat Thermoembonded Nonwoven 37.5" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 1 | cfu | FDA,0 |  |  | TM 610 | TM610 |  |  |  | left_only
| 3OP242 | 35gsm Flat Thermoembonded Nonwoven 37.5" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 |  |  |  | left_only
| 3OP242 | 35gsm Flat Thermoembonded Nonwoven 37.5" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 |  |  |  | left_only
| 3OP242 | 35gsm Flat Thermoembonded Nonwoven 37.5" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 |  |  |  | left_only
| 3OP242 | 35gsm Flat Thermoembonded Nonwoven 37.5" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | Non-Subculturable Growth (Done) |  |  |  |  | WI QAS 076 | WI_QAS_076 |  |  |  | left_only
3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | 3OP257 | Spunlaced Nonwoven 50gsm 100% Polyester - 37.5" | Spunlaced Nonwoven 50 GSM 100% Polyester - 37.5" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (COA) | 43 | in | FDA,0 | 41 | 45 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | 3OP257 | Spunlaced Nonwoven 50gsm 100% Polyester - 37.5" | Spunlaced Nonwoven 50 GSM 100% Polyester - 37.5" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 759 | Applicator Measurements | Roll width (COA) | 37.5 | in | FDA,2 | 37.25 | 37.75 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | 3OP257 | Spunlaced Nonwoven 50gsm 100% Polyester - 37.5" | Spunlaced Nonwoven 50 GSM 100% Polyester - 37.5" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Basis Weight |  | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | 3OP257 | Spunlaced Nonwoven 50gsm 100% Polyester - 37.5" | Spunlaced Nonwoven 50 GSM 100% Polyester - 37.5" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 756 | Basis Weight | Basis Weight | 51.02 | GSM | FDA,0 | 45 | 55 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | 3OP257 | Spunlaced Nonwoven 50gsm 100% Polyester - 37.5" | Spunlaced Nonwoven 50 GSM 100% Polyester - 37.5" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 757 | Caliper | Caliper | 0.52 | mm | FDA,2 | 0.51 | 0.63 | TM 102 | TM102 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | 3OP257 | Spunlaced Nonwoven 50gsm 100% Polyester - 37.5" | Spunlaced Nonwoven 50 GSM 100% Polyester - 37.5" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 758 | Tensile Strength | Dry CD | 2193.6 | g/in | FDA,0 | 1801 | 2821 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | 3OP257 | Spunlaced Nonwoven 50gsm 100% Polyester - 37.5" | Spunlaced Nonwoven 50 GSM 100% Polyester - 37.5" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 758 | Tensile Strength | Dry CD | 7519.45 | g/in | FDA,0 | 6675 | 9447 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | 3OP257 | Spunlaced Nonwoven 50gsm 100% Polyester - 37.5" | Spunlaced Nonwoven 50 GSM 100% Polyester - 37.5" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 748 | Visual Inspection | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | 3OP257 | Spunlaced Nonwoven 50gsm 100% Polyester - 37.5" | Spunlaced Nonwoven 50 GSM 100% Polyester - 37.5" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 748 | Visual Inspection | Color | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | 3OP257 | Spunlaced Nonwoven 50gsm 100% Polyester - 37.5" | Spunlaced Nonwoven 50 GSM 100% Polyester - 37.5" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 748 | Visual Inspection | Odor | No objectionable odor Wet or Dry [Done] |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | SMS Nonwovens Honeycomb Embossed- 35 gsm - 953 mm | Sms Nonwovens Honeycomb Embossed- 35 GSM - 953 Mm | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Color | White with no off colors or shading |  |  |  |  | Color (TM 208) | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | SMS Nonwovens Honeycomb Embossed- 35 gsm - 953 mm | Sms Nonwovens Honeycomb Embossed- 35 GSM - 953 Mm | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 759 | Applicator Measurements | Roll Width | 953 | mm | FDA,0 | 950 | 956 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | SMS Nonwovens Honeycomb Embossed- 35 gsm - 953 mm | Sms Nonwovens Honeycomb Embossed- 35 GSM - 953 Mm | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Appearance | Free of hard wrinkles, formation uniform/quilted big dot emboss pattern (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | SMS Nonwovens Honeycomb Embossed- 35 gsm - 953 mm | Sms Nonwovens Honeycomb Embossed- 35 GSM - 953 Mm | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | 315 |  | FDA,0 |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | SMS Nonwovens Honeycomb Embossed- 35 gsm - 953 mm | Sms Nonwovens Honeycomb Embossed- 35 GSM - 953 Mm | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Basis Weight | 32.08 | GSM | FDA,1 | 31.5 | 38.5 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | SMS Nonwovens Honeycomb Embossed- 35 gsm - 953 mm | Sms Nonwovens Honeycomb Embossed- 35 GSM - 953 Mm | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Caliper | 0.34 | mm | FDA,2 | 0.26 | 0.4 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | SMS Nonwovens Honeycomb Embossed- 35 gsm - 953 mm | Sms Nonwovens Honeycomb Embossed- 35 GSM - 953 Mm | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | SMS Nonwovens Honeycomb Embossed- 35 gsm - 953 mm | Sms Nonwovens Honeycomb Embossed- 35 GSM - 953 Mm | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Dry CD | 1254.21 | g/in | FDA,0 | 1100 | 1700 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | SMS Nonwovens Honeycomb Embossed- 35 gsm - 953 mm | Sms Nonwovens Honeycomb Embossed- 35 GSM - 953 Mm | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Dry CD | 2070.52 | g/in | FDA,0 | 2000 | 3000 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | SMS Nonwovens Honeycomb Embossed- 35 gsm - 953 mm | Sms Nonwovens Honeycomb Embossed- 35 GSM - 953 Mm | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Odor | May have slight odor when wet (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | SMS Nonwovens Honeycomb Embossed- 35 gsm - 953 mm | Sms Nonwovens Honeycomb Embossed- 35 GSM - 953 Mm | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter | 40 | in | FDA,0 | 40 |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | SMS Nonwovens Honeycomb Embossed- 35 gsm - 953 mm | Sms Nonwovens Honeycomb Embossed- 35 GSM - 953 Mm | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Roll Width | 37.5 | in | FDA,2 | 37.25 | 37.75 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | SMS Nonwovens Honeycomb Embossed- 35 gsm - 953 mm | Sms Nonwovens Honeycomb Embossed- 35 GSM - 953 Mm | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 756 | Basis Weight | Basis Weight | 34.44 | GSM | FDA,1 | 31.5 | 38.5 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | SMS Nonwovens Honeycomb Embossed- 35 gsm - 953 mm | Sms Nonwovens Honeycomb Embossed- 35 GSM - 953 Mm | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 757 | Caliper | Caliper | 0.29 | mm | FDA,2 | 0.26 | 0.4 | TM 102 | TM102 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | SMS Nonwovens Honeycomb Embossed- 35 gsm - 953 mm | Sms Nonwovens Honeycomb Embossed- 35 GSM - 953 Mm | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 758 | Tensile Strength | Dry CD | 1143.94 | g/in | FDA,0 | 1100 | 1700 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | SMS Nonwovens Honeycomb Embossed- 35 gsm - 953 mm | Sms Nonwovens Honeycomb Embossed- 35 GSM - 953 Mm | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 758 | Tensile Strength | Dry CD | 2220.05 | g/in | FDA,0 | 2000 | 3000 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | SMS Nonwovens Honeycomb Embossed- 35 gsm - 953 mm | Sms Nonwovens Honeycomb Embossed- 35 GSM - 953 Mm | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 748 | Visual Inspection | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | SMS Nonwovens Honeycomb Embossed- 35 gsm - 953 mm | Sms Nonwovens Honeycomb Embossed- 35 GSM - 953 Mm | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 748 | Visual Inspection | Color | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | SMS Nonwovens Honeycomb Embossed- 35 gsm - 953 mm | Sms Nonwovens Honeycomb Embossed- 35 GSM - 953 Mm | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 748 | Visual Inspection | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | SMS Nonwovens Honeycomb Embossed- 35 gsm - 953 mm | Sms Nonwovens Honeycomb Embossed- 35 GSM - 953 Mm | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | 0.95 | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | SMS Nonwovens Honeycomb Embossed- 35 gsm - 953 mm | Sms Nonwovens Honeycomb Embossed- 35 GSM - 953 Mm | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | SMS Nonwovens Honeycomb Embossed- 35 gsm - 953 mm | Sms Nonwovens Honeycomb Embossed- 35 GSM - 953 Mm | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | SMS Nonwovens Honeycomb Embossed- 35 gsm - 953 mm | Sms Nonwovens Honeycomb Embossed- 35 GSM - 953 Mm | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | SMS Nonwovens Honeycomb Embossed- 35 gsm - 953 mm | Sms Nonwovens Honeycomb Embossed- 35 GSM - 953 Mm | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | SMS Nonwovens Honeycomb Embossed- 35 gsm - 953 mm | Sms Nonwovens Honeycomb Embossed- 35 GSM - 953 Mm | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | SMS Nonwovens Honeycomb Embossed- 35 gsm - 953 mm | Sms Nonwovens Honeycomb Embossed- 35 GSM - 953 Mm | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 1 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | SMS Nonwovens Honeycomb Embossed- 35 gsm - 953 mm | Sms Nonwovens Honeycomb Embossed- 35 GSM - 953 Mm | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | 0.21053 | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | 3OP920 | SMS Nonwovens Honeycomb Embossed- 35 gsm - 953 mm | Sms Nonwovens Honeycomb Embossed- 35 GSM - 953 Mm | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | [5 empty] | Certificate of Analysis | COA | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis | COA | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | [5 empty] | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 759 | Applicator Measurements | Roll Diameter |  | in | FDA,0 | 42 | 46 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 759 | Applicator Measurements | Roll Diameter | 44 | in | FDA,0 | 42 | 46 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 759 | Applicator Measurements | Roll Diameter |  | in | FDA,0 | 42 | 46 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 759 | Applicator Measurements | Roll Width |  | in | FDA,3 | 37.375 | 37.625 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 759 | Applicator Measurements | Roll Width | 36 | in | FDA,3 | 35.875 | 36.125 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 759 | Applicator Measurements | Roll Width | 30 | in | FDA,3 | 29.875 | 30.125 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 756 | Basis Weight | Basis Weight | 41.12 | GSM | FDA,1 | 35.5 | 44.5 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 756 | Basis Weight | Basis Weight | 39.61 | GSM | FDA,1 | 35.5 | 44.5 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 756 | Basis Weight | Basis Weight |  | GSM | FDA,1 | 35.5 | 44.5 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 757 | Caliper | Caliper | 0.46 | mm | FDA,2 | 0.4 | 0.48 | TM 102 | TM102 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 757 | Caliper | Caliper |  | mm | FDA,2 | 0.4 | 0.48 | TM 102 | TM102 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 757 | Caliper | Caliper | 0.45 | mm | FDA,2 | 0.4 | 0.48 | TM 102 | TM102 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 758 | Tensile Strength | Wet CD | 740.59 | g/in | FDA,0 | 600 | 1200 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 758 | Tensile Strength | Wet CD | 521.87 | g/in | FDA,0 | 600 | 1200 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 758 | Tensile Strength | Wet CD |  | g/in | FDA,0 | 600 | 1200 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 758 | Tensile Strength | Wet MD |  | g/in | FDA,0 | 2200 | 3400 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 758 | Tensile Strength | Wet MD | 1610.35 | g/in | FDA,0 | 2200 | 3400 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 758 | Tensile Strength | Wet MD |  | g/in | FDA,0 | 2200 | 3400 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 748 | Visual Inspection | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 748 | Visual Inspection | Appearance | [5 empty] | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 748 | Visual Inspection | Appearance | [5 empty] | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 748 | Visual Inspection | Color | White with no off colors or shading [Done] |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 748 | Visual Inspection | Color | [5 empty] | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 748 | Visual Inspection | Color | [5 empty] | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Analytical | 748 | Visual Inspection | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 748 | Visual Inspection | Odor | [5 empty] | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 748 | Visual Inspection | Odor | [5 empty] | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight |  | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | [5 empty] | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | [5 empty] | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) |  | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | 3OQ222 | Applicator - 30" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 30" | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) |  | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ333 | Embossed Spunlaced Nonwoven - 37 1/2" | Embossed Spunlaced Nonwoven - 37 1/2" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | 1.32 | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ333 | Embossed Spunlaced Nonwoven - 37 1/2" | Embossed Spunlaced Nonwoven - 37 1/2" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ333 | Embossed Spunlaced Nonwoven - 37 1/2" | Embossed Spunlaced Nonwoven - 37 1/2" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ333 | Embossed Spunlaced Nonwoven - 37 1/2" | Embossed Spunlaced Nonwoven - 37 1/2" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ333 | Embossed Spunlaced Nonwoven - 37 1/2" | Embossed Spunlaced Nonwoven - 37 1/2" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ333 | Embossed Spunlaced Nonwoven - 37 1/2" | Embossed Spunlaced Nonwoven - 37 1/2" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ333 | Embossed Spunlaced Nonwoven - 37 1/2" | Embossed Spunlaced Nonwoven - 37 1/2" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ333 | Embossed Spunlaced Nonwoven - 37 1/2" | Embossed Spunlaced Nonwoven - 37 1/2" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | 3OQ333 | Embossed Spunlaced Nonwoven - 37 1/2" | Embossed Spunlaced Nonwoven - 37 1/2" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ444 | Applicator- 36" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 36" | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight |  | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ444 | Applicator- 36" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 36" | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | [5 empty] | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ444 | Applicator- 36" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 36" | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | [5 empty] | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ444 | Applicator- 36" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 36" | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ444 | Applicator- 36" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 36" | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ444 | Applicator- 36" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 36" | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) |  | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ444 | Applicator- 36" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 36" | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ444 | Applicator- 36" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 36" | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | 3OQ444 | Applicator- 36" Embossed Spunlaced Nonwoven | Embossed Spunlaced Nonwoven - 36" | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) |  | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ551/Applicator_15E | 3OQ551 | 40gsm, Hydroentangled Nonwoven, 42" | Hydroentangled Nonwoven, 42" 40 GSM | Applicator_15E | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (COA) | 45 | in | FDA,0 | 43 | 47 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ551/Applicator_15E | 3OQ551 | 40gsm, Hydroentangled Nonwoven, 42" | Hydroentangled Nonwoven, 42" 40 GSM | Applicator_15E | NY-Analytical | 759 | Applicator Measurements | Roll Width | 42 | in | FDA,3 | 41.875 | 42.125 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ551/Applicator_15E | 3OQ551 | 40gsm, Hydroentangled Nonwoven, 42" | Hydroentangled Nonwoven, 42" 40 GSM | Applicator_15E | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Basis Weight |  | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ551/Applicator_15E | 3OQ551 | 40gsm, Hydroentangled Nonwoven, 42" | Hydroentangled Nonwoven, 42" 40 GSM | Applicator_15E | NY-Analytical | 756 | Basis Weight | Basis Weight |  | GSM | FDA,0 | 35 | 45 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ551/Applicator_15E | 3OQ551 | 40gsm, Hydroentangled Nonwoven, 42" | Hydroentangled Nonwoven, 42" 40 GSM | Applicator_15E | NY-Analytical | 757 | Caliper | Caliper | 431.8 | um | FDA,0 | 340 | 540 | TM 102 | TM102 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ551/Applicator_15E | 3OQ551 | 40gsm, Hydroentangled Nonwoven, 42" | Hydroentangled Nonwoven, 42" 40 GSM | Applicator_15E | NY-Analytical | 758 | Tensile Strength | Wet CD CoA | 1073 | g/25mm | FDA,0 | 400 |  | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ551/Applicator_15E | 3OQ551 | 40gsm, Hydroentangled Nonwoven, 42" | Hydroentangled Nonwoven, 42" 40 GSM | Applicator_15E | NY-Analytical | 758 | Tensile Strength | Wet MD CoA | 2506 | g/25mm | FDA,0 | 1000 |  | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ551/Applicator_15E | 3OQ551 | 40gsm, Hydroentangled Nonwoven, 42" | Hydroentangled Nonwoven, 42" 40 GSM | Applicator_15E | NY-Analytical | 748 | Visual Inspection | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ551/Applicator_15E | 3OQ551 | 40gsm, Hydroentangled Nonwoven, 42" | Hydroentangled Nonwoven, 42" 40 GSM | Applicator_15E | NY-Analytical | 748 | Visual Inspection | Color | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ551/Applicator_15E | 3OQ551 | 40gsm, Hydroentangled Nonwoven, 42" | Hydroentangled Nonwoven, 42" 40 GSM | Applicator_15E | NY-Analytical | 748 | Visual Inspection | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | 3OQ850N | Applicator Thermal Bonded Nonwoven-37 1/2" | Thermal Bonded Nonwoven-37 1/2" | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Color | White with no off colors or shading |  |  |  |  | Color (TM 208) | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ850N/Applicator_01E | 3OQ850N | Applicator Thermal Bonded Nonwoven-37 1/2" | Thermal Bonded Nonwoven-37 1/2" | Applicator_01E | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | [5 empty] | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ850N/Applicator_01E | 3OQ850N | Applicator Thermal Bonded Nonwoven-37 1/2" | Thermal Bonded Nonwoven-37 1/2" | Applicator_01E | NY-Analytical | 759 | Applicator Measurements | Roll Diameter |  | in | FDA,0 | 44 | 47 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ850N/Applicator_01E | 3OQ850N | Applicator Thermal Bonded Nonwoven-37 1/2" | Thermal Bonded Nonwoven-37 1/2" | Applicator_01E | NY-Analytical | 759 | Applicator Measurements | Roll Width |  | in | FDA,3 | 35.875 | 36.125 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | 3OQ850N | Applicator Thermal Bonded Nonwoven-37 1/2" | Thermal Bonded Nonwoven-37 1/2" | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Appearance | Smooth, free of hard wrinkles and formation uniform (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | 3OQ850N | Applicator Thermal Bonded Nonwoven-37 1/2" | Thermal Bonded Nonwoven-37 1/2" | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | 315 |  | FDA,0 |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | 3OQ850N | Applicator Thermal Bonded Nonwoven-37 1/2" | Thermal Bonded Nonwoven-37 1/2" | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Basis Weight | 40.69 | GSM | FDA,1 | 38 | 42 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | 3OQ850N | Applicator Thermal Bonded Nonwoven-37 1/2" | Thermal Bonded Nonwoven-37 1/2" | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Caliper | 0.41 | mm |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | 3OQ850N | Applicator Thermal Bonded Nonwoven-37 1/2" | Thermal Bonded Nonwoven-37 1/2" | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | 3OQ850N | Applicator Thermal Bonded Nonwoven-37 1/2" | Thermal Bonded Nonwoven-37 1/2" | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Dry CD | 1098.41 | g/in | FDA,0 | 513 | 1130 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | 3OQ850N | Applicator Thermal Bonded Nonwoven-37 1/2" | Thermal Bonded Nonwoven-37 1/2" | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Dry CD | 2976.77 | g/in | FDA,0 | 2054 | 4108 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | 3OQ850N | Applicator Thermal Bonded Nonwoven-37 1/2" | Thermal Bonded Nonwoven-37 1/2" | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Odor | No objectionable odor detectable wet or dry (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | 3OQ850N | Applicator Thermal Bonded Nonwoven-37 1/2" | Thermal Bonded Nonwoven-37 1/2" | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter | 46 | in | FDA,0 | 44 | 47 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | 3OQ850N | Applicator Thermal Bonded Nonwoven-37 1/2" | Thermal Bonded Nonwoven-37 1/2" | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Roll Width | 37.5 | in | FDA,3 | 37.375 | 37.625 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ850N/Applicator_01E | 3OQ850N | Applicator Thermal Bonded Nonwoven-37 1/2" | Thermal Bonded Nonwoven-37 1/2" | Applicator_01E | NY-Analytical | 756 | Basis Weight | Basis Weight |  | GSM | FDA,0 | 38 | 42 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ850N/Applicator_01E | 3OQ850N | Applicator Thermal Bonded Nonwoven-37 1/2" | Thermal Bonded Nonwoven-37 1/2" | Applicator_01E | NY-Analytical | 757 | Caliper | Caliper |  | mm | FDA,2 | 0.35 | 0.45 | TM 102 | TM102 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ850N/Applicator_01E | 3OQ850N | Applicator Thermal Bonded Nonwoven-37 1/2" | Thermal Bonded Nonwoven-37 1/2" | Applicator_01E | NY-Analytical | 758 | Tensile Strength | Dry CD |  | g/in | FDA,0 | 513 | 1130 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ850N/Applicator_01E | 3OQ850N | Applicator Thermal Bonded Nonwoven-37 1/2" | Thermal Bonded Nonwoven-37 1/2" | Applicator_01E | NY-Analytical | 758 | Tensile Strength | Dry CD |  | g/in | FDA,0 | 2054 | 4108 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ850N/Applicator_01E | 3OQ850N | Applicator Thermal Bonded Nonwoven-37 1/2" | Thermal Bonded Nonwoven-37 1/2" | Applicator_01E | NY-Analytical | 748 | Visual Inspection | Appearance | [5 empty] | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ850N/Applicator_01E | 3OQ850N | Applicator Thermal Bonded Nonwoven-37 1/2" | Thermal Bonded Nonwoven-37 1/2" | Applicator_01E | NY-Analytical | 748 | Visual Inspection | Color | [5 empty] | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OQ850N/Applicator_01E | 3OQ850N | Applicator Thermal Bonded Nonwoven-37 1/2" | Thermal Bonded Nonwoven-37 1/2" | Applicator_01E | NY-Analytical | 748 | Visual Inspection | Odor | [5 empty] | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OS407/Applicator_25B | 3OS407 | SWB SUPER BRUSH FOAM TIP SWABS | Swb Super Brush Foam Tip Swabs | Applicator_25B | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Swab Measurements | TM109 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS407/Applicator_25B | 3OS407 | SWB SUPER BRUSH FOAM TIP SWABS | Swb Super Brush Foam Tip Swabs | Applicator_25B | NY-Analytical | 760 | Swab Measurements | Mitt Length | 0.99 | in | FDA,3 | 0.97 | 1.03 | TM 109 | TM109 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS407/Applicator_25B | 3OS407 | SWB SUPER BRUSH FOAM TIP SWABS | Swb Super Brush Foam Tip Swabs | Applicator_25B | NY-Analytical | 760 | Swab Measurements | Mitt Thickness | 0.314 | in | FDA,3 | 0.265 | 0.385 | TM 109 | TM109 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS407/Applicator_25B | 3OS407 | SWB SUPER BRUSH FOAM TIP SWABS | Swb Super Brush Foam Tip Swabs | Applicator_25B | NY-Analytical | 760 | Swab Measurements | Mitt Width | 0.49 | in | FDA,2 | 0.47 | 0.53 | TM 109 | TM109 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS407/Applicator_25B | 3OS407 | SWB SUPER BRUSH FOAM TIP SWABS | Swb Super Brush Foam Tip Swabs | Applicator_25B | NY-Analytical | 760 | Swab Measurements | Overall Swab Length | 4.77 | in | FDA,3 | 4.69 | 4.81 | TM 109 | TM109 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS407/Applicator_25B | 3OS407 | SWB SUPER BRUSH FOAM TIP SWABS | Swb Super Brush Foam Tip Swabs | Applicator_25B | NY-Analytical | 760 | Swab Measurements | Scrub Test | Meets specifications (Done) |  |  |  |  | TM 109 | TM109 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS407/Applicator_25B | 3OS407 | SWB SUPER BRUSH FOAM TIP SWABS | Swb Super Brush Foam Tip Swabs | Applicator_25B | NY-Analytical | 760 | Swab Measurements | Tuft Adhesion | 2.59 | lb | FDA,1 | 1.4 |  | TM 109 | TM109 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS407/Applicator_25B | 3OS407 | SWB SUPER BRUSH FOAM TIP SWABS | Swb Super Brush Foam Tip Swabs | Applicator_25B | NY-Analytical | 760 | Swab Measurements | Use Test-Solvents | Meets specifications (Done) |  |  |  |  | TM 109 | TM109 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS407/Applicator_25B | 3OS407 | SWB SUPER BRUSH FOAM TIP SWABS | Swb Super Brush Foam Tip Swabs | Applicator_25B | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | 1.2 | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS407/Applicator_25B | 3OS407 | SWB SUPER BRUSH FOAM TIP SWABS | Swb Super Brush Foam Tip Swabs | Applicator_25B | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS407/Applicator_25B | 3OS407 | SWB SUPER BRUSH FOAM TIP SWABS | Swb Super Brush Foam Tip Swabs | Applicator_25B | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS407/Applicator_25B | 3OS407 | SWB SUPER BRUSH FOAM TIP SWABS | Swb Super Brush Foam Tip Swabs | Applicator_25B | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS407/Applicator_25B | 3OS407 | SWB SUPER BRUSH FOAM TIP SWABS | Swb Super Brush Foam Tip Swabs | Applicator_25B | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS407/Applicator_25B | 3OS407 | SWB SUPER BRUSH FOAM TIP SWABS | Swb Super Brush Foam Tip Swabs | Applicator_25B | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) |  | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS407/Applicator_25B | 3OS407 | SWB SUPER BRUSH FOAM TIP SWABS | Swb Super Brush Foam Tip Swabs | Applicator_25B | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS407/Applicator_25B | 3OS407 | SWB SUPER BRUSH FOAM TIP SWABS | Swb Super Brush Foam Tip Swabs | Applicator_25B | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS407/Applicator_25B | 3OS407 | SWB SUPER BRUSH FOAM TIP SWABS | Swb Super Brush Foam Tip Swabs | Applicator_25B | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) |  | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS407/Applicator_25B | 3OS407 | SWB SUPER BRUSH FOAM TIP SWABS | Swb Super Brush Foam Tip Swabs | Applicator_25B | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | Gram Positive Rods (Done) |  |  |  |  | WI QAS 076 | WI_QAS_076 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS407/Applicator_25B | 3OS407 | SWB SUPER BRUSH FOAM TIP SWABS | Swb Super Brush Foam Tip Swabs | Applicator_25B | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Non-Subculturable Growth (Done) |  |  |  |  | WI QAS 076 | WI_QAS_076 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS419/Applicator_25C | 3OS419 | SWB FOAM TIP SWABSTICK FLAT6" | Swb Foam Tip Swabstick Flat6" | Applicator_25C | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Swab Measurements | TM109 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS419/Applicator_25C | 3OS419 | SWB FOAM TIP SWABSTICK FLAT6" | Swb Foam Tip Swabstick Flat6" | Applicator_25C | NY-Analytical | 760 | Swab Measurements | Mitt Length | 1.22 | in | FDA,2 | 1.16 | 1.36 | TM 109 | TM109 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS419/Applicator_25C | 3OS419 | SWB FOAM TIP SWABSTICK FLAT6" | Swb Foam Tip Swabstick Flat6" | Applicator_25C | NY-Analytical | 760 | Swab Measurements | Mitt Thickness | 0.47 | in | FDA,3 | 0.41 | 0.53 | TM 109 | TM109 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS419/Applicator_25C | 3OS419 | SWB FOAM TIP SWABSTICK FLAT6" | Swb Foam Tip Swabstick Flat6" | Applicator_25C | NY-Analytical | 760 | Swab Measurements | Mitt Width | 1.19 | in | FDA,2 | 1.1 | 1.3 | TM 109 | TM109 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS419/Applicator_25C | 3OS419 | SWB FOAM TIP SWABSTICK FLAT6" | Swb Foam Tip Swabstick Flat6" | Applicator_25C | NY-Analytical | 760 | Swab Measurements | Overall Swab Length | 6.03 | in | FDA,2 | 5.85 | 6.15 | TM 109 | TM109 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS419/Applicator_25C | 3OS419 | SWB FOAM TIP SWABSTICK FLAT6" | Swb Foam Tip Swabstick Flat6" | Applicator_25C | NY-Analytical | 760 | Swab Measurements | Scrub Test | Meets specifications (Done) |  |  |  |  | TM 109 | TM109 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS419/Applicator_25C | 3OS419 | SWB FOAM TIP SWABSTICK FLAT6" | Swb Foam Tip Swabstick Flat6" | Applicator_25C | NY-Analytical | 760 | Swab Measurements | Tuft Adhesion | 5.12 | lb | FDA,1 | 1.4 |  | TM 109 | TM109 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS419/Applicator_25C | 3OS419 | SWB FOAM TIP SWABSTICK FLAT6" | Swb Foam Tip Swabstick Flat6" | Applicator_25C | NY-Analytical | 760 | Swab Measurements | Use Test-Solvents | Meets specifications (Done) |  |  |  |  | TM 109 | TM109 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS419/Applicator_25C | 3OS419 | SWB FOAM TIP SWABSTICK FLAT6" | Swb Foam Tip Swabstick Flat6" | Applicator_25C | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | 1.26 | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS419/Applicator_25C | 3OS419 | SWB FOAM TIP SWABSTICK FLAT6" | Swb Foam Tip Swabstick Flat6" | Applicator_25C | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | [5 empty] | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS419/Applicator_25C | 3OS419 | SWB FOAM TIP SWABSTICK FLAT6" | Swb Foam Tip Swabstick Flat6" | Applicator_25C | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | [5 empty] | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS419/Applicator_25C | 3OS419 | SWB FOAM TIP SWABSTICK FLAT6" | Swb Foam Tip Swabstick Flat6" | Applicator_25C | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) |  | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS419/Applicator_25C | 3OS419 | SWB FOAM TIP SWABSTICK FLAT6" | Swb Foam Tip Swabstick Flat6" | Applicator_25C | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) |  | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS419/Applicator_25C | 3OS419 | SWB FOAM TIP SWABSTICK FLAT6" | Swb Foam Tip Swabstick Flat6" | Applicator_25C | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) |  | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS419/Applicator_25C | 3OS419 | SWB FOAM TIP SWABSTICK FLAT6" | Swb Foam Tip Swabstick Flat6" | Applicator_25C | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS419/Applicator_25C | 3OS419 | SWB FOAM TIP SWABSTICK FLAT6" | Swb Foam Tip Swabstick Flat6" | Applicator_25C | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS419/Applicator_25C | 3OS419 | SWB FOAM TIP SWABSTICK FLAT6" | Swb Foam Tip Swabstick Flat6" | Applicator_25C | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS419/Applicator_25C | 3OS419 | SWB FOAM TIP SWABSTICK FLAT6" | Swb Foam Tip Swabstick Flat6" | Applicator_25C | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | Gram Positive Rods (Done) |  |  |  |  | WI QAS 076 | WI_QAS_076 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | Swabstick | Swabstick | REDUCED TESTING_NY_3OS428 | NY-Analytical | 760 | Swab Measurements | AQL Sample Size | 315 | in |  |  |  | TM 109 | TM109 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | Swabstick | Swabstick | REDUCED TESTING_NY_3OS428 | NY-Analytical | 760 | Swab Measurements | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | TM 109 | TM109 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | Swabstick | Swabstick | REDUCED TESTING_NY_3OS428 | NY-Analytical | 760 | Swab Measurements | Scrub Test | Pass (Done) |  |  |  |  | TM 109 | TM109 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | Swabstick | Swabstick | REDUCED TESTING_NY_3OS428 | NY-Analytical | 760 | Swab Measurements | Tuft Adhesion | 2.4 | lb | FDA,1 | 1.4 |  | TM 109 | TM109 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | Swabstick | Swabstick | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | 2.5 | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | Swabstick | Swabstick | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | Swabstick | Swabstick | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | Swabstick | Swabstick | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | Swabstick | Swabstick | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | Swabstick | Swabstick | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | Swabstick | Swabstick | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | Swabstick | Swabstick | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS428/REDUCED TESTING_NY_3OS428 | 3OS428 | Swabstick | Swabstick | REDUCED TESTING_NY_3OS428 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS484/Applicator_26A | 3OS484 | SWB 8" DBL-TIP FOAM SWABSTICK | Swb 8" Dbl-Tip Foam Swabstick | Applicator_26A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Swab Measurements | TM109 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS484/Applicator_26A | 3OS484 | SWB 8" DBL-TIP FOAM SWABSTICK | Swb 8" Dbl-Tip Foam Swabstick | Applicator_26A | NY-Analytical | 760 | Swab Measurements | Overall Swab Length | 7.875 | in | FDA,2 | 7.82 | 7.94 | TM 109 | TM109 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS484/Applicator_26A | 3OS484 | SWB 8" DBL-TIP FOAM SWABSTICK | Swb 8" Dbl-Tip Foam Swabstick | Applicator_26A | NY-Analytical | 760 | Swab Measurements | Scrub Test | Meets specifications (Done) |  |  |  |  | TM 109 | TM109 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS484/Applicator_26A | 3OS484 | SWB 8" DBL-TIP FOAM SWABSTICK | Swb 8" Dbl-Tip Foam Swabstick | Applicator_26A | NY-Analytical | 760 | Swab Measurements | Tuft Diameter | 0.48 | in | FDA,2 | 0.42 | 0.53 | TM 109 | TM109 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS484/Applicator_26A | 3OS484 | SWB 8" DBL-TIP FOAM SWABSTICK | Swb 8" Dbl-Tip Foam Swabstick | Applicator_26A | NY-Analytical | 760 | Swab Measurements | Use Test-Solvents | Meets specifications (Done) |  |  |  |  | TM 109 | TM109 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS484/Applicator_26A | 3OS484 | SWB 8" DBL-TIP FOAM SWABSTICK | Swb 8" Dbl-Tip Foam Swabstick | Applicator_26A | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | 1.01 | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS484/Applicator_26A | 3OS484 | SWB 8" DBL-TIP FOAM SWABSTICK | Swb 8" Dbl-Tip Foam Swabstick | Applicator_26A | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS484/Applicator_26A | 3OS484 | SWB 8" DBL-TIP FOAM SWABSTICK | Swb 8" Dbl-Tip Foam Swabstick | Applicator_26A | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS484/Applicator_26A | 3OS484 | SWB 8" DBL-TIP FOAM SWABSTICK | Swb 8" Dbl-Tip Foam Swabstick | Applicator_26A | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | 1 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS484/Applicator_26A | 3OS484 | SWB 8" DBL-TIP FOAM SWABSTICK | Swb 8" Dbl-Tip Foam Swabstick | Applicator_26A | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | 0.20833 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS484/Applicator_26A | 3OS484 | SWB 8" DBL-TIP FOAM SWABSTICK | Swb 8" Dbl-Tip Foam Swabstick | Applicator_26A | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS484/Applicator_26A | 3OS484 | SWB 8" DBL-TIP FOAM SWABSTICK | Swb 8" Dbl-Tip Foam Swabstick | Applicator_26A | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS484/Applicator_26A | 3OS484 | SWB 8" DBL-TIP FOAM SWABSTICK | Swb 8" Dbl-Tip Foam Swabstick | Applicator_26A | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS484/Applicator_26A | 3OS484 | SWB 8" DBL-TIP FOAM SWABSTICK | Swb 8" Dbl-Tip Foam Swabstick | Applicator_26A | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OS484/Applicator_26A | 3OS484 | SWB 8" DBL-TIP FOAM SWABSTICK | Swb 8" Dbl-Tip Foam Swabstick | Applicator_26A | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | Gram Negative Rods (OOS) |  |  |  |  | WI QAS 076 | WI_QAS_076 | NY | RAW MATERIAL - NY | RM_SWAB | both
3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | 3OU252 | Nonwoven Rolls Spunlaced Nonwoven - 9" - 67.8gsm | Nonwoven Rolls Spunlaced Nonwoven - 9" - 67.8GSM | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Appearance | [5 empty] | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | 3OU252 | Nonwoven Rolls Spunlaced Nonwoven - 9" - 67.8gsm | Nonwoven Rolls Spunlaced Nonwoven - 9" - 67.8GSM | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size |  |  | FDA,0 |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | 3OU252 | Nonwoven Rolls Spunlaced Nonwoven - 9" - 67.8gsm | Nonwoven Rolls Spunlaced Nonwoven - 9" - 67.8GSM | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Basis Weight |  | GSM | FDA,1 | 64.4 | 71.2 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | 3OU252 | Nonwoven Rolls Spunlaced Nonwoven - 9" - 67.8gsm | Nonwoven Rolls Spunlaced Nonwoven - 9" - 67.8GSM | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Caliper |  | mm | FDA,2 | 0.28 | 0.42 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | 3OU252 | Nonwoven Rolls Spunlaced Nonwoven - 9" - 67.8gsm | Nonwoven Rolls Spunlaced Nonwoven - 9" - 67.8GSM | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | [5 empty] | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | 3OU252 | Nonwoven Rolls Spunlaced Nonwoven - 9" - 67.8gsm | Nonwoven Rolls Spunlaced Nonwoven - 9" - 67.8GSM | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Color | [5 empty] | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | 3OU252 | Nonwoven Rolls Spunlaced Nonwoven - 9" - 67.8gsm | Nonwoven Rolls Spunlaced Nonwoven - 9" - 67.8GSM | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Odor | [5 empty] | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | 3OU252 | Nonwoven Rolls Spunlaced Nonwoven - 9" - 67.8gsm | Nonwoven Rolls Spunlaced Nonwoven - 9" - 67.8GSM | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (COA) |  |  | FDA,0 | 41 | 45 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | 3OU252 | Nonwoven Rolls Spunlaced Nonwoven - 9" - 67.8gsm | Nonwoven Rolls Spunlaced Nonwoven - 9" - 67.8GSM | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Roll width (COA) |  | in | FDA,2 | 8.94 | 9.13 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | 3OU252 | Nonwoven Rolls Spunlaced Nonwoven - 9" - 67.8gsm | Nonwoven Rolls Spunlaced Nonwoven - 9" - 67.8GSM | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Wet CD |  | g/in | FDA,0 | 1500 | 2500 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | 3OU252 | Nonwoven Rolls Spunlaced Nonwoven - 9" - 67.8gsm | Nonwoven Rolls Spunlaced Nonwoven - 9" - 67.8GSM | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Wet MD |  | g/in | FDA,0 | 4500 | 6500 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU265/Applicator_10H | 3OU265 | Spunlaced Nonwoven 50 gsm 100% Polyester – 11 1/2” | Spunlaced Nonwoven 50 GSM 100% Polyester – 11 1/2” | Applicator_10H | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU265/Applicator_10H | 3OU265 | Spunlaced Nonwoven 50 gsm 100% Polyester – 11 1/2” | Spunlaced Nonwoven 50 GSM 100% Polyester – 11 1/2” | Applicator_10H | NY-Analytical | 759 | Applicator Measurements | Roll Diameter | 44 | in | FDA,0 | 41 | 45 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU265/Applicator_10H | 3OU265 | Spunlaced Nonwoven 50 gsm 100% Polyester – 11 1/2” | Spunlaced Nonwoven 50 GSM 100% Polyester – 11 1/2” | Applicator_10H | NY-Analytical | 759 | Applicator Measurements | Roll Width | 11.5 | in | FDA,2 | 11.25 | 11.75 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU265/Applicator_10H | 3OU265 | Spunlaced Nonwoven 50 gsm 100% Polyester – 11 1/2” | Spunlaced Nonwoven 50 GSM 100% Polyester – 11 1/2” | Applicator_10H | NY-Analytical | 756 | Basis Weight | Basis Weight | 50.6 | GSM | FDA,1 | 45 | 55 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU265/Applicator_10H | 3OU265 | Spunlaced Nonwoven 50 gsm 100% Polyester – 11 1/2” | Spunlaced Nonwoven 50 GSM 100% Polyester – 11 1/2” | Applicator_10H | NY-Analytical | 757 | Caliper | Caliper | 0.52 | mm | FDA,2 | 0.51 | 0.63 | TM 102 | TM102 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU265/Applicator_10H | 3OU265 | Spunlaced Nonwoven 50 gsm 100% Polyester – 11 1/2” | Spunlaced Nonwoven 50 GSM 100% Polyester – 11 1/2” | Applicator_10H | NY-Analytical | 758 | Tensile Strength | Dry CD | 2302.98 | g/in | FDA,0 | 1801 | 2821 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU265/Applicator_10H | 3OU265 | Spunlaced Nonwoven 50 gsm 100% Polyester – 11 1/2” | Spunlaced Nonwoven 50 GSM 100% Polyester – 11 1/2” | Applicator_10H | NY-Analytical | 758 | Tensile Strength | Dry CD | 8886.78 | g/in | FDA,0 | 6675 | 9447 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU265/Applicator_10H | 3OU265 | Spunlaced Nonwoven 50 gsm 100% Polyester – 11 1/2” | Spunlaced Nonwoven 50 GSM 100% Polyester – 11 1/2” | Applicator_10H | NY-Analytical | 748 | Visual Inspection | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU265/Applicator_10H | 3OU265 | Spunlaced Nonwoven 50 gsm 100% Polyester – 11 1/2” | Spunlaced Nonwoven 50 GSM 100% Polyester – 11 1/2” | Applicator_10H | NY-Analytical | 748 | Visual Inspection | Color | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU265/Applicator_10H | 3OU265 | Spunlaced Nonwoven 50 gsm 100% Polyester – 11 1/2” | Spunlaced Nonwoven 50 GSM 100% Polyester – 11 1/2” | Applicator_10H | NY-Analytical | 748 | Visual Inspection | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU265/Applicator_10H | 3OU265 | Spunlaced Nonwoven 50 gsm 100% Polyester – 11 1/2” | Spunlaced Nonwoven 50 GSM 100% Polyester – 11 1/2” | Applicator_10H | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | 1.15 | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU265/Applicator_10H | 3OU265 | Spunlaced Nonwoven 50 gsm 100% Polyester – 11 1/2” | Spunlaced Nonwoven 50 GSM 100% Polyester – 11 1/2” | Applicator_10H | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU265/Applicator_10H | 3OU265 | Spunlaced Nonwoven 50 gsm 100% Polyester – 11 1/2” | Spunlaced Nonwoven 50 GSM 100% Polyester – 11 1/2” | Applicator_10H | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU265/Applicator_10H | 3OU265 | Spunlaced Nonwoven 50 gsm 100% Polyester – 11 1/2” | Spunlaced Nonwoven 50 GSM 100% Polyester – 11 1/2” | Applicator_10H | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU265/Applicator_10H | 3OU265 | Spunlaced Nonwoven 50 gsm 100% Polyester – 11 1/2” | Spunlaced Nonwoven 50 GSM 100% Polyester – 11 1/2” | Applicator_10H | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU265/Applicator_10H | 3OU265 | Spunlaced Nonwoven 50 gsm 100% Polyester – 11 1/2” | Spunlaced Nonwoven 50 GSM 100% Polyester – 11 1/2” | Applicator_10H | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU265/Applicator_10H | 3OU265 | Spunlaced Nonwoven 50 gsm 100% Polyester – 11 1/2” | Spunlaced Nonwoven 50 GSM 100% Polyester – 11 1/2” | Applicator_10H | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU265/Applicator_10H | 3OU265 | Spunlaced Nonwoven 50 gsm 100% Polyester – 11 1/2” | Spunlaced Nonwoven 50 GSM 100% Polyester – 11 1/2” | Applicator_10H | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU265/Applicator_10H | 3OU265 | Spunlaced Nonwoven 50 gsm 100% Polyester – 11 1/2” | Spunlaced Nonwoven 50 GSM 100% Polyester – 11 1/2” | Applicator_10H | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU275/Applicator_36C | 3OU275 | 40 gsm, Hydroentangled Nonwoven 11 3/4” with Diamond Emboss | Hydroentangled Nonwoven 11 3/4” With Diamond Emboss 40 GSM | Applicator_36C | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU275/Applicator_36C | 3OU275 | 40 gsm, Hydroentangled Nonwoven 11 3/4” with Diamond Emboss | Hydroentangled Nonwoven 11 3/4” With Diamond Emboss 40 GSM | Applicator_36C | NY-Analytical | 759 | Applicator Measurements | Roll Diameter | 32 | in | FDA,0 | 30 | 34 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU275/Applicator_36C | 3OU275 | 40 gsm, Hydroentangled Nonwoven 11 3/4” with Diamond Emboss | Hydroentangled Nonwoven 11 3/4” With Diamond Emboss 40 GSM | Applicator_36C | NY-Analytical | 759 | Applicator Measurements | Roll Width | 11.75 | in | FDA,3 | 11.625 | 11.875 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU275/Applicator_36C | 3OU275 | 40 gsm, Hydroentangled Nonwoven 11 3/4” with Diamond Emboss | Hydroentangled Nonwoven 11 3/4” With Diamond Emboss 40 GSM | Applicator_36C | NY-Analytical | 756 | Basis Weight | Basis Weight | 37.67 | GSM | FDA,0 | 35 | 45 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU275/Applicator_36C | 3OU275 | 40 gsm, Hydroentangled Nonwoven 11 3/4” with Diamond Emboss | Hydroentangled Nonwoven 11 3/4” With Diamond Emboss 40 GSM | Applicator_36C | NY-Analytical | 757 | Caliper | Caliper | 458 | um | FDA,0 | 340 | 540 | TM 102 | TM102 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU275/Applicator_36C | 3OU275 | 40 gsm, Hydroentangled Nonwoven 11 3/4” with Diamond Emboss | Hydroentangled Nonwoven 11 3/4” With Diamond Emboss 40 GSM | Applicator_36C | NY-Analytical | 758 | Tensile Strength | Wet CD CoA | 1256.91 | g/in | FDA,0 | 400 |  | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU275/Applicator_36C | 3OU275 | 40 gsm, Hydroentangled Nonwoven 11 3/4” with Diamond Emboss | Hydroentangled Nonwoven 11 3/4” With Diamond Emboss 40 GSM | Applicator_36C | NY-Analytical | 758 | Tensile Strength | Wet MD CoA | 3006.85 | g/in | FDA,0 | 1000 |  | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU275/Applicator_36C | 3OU275 | 40 gsm, Hydroentangled Nonwoven 11 3/4” with Diamond Emboss | Hydroentangled Nonwoven 11 3/4” With Diamond Emboss 40 GSM | Applicator_36C | NY-Analytical | 748 | Visual Inspection | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU275/Applicator_36C | 3OU275 | 40 gsm, Hydroentangled Nonwoven 11 3/4” with Diamond Emboss | Hydroentangled Nonwoven 11 3/4” With Diamond Emboss 40 GSM | Applicator_36C | NY-Analytical | 748 | Visual Inspection | Color | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU275/Applicator_36C | 3OU275 | 40 gsm, Hydroentangled Nonwoven 11 3/4” with Diamond Emboss | Hydroentangled Nonwoven 11 3/4” With Diamond Emboss 40 GSM | Applicator_36C | NY-Analytical | 748 | Visual Inspection | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU351/Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | 3OU351 | 1.25 oz/sy (42.4 gsm) Ultrasoft Spunbond Nonwoven White | Ultrasoft Spunbond Nonwoven White 1.25 Oz/Sy (42.4 GSM) | Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA | COA | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU351/Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | 3OU351 | 1.25 oz/sy (42.4 gsm) Ultrasoft Spunbond Nonwoven White | Ultrasoft Spunbond Nonwoven White 1.25 Oz/Sy (42.4 GSM) | Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | COA Meets Specification (Done) |  |  |  |  | COA | COA | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU351/Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | 3OU351 | 1.25 oz/sy (42.4 gsm) Ultrasoft Spunbond Nonwoven White | Ultrasoft Spunbond Nonwoven White 1.25 Oz/Sy (42.4 GSM) | Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | NY-Analytical | 759 | Applicator Measurements | Roll Width | 5.968 | in | FDA,3 | 5.875 | 6.125 | TM 143 | TM143 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU351/Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | 3OU351 | 1.25 oz/sy (42.4 gsm) Ultrasoft Spunbond Nonwoven White | Ultrasoft Spunbond Nonwoven White 1.25 Oz/Sy (42.4 GSM) | Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | NY-Analytical | 397 | Appearance | Appearance/Matches Embossing Pattern | Free of Hard wrinkles. Formation uniform (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU351/Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | 3OU351 | 1.25 oz/sy (42.4 gsm) Ultrasoft Spunbond Nonwoven White | Ultrasoft Spunbond Nonwoven White 1.25 Oz/Sy (42.4 GSM) | Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | NY-Analytical | 395 | Color | Color | White with no off colors or shading (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OU351/Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | 3OU351 | 1.25 oz/sy (42.4 gsm) Ultrasoft Spunbond Nonwoven White | Ultrasoft Spunbond Nonwoven White 1.25 Oz/Sy (42.4 GSM) | Applicator_NY_1.25 oz/sy - 42.4gsm Spunbound Nonwoven | NY-Analytical | 1038 | KC Part Number | Verify KC Part Number | SPEC-08468 (Done) | [6 empty] | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Foam 10.5 mL | Foam 10.5 Ml | Applicatpr_NY_Foam 10.5ml | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Conformance | COA | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Foam 10.5 mL | Foam 10.5 Ml | Applicatpr_NY_Foam 10.5ml | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Bioburden_TAMC |  | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Foam 10.5 mL | Foam 10.5 Ml | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | 0.91 | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Foam 10.5 mL | Foam 10.5 Ml | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Foam 10.5 mL | Foam 10.5 Ml | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Foam 10.5 mL | Foam 10.5 Ml | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Foam 10.5 mL | Foam 10.5 Ml | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Foam 10.5 mL | Foam 10.5 Ml | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Foam 10.5 mL | Foam 10.5 Ml | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 1 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Foam 10.5 mL | Foam 10.5 Ml | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | 0.32967 | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Foam 10.5 mL | Foam 10.5 Ml | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Foam 10.5 mL | Foam 10.5 Ml | Applicatpr_NY_Foam 10.5ml | NY-Analytical | 397 | Appearance | Appearance | Meets specifications (Done) |  |  |  |  | Visual | Visual | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX105F/Applicatpr_NY_Foam 10.5ml | 3OX105F | Foam 10.5 mL | Foam 10.5 Ml | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Mold (OOS) |  |  |  |  | WI QAS 076 | WI_QAS_076 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | Handle, 26 Ml | Handle, 26 ml | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA | COA | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | Handle, 26 Ml | Handle, 26 ml | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | Handle, 26 Ml | Handle, 26 ml | NY-Analytical | 765 | Canister & Lid Inspection | Color | Natural (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | Handle, 26 Ml | Handle, 26 ml | NY-Analytical | 765 | Canister & Lid Inspection | Handle Internal Diameter | 0.9575 | in | FDA,3 | 0.951 | 0.957 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX465H/Microbiologu - Bioburden_Applicator_NY | 3OX465H | Handle, 26 ml | Handle, 26 Ml | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 765 | Canister & Lid Inspection | Handle Internal Diameter | 0.9545 | in | FDA,3 | 0.951 | 0.957 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | Handle, 26 Ml | Handle, 26 ml | NY-Analytical | 765 | Canister & Lid Inspection | Overall Length | 5.23 | in | FDA,3 | 5.2 | 5.25 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | Handle, 26 Ml | Handle, 26 ml | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | 87.86 | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | Handle, 26 Ml | Handle, 26 ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | Handle, 26 Ml | Handle, 26 ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | Handle, 26 Ml | Handle, 26 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | 3 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | Handle, 26 Ml | Handle, 26 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | 0.034145 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | Handle, 26 Ml | Handle, 26 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | Handle, 26 Ml | Handle, 26 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | Handle, 26 Ml | Handle, 26 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | Handle, 26 Ml | Handle, 26 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 13 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | Handle, 26 Ml | Handle, 26 ml | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | Non-Subculturable Growth (Done) |  |  |  |  | WI QAS 076 | WI_QAS_076 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX465H/Handle, 26 ml | 3OX465H | Handle, 26 ml | Handle, 26 Ml | Handle, 26 ml | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Gram Positive Rods (Done) |  |  |  |  | WI QAS 076 | WI_QAS_076 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swabstick, Cotton - Tipped | Swabstick, Cotton - Tipped | Swab_NY_Swabstick, Cotton-Tipped | NY-Analytical | 760 | Swab Measurements | Mitt Length |  | in | FDA,3 | 0.593 | 0.781 | Swab Measurements | TM109 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swabstick, Cotton - Tipped | Swabstick, Cotton - Tipped | Swab_NY_Swabstick, Cotton-Tipped | NY-Analytical | 760 | Swab Measurements | Mitt Width |  | in | FDA,3 | 0.188 | 0.28 | Swab Measurements | TM109 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swabstick, Cotton - Tipped | Swabstick, Cotton - Tipped | Swab_NY_Swabstick, Cotton-Tipped | NY-Analytical | 760 | Swab Measurements | Overall Swab Length |  | in | FDA,3 | 5.875 | 5.999 | Swab Measurements | TM109 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swabstick, Cotton - Tipped | Swabstick, Cotton - Tipped | Swab_NY_Swabstick, Cotton-Tipped | NY-Analytical | 760 | Swab Measurements | Stick Diameter |  | in | FDA,3 | 0.093 | 0.107 | Swab Measurements | TM109 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swabstick, Cotton - Tipped | Swabstick, Cotton - Tipped | Swab_NY_Swabstick, Cotton-Tipped | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | [5 empty] | Certificate of Conformance | COA | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swabstick, Cotton - Tipped | Swabstick, Cotton - Tipped | Swab_NY_Swabstick, Cotton-Tipped | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | [5 empty] | Bioburden_TAMC |  | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swabstick, Cotton - Tipped | Swabstick, Cotton - Tipped | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | 0.58 | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swabstick, Cotton - Tipped | Swabstick, Cotton - Tipped | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swabstick, Cotton - Tipped | Swabstick, Cotton - Tipped | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swabstick, Cotton - Tipped | Swabstick, Cotton - Tipped | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swabstick, Cotton - Tipped | Swabstick, Cotton - Tipped | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram) | 0 | CFU/g |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swabstick, Cotton - Tipped | Swabstick, Cotton - Tipped | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swabstick, Cotton - Tipped | Swabstick, Cotton - Tipped | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swabstick, Cotton - Tipped | Swabstick, Cotton - Tipped | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | 3OX465Q | Swabstick, Cotton - Tipped | Swabstick, Cotton - Tipped | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | Handle, 5.1 Ml | Handle, 5.1 ml | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA | COA | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | Handle, 5.1 Ml | Handle, 5.1 ml | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Meets Specifications (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | Handle, 5.1 Ml | Handle, 5.1 ml | NY-Analytical | 765 | Canister & Lid Inspection | Color | Natural (Done) |  |  |  |  | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | Handle, 5.1 Ml | Handle, 5.1 ml | NY-Analytical | 765 | Canister & Lid Inspection | Handle Internal Diameter | 0.6405 | in | FDA,3 | 0.639 | 0.645 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | Handle, 5.1 Ml | Handle, 5.1 ml | NY-Analytical | 765 | Canister & Lid Inspection | Overall Length | 3.156 | in | FDA,3 | 3.149 | 3.189 | Canister & Lid Inspection | TM143 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | Handle, 5.1 Ml | Handle, 5.1 ml | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | 1.5 | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | Handle, 5.1 Ml | Handle, 5.1 ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | Handle, 5.1 Ml | Handle, 5.1 ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | Handle, 5.1 Ml | Handle, 5.1 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | Handle, 5.1 Ml | Handle, 5.1 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | Handle, 5.1 Ml | Handle, 5.1 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | Handle, 5.1 Ml | Handle, 5.1 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | Handle, 5.1 Ml | Handle, 5.1 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3OX510H/Handle, 5.1 ml | 3OX510H | Handle, 5.1 ml | Handle, 5.1 Ml | Handle, 5.1 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | 35gsm Stora Flat - 36" | Stora Flat - 36" 35 GSM | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | 35gsm Stora Flat - 36" | Stora Flat - 36" 35 GSM | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 759 | Applicator Measurements | Roll Diameter | 39 | in | FDA,0 | 39 | 40 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | 35gsm Stora Flat - 36" | Stora Flat - 36" 35 GSM | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 759 | Applicator Measurements | Roll Width | 36 | in | FDA,3 | 35.875 | 36.125 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | 35gsm Stora Flat - 36" | Stora Flat - 36" 35 GSM | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 884 | Applicator_Inspection | Caliper | 0.0082 | in | FDA,3 | 0.006 | 0.009 | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | 35gsm Stora Flat - 36" | Stora Flat - 36" 35 GSM | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 756 | Basis Weight | Basis Weight | 34.44 | GSM | FDA,1 | 31.5 | 38.5 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | 35gsm Stora Flat - 36" | Stora Flat - 36" 35 GSM | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 758 | Tensile Strength | Wet CD | 842.87 | g/in | FDA,0 | 250 | 1250 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | 35gsm Stora Flat - 36" | Stora Flat - 36" 35 GSM | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 758 | Tensile Strength | Wet MD | 1755.59 | g/in | FDA,0 | 800 | 2600 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | 35gsm Stora Flat - 36" | Stora Flat - 36" 35 GSM | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 748 | Visual Inspection | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | 35gsm Stora Flat - 36" | Stora Flat - 36" 35 GSM | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 748 | Visual Inspection | Color | White with no off colors or shading [Done] |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | 35gsm Stora Flat - 36" | Stora Flat - 36" 35 GSM | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 748 | Visual Inspection | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | 35gsm Stora Flat - 36" | Stora Flat - 36" 35 GSM | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | 1.55 | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | 35gsm Stora Flat - 36" | Stora Flat - 36" 35 GSM | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | 35gsm Stora Flat - 36" | Stora Flat - 36" 35 GSM | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | 35gsm Stora Flat - 36" | Stora Flat - 36" 35 GSM | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | 1 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | 35gsm Stora Flat - 36" | Stora Flat - 36" 35 GSM | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | 0.12903 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | 35gsm Stora Flat - 36" | Stora Flat - 36" 35 GSM | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 2 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | 35gsm Stora Flat - 36" | Stora Flat - 36" 35 GSM | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | 35gsm Stora Flat - 36" | Stora Flat - 36" 35 GSM | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | 35gsm Stora Flat - 36" | Stora Flat - 36" 35 GSM | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP134/Applicator_NY_35gsm Stora Flat - 36" | 3XP134 | 35gsm Stora Flat - 36" | Stora Flat - 36" 35 GSM | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | Gram Negative Rods (OOS) |  |  |  |  | WI QAS 076 | WI_QAS_076 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP138/Applicator_36B | 3XP138 | 40 gsm, Hydroentangled Nonwoven 37 1/2" with Diamond Emboss | Hydroentangled Nonwoven 37 1/2" With Diamond Emboss 40 GSM | Applicator_36B | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP138/Applicator_36B | 3XP138 | 40 gsm, Hydroentangled Nonwoven 37 1/2" with Diamond Emboss | Hydroentangled Nonwoven 37 1/2" With Diamond Emboss 40 GSM | Applicator_36B | NY-Analytical | 759 | Applicator Measurements | Roll Diameter | 45 | in | FDA,0 | 43 | 47 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP138/Applicator_36B | 3XP138 | 40 gsm, Hydroentangled Nonwoven 37 1/2" with Diamond Emboss | Hydroentangled Nonwoven 37 1/2" With Diamond Emboss 40 GSM | Applicator_36B | NY-Analytical | 759 | Applicator Measurements | Roll Width | 37.5 | in | FDA,3 | 37.375 | 37.625 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP138/Applicator_36B | 3XP138 | 40 gsm, Hydroentangled Nonwoven 37 1/2" with Diamond Emboss | Hydroentangled Nonwoven 37 1/2" With Diamond Emboss 40 GSM | Applicator_36B | NY-Analytical | 756 | Basis Weight | Basis Weight | 40.04 | GSM | FDA,0 | 35 | 45 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP138/Applicator_36B | 3XP138 | 40 gsm, Hydroentangled Nonwoven 37 1/2" with Diamond Emboss | Hydroentangled Nonwoven 37 1/2" With Diamond Emboss 40 GSM | Applicator_36B | NY-Analytical | 757 | Caliper | Caliper | 419.8 | um | FDA,0 | 340 | 540 | TM 102 | TM102 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP138/Applicator_36B | 3XP138 | 40 gsm, Hydroentangled Nonwoven 37 1/2" with Diamond Emboss | Hydroentangled Nonwoven 37 1/2" With Diamond Emboss 40 GSM | Applicator_36B | NY-Analytical | 758 | Tensile Strength | Wet CD CoA | 1069 | g/in | FDA,0 | 400 |  | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP138/Applicator_36B | 3XP138 | 40 gsm, Hydroentangled Nonwoven 37 1/2" with Diamond Emboss | Hydroentangled Nonwoven 37 1/2" With Diamond Emboss 40 GSM | Applicator_36B | NY-Analytical | 758 | Tensile Strength | Wet MD CoA | 2099 | g/in | FDA,0 | 1000 |  | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP138/Applicator_36B | 3XP138 | 40 gsm, Hydroentangled Nonwoven 37 1/2" with Diamond Emboss | Hydroentangled Nonwoven 37 1/2" With Diamond Emboss 40 GSM | Applicator_36B | NY-Analytical | 748 | Visual Inspection | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP138/Applicator_36B | 3XP138 | 40 gsm, Hydroentangled Nonwoven 37 1/2" with Diamond Emboss | Hydroentangled Nonwoven 37 1/2" With Diamond Emboss 40 GSM | Applicator_36B | NY-Analytical | 748 | Visual Inspection | Color | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP138/Applicator_36B | 3XP138 | 40 gsm, Hydroentangled Nonwoven 37 1/2" with Diamond Emboss | Hydroentangled Nonwoven 37 1/2" With Diamond Emboss 40 GSM | Applicator_36B | NY-Analytical | 748 | Visual Inspection | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
| 3XP253 | 35gsm Flat Thermoembonded Nonwoven 30" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Applicator Measurements | TM142 |  |  |  | left_only
| 3XP253 | 35gsm Flat Thermoembonded Nonwoven 30" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 759 | Applicator Measurements | Roll Width | 30 | in | FDA,3 | 29.875 | 30.125 | Applicator Measurements | TM142 |  |  |  | left_only
| 3XP253 | 35gsm Flat Thermoembonded Nonwoven 30" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 756 | Basis Weight | Basis Weight | 37.03 | GSM | FDA,1 | 31.5 | 38.5 | TM 101 | TM101 |  |  |  | left_only
| 3XP253 | 35gsm Flat Thermoembonded Nonwoven 30" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 757 | Caliper | Caliper | 0.33 | mm | FDA,2 | 0.27 | 0.41 | TM 102 | TM102 |  |  |  | left_only
| 3XP253 | 35gsm Flat Thermoembonded Nonwoven 30" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 758 | Tensile Strength | Wet CD | 599.1 | g/in | FDA,0 | 300 |  | TM 103 | TM103 |  |  |  | left_only
| 3XP253 | 35gsm Flat Thermoembonded Nonwoven 30" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 758 | Tensile Strength | Wet MD | 3682.03 | g/in | FDA,0 | 2000 | 4400 | TM 103 | TM103 |  |  |  | left_only
| 3XP253 | 35gsm Flat Thermoembonded Nonwoven 30" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 748 | Visual Inspection | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 |  |  |  | left_only
| 3XP253 | 35gsm Flat Thermoembonded Nonwoven 30" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 748 | Visual Inspection | Color | White with no off colors or shading [Done] |  |  |  |  | TM 208 | TM208 |  |  |  | left_only
| 3XP253 | 35gsm Flat Thermoembonded Nonwoven 30" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 748 | Visual Inspection | Odor | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 |  |  |  | left_only
| 3XP253 | 35gsm Flat Thermoembonded Nonwoven 30" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | 0.98 | g | FDA,2 |  |  | TM 610 | TM610 |  |  |  | left_only
| 3XP253 | 35gsm Flat Thermoembonded Nonwoven 30" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA |  |  |  |  | TM 610 | TM610 |  |  |  | left_only
| 3XP253 | 35gsm Flat Thermoembonded Nonwoven 30" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | No Growth SDA |  |  |  |  | TM 610 | TM610 |  |  |  | left_only
| 3XP253 | 35gsm Flat Thermoembonded Nonwoven 30" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | 1 | cfu/device | FDA,1 |  |  | TM 610 | TM610 |  |  |  | left_only
| 3XP253 | 35gsm Flat Thermoembonded Nonwoven 30" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | 0.91837 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 |  |  |  | left_only
| 3XP253 | 35gsm Flat Thermoembonded Nonwoven 30" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 1 | cfu | FDA,0 |  |  | TM 610 | TM610 |  |  |  | left_only
| 3XP253 | 35gsm Flat Thermoembonded Nonwoven 30" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 |  |  |  | left_only
| 3XP253 | 35gsm Flat Thermoembonded Nonwoven 30" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 |  |  |  | left_only
| 3XP253 | 35gsm Flat Thermoembonded Nonwoven 30" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 |  |  |  | left_only
| 3XP253 | 35gsm Flat Thermoembonded Nonwoven 30" |  | Analytical - Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | Gram Positive Rods (Done) |  |  |  |  | WI QAS 076 | WI_QAS_076 |  |  |  | left_only
3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 3XP540 | Spunlaced Nonwoven 50gsm 100% Polyester - 42" | Spunlaced Nonwoven 50 GSM 100% Polyester - 42" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (COA) | 43 | in | FDA,0 | 41 | 45 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 3XP540 | Spunlaced Nonwoven 50gsm 100% Polyester - 42" | Spunlaced Nonwoven 50 GSM 100% Polyester - 42" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 759 | Applicator Measurements | Roll width (COA) | 42 | in | FDA,2 | 41.75 | 42.25 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 3XP540 | Spunlaced Nonwoven 50gsm 100% Polyester - 42" | Spunlaced Nonwoven 50 GSM 100% Polyester - 42" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | [5 empty] | Basis Weight |  | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 3XP540 | Spunlaced Nonwoven 50gsm 100% Polyester - 42" | Spunlaced Nonwoven 50 GSM 100% Polyester - 42" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 756 | Basis Weight | Basis Weight | 52.31 | GSM | FDA,0 | 45 | 55 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 3XP540 | Spunlaced Nonwoven 50gsm 100% Polyester - 42" | Spunlaced Nonwoven 50 GSM 100% Polyester - 42" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 757 | Caliper | Caliper | 0.42 | mm | FDA,2 | 0.51 | 0.63 | TM 102 | TM102 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 3XP540 | Spunlaced Nonwoven 50gsm 100% Polyester - 42" | Spunlaced Nonwoven 50 GSM 100% Polyester - 42" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 758 | Tensile Strength | Dry CD | 2372 | g/in | FDA,0 | 1801 | 2821 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 3XP540 | Spunlaced Nonwoven 50gsm 100% Polyester - 42" | Spunlaced Nonwoven 50 GSM 100% Polyester - 42" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 758 | Tensile Strength | Dry CD | 9752 | g/in | FDA,0 | 6675 | 9447 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 3XP540 | Spunlaced Nonwoven 50gsm 100% Polyester - 42" | Spunlaced Nonwoven 50 GSM 100% Polyester - 42" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 748 | Visual Inspection | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 3XP540 | Spunlaced Nonwoven 50gsm 100% Polyester - 42" | Spunlaced Nonwoven 50 GSM 100% Polyester - 42" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 748 | Visual Inspection | Color | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | 3XP540 | Spunlaced Nonwoven 50gsm 100% Polyester - 42" | Spunlaced Nonwoven 50 GSM 100% Polyester - 42" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 748 | Visual Inspection | Odor | No objectionable odor Wet or Dry [Done] |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Spunlaced Nonwoven 50gsm 100% Polyester - 38.75" | Spunlaced Nonwoven 50 GSM 100% Polyester - 38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 759 | Applicator Measurements | Roll Diameter | 43 | in | FDA,0 | 41 | 45 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Spunlaced Nonwoven 50gsm 100% Polyester - 38.75" | Spunlaced Nonwoven 50 GSM 100% Polyester - 38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 759 | Applicator Measurements | Roll Width | 38.75 | in | FDA,2 | 38.5 | 39 | Applicator Measurements | TM142 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Spunlaced Nonwoven 50gsm 100% Polyester - 38.75" | Spunlaced Nonwoven 50 GSM 100% Polyester - 38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Basis Weight |  | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Spunlaced Nonwoven 50gsm 100% Polyester - 38.75" | Spunlaced Nonwoven 50 GSM 100% Polyester - 38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 756 | Basis Weight | Basis Weight | 49.94 | GSM | FDA,1 | 45 | 55 | TM 101 | TM101 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Spunlaced Nonwoven 50gsm 100% Polyester - 38.75" | Spunlaced Nonwoven 50 GSM 100% Polyester - 38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 757 | Caliper | Caliper | 0.56 | mm | FDA,2 | 0.51 | 0.63 | TM 102 | TM102 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Spunlaced Nonwoven 50gsm 100% Polyester - 38.75" | Spunlaced Nonwoven 50 GSM 100% Polyester - 38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 758 | Tensile Strength | Dry CD | 2635 | g/in | FDA,0 | 1801 | 2821 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Spunlaced Nonwoven 50gsm 100% Polyester - 38.75" | Spunlaced Nonwoven 50 GSM 100% Polyester - 38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 758 | Tensile Strength | Dry CD | 8881 | g/in | FDA,0 | 6675 | 9447 | TM 103 | TM103 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Spunlaced Nonwoven 50gsm 100% Polyester - 38.75" | Spunlaced Nonwoven 50 GSM 100% Polyester - 38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 748 | Visual Inspection | Appearance | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Spunlaced Nonwoven 50gsm 100% Polyester - 38.75" | Spunlaced Nonwoven 50 GSM 100% Polyester - 38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 748 | Visual Inspection | Color | Meets specifications (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Spunlaced Nonwoven 50gsm 100% Polyester - 38.75" | Spunlaced Nonwoven 50 GSM 100% Polyester - 38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 748 | Visual Inspection | Odor | No objectionable odor Wet or Dry [Done] |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Spunlaced Nonwoven 50gsm 100% Polyester - 38.75" | Spunlaced Nonwoven 50 GSM 100% Polyester - 38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | 0.94 | g | FDA,2 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Spunlaced Nonwoven 50gsm 100% Polyester - 38.75" | Spunlaced Nonwoven 50 GSM 100% Polyester - 38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Spunlaced Nonwoven 50gsm 100% Polyester - 38.75" | Spunlaced Nonwoven 50 GSM 100% Polyester - 38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer |  |  |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Spunlaced Nonwoven 50gsm 100% Polyester - 38.75" | Spunlaced Nonwoven 50 GSM 100% Polyester - 38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Spunlaced Nonwoven 50gsm 100% Polyester - 38.75" | Spunlaced Nonwoven 50 GSM 100% Polyester - 38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Spunlaced Nonwoven 50gsm 100% Polyester - 38.75" | Spunlaced Nonwoven 50 GSM 100% Polyester - 38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Spunlaced Nonwoven 50gsm 100% Polyester - 38.75" | Spunlaced Nonwoven 50 GSM 100% Polyester - 38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | 0 | cfu/device | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Spunlaced Nonwoven 50gsm 100% Polyester - 38.75" | Spunlaced Nonwoven 50 GSM 100% Polyester - 38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | 0 | CFU/g | FDA,1 |  | 15 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | 3XP667 | Spunlaced Nonwoven 50gsm 100% Polyester - 38.75" | Spunlaced Nonwoven 50 GSM 100% Polyester - 38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_APPLICATOR | both
| 4YBBY12 | NDI 20X Concentrate |  | 4YBBY12 | JB-Analytical | 398 | pH | pH |  |  | FDA,1 | 10.8 | 11.3 | TM200 | TM200 |  |  |  | left_only
| 4YBBY12 | NDI 20X Concentrate |  | 4YBBY12 | JB-Analytical | 902 | Bulk Liquid Inspection | Appearance | [5 empty] | See Analyte for Method # | See_Analyte_for_Method |  |  |  | left_only
| 4YBBY12 | NDI 20X Concentrate |  | 4YBBY12 | JB-Analytical | 902 | Bulk Liquid Inspection | Color | [5 empty] | See Analyte for Method # | See_Analyte_for_Method |  |  |  | left_only
| 4YBBY12 | NDI 20X Concentrate |  | 4YBBY12 | JB-Analytical | 902 | Bulk Liquid Inspection | Contamination | [5 empty] | See Analyte for Method # | See_Analyte_for_Method |  |  |  | left_only
| 4YBBY12 | NDI 20X Concentrate |  | 4YBBY12 | JB-Analytical | 902 | Bulk Liquid Inspection | Odor | [5 empty] | See Analyte for Method # | See_Analyte_for_Method |  |  |  | left_only
| 4YBBY12 | NDI 20X Concentrate |  | Analytical - BL_JB 20X Sodium Benz Base Blend | JB-Analytical | 447 | pH as is | pH | 11.86 |  | FDA,2 | 11.2 | 12.6 | TM 200 | TM200 |  |  |  | left_only
| 4YBBY12 | NDI 20X Concentrate |  | Analytical - BL_JB 20X Sodium Benz Base Blend | JB-Analytical | 748 | Visual Inspection | Appearance | Slighlty Hazy (Done) |  |  |  |  | TM 208 | TM208 |  |  |  | left_only
| 4YBBY12 | NDI 20X Concentrate |  | Analytical - BL_JB 20X Sodium Benz Base Blend | JB-Analytical | 748 | Visual Inspection | Color | Colorless to Slight Yellow (Done) |  |  |  |  | TM 208 | TM208 |  |  |  | left_only
| 4YBBY12 | NDI 20X Concentrate |  | Analytical - BL_JB 20X Sodium Benz Base Blend | JB-Analytical | 748 | Visual Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 |  |  |  | left_only
| 4YBBY12 | NDI 20X Concentrate |  | Analytical - BL_JB 20X Sodium Benz Base Blend | JB-Analytical | 748 | Visual Inspection | Odor | Unscented (Done) |  |  |  |  | TM 208 | TM208 |  |  |  | left_only
| 4YBBY12 | NDI 20X Concentrate |  | Analytical - BL_JB 20X Sodium Benz Base Blend | JB-Analytical | 425 | % Sodium Benzoate | Sodium Benzoate | 20.47 | % | FDA,2 | 17.86 | 20.68 | TM 395 | TM395 |  |  |  | left_only
| 4YBBY12 | NDI 20X Concentrate |  | 4YBBY12 | JB-Analytical | 425 | % Sodium Benzoate | Sodium Benzoate |  | % | FDA,2 | 8.46 | 10.34 | TM 395 | TM395 |  |  |  | left_only
| 4YBBY12 | NDI 20X Concentrate |  | Analytical - BL_JB 20X Sodium Benz Base Blend | JB-Analytical | 425 | % Sodium Benzoate | Sodium Benzoate Final | 20.41 | % | FDA,2 | 17.86 | 20.68 | TM 395 | TM395 |  |  |  | left_only
| 4YBBY12 | NDI 20X Concentrate |  | 4YBBY12 | JB-Analytical | 425 | % Sodium Benzoate | Sodium Benzoate Final |  | % | FDA,2 | 8.46 | 10.34 | TM 395 | TM395 |  |  |  | left_only
| 4YBBY23 | NDI 20X Concentrate w/ Vitamin E |  | 4YBBY23 | JB-Analytical | 902 | Bulk Liquid Inspection | Appearance | [5 empty] | See Analyte for Method # | See_Analyte_for_Method |  |  |  | left_only
| 4YBBY23 | NDI 20X Concentrate w/ Vitamin E |  | Analytical - BL_JB 40X Natural Baby W/Vitamin E | JB-Analytical | 902 | Bulk Liquid Inspection | Appearance | Slightly Hazy Liquid (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method |  |  |  | left_only
| 4YBBY23 | NDI 20X Concentrate w/ Vitamin E |  | 4YBBY23 | JB-Analytical | 902 | Bulk Liquid Inspection | Color | [5 empty] | See Analyte for Method # | See_Analyte_for_Method |  |  |  | left_only
| 4YBBY23 | NDI 20X Concentrate w/ Vitamin E |  | Analytical - BL_JB 40X Natural Baby W/Vitamin E | JB-Analytical | 902 | Bulk Liquid Inspection | Color | Slightly Yellow Liquid (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method |  |  |  | left_only
| 4YBBY23 | NDI 20X Concentrate w/ Vitamin E |  | 4YBBY23 | JB-Analytical | 902 | Bulk Liquid Inspection | Contamination | [5 empty] | See Analyte for Method # | See_Analyte_for_Method |  |  |  | left_only
| 4YBBY23 | NDI 20X Concentrate w/ Vitamin E |  | Analytical - BL_JB 40X Natural Baby W/Vitamin E | JB-Analytical | 902 | Bulk Liquid Inspection | Contamination | None (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method |  |  |  | left_only
| 4YBBY23 | NDI 20X Concentrate w/ Vitamin E |  | 4YBBY23 | JB-Analytical | 902 | Bulk Liquid Inspection | Odor | [5 empty] | See Analyte for Method # | See_Analyte_for_Method |  |  |  | left_only
| 4YBBY23 | NDI 20X Concentrate w/ Vitamin E |  | Analytical - BL_JB 40X Natural Baby W/Vitamin E | JB-Analytical | 902 | Bulk Liquid Inspection | Odor | Unscented (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method |  |  |  | left_only
| 4YBBY23 | NDI 20X Concentrate w/ Vitamin E |  | Analytical - BL_JB 40X Natural Baby W/Vitamin E | JB-Analytical | 902 | Bulk Liquid Inspection | pH | 11.88 |  | FDA,2 | 11.2 | 12.6 | See Analyte for Method # | See_Analyte_for_Method |  |  |  | left_only
| 4YBBY23 | NDI 20X Concentrate w/ Vitamin E |  | 4YBBY23 | JB-Analytical | 398 | pH | pH |  |  | FDA,1 | 10.8 | 11.3 | TM 200 | TM200 |  |  |  | left_only
| 4YBBY23 | NDI 20X Concentrate w/ Vitamin E |  | 4YBBY23 | JB-Analytical | 425 | % Sodium Benzoate | Sodium Benzoate |  | % | FDA,2 | 8.46 | 10.34 | TM 395 | TM395 |  |  |  | left_only
| 4YBBY23 | NDI 20X Concentrate w/ Vitamin E |  | Analytical - BL_JB 40X Natural Baby W/Vitamin E | JB-Analytical | 425 | % Sodium Benzoate | Sodium Benzoate | 20.11 | % | FDA,2 | 17.86 | 20.68 | TM 395 | TM395 |  |  |  | left_only
| 4YBBY23 | NDI 20X Concentrate w/ Vitamin E |  | 4YBBY23 | JB-Analytical | 425 | % Sodium Benzoate | Sodium Benzoate Final | 0 | % | FDA,2 | 8.46 | 10.34 | TM 395 | TM395 |  |  |  | left_only
| 4YBBY23 | NDI 20X Concentrate w/ Vitamin E |  | Analytical - BL_JB 40X Natural Baby W/Vitamin E | JB-Analytical | 425 | % Sodium Benzoate | Sodium Benzoate Final | 20.13 | % | FDA,2 | 17.86 | 20.68 | TM 395 | TM395 |  |  |  | left_only
5AA347BT/Box_NY_0.020SBS_300_3.375-1.313-2.625 | 5AA347BT | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | Box_NY_0.020SBS_300_3.375-1.313-2.625 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | 315 |  | FDA,0 |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AA347BT/Box_NY_0.020SBS_300_3.375-1.313-2.625 | 5AA347BT | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | Box_NY_0.020SBS_300_3.375-1.313-2.625 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AA347BT/Box_NY_0.020SBS_300_3.375-1.313-2.625 | 5AA347BT | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | Box_NY_0.020SBS_300_3.375-1.313-2.625 | NY-Analytical | 772 | Box Inspection | Caliper | 0.019 | in | FDA,3 | 0.019 | 0.021 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AA347BT/Box_NY_0.020SBS_300_3.375-1.313-2.625 | 5AA347BT | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | Box_NY_0.020SBS_300_3.375-1.313-2.625 | NY-Analytical | 772 | Box Inspection | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AA347BT/Box_NY_0.020SBS_300_3.375-1.313-2.625 | 5AA347BT | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | Box_NY_0.020SBS_300_3.375-1.313-2.625 | NY-Analytical | 772 | Box Inspection | Carton Style | Open Top/Auto Bottom Tray (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AA347BT/Box_NY_Skip lot testing | 5AA347BT | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | Box_NY_Skip lot testing | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AA347BT/Box_NY_0.020SBS_300_3.375-1.313-2.625 | 5AA347BT | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | Box_NY_0.020SBS_300_3.375-1.313-2.625 | NY-Analytical | 772 | Box Inspection | Height | 2.625 | in | FDA,5 | 2.5938 | 2.6562 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AA347BT/Box_NY_0.020SBS_300_3.375-1.313-2.625 | 5AA347BT | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | Box_NY_0.020SBS_300_3.375-1.313-2.625 | NY-Analytical | 772 | Box Inspection | Length | 3.375 | in | FDA,5 | 3.3438 | 3.4062 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AA347BT/Box_NY_0.020SBS_300_3.375-1.313-2.625 | 5AA347BT | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | Box_NY_0.020SBS_300_3.375-1.313-2.625 | NY-Analytical | 772 | Box Inspection | Width | 1.3125 | in | FDA,5 | 1.2817 | 1.3442 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
| 5AA404S48 | Preparation H Medicated Wipes 48ct Box - NY |  | Analytical - Box_NY_0.018SBS_111_6-5.875-3.625 | NY-Analytical | 772 | Box Inspection | Artwork / Copy |  | in |  |  |  | Box Inspection | Box Inspection |  |  |  | left_only
| 5AA404S48 | Preparation H Medicated Wipes 48ct Box - NY |  | Analytical - Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection |  |  |  | left_only
| 5AA404S48 | Preparation H Medicated Wipes 48ct Box - NY |  | Analytical - Box_NY_0.018SBS_111_6-5.875-3.625 | NY-Analytical | 772 | Box Inspection | Caliper |  | in | FDA,3 | 0.017 | 0.019 | Box Inspection | Box Inspection |  |  |  | left_only
| 5AA404S48 | Preparation H Medicated Wipes 48ct Box - NY |  | Analytical - Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | NY-Analytical | 772 | Box Inspection | Caliper | 0.018 | in | FDA,3 | 0.017 | 0.019 | Box Inspection | Box Inspection |  |  |  | left_only
| 5AA404S48 | Preparation H Medicated Wipes 48ct Box - NY |  | Analytical - Box_NY_0.018SBS_111_6-5.875-3.625 | NY-Analytical | 772 | Box Inspection | Carton Material | [5 empty] | Box Inspection | Box Inspection |  |  |  | left_only
| 5AA404S48 | Preparation H Medicated Wipes 48ct Box - NY |  | Analytical - Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | NY-Analytical | 772 | Box Inspection | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | Box Inspection |  |  |  | left_only
| 5AA404S48 | Preparation H Medicated Wipes 48ct Box - NY |  | Analytical - Box_NY_0.018SBS_111_6-5.875-3.625 | NY-Analytical | 772 | Box Inspection | Carton Style | [5 empty] | Box Inspection | Box Inspection |  |  |  | left_only
| 5AA404S48 | Preparation H Medicated Wipes 48ct Box - NY |  | Analytical - Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | NY-Analytical | 772 | Box Inspection | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | Box Inspection |  |  |  | left_only
| 5AA404S48 | Preparation H Medicated Wipes 48ct Box - NY |  | Analytical - Box_NY_0.018SBS_111_6-5.875-3.625 | NY-Analytical | 772 | Box Inspection | Height |  | in | FDA,5 | 3.5938 | 3.6562 | Box Inspection | Box Inspection |  |  |  | left_only
| 5AA404S48 | Preparation H Medicated Wipes 48ct Box - NY |  | Analytical - Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | NY-Analytical | 772 | Box Inspection | Height | 6.0625 | in | FDA,5 | 6.0312 | 6.0938 | Box Inspection | Box Inspection |  |  |  | left_only
| 5AA404S48 | Preparation H Medicated Wipes 48ct Box - NY |  | Analytical - Box_NY_0.018SBS_111_6-5.875-3.625 | NY-Analytical | 772 | Box Inspection | Length |  | in | FDA,5 | 5.9688 | 6.0312 | Box Inspection | Box Inspection |  |  |  | left_only
| 5AA404S48 | Preparation H Medicated Wipes 48ct Box - NY |  | Analytical - Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | NY-Analytical | 772 | Box Inspection | Length | 7.1875 | in | FDA,5 | 7.1562 | 7.2188 | Box Inspection | Box Inspection |  |  |  | left_only
| 5AA404S48 | Preparation H Medicated Wipes 48ct Box - NY |  | Analytical - Box_NY_0.018SBS_111_6-5.875-3.625 | NY-Analytical | 772 | Box Inspection | Width |  | in | FDA,5 | 5.8438 | 5.9062 | Box Inspection | Box Inspection |  |  |  | left_only
| 5AA404S48 | Preparation H Medicated Wipes 48ct Box - NY |  | Analytical - Box_NY_0.018SBS_100_7.1875-3.6875-6.0625 | NY-Analytical | 772 | Box Inspection | Width | 3.6848 | in | FDA,5 | 3.6562 | 3.7188 | Box Inspection | Box Inspection |  |  |  | left_only
5AA617S96/Box_NY_0.018SBS_47B_3.625-3.875-6 | 5AA617S96 | Preparation H Medicated Wipes 96ct Box | Prep H Medicated Wipes 96Ct. Box | Box_NY_0.018SBS_47B_3.625-3.875-6 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AA617S96/Box_NY_0.018SBS_47B_3.625-3.875-6 | 5AA617S96 | Preparation H Medicated Wipes 96ct Box | Prep H Medicated Wipes 96Ct. Box | Box_NY_0.018SBS_47B_3.625-3.875-6 | NY-Analytical | 772 | Box Inspection | Caliper | 0.019 | in | FDA,3 | 0.017 | 0.019 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AA617S96/Box_NY_0.018SBS_47B_3.625-3.875-6 | 5AA617S96 | Preparation H Medicated Wipes 96ct Box | Prep H Medicated Wipes 96Ct. Box | Box_NY_0.018SBS_47B_3.625-3.875-6 | NY-Analytical | 772 | Box Inspection | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AA617S96/Box_NY_0.018SBS_47B_3.625-3.875-6 | 5AA617S96 | Preparation H Medicated Wipes 96ct Box | Prep H Medicated Wipes 96Ct. Box | Box_NY_0.018SBS_47B_3.625-3.875-6 | NY-Analytical | 772 | Box Inspection | Carton Style | Straight Tuck End - French Style (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AA617S96/Box_NY_0.018SBS_47B_3.625-3.875-6 | 5AA617S96 | Preparation H Medicated Wipes 96ct Box | Prep H Medicated Wipes 96Ct. Box | Box_NY_0.018SBS_47B_3.625-3.875-6 | NY-Analytical | 772 | Box Inspection | Height | 6 | in | FDA,5 | 5.9688 | 6.0312 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AA617S96/Box_NY_0.018SBS_47B_3.625-3.875-6 | 5AA617S96 | Preparation H Medicated Wipes 96ct Box | Prep H Medicated Wipes 96Ct. Box | Box_NY_0.018SBS_47B_3.625-3.875-6 | NY-Analytical | 772 | Box Inspection | Length | 3.625 | in | FDA,5 | 3.5938 | 3.6562 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AA617S96/Box_NY_0.018SBS_47B_3.625-3.875-6 | 5AA617S96 | Preparation H Medicated Wipes 96ct Box | Prep H Medicated Wipes 96Ct. Box | Box_NY_0.018SBS_47B_3.625-3.875-6 | NY-Analytical | 772 | Box Inspection | Width | 3.875 | in | FDA,5 | 3.8438 | 3.9062 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | 5AA632S60 | Preparation H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Box_NY_0.020SBS_101_6-3.625-7 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | 3 |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | 5AA632S60 | Preparation H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Box_NY_0.020SBS_101_6-3.625-7 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | 5AA632S60 | Preparation H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Box_NY_0.020SBS_101_6-3.625-7 | NY-Analytical | 772 | Box Inspection | Caliper | 0.021 | in | FDA,3 | 0.019 | 0.021 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | 5AA632S60 | Preparation H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Box_NY_0.020SBS_101_6-3.625-7 | NY-Analytical | 772 | Box Inspection | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | 5AA632S60 | Preparation H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Box_NY_0.020SBS_101_6-3.625-7 | NY-Analytical | 772 | Box Inspection | Carton Style | Tuck Top / Auto Bottom - French Style (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | 5AA632S60 | Preparation H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Box_NY_0.020SBS_101_6-3.625-7 | NY-Analytical | 772 | Box Inspection | Height | 7 | in | FDA,5 | 6.9688 | 7.0312 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | 5AA632S60 | Preparation H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Box_NY_0.020SBS_101_6-3.625-7 | NY-Analytical | 772 | Box Inspection | Length | 6 | in | FDA,5 | 5.9688 | 6.0312 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | 5AA632S60 | Preparation H Medicated Hemorrhoidal Wipes | Prep H Medicated Hemorrhoidal Wipes | Box_NY_0.020SBS_101_6-3.625-7 | NY-Analytical | 772 | Box Inspection | Width | 3.625 | in | FDA,5 | 3.5938 | 3.6562 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | 5AA911TY | Dove Baby Wipes Rich Moisture 30ct | Dove Baby Wipes Rich Moisture 30Ct. | Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | NY-Analytical | 772 | Box Inspection | Carton Style | Tuck Top / Auto Bottom w/perf |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | 5AA911TY | Dove Baby Wipes Rich Moisture 30ct | Dove Baby Wipes Rich Moisture 30Ct. | Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | 80 |  | FDA,0 |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | 5AA911TY | Dove Baby Wipes Rich Moisture 30ct | Dove Baby Wipes Rich Moisture 30Ct. | Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | 5AA911TY | Dove Baby Wipes Rich Moisture 30ct | Dove Baby Wipes Rich Moisture 30Ct. | Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | NY-Analytical | 772 | Box Inspection | Caliper | 0.0195 | in | FDA,3 | 0.019 | 0.021 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | 5AA911TY | Dove Baby Wipes Rich Moisture 30ct | Dove Baby Wipes Rich Moisture 30Ct. | Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | NY-Analytical | 772 | Box Inspection | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | 5AA911TY | Dove Baby Wipes Rich Moisture 30ct | Dove Baby Wipes Rich Moisture 30Ct. | Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | NY-Analytical | 772 | Box Inspection | Height | 7.94 | in | FDA,5 | 7.9062 | 7.9688 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | 5AA911TY | Dove Baby Wipes Rich Moisture 30ct | Dove Baby Wipes Rich Moisture 30Ct. | Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | NY-Analytical | 772 | Box Inspection | Length | 4.51 | in | FDA,5 | 4.4688 | 4.5312 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | 5AA911TY | Dove Baby Wipes Rich Moisture 30ct | Dove Baby Wipes Rich Moisture 30Ct. | Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | NY-Analytical | 772 | Box Inspection | Width | 7.94 | in | FDA,5 | 7.9062 | 7.9688 | Certificate of Conformance | COA | NY | RAW MATERIAL - NY | RM_BOX | both
5AB123CP/Label_NY_Skip lot testing | 5AB123CP | Clear Hang Tab w/Hang Tite 203A Clear Adhesive | Clear Hang Tab W/Hang Tite 203A Clear Adhesive | Label_NY_Skip lot testing | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Certificate of Conformance | COA | NY | RAW MATERIAL - NY | RM_BOX | both
5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | 5AB123CP | Clear Hang Tab w/Hang Tite 203A Clear Adhesive | Clear Hang Tab W/Hang Tite 203A Clear Adhesive | Label_NY_5AB123CP_1.9375-1.875 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | [5 empty] | Label Inspection |  | NY | RAW MATERIAL - NY | RM_BOX | both
5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | 5AB123CP | Clear Hang Tab w/Hang Tite 203A Clear Adhesive | Clear Hang Tab W/Hang Tite 203A Clear Adhesive | Label_NY_5AB123CP_1.9375-1.875 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_BOX | both
5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | 5AB123CP | Clear Hang Tab w/Hang Tite 203A Clear Adhesive | Clear Hang Tab W/Hang Tite 203A Clear Adhesive | Label_NY_5AB123CP_1.9375-1.875 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_BOX | both
5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | 5AB123CP | Clear Hang Tab w/Hang Tite 203A Clear Adhesive | Clear Hang Tab W/Hang Tite 203A Clear Adhesive | Label_NY_5AB123CP_1.9375-1.875 | NY-Analytical | 770 | Label Inspection | Length | 1.875 | in | FDA,4 | 1.875 | 2 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_BOX | both
5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | 5AB123CP | Clear Hang Tab w/Hang Tite 203A Clear Adhesive | Clear Hang Tab W/Hang Tite 203A Clear Adhesive | Label_NY_5AB123CP_1.9375-1.875 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_BOX | both
5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | 5AB123CP | Clear Hang Tab w/Hang Tite 203A Clear Adhesive | Clear Hang Tab W/Hang Tite 203A Clear Adhesive | Label_NY_5AB123CP_1.9375-1.875 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_BOX | both
5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | 5AB123CP | Clear Hang Tab w/Hang Tite 203A Clear Adhesive | Clear Hang Tab W/Hang Tite 203A Clear Adhesive | Label_NY_5AB123CP_1.9375-1.875 | NY-Analytical | 770 | Label Inspection | Width | 1.9375 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_BOX | both
5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | 5AB123ST | PDIPrevantics Device Swab Box | Prevantics Device Swab Box | Box_NY_0.018SBS_107_10.25-2.375-2.75 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | 5AB123ST | PDIPrevantics Device Swab Box | Prevantics Device Swab Box | Box_NY_0.018SBS_107_10.25-2.375-2.75 | NY-Analytical | 772 | Box Inspection | Caliper | 0.018 | in | FDA,3 | 0.017 | 0.019 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | 5AB123ST | PDIPrevantics Device Swab Box | Prevantics Device Swab Box | Box_NY_0.018SBS_107_10.25-2.375-2.75 | NY-Analytical | 772 | Box Inspection | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | 5AB123ST | PDIPrevantics Device Swab Box | Prevantics Device Swab Box | Box_NY_0.018SBS_107_10.25-2.375-2.75 | NY-Analytical | 772 | Box Inspection | Carton Style | TT/AB w/friction lock (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | 5AB123ST | PDIPrevantics Device Swab Box | Prevantics Device Swab Box | Box_NY_0.018SBS_107_10.25-2.375-2.75 | NY-Analytical | 772 | Box Inspection | Height | 2.7515 | in | FDA,5 | 2.7188 | 2.7812 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | 5AB123ST | PDIPrevantics Device Swab Box | Prevantics Device Swab Box | Box_NY_0.018SBS_107_10.25-2.375-2.75 | NY-Analytical | 772 | Box Inspection | Length | 10.22 | in | FDA,5 | 10.2188 | 10.2812 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | 5AB123ST | PDIPrevantics Device Swab Box | Prevantics Device Swab Box | Box_NY_0.018SBS_107_10.25-2.375-2.75 | NY-Analytical | 772 | Box Inspection | Width | 2.351 | in | FDA,5 | 2.3438 | 2.4062 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | 5AD02424 | Wet Nap Antibacterial Hand Wipes - Fresh | Wet Nap AntibaCt. erial Hand Wipes - Fresh | Box_NY_0.018SBS_107_4-2.03125-4.5 | NY-Analytical | 772 | Box Inspection | Carton Style | Tuck Top Auto Bottom w/ Tab Lock |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | 5AD02424 | Wet Nap Antibacterial Hand Wipes - Fresh | Wet Nap AntibaCt. erial Hand Wipes - Fresh | Box_NY_0.018SBS_107_4-2.03125-4.5 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | 5AD02424 | Wet Nap Antibacterial Hand Wipes - Fresh | Wet Nap AntibaCt. erial Hand Wipes - Fresh | Box_NY_0.018SBS_107_4-2.03125-4.5 | NY-Analytical | 772 | Box Inspection | Caliper | 0.0175 | in | FDA,3 | 0.017 | 0.019 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | 5AD02424 | Wet Nap Antibacterial Hand Wipes - Fresh | Wet Nap AntibaCt. erial Hand Wipes - Fresh | Box_NY_0.018SBS_107_4-2.03125-4.5 | NY-Analytical | 772 | Box Inspection | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | 5AD02424 | Wet Nap Antibacterial Hand Wipes - Fresh | Wet Nap AntibaCt. erial Hand Wipes - Fresh | Box_NY_0.018SBS_107_4-2.03125-4.5 | NY-Analytical | 772 | Box Inspection | Height | 4.5 | in | FDA,5 | 4.4688 | 4.5312 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | 5AD02424 | Wet Nap Antibacterial Hand Wipes - Fresh | Wet Nap AntibaCt. erial Hand Wipes - Fresh | Box_NY_0.018SBS_107_4-2.03125-4.5 | NY-Analytical | 772 | Box Inspection | Length | 4 | in | FDA,5 | 3.9688 | 4.0312 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | 5AD02424 | Wet Nap Antibacterial Hand Wipes - Fresh | Wet Nap AntibaCt. erial Hand Wipes - Fresh | Box_NY_0.018SBS_107_4-2.03125-4.5 | NY-Analytical | 772 | Box Inspection | Width | 2.0312 | in | FDA,5 | 2 | 2.0625 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | 5AH45016 | CVS Health Cleansing Cloths 16ct Box | CVS Health Cleansing Cloths 16Ct. Box | Box_NY_0.018SBS_101_3.875-2-3.5 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | 200 |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | 5AH45016 | CVS Health Cleansing Cloths 16ct Box | CVS Health Cleansing Cloths 16Ct. Box | Box_NY_0.018SBS_101_3.875-2-3.5 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | 5AH45016 | CVS Health Cleansing Cloths 16ct Box | CVS Health Cleansing Cloths 16Ct. Box | Box_NY_0.018SBS_101_3.875-2-3.5 | NY-Analytical | 772 | Box Inspection | Caliper | 0.018 | in | FDA,3 | 0.017 | 0.019 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | 5AH45016 | CVS Health Cleansing Cloths 16ct Box | CVS Health Cleansing Cloths 16Ct. Box | Box_NY_0.018SBS_101_3.875-2-3.5 | NY-Analytical | 772 | Box Inspection | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | 5AH45016 | CVS Health Cleansing Cloths 16ct Box | CVS Health Cleansing Cloths 16Ct. Box | Box_NY_0.018SBS_101_3.875-2-3.5 | NY-Analytical | 772 | Box Inspection | Carton Style | French Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | 5AH45016 | CVS Health Cleansing Cloths 16ct Box | CVS Health Cleansing Cloths 16Ct. Box | Box_NY_0.018SBS_101_3.875-2-3.5 | NY-Analytical | 772 | Box Inspection | Height | 3.5 | in | FDA,5 | 3.4688 | 3.5312 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | 5AH45016 | CVS Health Cleansing Cloths 16ct Box | CVS Health Cleansing Cloths 16Ct. Box | Box_NY_0.018SBS_101_3.875-2-3.5 | NY-Analytical | 772 | Box Inspection | Length | 3.875 | in | FDA,5 | 3.8438 | 3.9062 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | 5AH45016 | CVS Health Cleansing Cloths 16ct Box | CVS Health Cleansing Cloths 16Ct. Box | Box_NY_0.018SBS_101_3.875-2-3.5 | NY-Analytical | 772 | Box Inspection | Width | 2 | in | FDA,5 | 1.9688 | 2.0312 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AH58195/Box_NY_0.020SBS_127_6.25-2.5625-4.625 | 5AH58195 | PDI Sani Cloth Bleach | Sani-Cloth Bleach | Box_NY_0.020SBS_127_6.25-2.5625-4.625 | NY-Analytical | 772 | Box Inspection | Carton Style | Straight Tuck w/Partition Dispenser Box |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AH58195/Box_NY_0.020SBS_127_6.25-2.5625-4.625 | 5AH58195 | PDI Sani Cloth Bleach | Sani-Cloth Bleach | Box_NY_0.020SBS_127_6.25-2.5625-4.625 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AH58195/Box_NY_0.020SBS_127_6.25-2.5625-4.625 | 5AH58195 | PDI Sani Cloth Bleach | Sani-Cloth Bleach | Box_NY_0.020SBS_127_6.25-2.5625-4.625 | NY-Analytical | 772 | Box Inspection | Caliper | 0.0195 | in | FDA,3 | 0.019 | 0.021 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AH58195/Box_NY_0.020SBS_127_6.25-2.5625-4.625 | 5AH58195 | PDI Sani Cloth Bleach | Sani-Cloth Bleach | Box_NY_0.020SBS_127_6.25-2.5625-4.625 | NY-Analytical | 772 | Box Inspection | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AH58195/Box_NY_0.020SBS_127_6.25-2.5625-4.625 | 5AH58195 | PDI Sani Cloth Bleach | Sani-Cloth Bleach | Box_NY_0.020SBS_127_6.25-2.5625-4.625 | NY-Analytical | 772 | Box Inspection | Height | 4.625 | in | FDA,5 | 4.5938 | 4.6562 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AH58195/Box_NY_0.020SBS_127_6.25-2.5625-4.625 | 5AH58195 | PDI Sani Cloth Bleach | Sani-Cloth Bleach | Box_NY_0.020SBS_127_6.25-2.5625-4.625 | NY-Analytical | 772 | Box Inspection | Length | 6.25 | in | FDA,5 | 6.2188 | 6.2812 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AH58195/Box_NY_0.020SBS_127_6.25-2.5625-4.625 | 5AH58195 | PDI Sani Cloth Bleach | Sani-Cloth Bleach | Box_NY_0.020SBS_127_6.25-2.5625-4.625 | NY-Analytical | 772 | Box Inspection | Width | 2.5625 | in | FDA,5 | 2.5312 | 2.5938 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AJ130BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | 5AJ130BT | Cetaphil Gentle Skin Cleansing Cloths Boot | Cetaphil Gentle Skin Cleansing Cloths Boot | Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | NY-Analytical | 772 | Box Inspection | AQL Sample Size |  |  | FDA,0 |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AJ130BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | 5AJ130BT | Cetaphil Gentle Skin Cleansing Cloths Boot | Cetaphil Gentle Skin Cleansing Cloths Boot | Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | [5 empty] | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AJ130BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | 5AJ130BT | Cetaphil Gentle Skin Cleansing Cloths Boot | Cetaphil Gentle Skin Cleansing Cloths Boot | Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | NY-Analytical | 772 | Box Inspection | Caliper |  | in | FDA,3 | 0.019 | 0.021 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AJ130BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | 5AJ130BT | Cetaphil Gentle Skin Cleansing Cloths Boot | Cetaphil Gentle Skin Cleansing Cloths Boot | Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | NY-Analytical | 772 | Box Inspection | Carton Material | [5 empty] | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AJ130BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | 5AJ130BT | Cetaphil Gentle Skin Cleansing Cloths Boot | Cetaphil Gentle Skin Cleansing Cloths Boot | Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | NY-Analytical | 772 | Box Inspection | Carton Style | [5 empty] | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AJ130BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | 5AJ130BT | Cetaphil Gentle Skin Cleansing Cloths Boot | Cetaphil Gentle Skin Cleansing Cloths Boot | Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | NY-Analytical | 772 | Box Inspection | Height |  | in | FDA,5 | 2.5938 | 2.6562 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AJ130BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | 5AJ130BT | Cetaphil Gentle Skin Cleansing Cloths Boot | Cetaphil Gentle Skin Cleansing Cloths Boot | Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | NY-Analytical | 772 | Box Inspection | Length |  | in | FDA,5 | 3.2812 | 3.3438 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AJ130BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | 5AJ130BT | Cetaphil Gentle Skin Cleansing Cloths Boot | Cetaphil Gentle Skin Cleansing Cloths Boot | Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | NY-Analytical | 772 | Box Inspection | Width |  | in | FDA,5 | 1.4062 | 1.4688 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | 5AP07530 | L 'Oreal Revitalift Makeup Removing Cleans Twlts. | L'Oreal Revitalift Makeup Removing Cleans Twlts. | Box_NY_0.020BC_110_6.375-3.4375-4.75 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | 125 |  | FDA,0 |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | 5AP07530 | L 'Oreal Revitalift Makeup Removing Cleans Twlts. | L'Oreal Revitalift Makeup Removing Cleans Twlts. | Box_NY_0.020BC_110_6.375-3.4375-4.75 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | 5AP07530 | L 'Oreal Revitalift Makeup Removing Cleans Twlts. | L'Oreal Revitalift Makeup Removing Cleans Twlts. | Box_NY_0.020BC_110_6.375-3.4375-4.75 | NY-Analytical | 772 | Box Inspection | Caliper | 0.02 | in | FDA,3 | 0.019 | 0.021 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | 5AP07530 | L 'Oreal Revitalift Makeup Removing Cleans Twlts. | L'Oreal Revitalift Makeup Removing Cleans Twlts. | Box_NY_0.020BC_110_6.375-3.4375-4.75 | NY-Analytical | 772 | Box Inspection | Carton Material | BC (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | 5AP07530 | L 'Oreal Revitalift Makeup Removing Cleans Twlts. | L'Oreal Revitalift Makeup Removing Cleans Twlts. | Box_NY_0.020BC_110_6.375-3.4375-4.75 | NY-Analytical | 772 | Box Inspection | Carton Style | TT/AB w/finger holes (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | 5AP07530 | L 'Oreal Revitalift Makeup Removing Cleans Twlts. | L'Oreal Revitalift Makeup Removing Cleans Twlts. | Box_NY_0.020BC_110_6.375-3.4375-4.75 | NY-Analytical | 772 | Box Inspection | Height | 4.771 | in | FDA,5 | 4.7188 | 4.7812 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | 5AP07530 | L 'Oreal Revitalift Makeup Removing Cleans Twlts. | L'Oreal Revitalift Makeup Removing Cleans Twlts. | Box_NY_0.020BC_110_6.375-3.4375-4.75 | NY-Analytical | 772 | Box Inspection | Length | 6.395 | in | FDA,5 | 6.3438 | 6.4062 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | 5AP07530 | L 'Oreal Revitalift Makeup Removing Cleans Twlts. | L'Oreal Revitalift Makeup Removing Cleans Twlts. | Box_NY_0.020BC_110_6.375-3.4375-4.75 | NY-Analytical | 772 | Box Inspection | Width | 3.4225 | in | FDA,5 | 3.4062 | 3.4688 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | 5AP209SL | 3X25ct L 'Oreal Tray Sleeve | L'Oreal Tray Sleeve 3X25Ct. | Box_NY_0.020BC_220_5.09375-3.6875-4.625 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | 30 |  | FDA,0 |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | 5AP209SL | 3X25ct L 'Oreal Tray Sleeve | L'Oreal Tray Sleeve 3X25Ct. | Box_NY_0.020BC_220_5.09375-3.6875-4.625 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | 5AP209SL | 3X25ct L 'Oreal Tray Sleeve | L'Oreal Tray Sleeve 3X25Ct. | Box_NY_0.020BC_220_5.09375-3.6875-4.625 | NY-Analytical | 772 | Box Inspection | Caliper | 0.0195 | in | FDA,3 | 0.019 | 0.021 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | 5AP209SL | 3X25ct L 'Oreal Tray Sleeve | L'Oreal Tray Sleeve 3X25Ct. | Box_NY_0.020BC_220_5.09375-3.6875-4.625 | NY-Analytical | 772 | Box Inspection | Carton Material | BC (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | 5AP209SL | 3X25ct L 'Oreal Tray Sleeve | L'Oreal Tray Sleeve 3X25Ct. | Box_NY_0.020BC_220_5.09375-3.6875-4.625 | NY-Analytical | 772 | Box Inspection | Carton Style | Diecut Glued Sleeve w/ 2-1" fingerholes (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | 5AP209SL | 3X25ct L 'Oreal Tray Sleeve | L'Oreal Tray Sleeve 3X25Ct. | Box_NY_0.020BC_220_5.09375-3.6875-4.625 | NY-Analytical | 772 | Box Inspection | Height | 4.61 | in | FDA,5 | 4.5938 | 4.6562 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | 5AP209SL | 3X25ct L 'Oreal Tray Sleeve | L'Oreal Tray Sleeve 3X25Ct. | Box_NY_0.020BC_220_5.09375-3.6875-4.625 | NY-Analytical | 772 | Box Inspection | Length | 5.07 | in | FDA,5 | 5.0625 | 5.125 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | 5AP209SL | 3X25ct L 'Oreal Tray Sleeve | L'Oreal Tray Sleeve 3X25Ct. | Box_NY_0.020BC_220_5.09375-3.6875-4.625 | NY-Analytical | 772 | Box Inspection | Width | 3.7 | in | FDA,5 | 3.6562 | 3.7188 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | 5AP227TY | Garnier Skin Active Tray | Garnier Skin ACt. ive Tray | Box_NY_0.020SBS_300_4.625-5-3.625 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | 125 |  | FDA,0 |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | 5AP227TY | Garnier Skin Active Tray | Garnier Skin ACt. ive Tray | Box_NY_0.020SBS_300_4.625-5-3.625 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | 5AP227TY | Garnier Skin Active Tray | Garnier Skin ACt. ive Tray | Box_NY_0.020SBS_300_4.625-5-3.625 | NY-Analytical | 772 | Box Inspection | Caliper | 0.02 | in | FDA,3 | 0.019 | 0.021 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | 5AP227TY | Garnier Skin Active Tray | Garnier Skin ACt. ive Tray | Box_NY_0.020SBS_300_4.625-5-3.625 | NY-Analytical | 772 | Box Inspection | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | 5AP227TY | Garnier Skin Active Tray | Garnier Skin ACt. ive Tray | Box_NY_0.020SBS_300_4.625-5-3.625 | NY-Analytical | 772 | Box Inspection | Carton Style | Open Top/Auto Bottom Tray (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | 5AP227TY | Garnier Skin Active Tray | Garnier Skin ACt. ive Tray | Box_NY_0.020SBS_300_4.625-5-3.625 | NY-Analytical | 772 | Box Inspection | Height | 3.61 | in | FDA,5 | 3.5938 | 3.6562 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | 5AP227TY | Garnier Skin Active Tray | Garnier Skin ACt. ive Tray | Box_NY_0.020SBS_300_4.625-5-3.625 | NY-Analytical | 772 | Box Inspection | Length | 4.6 | in | FDA,5 | 4.5938 | 4.6562 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | 5AP227TY | Garnier Skin Active Tray | Garnier Skin ACt. ive Tray | Box_NY_0.020SBS_300_4.625-5-3.625 | NY-Analytical | 772 | Box Inspection | Width | 5.01 | in | FDA,5 | 4.9688 | 5.0312 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AP308TY/Box_63A | 5AP308TY | Plastic Wipes Try | Plastic Wipes Try | Box_63A | NY-Analytical | 772 | Box Inspection | Appearance | Meets specification (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AP308TY/Box_63A | 5AP308TY | Plastic Wipes Try | Plastic Wipes Try | Box_63A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AP308TY/Box_63A | 5AP308TY | Plastic Wipes Try | Plastic Wipes Try | Box_63A | NY-Analytical | 772 | Box Inspection | Color | Clear (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AP308TY/Box_63A | 5AP308TY | Plastic Wipes Try | Plastic Wipes Try | Box_63A | NY-Analytical | 772 | Box Inspection | Flange - Sides | 0.075 | in | FDA,3 | 0.045 | 0.135 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AP308TY/Box_63A | 5AP308TY | Plastic Wipes Try | Plastic Wipes Try | Box_63A | NY-Analytical | 772 | Box Inspection | Flange - Top & Bottom | 0.099 | in | FDA,3 | 0.045 | 0.135 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AP308TY/Box_63A | 5AP308TY | Plastic Wipes Try | Plastic Wipes Try | Box_63A | NY-Analytical | 772 | Box Inspection | Length | 7.2 | in | FDA,3 | 7.156 | 7.22 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AP308TY/Box_63A | 5AP308TY | Plastic Wipes Try | Plastic Wipes Try | Box_63A | NY-Analytical | 772 | Box Inspection | Odor | None Detectable (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AP308TY/Box_63A | 5AP308TY | Plastic Wipes Try | Plastic Wipes Try | Box_63A | NY-Analytical | 772 | Box Inspection | Total Height | 2.95 | in | FDA,3 | 2.899 | 2.963 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AP308TY/Box_63A | 5AP308TY | Plastic Wipes Try | Plastic Wipes Try | Box_63A | NY-Analytical | 772 | Box Inspection | Width | 6.1 | in | FDA,3 | 6.093 | 6.157 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | 5AU25206 | L'Oreal Paris Sublime Bronze Self Tan Twlt | L'Oreal Paris Sublime Bronze Self Tan Twlt | Box_NY_0.018SBS_100_101.6-44.45-101.6 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | 200 |  | FDA,0 |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | 5AU25206 | L'Oreal Paris Sublime Bronze Self Tan Twlt | L'Oreal Paris Sublime Bronze Self Tan Twlt | Box_NY_0.018SBS_100_101.6-44.45-101.6 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | 5AU25206 | L'Oreal Paris Sublime Bronze Self Tan Twlt | L'Oreal Paris Sublime Bronze Self Tan Twlt | Box_NY_0.018SBS_100_101.6-44.45-101.6 | NY-Analytical | 772 | Box Inspection | Caliper | 0.018 | in | FDA,3 | 0.017 | 0.019 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | 5AU25206 | L'Oreal Paris Sublime Bronze Self Tan Twlt | L'Oreal Paris Sublime Bronze Self Tan Twlt | Box_NY_0.018SBS_100_101.6-44.45-101.6 | NY-Analytical | 772 | Box Inspection | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | 5AU25206 | L'Oreal Paris Sublime Bronze Self Tan Twlt | L'Oreal Paris Sublime Bronze Self Tan Twlt | Box_NY_0.018SBS_100_101.6-44.45-101.6 | NY-Analytical | 772 | Box Inspection | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | 5AU25206 | L'Oreal Paris Sublime Bronze Self Tan Twlt | L'Oreal Paris Sublime Bronze Self Tan Twlt | Box_NY_0.018SBS_100_101.6-44.45-101.6 | NY-Analytical | 772 | Box Inspection | Height | 102 | mm | FDA,1 | 100.8 | 102.4 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | 5AU25206 | L'Oreal Paris Sublime Bronze Self Tan Twlt | L'Oreal Paris Sublime Bronze Self Tan Twlt | Box_NY_0.018SBS_100_101.6-44.45-101.6 | NY-Analytical | 772 | Box Inspection | Length | 102 | mm | FDA,1 | 100.8 | 102.4 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | 5AU25206 | L'Oreal Paris Sublime Bronze Self Tan Twlt | L'Oreal Paris Sublime Bronze Self Tan Twlt | Box_NY_0.018SBS_100_101.6-44.45-101.6 | NY-Analytical | 772 | Box Inspection | Width | 45 | mm | FDA,1 | 43.7 | 45.2 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | 5AU252IN | L' Oreal Paris Sublime Bronze Self Tan Twlts | L'Oreal Paris Sublime Bronze Self Tan Twlts | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | 125 |  | FDA,0 |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | 5AU252IN | L' Oreal Paris Sublime Bronze Self Tan Twlts | L'Oreal Paris Sublime Bronze Self Tan Twlts | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Does not meet specifications (OOS) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | 5AU252IN | L' Oreal Paris Sublime Bronze Self Tan Twlts | L'Oreal Paris Sublime Bronze Self Tan Twlts | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 772 | Box Inspection | Caliper | 0.019 | in | FDA,3 | 0.019 | 0.021 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | 5AU252IN | L' Oreal Paris Sublime Bronze Self Tan Twlts | L'Oreal Paris Sublime Bronze Self Tan Twlts | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 772 | Box Inspection | Carton Material | TBC (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | 5AU252IN | L' Oreal Paris Sublime Bronze Self Tan Twlts | L'Oreal Paris Sublime Bronze Self Tan Twlts | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 772 | Box Inspection | Carton Style | RT (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | 5AU252IN | L' Oreal Paris Sublime Bronze Self Tan Twlts | L'Oreal Paris Sublime Bronze Self Tan Twlts | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | [5 empty] | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | 5AU252IN | L' Oreal Paris Sublime Bronze Self Tan Twlts | L'Oreal Paris Sublime Bronze Self Tan Twlts | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 772 | Box Inspection | Height | 4.2812 | in | FDA,5 | 4.2812 | 4.3438 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | 5AU252IN | L' Oreal Paris Sublime Bronze Self Tan Twlts | L'Oreal Paris Sublime Bronze Self Tan Twlts | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 772 | Box Inspection | Length | 4.125 | in | FDA,5 | 4.0625 | 4.125 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | 5AU252IN | L' Oreal Paris Sublime Bronze Self Tan Twlts | L'Oreal Paris Sublime Bronze Self Tan Twlts | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 772 | Box Inspection | Width | 3.625 | in | FDA,5 | 3.5938 | 3.6562 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | 5AU31450 | Preparation H Totables Dispenser | Prep H Totables Dispenser | Box_NY_0.020SBS_169_6-3.125-5.9375 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | 5AU31450 | Preparation H Totables Dispenser | Prep H Totables Dispenser | Box_NY_0.020SBS_169_6-3.125-5.9375 | NY-Analytical | 772 | Box Inspection | Caliper | 0.019 | in | FDA,3 | 0.019 | 0.021 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | 5AU31450 | Preparation H Totables Dispenser | Prep H Totables Dispenser | Box_NY_0.020SBS_169_6-3.125-5.9375 | NY-Analytical | 772 | Box Inspection | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | 5AU31450 | Preparation H Totables Dispenser | Prep H Totables Dispenser | Box_NY_0.020SBS_169_6-3.125-5.9375 | NY-Analytical | 772 | Box Inspection | Carton Style | Dispenser (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | 5AU31450 | Preparation H Totables Dispenser | Prep H Totables Dispenser | Box_NY_0.020SBS_169_6-3.125-5.9375 | NY-Analytical | 772 | Box Inspection | Height | 5.9375 | in | FDA,5 | 5.9062 | 5.9688 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | 5AU31450 | Preparation H Totables Dispenser | Prep H Totables Dispenser | Box_NY_0.020SBS_169_6-3.125-5.9375 | NY-Analytical | 772 | Box Inspection | Length | 6 | in | FDA,5 | 5.9688 | 6.0312 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | 5AU31450 | Preparation H Totables Dispenser | Prep H Totables Dispenser | Box_NY_0.020SBS_169_6-3.125-5.9375 | NY-Analytical | 772 | Box Inspection | Width | 3.125 | in | FDA,5 | 3.0938 | 3.1562 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU44410/Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | 5AU44410 | Prep H Totables 10ct Box | Prep H Totables 10Ct. Box | Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU44410/Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | 5AU44410 | Prep H Totables 10ct Box | Prep H Totables 10Ct. Box | Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Caliper | 0.0155 | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU44410/Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | 5AU44410 | Prep H Totables 10ct Box | Prep H Totables 10Ct. Box | Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU44410/Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | 5AU44410 | Prep H Totables 10ct Box | Prep H Totables 10Ct. Box | Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Carton Style | Straight Tuck End w/1" glued double header (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU44410/Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | 5AU44410 | Prep H Totables 10ct Box | Prep H Totables 10Ct. Box | Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Height | 3.8125 | in | FDA,5 | 3.7812 | 3.8438 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU44410/Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | 5AU44410 | Prep H Totables 10ct Box | Prep H Totables 10Ct. Box | Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Length | 2.9375 | in | FDA,5 | 2.9062 | 2.9688 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU44410/Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | 5AU44410 | Prep H Totables 10ct Box | Prep H Totables 10Ct. Box | Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Width | 1.875 | in | FDA,5 | 1.8438 | 1.9062 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | 5AU57314 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | Box_NY_0.016CCSBS_100_3-2.5-4 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | 200 |  | FDA,0 |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | 5AU57314 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | Box_NY_0.016CCSBS_100_3-2.5-4 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | 5AU57314 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | Box_NY_0.016CCSBS_100_3-2.5-4 | NY-Analytical | 772 | Box Inspection | Caliper | 0.0155 | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | 5AU57314 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | Box_NY_0.016CCSBS_100_3-2.5-4 | NY-Analytical | 772 | Box Inspection | Carton Material | Clay Coated SBS (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | 5AU57314 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | Box_NY_0.016CCSBS_100_3-2.5-4 | NY-Analytical | 772 | Box Inspection | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | 5AU57314 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | Box_NY_0.016CCSBS_100_3-2.5-4 | NY-Analytical | 772 | Box Inspection | Height | 4 | in | FDA,5 | 3.9688 | 4.0312 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | 5AU57314 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | Box_NY_0.016CCSBS_100_3-2.5-4 | NY-Analytical | 772 | Box Inspection | Length | 3 | in | FDA,5 | 2.9688 | 3.0312 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | 5AU57314 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | Box_NY_0.016CCSBS_100_3-2.5-4 | NY-Analytical | 772 | Box Inspection | Width | 2.5 | in | FDA,5 | 2.4688 | 2.5312 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
| 5AU66522 | Preparation H Rapid Relief w/Lidocaine Totables |  | Analytical - Box_NY_0.016BMPB_122_3.625-3-3.8125 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection |  |  |  | left_only
| 5AU66522 | Preparation H Rapid Relief w/Lidocaine Totables |  | Analytical - Box_NY_0.016BMPB_122_3.625-3-3.8125 | NY-Analytical | 772 | Box Inspection | Caliper | 0.017 | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Box Inspection |  |  |  | left_only
| 5AU66522 | Preparation H Rapid Relief w/Lidocaine Totables |  | Analytical - Box_NY_0.016BMPB_122_3.625-3-3.8125 | NY-Analytical | 772 | Box Inspection | Carton Material | Bright Metalized Poly Board (Done) |  |  |  |  | Box Inspection | Box Inspection |  |  |  | left_only
| 5AU66522 | Preparation H Rapid Relief w/Lidocaine Totables |  | Analytical - Box_NY_0.016BMPB_122_3.625-3-3.8125 | NY-Analytical | 772 | Box Inspection | Carton Style | Straight Tuck End (Done) |  |  |  |  | Box Inspection | Box Inspection |  |  |  | left_only
| 5AU66522 | Preparation H Rapid Relief w/Lidocaine Totables |  | Analytical - Box_NY_0.016BMPB_122_3.625-3-3.8125 | NY-Analytical | 772 | Box Inspection | Height | 3.812 | in | FDA,5 | 3.7812 | 3.8438 | Box Inspection | Box Inspection |  |  |  | left_only
| 5AU66522 | Preparation H Rapid Relief w/Lidocaine Totables |  | Analytical - Box_NY_0.016BMPB_122_3.625-3-3.8125 | NY-Analytical | 772 | Box Inspection | Length | 3.625 | in | FDA,5 | 3.5938 | 3.6562 | Box Inspection | Box Inspection |  |  |  | left_only
| 5AU66522 | Preparation H Rapid Relief w/Lidocaine Totables |  | Analytical - Box_NY_0.016BMPB_122_3.625-3-3.8125 | NY-Analytical | 772 | Box Inspection | Width | 3 | in | FDA,5 | 2.9688 | 3.0312 | Box Inspection | Box Inspection |  |  |  | left_only
5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | 5BU19101 | Cetaphil Gentle Skin Cleans. Cloths Dispenser | Cetaphil Gentle Skin Cleans. Cloths Dispenser | Box_NY_0.020SBS_129_5.75-3-3.875 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | 5BU19101 | Cetaphil Gentle Skin Cleans. Cloths Dispenser | Cetaphil Gentle Skin Cleans. Cloths Dispenser | Box_NY_0.020SBS_129_5.75-3-3.875 | NY-Analytical | 772 | Box Inspection | Caliper | 0.0207 | in | FDA,3 | 0.019 | 0.021 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | 5BU19101 | Cetaphil Gentle Skin Cleans. Cloths Dispenser | Cetaphil Gentle Skin Cleans. Cloths Dispenser | Box_NY_0.020SBS_129_5.75-3-3.875 | NY-Analytical | 772 | Box Inspection | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | 5BU19101 | Cetaphil Gentle Skin Cleans. Cloths Dispenser | Cetaphil Gentle Skin Cleans. Cloths Dispenser | Box_NY_0.020SBS_129_5.75-3-3.875 | NY-Analytical | 772 | Box Inspection | Carton Style | Straight Tuck End French Style Dipsenser w/False Bottom |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | 5BU19101 | Cetaphil Gentle Skin Cleans. Cloths Dispenser | Cetaphil Gentle Skin Cleans. Cloths Dispenser | Box_NY_0.020SBS_129_5.75-3-3.875 | NY-Analytical | 772 | Box Inspection | Height | 3.875 | in | FDA,5 | 3.8438 | 3.9062 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | 5BU19101 | Cetaphil Gentle Skin Cleans. Cloths Dispenser | Cetaphil Gentle Skin Cleans. Cloths Dispenser | Box_NY_0.020SBS_129_5.75-3-3.875 | NY-Analytical | 772 | Box Inspection | Length | 5.75 | in | FDA,5 | 5.7188 | 5.7812 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | 5BU19101 | Cetaphil Gentle Skin Cleans. Cloths Dispenser | Cetaphil Gentle Skin Cleans. Cloths Dispenser | Box_NY_0.020SBS_129_5.75-3-3.875 | NY-Analytical | 772 | Box Inspection | Width | 3 | in | FDA,5 | 2.9688 | 3.0312 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU26595/Box_NY_0.024SBS_102_4.390625-4.390625-12 | 5BU26595 | PDI Sani Cloth Bleach 40ct Box | Sani-Cloth Bleach 40Ct. Box | Box_NY_0.024SBS_102_4.390625-4.390625-12 | NY-Analytical | 772 | Box Inspection | Carton Style | Tuck Top / Auto Bottom Dispenser Box |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU26595/Box_NY_0.024SBS_102_4.390625-4.390625-12 | 5BU26595 | PDI Sani Cloth Bleach 40ct Box | Sani-Cloth Bleach 40Ct. Box | Box_NY_0.024SBS_102_4.390625-4.390625-12 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU26595/Box_NY_0.024SBS_102_4.390625-4.390625-12 | 5BU26595 | PDI Sani Cloth Bleach 40ct Box | Sani-Cloth Bleach 40Ct. Box | Box_NY_0.024SBS_102_4.390625-4.390625-12 | NY-Analytical | 772 | Box Inspection | Caliper | 0.024 | in | FDA,3 | 0.023 | 0.025 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU26595/Box_NY_0.024SBS_102_4.390625-4.390625-12 | 5BU26595 | PDI Sani Cloth Bleach 40ct Box | Sani-Cloth Bleach 40Ct. Box | Box_NY_0.024SBS_102_4.390625-4.390625-12 | NY-Analytical | 772 | Box Inspection | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU26595/Box_NY_0.024SBS_102_4.390625-4.390625-12 | 5BU26595 | PDI Sani Cloth Bleach 40ct Box | Sani-Cloth Bleach 40Ct. Box | Box_NY_0.024SBS_102_4.390625-4.390625-12 | NY-Analytical | 772 | Box Inspection | Height | 12 | in | FDA,5 | 11.9688 | 12.0312 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU26595/Box_NY_0.024SBS_102_4.390625-4.390625-12 | 5BU26595 | PDI Sani Cloth Bleach 40ct Box | Sani-Cloth Bleach 40Ct. Box | Box_NY_0.024SBS_102_4.390625-4.390625-12 | NY-Analytical | 772 | Box Inspection | Length | 4.41 | in | FDA,5 | 4.3594 | 4.4219 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU26595/Box_NY_0.024SBS_102_4.390625-4.390625-12 | 5BU26595 | PDI Sani Cloth Bleach 40ct Box | Sani-Cloth Bleach 40Ct. Box | Box_NY_0.024SBS_102_4.390625-4.390625-12 | NY-Analytical | 772 | Box Inspection | Width | 4.39 | in | FDA,5 | 4.3594 | 4.4219 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU27500/Box_NY_0.024SBS_102_4.3906-4.3906-12 | 5BU27500 | PDI Sani-Cloth AF3 Box | Sani-Cloth AF3 Box | Box_NY_0.024SBS_102_4.3906-4.3906-12 | NY-Analytical | 772 | Box Inspection | Carton Style | Tuck Top / Auto Bottom Dispenser Box |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU27500/Box_NY_0.024SBS_102_4.3906-4.3906-12 | 5BU27500 | PDI Sani-Cloth AF3 Box | Sani-Cloth AF3 Box | Box_NY_0.024SBS_102_4.3906-4.3906-12 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU27500/Box_NY_0.024SBS_102_4.3906-4.3906-12 | 5BU27500 | PDI Sani-Cloth AF3 Box | Sani-Cloth AF3 Box | Box_NY_0.024SBS_102_4.3906-4.3906-12 | NY-Analytical | 772 | Box Inspection | Caliper | 0.024 | in | FDA,3 | 0.023 | 0.025 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU27500/Box_NY_0.024SBS_102_4.3906-4.3906-12 | 5BU27500 | PDI Sani-Cloth AF3 Box | Sani-Cloth AF3 Box | Box_NY_0.024SBS_102_4.3906-4.3906-12 | NY-Analytical | 772 | Box Inspection | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU27500/Box_NY_0.024SBS_102_4.3906-4.3906-12 | 5BU27500 | PDI Sani-Cloth AF3 Box | Sani-Cloth AF3 Box | Box_NY_0.024SBS_102_4.3906-4.3906-12 | NY-Analytical | 772 | Box Inspection | Height | 12 | in | FDA,5 | 11.9688 | 12.0312 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU27500/Box_NY_0.024SBS_102_4.3906-4.3906-12 | 5BU27500 | PDI Sani-Cloth AF3 Box | Sani-Cloth AF3 Box | Box_NY_0.024SBS_102_4.3906-4.3906-12 | NY-Analytical | 772 | Box Inspection | Length | 4.4 | in | FDA,5 | 4.3594 | 4.4219 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU27500/Box_NY_0.024SBS_102_4.3906-4.3906-12 | 5BU27500 | PDI Sani-Cloth AF3 Box | Sani-Cloth AF3 Box | Box_NY_0.024SBS_102_4.3906-4.3906-12 | NY-Analytical | 772 | Box Inspection | Width | 4.41 | in | FDA,5 | 4.3594 | 4.4219 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU37230/Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | 5BU37230 | Harmon Smart Values Flushable Cleans Cloth. | Harmon Smart Values Flushable Cleans Cloth. | Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | 125 |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU37230/Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | 5BU37230 | Harmon Smart Values Flushable Cleans Cloth. | Harmon Smart Values Flushable Cleans Cloth. | Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU37230/Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | 5BU37230 | Harmon Smart Values Flushable Cleans Cloth. | Harmon Smart Values Flushable Cleans Cloth. | Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | NY-Analytical | 772 | Box Inspection | Caliper | 0.017 | in | FDA,3 | 0.017 | 0.019 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU37230/Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | 5BU37230 | Harmon Smart Values Flushable Cleans Cloth. | Harmon Smart Values Flushable Cleans Cloth. | Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | NY-Analytical | 772 | Box Inspection | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU37230/Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | 5BU37230 | Harmon Smart Values Flushable Cleans Cloth. | Harmon Smart Values Flushable Cleans Cloth. | Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | NY-Analytical | 772 | Box Inspection | Carton Style | Tuck Top / Auto Bottom - French Style (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU37230/Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | 5BU37230 | Harmon Smart Values Flushable Cleans Cloth. | Harmon Smart Values Flushable Cleans Cloth. | Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | NY-Analytical | 772 | Box Inspection | Height | 2.93 | in | FDA,5 | 2.9062 | 2.9688 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU37230/Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | 5BU37230 | Harmon Smart Values Flushable Cleans Cloth. | Harmon Smart Values Flushable Cleans Cloth. | Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | NY-Analytical | 772 | Box Inspection | Length | 3.84 | in | FDA,5 | 3.7812 | 3.8438 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU37230/Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | 5BU37230 | Harmon Smart Values Flushable Cleans Cloth. | Harmon Smart Values Flushable Cleans Cloth. | Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | NY-Analytical | 772 | Box Inspection | Width | 3.38 | in | FDA,5 | 3.3438 | 3.4062 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | 5BU55510 | Preparation H Rapid Relief w/Lidocaine Totables | Prep H Rapid Relief W/Lidocaine Totables | Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | 5BU55510 | Preparation H Rapid Relief w/Lidocaine Totables | Prep H Rapid Relief W/Lidocaine Totables | Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Caliper | 0.0155 | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | 5BU55510 | Preparation H Rapid Relief w/Lidocaine Totables | Prep H Rapid Relief W/Lidocaine Totables | Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Carton Material | Bright Metalized Poly Board (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | 5BU55510 | Preparation H Rapid Relief w/Lidocaine Totables | Prep H Rapid Relief W/Lidocaine Totables | Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Carton Style | Straight Tuck End w/glued header (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | 5BU55510 | Preparation H Rapid Relief w/Lidocaine Totables | Prep H Rapid Relief W/Lidocaine Totables | Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Height | 3.8125 | in | FDA,5 | 3.7812 | 3.8438 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | 5BU55510 | Preparation H Rapid Relief w/Lidocaine Totables | Prep H Rapid Relief W/Lidocaine Totables | Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Length | 2.9375 | in | FDA,5 | 2.9062 | 2.9688 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | 5BU55510 | Preparation H Rapid Relief w/Lidocaine Totables | Prep H Rapid Relief W/Lidocaine Totables | Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Width | 1.875 | in | FDA,5 | 1.8438 | 1.9062 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | 5BU71230 | Preparation H Rapid Relief Totables w/Lidocaine | Prep H Rapid Relief Totables W/Lidocaine | Box_NY_0.018SBS_122_4.6875-3-3.8125 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | 5BU71230 | Preparation H Rapid Relief Totables w/Lidocaine | Prep H Rapid Relief Totables W/Lidocaine | Box_NY_0.018SBS_122_4.6875-3-3.8125 | NY-Analytical | 772 | Box Inspection | Caliper | 0.018 | in | FDA,3 | 0.017 | 0.019 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | 5BU71230 | Preparation H Rapid Relief Totables w/Lidocaine | Prep H Rapid Relief Totables W/Lidocaine | Box_NY_0.018SBS_122_4.6875-3-3.8125 | NY-Analytical | 772 | Box Inspection | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | 5BU71230 | Preparation H Rapid Relief Totables w/Lidocaine | Prep H Rapid Relief Totables W/Lidocaine | Box_NY_0.018SBS_122_4.6875-3-3.8125 | NY-Analytical | 772 | Box Inspection | Carton Style | Straight Tuck End (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | 5BU71230 | Preparation H Rapid Relief Totables w/Lidocaine | Prep H Rapid Relief Totables W/Lidocaine | Box_NY_0.018SBS_122_4.6875-3-3.8125 | NY-Analytical | 772 | Box Inspection | Height | 3.82 | in | FDA,5 | 3.7812 | 3.8438 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | 5BU71230 | Preparation H Rapid Relief Totables w/Lidocaine | Prep H Rapid Relief Totables W/Lidocaine | Box_NY_0.018SBS_122_4.6875-3-3.8125 | NY-Analytical | 772 | Box Inspection | Length | 4.7 | in | FDA,5 | 4.6562 | 4.7188 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | 5BU71230 | Preparation H Rapid Relief Totables w/Lidocaine | Prep H Rapid Relief Totables W/Lidocaine | Box_NY_0.018SBS_122_4.6875-3-3.8125 | NY-Analytical | 772 | Box Inspection | Width | 3.01 | in | FDA,5 | 2.9688 | 3.0312 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | 5CA284TR | Preparation H 48ct | Prep H 48Ct. | Box_NY_0.018SBS_111_6-5.3125-3.6875 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | 5CA284TR | Preparation H 48ct | Prep H 48Ct. | Box_NY_0.018SBS_111_6-5.3125-3.6875 | NY-Analytical | 772 | Box Inspection | Caliper | 0.0175 | in | FDA,3 | 0.017 | 0.019 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | 5CA284TR | Preparation H 48ct | Prep H 48Ct. | Box_NY_0.018SBS_111_6-5.3125-3.6875 | NY-Analytical | 772 | Box Inspection | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | 5CA284TR | Preparation H 48ct | Prep H 48Ct. | Box_NY_0.018SBS_111_6-5.3125-3.6875 | NY-Analytical | 772 | Box Inspection | Carton Style | Tuck Top / Auto Bottom French Style w/Perfs (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | 5CA284TR | Preparation H 48ct | Prep H 48Ct. | Box_NY_0.018SBS_111_6-5.3125-3.6875 | NY-Analytical | 772 | Box Inspection | Height | 3.6875 | in | FDA,5 | 3.6562 | 3.7188 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | 5CA284TR | Preparation H 48ct | Prep H 48Ct. | Box_NY_0.018SBS_111_6-5.3125-3.6875 | NY-Analytical | 772 | Box Inspection | Length | 6 | in | FDA,5 | 5.9688 | 6.0312 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | 5CA284TR | Preparation H 48ct | Prep H 48Ct. | Box_NY_0.018SBS_111_6-5.3125-3.6875 | NY-Analytical | 772 | Box Inspection | Width | 5.3125 | in | FDA,5 | 5.2812 | 5.3438 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 5CB19600 | PDI Prevantics Device Swab 100ct Box | Prevantics Device Swab 100Ct. Box | Box_NY_0.018SBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | 125 |  | FDA,0 |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 5CB19600 | PDI Prevantics Device Swab 100ct Box | Prevantics Device Swab 100Ct. Box | Box_NY_0.018SBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 5CB19600 | PDI Prevantics Device Swab 100ct Box | Prevantics Device Swab 100Ct. Box | Box_NY_0.018SBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Caliper | 0.018 | in | FDA,3 | 0.017 | 0.019 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 5CB19600 | PDI Prevantics Device Swab 100ct Box | Prevantics Device Swab 100Ct. Box | Box_NY_0.018SBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 5CB19600 | PDI Prevantics Device Swab 100ct Box | Prevantics Device Swab 100Ct. Box | Box_NY_0.018SBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 5CB19600 | PDI Prevantics Device Swab 100ct Box | Prevantics Device Swab 100Ct. Box | Box_NY_0.018SBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Height | 2.0938 | in | FDA,5 | 2.0625 | 2.125 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 5CB19600 | PDI Prevantics Device Swab 100ct Box | Prevantics Device Swab 100Ct. Box | Box_NY_0.018SBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Length | 5.125 | in | FDA,5 | 5.0938 | 5.1562 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | 5CB19600 | PDI Prevantics Device Swab 100ct Box | Prevantics Device Swab 100Ct. Box | Box_NY_0.018SBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Width | 2.25 | in | FDA,5 | 2.2188 | 2.2812 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | 5CB51200 | PDI Povidone-Iodine Prep Pad | Povidone-Iodine Prep Pad | Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | NY-Analytical | 772 | Box Inspection | Carton Material | UCSBS |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | 5CB51200 | PDI Povidone-Iodine Prep Pad | Povidone-Iodine Prep Pad | Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | 5CB51200 | PDI Povidone-Iodine Prep Pad | Povidone-Iodine Prep Pad | Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | NY-Analytical | 772 | Box Inspection | Caliper | 0.0155 | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | 5CB51200 | PDI Povidone-Iodine Prep Pad | Povidone-Iodine Prep Pad | Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | NY-Analytical | 772 | Box Inspection | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | 5CB51200 | PDI Povidone-Iodine Prep Pad | Povidone-Iodine Prep Pad | Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | NY-Analytical | 772 | Box Inspection | Height | 2.375 | in | FDA,5 | 2.3438 | 2.4062 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | 5CB51200 | PDI Povidone-Iodine Prep Pad | Povidone-Iodine Prep Pad | Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | NY-Analytical | 772 | Box Inspection | Length | 6.125 | in | FDA,5 | 6.0938 | 6.1562 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | 5CB51200 | PDI Povidone-Iodine Prep Pad | Povidone-Iodine Prep Pad | Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | NY-Analytical | 772 | Box Inspection | Width | 2.75 | in | FDA,5 | 2.7188 | 2.7812 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | 5CC80100 | Philips Sali-Wipes Sterile saline Wipes. | Philips Sali-Wipes Sterile Saline Wipes. | Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Carton Material | UCSBS |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | 5CC80100 | Philips Sali-Wipes Sterile saline Wipes. | Philips Sali-Wipes Sterile Saline Wipes. | Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | 80 |  | FDA,0 |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | 5CC80100 | Philips Sali-Wipes Sterile saline Wipes. | Philips Sali-Wipes Sterile Saline Wipes. | Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | 5CC80100 | Philips Sali-Wipes Sterile saline Wipes. | Philips Sali-Wipes Sterile Saline Wipes. | Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Caliper | 0.015 | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | 5CC80100 | Philips Sali-Wipes Sterile saline Wipes. | Philips Sali-Wipes Sterile Saline Wipes. | Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | 5CC80100 | Philips Sali-Wipes Sterile saline Wipes. | Philips Sali-Wipes Sterile Saline Wipes. | Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Height | 3.07 | in | FDA,5 | 3.0312 | 3.0938 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | 5CC80100 | Philips Sali-Wipes Sterile saline Wipes. | Philips Sali-Wipes Sterile Saline Wipes. | Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Length | 5.07 | in | FDA,5 | 5.0312 | 5.0938 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | 5CC80100 | Philips Sali-Wipes Sterile saline Wipes. | Philips Sali-Wipes Sterile Saline Wipes. | Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Width | 1.73 | in | FDA,5 | 1.7188 | 1.7812 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 5CD35185 | PDI Hygea BZK Antiseptic Towelette | PDI Hygea BZK Antiseptic Towelette | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 772 | Box Inspection | Carton Material | UCSBS |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 5CD35185 | PDI Hygea BZK Antiseptic Towelette | PDI Hygea BZK Antiseptic Towelette | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | 125 |  | FDA,0 |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 5CD35185 | PDI Hygea BZK Antiseptic Towelette | PDI Hygea BZK Antiseptic Towelette | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 5CD35185 | PDI Hygea BZK Antiseptic Towelette | PDI Hygea BZK Antiseptic Towelette | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 772 | Box Inspection | Caliper | 0.016 | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 5CD35185 | PDI Hygea BZK Antiseptic Towelette | PDI Hygea BZK Antiseptic Towelette | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 772 | Box Inspection | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 5CD35185 | PDI Hygea BZK Antiseptic Towelette | PDI Hygea BZK Antiseptic Towelette | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 772 | Box Inspection | Height | 6.078 | in | FDA,5 | 6.0312 | 6.0938 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 5CD35185 | PDI Hygea BZK Antiseptic Towelette | PDI Hygea BZK Antiseptic Towelette | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 772 | Box Inspection | Length | 6.5 | in | FDA,5 | 6.4688 | 6.5312 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 5CD35185 | PDI Hygea BZK Antiseptic Towelette | PDI Hygea BZK Antiseptic Towelette | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 772 | Box Inspection | Width | 2.312 | in | FDA,5 | 2.2812 | 2.3438 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 5CS42050 | PDI Povidone-Iodine Scrub 1's CDA | Povidone-Iodine Scrub 1'S Cda | Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | NY-Analytical | 772 | Box Inspection | Carton Material | UCSBS |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 5CS42050 | PDI Povidone-Iodine Scrub 1's CDA | Povidone-Iodine Scrub 1'S Cda | Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | 80 |  | FDA,0 |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 5CS42050 | PDI Povidone-Iodine Scrub 1's CDA | Povidone-Iodine Scrub 1'S Cda | Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 5CS42050 | PDI Povidone-Iodine Scrub 1's CDA | Povidone-Iodine Scrub 1'S Cda | Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | NY-Analytical | 772 | Box Inspection | Caliper | 0.016 | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 5CS42050 | PDI Povidone-Iodine Scrub 1's CDA | Povidone-Iodine Scrub 1'S Cda | Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | NY-Analytical | 772 | Box Inspection | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 5CS42050 | PDI Povidone-Iodine Scrub 1's CDA | Povidone-Iodine Scrub 1'S Cda | Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | NY-Analytical | 772 | Box Inspection | Height | 5.81 | in | FDA,5 | 5.7812 | 5.8438 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 5CS42050 | PDI Povidone-Iodine Scrub 1's CDA | Povidone-Iodine Scrub 1'S Cda | Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | NY-Analytical | 772 | Box Inspection | Length | 6.51 | in | FDA,5 | 6.4688 | 6.5312 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | 5CS42050 | PDI Povidone-Iodine Scrub 1's CDA | Povidone-Iodine Scrub 1'S Cda | Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | NY-Analytical | 772 | Box Inspection | Width | 3.32 | in | FDA,5 | 3.2812 | 3.3438 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5DA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | 5DA941TY | Target Up & Up Extra Large Cleans. Washcloths | Target Up & Up Extra Large Cleans. Washcloths | Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5DA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | 5DA941TY | Target Up & Up Extra Large Cleans. Washcloths | Target Up & Up Extra Large Cleans. Washcloths | Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | NY-Analytical | 772 | Box Inspection | Back Height | 8.625 | in | FDA,5 | 8.5938 | 8.6562 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5DA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | 5DA941TY | Target Up & Up Extra Large Cleans. Washcloths | Target Up & Up Extra Large Cleans. Washcloths | Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | NY-Analytical | 772 | Box Inspection | Caliper | 0.024 | in | FDA,3 | 0.023 | 0.025 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5DA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | 5DA941TY | Target Up & Up Extra Large Cleans. Washcloths | Target Up & Up Extra Large Cleans. Washcloths | Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | NY-Analytical | 772 | Box Inspection | Carton Material | WCK (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5DA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | 5DA941TY | Target Up & Up Extra Large Cleans. Washcloths | Target Up & Up Extra Large Cleans. Washcloths | Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | NY-Analytical | 772 | Box Inspection | Carton Style | Tuck Top / Auto Bottom French Style w/Perfs (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5DA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | 5DA941TY | Target Up & Up Extra Large Cleans. Washcloths | Target Up & Up Extra Large Cleans. Washcloths | Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | NY-Analytical | 772 | Box Inspection | Front Height | 3.5 | in | FDA,5 | 3.4688 | 3.5312 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5DA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | 5DA941TY | Target Up & Up Extra Large Cleans. Washcloths | Target Up & Up Extra Large Cleans. Washcloths | Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | NY-Analytical | 772 | Box Inspection | Length | 5.24 | in | FDA,5 | 5.2188 | 5.2812 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5DA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | 5DA941TY | Target Up & Up Extra Large Cleans. Washcloths | Target Up & Up Extra Large Cleans. Washcloths | Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | NY-Analytical | 772 | Box Inspection | Width | 6.26 | in | FDA,5 | 6.2188 | 6.2812 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | 5DH12824 | CVS Health Instant Hand Sanitizer Wipes | CVS Health Instant Hand Sanitizer Wipes | Box_NY_0.018SBS_100_4-2.125-6.125 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | 30 |  | FDA,0 |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | 5DH12824 | CVS Health Instant Hand Sanitizer Wipes | CVS Health Instant Hand Sanitizer Wipes | Box_NY_0.018SBS_100_4-2.125-6.125 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | 5DH12824 | CVS Health Instant Hand Sanitizer Wipes | CVS Health Instant Hand Sanitizer Wipes | Box_NY_0.018SBS_100_4-2.125-6.125 | NY-Analytical | 772 | Box Inspection | Caliper | 0.017 | in | FDA,3 | 0.017 | 0.019 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | 5DH12824 | CVS Health Instant Hand Sanitizer Wipes | CVS Health Instant Hand Sanitizer Wipes | Box_NY_0.018SBS_100_4-2.125-6.125 | NY-Analytical | 772 | Box Inspection | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | 5DH12824 | CVS Health Instant Hand Sanitizer Wipes | CVS Health Instant Hand Sanitizer Wipes | Box_NY_0.018SBS_100_4-2.125-6.125 | NY-Analytical | 772 | Box Inspection | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | 5DH12824 | CVS Health Instant Hand Sanitizer Wipes | CVS Health Instant Hand Sanitizer Wipes | Box_NY_0.018SBS_100_4-2.125-6.125 | NY-Analytical | 772 | Box Inspection | Height | 6.125 | in | FDA,5 | 6.0938 | 6.1562 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | 5DH12824 | CVS Health Instant Hand Sanitizer Wipes | CVS Health Instant Hand Sanitizer Wipes | Box_NY_0.018SBS_100_4-2.125-6.125 | NY-Analytical | 772 | Box Inspection | Length | 4 | in | FDA,5 | 3.9688 | 4.0312 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | 5DH12824 | CVS Health Instant Hand Sanitizer Wipes | CVS Health Instant Hand Sanitizer Wipes | Box_NY_0.018SBS_100_4-2.125-6.125 | NY-Analytical | 772 | Box Inspection | Width | 2.125 | in | FDA,5 | 2.0938 | 2.1562 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | 5EH21130 | CVS Total Home Antibac. Wipes 30ct Box | CVS Total Home Antibac. Wipes 30Ct. Box | Box_NY_0.016SBS_100_4.75-2.375-3.25 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | 80 |  | FDA,0 |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | 5EH21130 | CVS Total Home Antibac. Wipes 30ct Box | CVS Total Home Antibac. Wipes 30Ct. Box | Box_NY_0.016SBS_100_4.75-2.375-3.25 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | 5EH21130 | CVS Total Home Antibac. Wipes 30ct Box | CVS Total Home Antibac. Wipes 30Ct. Box | Box_NY_0.016SBS_100_4.75-2.375-3.25 | NY-Analytical | 772 | Box Inspection | Caliper | 0.015 | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | 5EH21130 | CVS Total Home Antibac. Wipes 30ct Box | CVS Total Home Antibac. Wipes 30Ct. Box | Box_NY_0.016SBS_100_4.75-2.375-3.25 | NY-Analytical | 772 | Box Inspection | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | 5EH21130 | CVS Total Home Antibac. Wipes 30ct Box | CVS Total Home Antibac. Wipes 30Ct. Box | Box_NY_0.016SBS_100_4.75-2.375-3.25 | NY-Analytical | 772 | Box Inspection | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | 5EH21130 | CVS Total Home Antibac. Wipes 30ct Box | CVS Total Home Antibac. Wipes 30Ct. Box | Box_NY_0.016SBS_100_4.75-2.375-3.25 | NY-Analytical | 772 | Box Inspection | Height | 3.26 | in | FDA,5 | 3.2188 | 3.2812 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | 5EH21130 | CVS Total Home Antibac. Wipes 30ct Box | CVS Total Home Antibac. Wipes 30Ct. Box | Box_NY_0.016SBS_100_4.75-2.375-3.25 | NY-Analytical | 772 | Box Inspection | Length | 4.75 | in | FDA,5 | 4.7188 | 4.7812 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | 5EH21130 | CVS Total Home Antibac. Wipes 30ct Box | CVS Total Home Antibac. Wipes 30Ct. Box | Box_NY_0.016SBS_100_4.75-2.375-3.25 | NY-Analytical | 772 | Box Inspection | Width | 2.36 | in | FDA,5 | 2.3438 | 2.4062 | Certificate of Analysis | COA | NY | RAW MATERIAL - NY | RM_BOX | both
5FD41900/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | 5FD41900 | PDI Castile Soap Towelette Box | PDI Castile Soap Towelette Box | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | 5FH07050 | EZ EM Mammo Wipes | EZ Em Mammo Wipes | Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | 5FH07050 | EZ EM Mammo Wipes | EZ Em Mammo Wipes | Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | NY-Analytical | 772 | Box Inspection | Caliper | 0.02 | in | FDA,3 | 0.019 | 0.021 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | 5FH07050 | EZ EM Mammo Wipes | EZ Em Mammo Wipes | Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | NY-Analytical | 772 | Box Inspection | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | 5FH07050 | EZ EM Mammo Wipes | EZ Em Mammo Wipes | Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | NY-Analytical | 772 | Box Inspection | Carton Style | Straight Tuck End w/Glued Internal Dividers&15/16" Dummy Platform Bottom (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | 5FH07050 | EZ EM Mammo Wipes | EZ Em Mammo Wipes | Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | NY-Analytical | 772 | Box Inspection | Height | 6 | in | FDA,5 | 5.9844 | 6.0469 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | 5FH07050 | EZ EM Mammo Wipes | EZ Em Mammo Wipes | Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | NY-Analytical | 772 | Box Inspection | Length | 6.14 | in | FDA,5 | 6.0938 | 6.1562 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | 5FH07050 | EZ EM Mammo Wipes | EZ Em Mammo Wipes | Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | NY-Analytical | 772 | Box Inspection | Width | 2.44 | in | FDA,5 | 2.4062 | 2.4688 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5GB10300/Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | 5GB10300 | PDI CHG Prep Pad Box | PDI CHG Prep Pad Box | Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | NY-Analytical | 772 | Box Inspection | Carton Material | UCSBS |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5GB10300/Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | 5GB10300 | PDI CHG Prep Pad Box | PDI CHG Prep Pad Box | Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5GB10300/Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | 5GB10300 | PDI CHG Prep Pad Box | PDI CHG Prep Pad Box | Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | NY-Analytical | 772 | Box Inspection | Caliper | 0.015 | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5GB10300/Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | 5GB10300 | PDI CHG Prep Pad Box | PDI CHG Prep Pad Box | Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | NY-Analytical | 772 | Box Inspection | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5GB10300/Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | 5GB10300 | PDI CHG Prep Pad Box | PDI CHG Prep Pad Box | Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | NY-Analytical | 772 | Box Inspection | Height | 2.375 | in | FDA,5 | 2.3438 | 2.4062 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5GB10300/Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | 5GB10300 | PDI CHG Prep Pad Box | PDI CHG Prep Pad Box | Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | NY-Analytical | 772 | Box Inspection | Length | 6.125 | in | FDA,5 | 6.0938 | 6.1562 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5GB10300/Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | 5GB10300 | PDI CHG Prep Pad Box | PDI CHG Prep Pad Box | Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | NY-Analytical | 772 | Box Inspection | Width | 2.75 | in | FDA,5 | 2.7188 | 2.7812 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 5HB59800 | PDI Electrode Prep Pad | PDI EleCt. rode Prep Pad | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | 125 |  | FDA,0 |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 5HB59800 | PDI Electrode Prep Pad | PDI EleCt. rode Prep Pad | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 5HB59800 | PDI Electrode Prep Pad | PDI EleCt. rode Prep Pad | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | Back Flip | 5.158 | in | FDA,5 | 4.6875 | 5.3125 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 5HB59800 | PDI Electrode Prep Pad | PDI EleCt. rode Prep Pad | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | Caliper | 0.001 | in | FDA,4 | 0.0008 | 0.0028 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 5HB59800 | PDI Electrode Prep Pad | PDI EleCt. rode Prep Pad | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | Carton Material | LDPE (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 5HB59800 | PDI Electrode Prep Pad | PDI EleCt. rode Prep Pad | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | Carton Style | Poly Bag w/fold over top on 2” wicketts (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 5HB59800 | PDI Electrode Prep Pad | PDI EleCt. rode Prep Pad | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | Gussett | 3.8 | in | FDA,5 | 3.6875 | 4.3125 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 5HB59800 | PDI Electrode Prep Pad | PDI EleCt. rode Prep Pad | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | Height | 5.76 | in | FDA,5 | 5.1875 | 5.8125 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5HB59800/Box_NY_0.0018LDPE_407_12-5.5-5-4 | 5HB59800 | PDI Electrode Prep Pad | PDI EleCt. rode Prep Pad | Box_NY_0.0018LDPE_407_12-5.5-5-4 | NY-Analytical | 772 | Box Inspection | Width | 11.812 | in | FDA,5 | 11.6875 | 12.3125 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | 5HC12400 | PDI Povidone-Iodine Prep Pad | Povidone-Iodine Prep Pad | Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | NY-Analytical | 772 | Box Inspection | Carton Material | UCSBS |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | 5HC12400 | PDI Povidone-Iodine Prep Pad | Povidone-Iodine Prep Pad | Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | 20 |  | FDA,0 |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | 5HC12400 | PDI Povidone-Iodine Prep Pad | Povidone-Iodine Prep Pad | Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | 5HC12400 | PDI Povidone-Iodine Prep Pad | Povidone-Iodine Prep Pad | Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | NY-Analytical | 772 | Box Inspection | Caliper | 0.0169 | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | 5HC12400 | PDI Povidone-Iodine Prep Pad | Povidone-Iodine Prep Pad | Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | NY-Analytical | 772 | Box Inspection | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | 5HC12400 | PDI Povidone-Iodine Prep Pad | Povidone-Iodine Prep Pad | Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | NY-Analytical | 772 | Box Inspection | Height | 3.0625 | in | FDA,5 | 3.0312 | 3.0938 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | 5HC12400 | PDI Povidone-Iodine Prep Pad | Povidone-Iodine Prep Pad | Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | NY-Analytical | 772 | Box Inspection | Length | 5.0625 | in | FDA,5 | 5.0312 | 5.0938 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | 5HC12400 | PDI Povidone-Iodine Prep Pad | Povidone-Iodine Prep Pad | Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | NY-Analytical | 772 | Box Inspection | Width | 4.25 | in | FDA,5 | 4.2188 | 4.2812 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | 5HC22324 | PDI Hygea Sterile Saline Wipes | PDI Hygea Sterile Saline Wipes | Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | 5HC22324 | PDI Hygea Sterile Saline Wipes | PDI Hygea Sterile Saline Wipes | Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Caliper | 0.0198 | in | FDA,3 | 0.019 | 0.021 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | 5HC22324 | PDI Hygea Sterile Saline Wipes | PDI Hygea Sterile Saline Wipes | Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Carton Material | WCCN (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | 5HC22324 | PDI Hygea Sterile Saline Wipes | PDI Hygea Sterile Saline Wipes | Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | 5HC22324 | PDI Hygea Sterile Saline Wipes | PDI Hygea Sterile Saline Wipes | Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Height | 3.0625 | in | FDA,5 | 3.0312 | 3.0938 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | 5HC22324 | PDI Hygea Sterile Saline Wipes | PDI Hygea Sterile Saline Wipes | Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Length | 5.0625 | in | FDA,5 | 5.0312 | 5.0938 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | 5HC22324 | PDI Hygea Sterile Saline Wipes | PDI Hygea Sterile Saline Wipes | Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Width | 1.75 | in | FDA,5 | 1.7188 | 1.7812 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | 5JB16400 | PDI Adhesive Tape Remover | PDI Adhesive Tape Remover | Box_NY_0.016UCSBS_100_4-3.125-2.375 | NY-Analytical | 772 | Box Inspection | Carton Material | UCSBS |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | 5JB16400 | PDI Adhesive Tape Remover | PDI Adhesive Tape Remover | Box_NY_0.016UCSBS_100_4-3.125-2.375 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | 5JB16400 | PDI Adhesive Tape Remover | PDI Adhesive Tape Remover | Box_NY_0.016UCSBS_100_4-3.125-2.375 | NY-Analytical | 772 | Box Inspection | Caliper | 0.015 | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | 5JB16400 | PDI Adhesive Tape Remover | PDI Adhesive Tape Remover | Box_NY_0.016UCSBS_100_4-3.125-2.375 | NY-Analytical | 772 | Box Inspection | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | 5JB16400 | PDI Adhesive Tape Remover | PDI Adhesive Tape Remover | Box_NY_0.016UCSBS_100_4-3.125-2.375 | NY-Analytical | 772 | Box Inspection | Height | 2.375 | in | FDA,5 | 2.3438 | 2.4062 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | 5JB16400 | PDI Adhesive Tape Remover | PDI Adhesive Tape Remover | Box_NY_0.016UCSBS_100_4-3.125-2.375 | NY-Analytical | 772 | Box Inspection | Length | 4 | in | FDA,5 | 3.9688 | 4.0312 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | 5JB16400 | PDI Adhesive Tape Remover | PDI Adhesive Tape Remover | Box_NY_0.016UCSBS_100_4-3.125-2.375 | NY-Analytical | 772 | Box Inspection | Width | 3.125 | in | FDA,5 | 3.0938 | 3.1562 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | 5LB60302 | PDI Alcohol Prep Pad | PDI Alcohol Prep Pad | Box_NY_0.016UCSBS_100_6.125-4-2.375 | NY-Analytical | 772 | Box Inspection | Carton Material | UCSBS |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | 5LB60302 | PDI Alcohol Prep Pad | PDI Alcohol Prep Pad | Box_NY_0.016UCSBS_100_6.125-4-2.375 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | 80 |  | FDA,0 |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | 5LB60302 | PDI Alcohol Prep Pad | PDI Alcohol Prep Pad | Box_NY_0.016UCSBS_100_6.125-4-2.375 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | 5LB60302 | PDI Alcohol Prep Pad | PDI Alcohol Prep Pad | Box_NY_0.016UCSBS_100_6.125-4-2.375 | NY-Analytical | 772 | Box Inspection | Caliper | 0.016 | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | 5LB60302 | PDI Alcohol Prep Pad | PDI Alcohol Prep Pad | Box_NY_0.016UCSBS_100_6.125-4-2.375 | NY-Analytical | 772 | Box Inspection | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | 5LB60302 | PDI Alcohol Prep Pad | PDI Alcohol Prep Pad | Box_NY_0.016UCSBS_100_6.125-4-2.375 | NY-Analytical | 772 | Box Inspection | Height | 2.375 | in | FDA,5 | 2.3438 | 2.4062 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | 5LB60302 | PDI Alcohol Prep Pad | PDI Alcohol Prep Pad | Box_NY_0.016UCSBS_100_6.125-4-2.375 | NY-Analytical | 772 | Box Inspection | Length | 6.0938 | in | FDA,5 | 6.0938 | 6.1562 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | 5LB60302 | PDI Alcohol Prep Pad | PDI Alcohol Prep Pad | Box_NY_0.016UCSBS_100_6.125-4-2.375 | NY-Analytical | 772 | Box Inspection | Width | 4 | in | FDA,5 | 3.9688 | 4.0312 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | 5MS48050 | PDI Providone-Iodine Cleansing Scrub Box | Povidone-Iodine Cleansing Scrub Box | Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | NY-Analytical | 772 | Box Inspection | Carton Material | UCSBS |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | 5MS48050 | PDI Providone-Iodine Cleansing Scrub Box | Povidone-Iodine Cleansing Scrub Box | Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | 80 |  | FDA,0 |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | 5MS48050 | PDI Providone-Iodine Cleansing Scrub Box | Povidone-Iodine Cleansing Scrub Box | Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | 5MS48050 | PDI Providone-Iodine Cleansing Scrub Box | Povidone-Iodine Cleansing Scrub Box | Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | NY-Analytical | 772 | Box Inspection | Caliper | 0.0155 | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | 5MS48050 | PDI Providone-Iodine Cleansing Scrub Box | Povidone-Iodine Cleansing Scrub Box | Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | NY-Analytical | 772 | Box Inspection | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | 5MS48050 | PDI Providone-Iodine Cleansing Scrub Box | Povidone-Iodine Cleansing Scrub Box | Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | NY-Analytical | 772 | Box Inspection | Height | 5.8125 | in | FDA,5 | 5.7812 | 5.8438 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | 5MS48050 | PDI Providone-Iodine Cleansing Scrub Box | Povidone-Iodine Cleansing Scrub Box | Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | NY-Analytical | 772 | Box Inspection | Length | 6.5 | in | FDA,5 | 6.4688 | 6.5312 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5MS48050/Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | 5MS48050 | PDI Providone-Iodine Cleansing Scrub Box | Povidone-Iodine Cleansing Scrub Box | Box_NY_0.016UCSBS_100_6.5-3.125-5.8125 | NY-Analytical | 772 | Box Inspection | Width | 3.3125 | in | FDA,5 | 3.2812 | 3.3438 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | 5NB40600 | PDI Povidone Iodine Prep Med | Povidone-Iodine Prep Med | Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Carton Material | UCSBS |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | 5NB40600 | PDI Povidone Iodine Prep Med | Povidone-Iodine Prep Med | Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | 125 |  | FDA,0 |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | 5NB40600 | PDI Povidone Iodine Prep Med | Povidone-Iodine Prep Med | Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | 5NB40600 | PDI Povidone Iodine Prep Med | Povidone-Iodine Prep Med | Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Caliper | 0.0155 | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | 5NB40600 | PDI Povidone Iodine Prep Med | Povidone-Iodine Prep Med | Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | 5NB40600 | PDI Povidone Iodine Prep Med | Povidone-Iodine Prep Med | Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Height | 2.07 | in | FDA,5 | 2.0625 | 2.125 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | 5NB40600 | PDI Povidone Iodine Prep Med | Povidone-Iodine Prep Med | Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Length | 5.13 | in | FDA,5 | 5.0938 | 5.1562 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | 5NB40600 | PDI Povidone Iodine Prep Med | Povidone-Iodine Prep Med | Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Width | 2.24 | in | FDA,5 | 2.2188 | 2.2812 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | 5NU87250 | PDI Super Sani-Cloth 50ct. Box | Super Sani-Cloth 50Ct. Box | Box_NY_0.024SBS_127_4.375-4.375-12 | NY-Analytical | 772 | Box Inspection | Carton Style | Tuck Top / Auto Bottom Dispenser Box |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | 5NU87250 | PDI Super Sani-Cloth 50ct. Box | Super Sani-Cloth 50Ct. Box | Box_NY_0.024SBS_127_4.375-4.375-12 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | 5NU87250 | PDI Super Sani-Cloth 50ct. Box | Super Sani-Cloth 50Ct. Box | Box_NY_0.024SBS_127_4.375-4.375-12 | NY-Analytical | 772 | Box Inspection | Caliper | 0.024 | in | FDA,3 | 0.023 | 0.025 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | 5NU87250 | PDI Super Sani-Cloth 50ct. Box | Super Sani-Cloth 50Ct. Box | Box_NY_0.024SBS_127_4.375-4.375-12 | NY-Analytical | 772 | Box Inspection | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | 5NU87250 | PDI Super Sani-Cloth 50ct. Box | Super Sani-Cloth 50Ct. Box | Box_NY_0.024SBS_127_4.375-4.375-12 | NY-Analytical | 772 | Box Inspection | Height | 12 | in | FDA,5 | 11.9688 | 12.0312 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | 5NU87250 | PDI Super Sani-Cloth 50ct. Box | Super Sani-Cloth 50Ct. Box | Box_NY_0.024SBS_127_4.375-4.375-12 | NY-Analytical | 772 | Box Inspection | Length | 4.39 | in | FDA,5 | 4.3438 | 4.4062 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | 5NU87250 | PDI Super Sani-Cloth 50ct. Box | Super Sani-Cloth 50Ct. Box | Box_NY_0.024SBS_127_4.375-4.375-12 | NY-Analytical | 772 | Box Inspection | Width | 4.39 | in | FDA,5 | 4.3438 | 4.4062 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OD125PB/Box_80A | 5OD125PB | BAG WET NAP 8.25 X 6.5 X 2 | Bag Wet Nap 8.25 X 6.5 X 2 | Box_80A | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OD125PB/Box_80A | 5OD125PB | BAG WET NAP 8.25 X 6.5 X 2 | Bag Wet Nap 8.25 X 6.5 X 2 | Box_80A | NY-Analytical | 772 | Box Inspection | Caliper | 0.0005 | in | FDA,4 | 0.0005 | 0.0025 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OD125PB/Box_80A | 5OD125PB | BAG WET NAP 8.25 X 6.5 X 2 | Bag Wet Nap 8.25 X 6.5 X 2 | Box_80A | NY-Analytical | 772 | Box Inspection | Carton Material | LDPE (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OD125PB/Box_80A | 5OD125PB | BAG WET NAP 8.25 X 6.5 X 2 | Bag Wet Nap 8.25 X 6.5 X 2 | Box_80A | NY-Analytical | 772 | Box Inspection | Carton Style | Poly Bag w/fold over top on 4” wicketts (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OD125PB/Box_80A | 5OD125PB | BAG WET NAP 8.25 X 6.5 X 2 | Bag Wet Nap 8.25 X 6.5 X 2 | Box_80A | NY-Analytical | 772 | Box Inspection | Length | 6.5 | in | FDA,5 | 6.0625 | 6.5 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OD125PB/Box_80A | 5OD125PB | BAG WET NAP 8.25 X 6.5 X 2 | Bag Wet Nap 8.25 X 6.5 X 2 | Box_80A | NY-Analytical | 772 | Box Inspection | Outside Flap | 2.25 | in | FDA,5 | 2.2188 | 2.2812 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OD125PB/Box_80A | 5OD125PB | BAG WET NAP 8.25 X 6.5 X 2 | Bag Wet Nap 8.25 X 6.5 X 2 | Box_80A | NY-Analytical | 772 | Box Inspection | Width | 8.25 | in | FDA,5 | 8.25 | 8.5 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OD239PD/Box_51A | 5OD239PD | PAD 15-3/8" X 5-1/2" GROUP176 | Pad 15-3/8" X 5-1/2" Group176 | Box_51A | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OD239PD/Box_51A | 5OD239PD | PAD 15-3/8" X 5-1/2" GROUP176 | Pad 15-3/8" X 5-1/2" Group176 | Box_51A | NY-Analytical | 772 | Box Inspection | Caliper | 0.0175 | in | FDA,3 |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OD239PD/Box_51A | 5OD239PD | PAD 15-3/8" X 5-1/2" GROUP176 | Pad 15-3/8" X 5-1/2" Group176 | Box_51A | NY-Analytical | 772 | Box Inspection | Carton Material | Chip Board (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OD239PD/Box_51A | 5OD239PD | PAD 15-3/8" X 5-1/2" GROUP176 | Pad 15-3/8" X 5-1/2" Group176 | Box_51A | NY-Analytical | 772 | Box Inspection | Carton Style | Layer Pad (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OD239PD/Box_51A | 5OD239PD | PAD 15-3/8" X 5-1/2" GROUP176 | Pad 15-3/8" X 5-1/2" Group176 | Box_51A | NY-Analytical | 772 | Box Inspection | Length | 15.375 | in | FDA,5 | 15.3438 | 15.4062 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OD239PD/Box_51A | 5OD239PD | PAD 15-3/8" X 5-1/2" GROUP176 | Pad 15-3/8" X 5-1/2" Group176 | Box_51A | NY-Analytical | 772 | Box Inspection | Width | 5.5 | in | FDA,5 | 5.4688 | 5.5312 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OH382PD/Box_50D | 5OH382PD | PAD 11-5/8"X 9-1/2" GR-182 | Pad 11-5/8"X 9-1/2" Gr-182 | Box_50D | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OH382PD/Box_50D | 5OH382PD | PAD 11-5/8"X 9-1/2" GR-182 | Pad 11-5/8"X 9-1/2" Gr-182 | Box_50D | NY-Analytical | 772 | Box Inspection | Caliper | 0.0018 | in | FDA,3 |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OH382PD/Box_50D | 5OH382PD | PAD 11-5/8"X 9-1/2" GR-182 | Pad 11-5/8"X 9-1/2" Gr-182 | Box_50D | NY-Analytical | 772 | Box Inspection | Carton Material | Any Board (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OH382PD/Box_50D | 5OH382PD | PAD 11-5/8"X 9-1/2" GR-182 | Pad 11-5/8"X 9-1/2" Gr-182 | Box_50D | NY-Analytical | 772 | Box Inspection | Carton Style | Cut Pad (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OH382PD/Box_50D | 5OH382PD | PAD 11-5/8"X 9-1/2" GR-182 | Pad 11-5/8"X 9-1/2" Gr-182 | Box_50D | NY-Analytical | 772 | Box Inspection | Length | 11.5938 | in | FDA,5 | 11.5938 | 11.6562 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OH382PD/Box_50D | 5OH382PD | PAD 11-5/8"X 9-1/2" GR-182 | Pad 11-5/8"X 9-1/2" Gr-182 | Box_50D | NY-Analytical | 772 | Box Inspection | Width | 9.4688 | in | FDA,5 | 9.4688 | 9.5312 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | 5OP067BT | Ponds Original Fresh Twlts. 28ct - Canada | Pond's Original Fresh Twlts. 28Ct. - Canada | Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | 50 |  | FDA,0 |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | 5OP067BT | Ponds Original Fresh Twlts. 28ct - Canada | Pond's Original Fresh Twlts. 28Ct. - Canada | Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | 5OP067BT | Ponds Original Fresh Twlts. 28ct - Canada | Pond's Original Fresh Twlts. 28Ct. - Canada | Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Caliper | 0.016 | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | 5OP067BT | Ponds Original Fresh Twlts. 28ct - Canada | Pond's Original Fresh Twlts. 28Ct. - Canada | Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Carton Material | SBS C1S (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | 5OP067BT | Ponds Original Fresh Twlts. 28ct - Canada | Pond's Original Fresh Twlts. 28Ct. - Canada | Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Carton Style | Open Top/Tuck Bottom (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | 5OP067BT | Ponds Original Fresh Twlts. 28ct - Canada | Pond's Original Fresh Twlts. 28Ct. - Canada | Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Height | 2.44 | in | FDA,5 | 2.4062 | 2.4688 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | 5OP067BT | Ponds Original Fresh Twlts. 28ct - Canada | Pond's Original Fresh Twlts. 28Ct. - Canada | Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Length | 3.2 | in | FDA,5 | 3.1488 | 3.2113 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | 5OP067BT | Ponds Original Fresh Twlts. 28ct - Canada | Pond's Original Fresh Twlts. 28Ct. - Canada | Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Width | 2.19 | in | FDA,5 | 2.1562 | 2.2188 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP073BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | 5OP073BT | Ponds Luminous Clean Twlts. | Pond's Luminous Clean Twlts. | Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | 80 |  | FDA,0 |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP073BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | 5OP073BT | Ponds Luminous Clean Twlts. | Pond's Luminous Clean Twlts. | Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP073BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | 5OP073BT | Ponds Luminous Clean Twlts. | Pond's Luminous Clean Twlts. | Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Caliper | 0.016 | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP073BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | 5OP073BT | Ponds Luminous Clean Twlts. | Pond's Luminous Clean Twlts. | Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Carton Material | SBS C1S (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP073BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | 5OP073BT | Ponds Luminous Clean Twlts. | Pond's Luminous Clean Twlts. | Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Carton Style | Open Top/Tuck Bottom (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP073BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | 5OP073BT | Ponds Luminous Clean Twlts. | Pond's Luminous Clean Twlts. | Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Height | 2.42 | in | FDA,5 | 2.4062 | 2.4688 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP073BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | 5OP073BT | Ponds Luminous Clean Twlts. | Pond's Luminous Clean Twlts. | Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Length | 3.2 | in | FDA,5 | 3.1488 | 3.2113 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP073BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | 5OP073BT | Ponds Luminous Clean Twlts. | Pond's Luminous Clean Twlts. | Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Width | 2.17 | in | FDA,5 | 2.1562 | 2.2188 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | 5OP092BT | Ponds Original Fresh 15ct Twlts. | Pond's Original Fresh 15Ct. Twlts. | Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | 315 |  | FDA,0 |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | 5OP092BT | Ponds Original Fresh 15ct Twlts. | Pond's Original Fresh 15Ct. Twlts. | Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | 5OP092BT | Ponds Original Fresh 15ct Twlts. | Pond's Original Fresh 15Ct. Twlts. | Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | NY-Analytical | 772 | Box Inspection | Caliper | 0.016 | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | 5OP092BT | Ponds Original Fresh 15ct Twlts. | Pond's Original Fresh 15Ct. Twlts. | Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | NY-Analytical | 772 | Box Inspection | Carton Material | SBS C1S (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | 5OP092BT | Ponds Original Fresh 15ct Twlts. | Pond's Original Fresh 15Ct. Twlts. | Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | NY-Analytical | 772 | Box Inspection | Carton Style | Open Top/Tuck Bottom (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | 5OP092BT | Ponds Original Fresh 15ct Twlts. | Pond's Original Fresh 15Ct. Twlts. | Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | NY-Analytical | 772 | Box Inspection | Height | 2.6875 | in | FDA,5 | 2.6562 | 2.7188 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | 5OP092BT | Ponds Original Fresh 15ct Twlts. | Pond's Original Fresh 15Ct. Twlts. | Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | NY-Analytical | 772 | Box Inspection | Length | 3.1875 | in | FDA,5 | 3.1488 | 3.2113 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | 5OP092BT | Ponds Original Fresh 15ct Twlts. | Pond's Original Fresh 15Ct. Twlts. | Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | NY-Analytical | 772 | Box Inspection | Width | 1.4375 | in | FDA,5 | 1.4062 | 1.4688 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | 5OP092PB | Poly Bag w/fold over top on 2" wicket | Poly Bag W/Fold Over Top On 2" Wicket | Box_NY_LDPE_407_13.5-6.5-6-4 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | 80 |  | FDA,0 |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | 5OP092PB | Poly Bag w/fold over top on 2" wicket | Poly Bag W/Fold Over Top On 2" Wicket | Box_NY_LDPE_407_13.5-6.5-6-4 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | 5OP092PB | Poly Bag w/fold over top on 2" wicket | Poly Bag W/Fold Over Top On 2" Wicket | Box_NY_LDPE_407_13.5-6.5-6-4 | NY-Analytical | 772 | Box Inspection | Back Flip | 5.75 | in | FDA,5 | 5.6875 | 6.3125 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | 5OP092PB | Poly Bag w/fold over top on 2" wicket | Poly Bag W/Fold Over Top On 2" Wicket | Box_NY_LDPE_407_13.5-6.5-6-4 | NY-Analytical | 772 | Box Inspection | Carton Material | LDPE (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | 5OP092PB | Poly Bag w/fold over top on 2" wicket | Poly Bag W/Fold Over Top On 2" Wicket | Box_NY_LDPE_407_13.5-6.5-6-4 | NY-Analytical | 772 | Box Inspection | Carton Style | Poly Bag w/fold over top on 2” wicketts (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | 5OP092PB | Poly Bag w/fold over top on 2" wicket | Poly Bag W/Fold Over Top On 2" Wicket | Box_NY_LDPE_407_13.5-6.5-6-4 | NY-Analytical | 772 | Box Inspection | Gussett | 4 | in | FDA,5 | 3.6875 | 4.3125 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | 5OP092PB | Poly Bag w/fold over top on 2" wicket | Poly Bag W/Fold Over Top On 2" Wicket | Box_NY_LDPE_407_13.5-6.5-6-4 | NY-Analytical | 772 | Box Inspection | Height | 6.5 | in | FDA,5 | 6.1875 | 6.8125 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | 5OP092PB | Poly Bag w/fold over top on 2" wicket | Poly Bag W/Fold Over Top On 2" Wicket | Box_NY_LDPE_407_13.5-6.5-6-4 | NY-Analytical | 772 | Box Inspection | Width | 13.1875 | in | FDA,5 | 13.1875 | 13.8125 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP308PD/Box_50C | 5OP308PD | AQUA LAYER PAD | Aqua Layer Pad | Box_50C | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP308PD/Box_50C | 5OP308PD | AQUA LAYER PAD | Aqua Layer Pad | Box_50C | NY-Analytical | 772 | Box Inspection | Caliper | 0.02 | in | FDA,3 | 0.018 | 0.02 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP308PD/Box_50C | 5OP308PD | AQUA LAYER PAD | Aqua Layer Pad | Box_50C | NY-Analytical | 772 | Box Inspection | Carton Material | Chip Board (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP308PD/Box_50C | 5OP308PD | AQUA LAYER PAD | Aqua Layer Pad | Box_50C | NY-Analytical | 772 | Box Inspection | Carton Style | Cut Pad (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP308PD/Box_50C | 5OP308PD | AQUA LAYER PAD | Aqua Layer Pad | Box_50C | NY-Analytical | 772 | Box Inspection | Length | 7 | in | FDA,5 | 6.9688 | 7.0312 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5OP308PD/Box_50C | 5OP308PD | AQUA LAYER PAD | Aqua Layer Pad | Box_50C | NY-Analytical | 772 | Box Inspection | Width | 12.28 | in | FDA,5 | 12.2188 | 12.2812 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5PB114IS/Box_NY_NA50#_402_2.5-3.25 | 5PB114IS | PDI Prevantics Swab Insert | Prevantics Swab Insert | Box_NY_NA50#_402_2.5-3.25 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | 500 |  | FDA,0 |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5PB114IS/Box_NY_NA50#_402_2.5-3.25 | 5PB114IS | PDI Prevantics Swab Insert | Prevantics Swab Insert | Box_NY_NA50#_402_2.5-3.25 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5PB114IS/Box_NY_NA50#_402_2.5-3.25 | 5PB114IS | PDI Prevantics Swab Insert | Prevantics Swab Insert | Box_NY_NA50#_402_2.5-3.25 | NY-Analytical | 772 | Box Inspection | Carton Material | 50# White Paper Stock (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5PB114IS/Box_NY_NA50#_402_2.5-3.25 | 5PB114IS | PDI Prevantics Swab Insert | Prevantics Swab Insert | Box_NY_NA50#_402_2.5-3.25 | NY-Analytical | 772 | Box Inspection | Carton Style | Insert (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5PB114IS/Box_NY_NA50#_402_2.5-3.25 | 5PB114IS | PDI Prevantics Swab Insert | Prevantics Swab Insert | Box_NY_NA50#_402_2.5-3.25 | NY-Analytical | 772 | Box Inspection | Length | 2.5 | in | FDA,5 | 2.4688 | 2.5312 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5PB114IS/Box_NY_NA50#_402_2.5-3.25 | 5PB114IS | PDI Prevantics Swab Insert | Prevantics Swab Insert | Box_NY_NA50#_402_2.5-3.25 | NY-Analytical | 772 | Box Inspection | Width | 3.25 | in | FDA,5 | 3.2188 | 3.2812 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | 5PS40750 | PDI Prevantics Skin Swabstick 50ct Box | Prevantics Skin Swabstick 50Ct. Box | Box_NY_0.018SBS_100_7.125-3-6.625 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | 5PS40750 | PDI Prevantics Skin Swabstick 50ct Box | Prevantics Skin Swabstick 50Ct. Box | Box_NY_0.018SBS_100_7.125-3-6.625 | NY-Analytical | 772 | Box Inspection | Caliper | 0.0175 | in | FDA,3 | 0.017 | 0.019 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | 5PS40750 | PDI Prevantics Skin Swabstick 50ct Box | Prevantics Skin Swabstick 50Ct. Box | Box_NY_0.018SBS_100_7.125-3-6.625 | NY-Analytical | 772 | Box Inspection | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | 5PS40750 | PDI Prevantics Skin Swabstick 50ct Box | Prevantics Skin Swabstick 50Ct. Box | Box_NY_0.018SBS_100_7.125-3-6.625 | NY-Analytical | 772 | Box Inspection | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | 5PS40750 | PDI Prevantics Skin Swabstick 50ct Box | Prevantics Skin Swabstick 50Ct. Box | Box_NY_0.018SBS_100_7.125-3-6.625 | NY-Analytical | 772 | Box Inspection | Height | 6.625 | in | FDA,5 | 6.5938 | 6.6562 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | 5PS40750 | PDI Prevantics Skin Swabstick 50ct Box | Prevantics Skin Swabstick 50Ct. Box | Box_NY_0.018SBS_100_7.125-3-6.625 | NY-Analytical | 772 | Box Inspection | Length | 7.125 | in | FDA,5 | 7.0938 | 7.1562 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | 5PS40750 | PDI Prevantics Skin Swabstick 50ct Box | Prevantics Skin Swabstick 50Ct. Box | Box_NY_0.018SBS_100_7.125-3-6.625 | NY-Analytical | 772 | Box Inspection | Width | 3 | in | FDA,5 | 2.9688 | 3.0312 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5QS273IS/Box_NY_NA50#_402_5-3.25 | 5QS273IS | PDI Prevantics Swabstick Insert | Prevantics Swabstick Insert | Box_NY_NA50#_402_5-3.25 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | 500 |  | FDA,0 |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5QS273IS/Box_NY_NA50#_402_5-3.25 | 5QS273IS | PDI Prevantics Swabstick Insert | Prevantics Swabstick Insert | Box_NY_NA50#_402_5-3.25 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5QS273IS/Box_NY_NA50#_402_5-3.25 | 5QS273IS | PDI Prevantics Swabstick Insert | Prevantics Swabstick Insert | Box_NY_NA50#_402_5-3.25 | NY-Analytical | 772 | Box Inspection | Carton Material | 50# White Paper Stock (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5QS273IS/Box_NY_NA50#_402_5-3.25 | 5QS273IS | PDI Prevantics Swabstick Insert | Prevantics Swabstick Insert | Box_NY_NA50#_402_5-3.25 | NY-Analytical | 772 | Box Inspection | Carton Style | Insert (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5QS273IS/Box_NY_NA50#_402_5-3.25 | 5QS273IS | PDI Prevantics Swabstick Insert | Prevantics Swabstick Insert | Box_NY_NA50#_402_5-3.25 | NY-Analytical | 772 | Box Inspection | Length | 5 | in | FDA,5 | 4.9688 | 5.0312 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5QS273IS/Box_NY_NA50#_402_5-3.25 | 5QS273IS | PDI Prevantics Swabstick Insert | Prevantics Swabstick Insert | Box_NY_NA50#_402_5-3.25 | NY-Analytical | 772 | Box Inspection | Width | 3.2188 | in | FDA,5 | 3.2188 | 3.2812 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | 5QS41950 | PDI Prevantics Maxi Swabstick Box | Prevantics Maxi Swabstick Box | Box_NY_0.018SBS_100_6.625-4.125-8.125 | NY-Analytical | 772 | Box Inspection | Artwork / Copy | Meets specifications (Done) | in |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | 5QS41950 | PDI Prevantics Maxi Swabstick Box | Prevantics Maxi Swabstick Box | Box_NY_0.018SBS_100_6.625-4.125-8.125 | NY-Analytical | 772 | Box Inspection | Caliper | 0.017 | in | FDA,3 | 0.017 | 0.019 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | 5QS41950 | PDI Prevantics Maxi Swabstick Box | Prevantics Maxi Swabstick Box | Box_NY_0.018SBS_100_6.625-4.125-8.125 | NY-Analytical | 772 | Box Inspection | Carton Material | SBS (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | 5QS41950 | PDI Prevantics Maxi Swabstick Box | Prevantics Maxi Swabstick Box | Box_NY_0.018SBS_100_6.625-4.125-8.125 | NY-Analytical | 772 | Box Inspection | Carton Style | Tuck Top / Auto Bottom (Done) |  |  |  |  | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | 5QS41950 | PDI Prevantics Maxi Swabstick Box | Prevantics Maxi Swabstick Box | Box_NY_0.018SBS_100_6.625-4.125-8.125 | NY-Analytical | 772 | Box Inspection | Height | 8.125 | in | FDA,5 | 8.0938 | 8.1562 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | 5QS41950 | PDI Prevantics Maxi Swabstick Box | Prevantics Maxi Swabstick Box | Box_NY_0.018SBS_100_6.625-4.125-8.125 | NY-Analytical | 772 | Box Inspection | Length | 6.625 | in | FDA,5 | 6.5938 | 6.6562 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | 5QS41950 | PDI Prevantics Maxi Swabstick Box | Prevantics Maxi Swabstick Box | Box_NY_0.018SBS_100_6.625-4.125-8.125 | NY-Analytical | 772 | Box Inspection | Width | 4.125 | in | FDA,5 | 4.0938 | 4.1562 | Box Inspection | Box Inspection | NY | RAW MATERIAL - NY | RM_BOX | both
6AA34724/Shipper_NY_Skip lot testing | 6AA34724 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | Shipper_NY_Skip lot testing | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Shipper Inspection |  | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | 6AA34724 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 20 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | 6AA34724 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | 6AA34724 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | 6AA34724 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | 6AA34724 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | Height | 6.1875 | in | FDA,4 | 6.0625 | 6.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | 6AA34724 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | Length | 8.4375 | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | 6AA34724 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | 6AA34724 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | 6AA34724 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | Width | 7.1875 | in | FDA,4 | 6.8125 | 7.1875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | 6AA617S96 | Prep H 96ct Shipper | Prep H 96Ct. Shipper | SHIP_NY_44K_100_12.75-11.625-7.875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | 6AA617S96 | Prep H 96ct Shipper | Prep H 96Ct. Shipper | SHIP_NY_44K_100_12.75-11.625-7.875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | 6AA617S96 | Prep H 96ct Shipper | Prep H 96Ct. Shipper | SHIP_NY_44K_100_12.75-11.625-7.875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | 6AA617S96 | Prep H 96ct Shipper | Prep H 96Ct. Shipper | SHIP_NY_44K_100_12.75-11.625-7.875 | NY-Analytical | 731 | Shipper Inspection | Height | 7.75 | in | FDA,4 | 7.6875 | 8.0625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | 6AA617S96 | Prep H 96ct Shipper | Prep H 96Ct. Shipper | SHIP_NY_44K_100_12.75-11.625-7.875 | NY-Analytical | 731 | Shipper Inspection | Length | 12.75 | in | FDA,4 | 12.5625 | 12.9375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | 6AA617S96 | Prep H 96ct Shipper | Prep H 96Ct. Shipper | SHIP_NY_44K_100_12.75-11.625-7.875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | 6AA617S96 | Prep H 96ct Shipper | Prep H 96Ct. Shipper | SHIP_NY_44K_100_12.75-11.625-7.875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | 6AA617S96 | Prep H 96ct Shipper | Prep H 96Ct. Shipper | SHIP_NY_44K_100_12.75-11.625-7.875 | NY-Analytical | 731 | Shipper Inspection | Width | 11.5 | in | FDA,4 | 11.4375 | 11.8125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | 6AA632TY | Preparation H 180ct Tray | Prep H 180Ct. Tray | SHIP_NY_44K_231_18.5-7.375-14 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | 6AA632TY | Preparation H 180ct Tray | Prep H 180Ct. Tray | SHIP_NY_44K_231_18.5-7.375-14 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | 6AA632TY | Preparation H 180ct Tray | Prep H 180Ct. Tray | SHIP_NY_44K_231_18.5-7.375-14 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | 6AA632TY | Preparation H 180ct Tray | Prep H 180Ct. Tray | SHIP_NY_44K_231_18.5-7.375-14 | NY-Analytical | 731 | Shipper Inspection | Height | 14 | in | FDA,4 | 13.875 | 14.125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | 6AA632TY | Preparation H 180ct Tray | Prep H 180Ct. Tray | SHIP_NY_44K_231_18.5-7.375-14 | NY-Analytical | 731 | Shipper Inspection | Length | 18.5 | in | FDA,4 | 18.375 | 18.625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | 6AA632TY | Preparation H 180ct Tray | Prep H 180Ct. Tray | SHIP_NY_44K_231_18.5-7.375-14 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | 6AA632TY | Preparation H 180ct Tray | Prep H 180Ct. Tray | SHIP_NY_44K_231_18.5-7.375-14 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | D/C (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | 6AA632TY | Preparation H 180ct Tray | Prep H 180Ct. Tray | SHIP_NY_44K_231_18.5-7.375-14 | NY-Analytical | 731 | Shipper Inspection | Width | 7.5 | in | FDA,4 | 7.25 | 7.5 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 6AA709S48 | Preparation H for Women Shipper | Prep H For Women Shipper | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 200 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 6AA709S48 | Preparation H for Women Shipper | Prep H For Women Shipper | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 6AA709S48 | Preparation H for Women Shipper | Prep H For Women Shipper | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 40 | Lbs./In. | FDA,0 | 40 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 6AA709S48 | Preparation H for Women Shipper | Prep H For Women Shipper | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 6AA709S48 | Preparation H for Women Shipper | Prep H For Women Shipper | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | Height | 6.125 | in | FDA,4 | 5.9375 | 6.3125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 6AA709S48 | Preparation H for Women Shipper | Prep H For Women Shipper | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | Length | 15 | in | FDA,4 | 14.8125 | 15.1875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 6AA709S48 | Preparation H for Women Shipper | Prep H For Women Shipper | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 6AA709S48 | Preparation H for Women Shipper | Prep H For Women Shipper | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA709S48/SHIP_NY_40K_100_15-11-6.125 | 6AA709S48 | Preparation H for Women Shipper | Prep H For Women Shipper | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | Width | 11 | in | FDA,4 | 10.8125 | 11.1875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 6AA73530 | Baby Dove Wipes Sensitive Moisture Shipper | Baby Dove Wipes Sensitive Moisture Shipper | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 6AA73530 | Baby Dove Wipes Sensitive Moisture Shipper | Baby Dove Wipes Sensitive Moisture Shipper | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 6AA73530 | Baby Dove Wipes Sensitive Moisture Shipper | Baby Dove Wipes Sensitive Moisture Shipper | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 6AA73530 | Baby Dove Wipes Sensitive Moisture Shipper | Baby Dove Wipes Sensitive Moisture Shipper | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 6AA73530 | Baby Dove Wipes Sensitive Moisture Shipper | Baby Dove Wipes Sensitive Moisture Shipper | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | Height | 8.25 | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 6AA73530 | Baby Dove Wipes Sensitive Moisture Shipper | Baby Dove Wipes Sensitive Moisture Shipper | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | Length | 10 | in | FDA,4 | 9.8125 | 10.1875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 6AA73530 | Baby Dove Wipes Sensitive Moisture Shipper | Baby Dove Wipes Sensitive Moisture Shipper | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 6AA73530 | Baby Dove Wipes Sensitive Moisture Shipper | Baby Dove Wipes Sensitive Moisture Shipper | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | 6AA73530 | Baby Dove Wipes Sensitive Moisture Shipper | Baby Dove Wipes Sensitive Moisture Shipper | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | Width | 8.375 | in | FDA,4 | 8.1875 | 8.5625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA76930/SHIP_NY_44K_100_9-9-4.75 | 6AA76930 | Baby Dove Wipes Rich Moisture Shipper | Baby Dove Wipes Rich Moisture Shipper | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA76930/SHIP_NY_44K_100_9-9-4.75 | 6AA76930 | Baby Dove Wipes Rich Moisture Shipper | Baby Dove Wipes Rich Moisture Shipper | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA76930/SHIP_NY_44K_100_9-9-4.75 | 6AA76930 | Baby Dove Wipes Rich Moisture Shipper | Baby Dove Wipes Rich Moisture Shipper | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA76930/SHIP_NY_44K_100_9-9-4.75 | 6AA76930 | Baby Dove Wipes Rich Moisture Shipper | Baby Dove Wipes Rich Moisture Shipper | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA76930/SHIP_NY_44K_100_9-9-4.75 | 6AA76930 | Baby Dove Wipes Rich Moisture Shipper | Baby Dove Wipes Rich Moisture Shipper | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | Height | 4.75 | in | FDA,4 | 4.625 | 4.875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA76930/SHIP_NY_44K_100_9-9-4.75 | 6AA76930 | Baby Dove Wipes Rich Moisture Shipper | Baby Dove Wipes Rich Moisture Shipper | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | Length | 9 | in | FDA,4 | 8.875 | 9.125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA76930/SHIP_NY_44K_100_9-9-4.75 | 6AA76930 | Baby Dove Wipes Rich Moisture Shipper | Baby Dove Wipes Rich Moisture Shipper | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA76930/SHIP_NY_44K_100_9-9-4.75 | 6AA76930 | Baby Dove Wipes Rich Moisture Shipper | Baby Dove Wipes Rich Moisture Shipper | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AA76930/SHIP_NY_44K_100_9-9-4.75 | 6AA76930 | Baby Dove Wipes Rich Moisture Shipper | Baby Dove Wipes Rich Moisture Shipper | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | Width | 9 | in | FDA,4 | 8.875 | 9.125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 6AB16400 | PDI Adhesive Tape Remover | PDI Adhesive Tape Remover | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 6AB16400 | PDI Adhesive Tape Remover | PDI Adhesive Tape Remover | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 6AB16400 | PDI Adhesive Tape Remover | PDI Adhesive Tape Remover | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 6AB16400 | PDI Adhesive Tape Remover | PDI Adhesive Tape Remover | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 6AB16400 | PDI Adhesive Tape Remover | PDI Adhesive Tape Remover | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | Height | 5.25 | in | FDA,4 | 4.9375 | 5.3125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 6AB16400 | PDI Adhesive Tape Remover | PDI Adhesive Tape Remover | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | Length | 16.125 | in | FDA,4 | 15.8125 | 16.1875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 6AB16400 | PDI Adhesive Tape Remover | PDI Adhesive Tape Remover | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 6AB16400 | PDI Adhesive Tape Remover | PDI Adhesive Tape Remover | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | 6AB16400 | PDI Adhesive Tape Remover | PDI Adhesive Tape Remover | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | Width | 4.25 | in | FDA,4 | 4.0625 | 4.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AD11055/SHIP_NY_32K_100_10.625-6.375-11.25 | 6AD11055 | COR WET NAP/BULK 32 ECT | Wet Nap/Bulk 32 ECt. | SHIP_NY_32K_100_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AD11055/SHIP_NY_32K_100_10.625-6.375-11.25 | 6AD11055 | COR WET NAP/BULK 32 ECT | Wet Nap/Bulk 32 ECt. | SHIP_NY_32K_100_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AD11055/SHIP_NY_32K_100_10.625-6.375-11.25 | 6AD11055 | COR WET NAP/BULK 32 ECT | Wet Nap/Bulk 32 ECt. | SHIP_NY_32K_100_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AD11055/SHIP_NY_32K_100_10.625-6.375-11.25 | 6AD11055 | COR WET NAP/BULK 32 ECT | Wet Nap/Bulk 32 ECt. | SHIP_NY_32K_100_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Height | 11.375 | in | FDA,4 | 11.0625 | 11.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AD11055/SHIP_NY_32K_100_10.625-6.375-11.25 | 6AD11055 | COR WET NAP/BULK 32 ECT | Wet Nap/Bulk 32 ECt. | SHIP_NY_32K_100_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Length | 10.625 | in | FDA,4 | 10.4375 | 10.8125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AD11055/SHIP_NY_32K_100_10.625-6.375-11.25 | 6AD11055 | COR WET NAP/BULK 32 ECT | Wet Nap/Bulk 32 ECt. | SHIP_NY_32K_100_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AD11055/SHIP_NY_32K_100_10.625-6.375-11.25 | 6AD11055 | COR WET NAP/BULK 32 ECT | Wet Nap/Bulk 32 ECt. | SHIP_NY_32K_100_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AD11055/SHIP_NY_32K_100_10.625-6.375-11.25 | 6AD11055 | COR WET NAP/BULK 32 ECT | Wet Nap/Bulk 32 ECt. | SHIP_NY_32K_100_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Width | 6.375 | in | FDA,4 | 6.1875 | 6.5625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AD41900/SHIP_NY_32K_100_13.25-12-6.125 | 6AD41900 | PDI Castile Soap Towelette Shipper | PDI Castile Soap Towelette Shipper | SHIP_NY_32K_100_13.25-12-6.125 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AD41900/SHIP_NY_32K_100_13.25-12-6.125 | 6AD41900 | PDI Castile Soap Towelette Shipper | PDI Castile Soap Towelette Shipper | SHIP_NY_32K_100_13.25-12-6.125 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AD41900/SHIP_NY_32K_100_13.25-12-6.125 | 6AD41900 | PDI Castile Soap Towelette Shipper | PDI Castile Soap Towelette Shipper | SHIP_NY_32K_100_13.25-12-6.125 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AD41900/SHIP_NY_32K_100_13.25-12-6.125 | 6AD41900 | PDI Castile Soap Towelette Shipper | PDI Castile Soap Towelette Shipper | SHIP_NY_32K_100_13.25-12-6.125 | NY-Analytical | 731 | Shipper Inspection | Height | 6.25 | in | FDA,4 | 5.9375 | 6.3125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AD41900/SHIP_NY_32K_100_13.25-12-6.125 | 6AD41900 | PDI Castile Soap Towelette Shipper | PDI Castile Soap Towelette Shipper | SHIP_NY_32K_100_13.25-12-6.125 | NY-Analytical | 731 | Shipper Inspection | Length | 13.25 | in | FDA,4 | 13.0625 | 13.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AD41900/SHIP_NY_32K_100_13.25-12-6.125 | 6AD41900 | PDI Castile Soap Towelette Shipper | PDI Castile Soap Towelette Shipper | SHIP_NY_32K_100_13.25-12-6.125 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AD41900/SHIP_NY_32K_100_13.25-12-6.125 | 6AD41900 | PDI Castile Soap Towelette Shipper | PDI Castile Soap Towelette Shipper | SHIP_NY_32K_100_13.25-12-6.125 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AD41900/SHIP_NY_32K_100_13.25-12-6.125 | 6AD41900 | PDI Castile Soap Towelette Shipper | PDI Castile Soap Towelette Shipper | SHIP_NY_32K_100_13.25-12-6.125 | NY-Analytical | 731 | Shipper Inspection | Width | 11.875 | in | FDA,4 | 11.8125 | 12.1875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AH58195/SHIP_NY_32K_100_13.4375-6.25-9.6875 | 6AH58195 | PDI Sani Cloth Bleach Shipper | Sani-Cloth Bleach Shipper | SHIP_NY_32K_100_13.4375-6.25-9.6875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AH58195/SHIP_NY_32K_100_13.4375-6.25-9.6875 | 6AH58195 | PDI Sani Cloth Bleach Shipper | Sani-Cloth Bleach Shipper | SHIP_NY_32K_100_13.4375-6.25-9.6875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AH58195/SHIP_NY_32K_100_13.4375-6.25-9.6875 | 6AH58195 | PDI Sani Cloth Bleach Shipper | Sani-Cloth Bleach Shipper | SHIP_NY_32K_100_13.4375-6.25-9.6875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AH58195/SHIP_NY_32K_100_13.4375-6.25-9.6875 | 6AH58195 | PDI Sani Cloth Bleach Shipper | Sani-Cloth Bleach Shipper | SHIP_NY_32K_100_13.4375-6.25-9.6875 | NY-Analytical | 731 | Shipper Inspection | Height | 9.75 | in | FDA,4 | 9.5 | 9.875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AH58195/SHIP_NY_32K_100_13.4375-6.25-9.6875 | 6AH58195 | PDI Sani Cloth Bleach Shipper | Sani-Cloth Bleach Shipper | SHIP_NY_32K_100_13.4375-6.25-9.6875 | NY-Analytical | 731 | Shipper Inspection | Length | 13.4375 | in | FDA,4 | 13.25 | 13.625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AH58195/SHIP_NY_32K_100_13.4375-6.25-9.6875 | 6AH58195 | PDI Sani Cloth Bleach Shipper | Sani-Cloth Bleach Shipper | SHIP_NY_32K_100_13.4375-6.25-9.6875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AH58195/SHIP_NY_32K_100_13.4375-6.25-9.6875 | 6AH58195 | PDI Sani Cloth Bleach Shipper | Sani-Cloth Bleach Shipper | SHIP_NY_32K_100_13.4375-6.25-9.6875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AH58195/SHIP_NY_32K_100_13.4375-6.25-9.6875 | 6AH58195 | PDI Sani Cloth Bleach Shipper | Sani-Cloth Bleach Shipper | SHIP_NY_32K_100_13.4375-6.25-9.6875 | NY-Analytical | 731 | Shipper Inspection | Width | 6.25 | in | FDA,4 | 6.0625 | 6.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 6AJ19025 | Cetaphil - Asia 25ct Shipper | Cetaphil - Asia 25Ct. Shipper | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 6AJ19025 | Cetaphil - Asia 25ct Shipper | Cetaphil - Asia 25Ct. Shipper | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 6AJ19025 | Cetaphil - Asia 25ct Shipper | Cetaphil - Asia 25Ct. Shipper | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 6AJ19025 | Cetaphil - Asia 25ct Shipper | Cetaphil - Asia 25Ct. Shipper | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 6AJ19025 | Cetaphil - Asia 25ct Shipper | Cetaphil - Asia 25Ct. Shipper | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | Height | 6.75 | in | FDA,4 | 6.5625 | 6.9375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 6AJ19025 | Cetaphil - Asia 25ct Shipper | Cetaphil - Asia 25Ct. Shipper | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | Length | 12 | in | FDA,4 | 11.8125 | 12.1875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 6AJ19025 | Cetaphil - Asia 25ct Shipper | Cetaphil - Asia 25Ct. Shipper | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 6AJ19025 | Cetaphil - Asia 25ct Shipper | Cetaphil - Asia 25Ct. Shipper | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | 6AJ19025 | Cetaphil - Asia 25ct Shipper | Cetaphil - Asia 25Ct. Shipper | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | Width | 7.25 | in | FDA,4 | 7.0625 | 7.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AM30472/SHIP_NY_44K_100_15.5-9-12 | 6AM30472 | Sani Professional No-Rinse Sanitizing Multi-Surface Wipes | Sani Professional No-Rinse Sanitizing Multi-Surface Wipes | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AM30472/SHIP_NY_44K_100_15.5-9-12 | 6AM30472 | Sani Professional No-Rinse Sanitizing Multi-Surface Wipes | Sani Professional No-Rinse Sanitizing Multi-Surface Wipes | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AM30472/SHIP_NY_44K_100_15.5-9-12 | 6AM30472 | Sani Professional No-Rinse Sanitizing Multi-Surface Wipes | Sani Professional No-Rinse Sanitizing Multi-Surface Wipes | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AM30472/SHIP_NY_44K_100_15.5-9-12 | 6AM30472 | Sani Professional No-Rinse Sanitizing Multi-Surface Wipes | Sani Professional No-Rinse Sanitizing Multi-Surface Wipes | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AM30472/SHIP_NY_44K_100_15.5-9-12 | 6AM30472 | Sani Professional No-Rinse Sanitizing Multi-Surface Wipes | Sani Professional No-Rinse Sanitizing Multi-Surface Wipes | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | Height | 12.125 | in | FDA,4 | 11.8125 | 12.1875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AM30472/SHIP_NY_44K_100_15.5-9-12 | 6AM30472 | Sani Professional No-Rinse Sanitizing Multi-Surface Wipes | Sani Professional No-Rinse Sanitizing Multi-Surface Wipes | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | Length | 15.375 | in | FDA,4 | 15.3125 | 15.6875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AM30472/SHIP_NY_44K_100_15.5-9-12 | 6AM30472 | Sani Professional No-Rinse Sanitizing Multi-Surface Wipes | Sani Professional No-Rinse Sanitizing Multi-Surface Wipes | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AM30472/SHIP_NY_44K_100_15.5-9-12 | 6AM30472 | Sani Professional No-Rinse Sanitizing Multi-Surface Wipes | Sani Professional No-Rinse Sanitizing Multi-Surface Wipes | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AM30472/SHIP_NY_44K_100_15.5-9-12 | 6AM30472 | Sani Professional No-Rinse Sanitizing Multi-Surface Wipes | Sani Professional No-Rinse Sanitizing Multi-Surface Wipes | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | Width | 8.875 | in | FDA,4 | 8.8125 | 9.1875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP03707/SHIP_NY_32K_100_13-7.3125-4 | 6AP03707 | Simple Kind To Skin Facial Shipper USA | Simple Kind To Skin Facial Shipper USa | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP03707/SHIP_NY_32K_100_13-7.3125-4 | 6AP03707 | Simple Kind To Skin Facial Shipper USA | Simple Kind To Skin Facial Shipper USa | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP03707/SHIP_NY_32K_100_13-7.3125-4 | 6AP03707 | Simple Kind To Skin Facial Shipper USA | Simple Kind To Skin Facial Shipper USa | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP03707/SHIP_NY_32K_100_13-7.3125-4 | 6AP03707 | Simple Kind To Skin Facial Shipper USA | Simple Kind To Skin Facial Shipper USa | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP03707/SHIP_NY_32K_100_13-7.3125-4 | 6AP03707 | Simple Kind To Skin Facial Shipper USA | Simple Kind To Skin Facial Shipper USa | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | Height | 4.1875 | in | FDA,4 | 3.8125 | 4.1875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP03707/SHIP_NY_32K_100_13-7.3125-4 | 6AP03707 | Simple Kind To Skin Facial Shipper USA | Simple Kind To Skin Facial Shipper USa | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | Length | 13.062 | in | FDA,4 | 12.8125 | 13.1875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP03707/SHIP_NY_32K_100_13-7.3125-4 | 6AP03707 | Simple Kind To Skin Facial Shipper USA | Simple Kind To Skin Facial Shipper USa | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP03707/SHIP_NY_32K_100_13-7.3125-4 | 6AP03707 | Simple Kind To Skin Facial Shipper USA | Simple Kind To Skin Facial Shipper USa | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP03707/SHIP_NY_32K_100_13-7.3125-4 | 6AP03707 | Simple Kind To Skin Facial Shipper USA | Simple Kind To Skin Facial Shipper USa | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | Width | 7.5 | in | FDA,4 | 7.125 | 7.5 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | 6AP06528 | Ponds Evening Soothe 28ct Shipper | Pond's Evening Soothe 28Ct. Shipper | SHIP_NY_32K_100_13.25-7.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | 6AP06528 | Ponds Evening Soothe 28ct Shipper | Pond's Evening Soothe 28Ct. Shipper | SHIP_NY_32K_100_13.25-7.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | 6AP06528 | Ponds Evening Soothe 28ct Shipper | Pond's Evening Soothe 28Ct. Shipper | SHIP_NY_32K_100_13.25-7.75-5.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | 6AP06528 | Ponds Evening Soothe 28ct Shipper | Pond's Evening Soothe 28Ct. Shipper | SHIP_NY_32K_100_13.25-7.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Height | 5.9375 | in | FDA,4 | 5.5625 | 5.9375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | 6AP06528 | Ponds Evening Soothe 28ct Shipper | Pond's Evening Soothe 28Ct. Shipper | SHIP_NY_32K_100_13.25-7.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Length | 13.25 | in | FDA,4 | 13.0625 | 13.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | 6AP06528 | Ponds Evening Soothe 28ct Shipper | Pond's Evening Soothe 28Ct. Shipper | SHIP_NY_32K_100_13.25-7.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | 6AP06528 | Ponds Evening Soothe 28ct Shipper | Pond's Evening Soothe 28Ct. Shipper | SHIP_NY_32K_100_13.25-7.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | 6AP06528 | Ponds Evening Soothe 28ct Shipper | Pond's Evening Soothe 28Ct. Shipper | SHIP_NY_32K_100_13.25-7.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Width | 7.75 | in | FDA,4 | 7.5625 | 7.9375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | 6AP06550 | Pond's Moisture Clean Twlts. Evening Soothe | Pond's Moisture Clean Twlts. Evening Soothe | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 20 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | 6AP06550 | Pond's Moisture Clean Twlts. Evening Soothe | Pond's Moisture Clean Twlts. Evening Soothe | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | 6AP06550 | Pond's Moisture Clean Twlts. Evening Soothe | Pond's Moisture Clean Twlts. Evening Soothe | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | 6AP06550 | Pond's Moisture Clean Twlts. Evening Soothe | Pond's Moisture Clean Twlts. Evening Soothe | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | 6AP06550 | Pond's Moisture Clean Twlts. Evening Soothe | Pond's Moisture Clean Twlts. Evening Soothe | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | Height | 10.375 | in | FDA,4 | 10.375 | 10.625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | 6AP06550 | Pond's Moisture Clean Twlts. Evening Soothe | Pond's Moisture Clean Twlts. Evening Soothe | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | Length | 17.5 | in | FDA,4 | 17.25 | 17.5 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | 6AP06550 | Pond's Moisture Clean Twlts. Evening Soothe | Pond's Moisture Clean Twlts. Evening Soothe | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | 6AP06550 | Pond's Moisture Clean Twlts. Evening Soothe | Pond's Moisture Clean Twlts. Evening Soothe | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | 6AP06550 | Pond's Moisture Clean Twlts. Evening Soothe | Pond's Moisture Clean Twlts. Evening Soothe | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | Width | 11.9375 | in | FDA,4 | 11.9375 | 12.1875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 6AP06728 | Ponds Original Fresh 28ct Shipper - Canada | Pond's Original Fresh 28Ct. Shipper - Canada | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 6AP06728 | Ponds Original Fresh 28ct Shipper - Canada | Pond's Original Fresh 28Ct. Shipper - Canada | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 6AP06728 | Ponds Original Fresh 28ct Shipper - Canada | Pond's Original Fresh 28Ct. Shipper - Canada | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 6AP06728 | Ponds Original Fresh 28ct Shipper - Canada | Pond's Original Fresh 28Ct. Shipper - Canada | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 6AP06728 | Ponds Original Fresh 28ct Shipper - Canada | Pond's Original Fresh 28Ct. Shipper - Canada | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | Height | 6 | in | FDA,4 | 5.875 | 6.25 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 6AP06728 | Ponds Original Fresh 28ct Shipper - Canada | Pond's Original Fresh 28Ct. Shipper - Canada | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | Length | 7 | in | FDA,4 | 6.9375 | 7.3125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 6AP06728 | Ponds Original Fresh 28ct Shipper - Canada | Pond's Original Fresh 28Ct. Shipper - Canada | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 6AP06728 | Ponds Original Fresh 28ct Shipper - Canada | Pond's Original Fresh 28Ct. Shipper - Canada | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | 6AP06728 | Ponds Original Fresh 28ct Shipper - Canada | Pond's Original Fresh 28Ct. Shipper - Canada | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | Width | 7.125 | in | FDA,4 | 6.8125 | 7.1875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
| 6AP07207 | 6AP07207 |  | Analytical-Simple Kind To Skin facial 7ct Shipper Canada | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 |  |  |  | left_only
| 6AP07207 | 6AP07207 |  | Analytical-Simple Kind To Skin facial 7ct Shipper Canada | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 |  |  |  | left_only
| 6AP07207 | 6AP07207 |  | Analytical-Simple Kind To Skin facial 7ct Shipper Canada | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 |  |  |  | left_only
| 6AP07207 | 6AP07207 |  | Analytical-Simple Kind To Skin facial 7ct Shipper Canada | NY-Analytical | 731 | Shipper Inspection | Height | 4 | in | FDA,4 | 3.9375 | 4.0625 | TM 141 / Visual | TM141 |  |  |  | left_only
| 6AP07207 | 6AP07207 |  | Analytical-Simple Kind To Skin facial 7ct Shipper Canada | NY-Analytical | 731 | Shipper Inspection | Length | 13 | in | FDA,4 | 12.9375 | 13.0625 | TM 141 / Visual | TM141 |  |  |  | left_only
| 6AP07207 | 6AP07207 |  | Analytical-Simple Kind To Skin facial 7ct Shipper Canada | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 |  |  |  | left_only
| 6AP07207 | 6AP07207 |  | Analytical-Simple Kind To Skin facial 7ct Shipper Canada | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 |  |  |  | left_only
| 6AP07207 | 6AP07207 |  | Analytical-Simple Kind To Skin facial 7ct Shipper Canada | NY-Analytical | 731 | Shipper Inspection | Width | 7.375 | in | FDA,4 | 7.25 | 7.375 | TM 141 / Visual | TM141 |  |  |  | left_only
6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 6AP09215 | Ponds Original Fresh 15ct Shipper | Pond's Original Fresh 15Ct. Shipper | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 6AP09215 | Ponds Original Fresh 15ct Shipper | Pond's Original Fresh 15Ct. Shipper | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 6AP09215 | Ponds Original Fresh 15ct Shipper | Pond's Original Fresh 15Ct. Shipper | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 6AP09215 | Ponds Original Fresh 15ct Shipper | Pond's Original Fresh 15Ct. Shipper | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 6AP09215 | Ponds Original Fresh 15ct Shipper | Pond's Original Fresh 15Ct. Shipper | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | Height | 5.75 | in | FDA,4 | 5.5625 | 5.9375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 6AP09215 | Ponds Original Fresh 15ct Shipper | Pond's Original Fresh 15Ct. Shipper | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | Length | 13.5 | in | FDA,4 | 13.25 | 13.625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 6AP09215 | Ponds Original Fresh 15ct Shipper | Pond's Original Fresh 15Ct. Shipper | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 6AP09215 | Ponds Original Fresh 15ct Shipper | Pond's Original Fresh 15Ct. Shipper | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | 6AP09215 | Ponds Original Fresh 15ct Shipper | Pond's Original Fresh 15Ct. Shipper | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | Width | 9.25 | in | FDA,4 | 9.0625 | 9.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 6AP09384 | Noxzema Anti-Blemish Pads 90ct Shipper | Noxzema Anti-Blemish Pads 90Ct. Shipper | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 6AP09384 | Noxzema Anti-Blemish Pads 90ct Shipper | Noxzema Anti-Blemish Pads 90Ct. Shipper | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 6AP09384 | Noxzema Anti-Blemish Pads 90ct Shipper | Noxzema Anti-Blemish Pads 90Ct. Shipper | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 6AP09384 | Noxzema Anti-Blemish Pads 90ct Shipper | Noxzema Anti-Blemish Pads 90Ct. Shipper | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 6AP09384 | Noxzema Anti-Blemish Pads 90ct Shipper | Noxzema Anti-Blemish Pads 90Ct. Shipper | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | Height | 3.8125 | in | FDA,4 | 3.5625 | 3.8125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 6AP09384 | Noxzema Anti-Blemish Pads 90ct Shipper | Noxzema Anti-Blemish Pads 90Ct. Shipper | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | Length | 8.5 | in | FDA,4 | 8.375 | 8.625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 6AP09384 | Noxzema Anti-Blemish Pads 90ct Shipper | Noxzema Anti-Blemish Pads 90Ct. Shipper | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 6AP09384 | Noxzema Anti-Blemish Pads 90ct Shipper | Noxzema Anti-Blemish Pads 90Ct. Shipper | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | 6AP09384 | Noxzema Anti-Blemish Pads 90ct Shipper | Noxzema Anti-Blemish Pads 90Ct. Shipper | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | Width | 5.75 | in | FDA,4 | 5.625 | 5.875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP11325/SHIP_NY_44K_100_10.625-9.75-7.75 | 6AP11325 | Garnier Skin Active MUR Shipper | Garnier Skin ACt. ive Mur Shipper | SHIP_NY_44K_100_10.625-9.75-7.75 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP11325/SHIP_NY_44K_100_10.625-9.75-7.75 | 6AP11325 | Garnier Skin Active MUR Shipper | Garnier Skin ACt. ive Mur Shipper | SHIP_NY_44K_100_10.625-9.75-7.75 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP11325/SHIP_NY_44K_100_10.625-9.75-7.75 | 6AP11325 | Garnier Skin Active MUR Shipper | Garnier Skin ACt. ive Mur Shipper | SHIP_NY_44K_100_10.625-9.75-7.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP11325/SHIP_NY_44K_100_10.625-9.75-7.75 | 6AP11325 | Garnier Skin Active MUR Shipper | Garnier Skin ACt. ive Mur Shipper | SHIP_NY_44K_100_10.625-9.75-7.75 | NY-Analytical | 731 | Shipper Inspection | Height | 7.75 | in | FDA,4 | 7.5625 | 7.9375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP11325/SHIP_NY_44K_100_10.625-9.75-7.75 | 6AP11325 | Garnier Skin Active MUR Shipper | Garnier Skin ACt. ive Mur Shipper | SHIP_NY_44K_100_10.625-9.75-7.75 | NY-Analytical | 731 | Shipper Inspection | Length | 10.625 | in | FDA,4 | 10.4375 | 10.8125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP11325/SHIP_NY_44K_100_10.625-9.75-7.75 | 6AP11325 | Garnier Skin Active MUR Shipper | Garnier Skin ACt. ive Mur Shipper | SHIP_NY_44K_100_10.625-9.75-7.75 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP11325/SHIP_NY_44K_100_10.625-9.75-7.75 | 6AP11325 | Garnier Skin Active MUR Shipper | Garnier Skin ACt. ive Mur Shipper | SHIP_NY_44K_100_10.625-9.75-7.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP11325/SHIP_NY_44K_100_10.625-9.75-7.75 | 6AP11325 | Garnier Skin Active MUR Shipper | Garnier Skin ACt. ive Mur Shipper | SHIP_NY_44K_100_10.625-9.75-7.75 | NY-Analytical | 731 | Shipper Inspection | Width | 9.75 | in | FDA,4 | 9.5625 | 9.9375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | 6AP13472 | PDI Sani-Hands Shipper | Sani-Hands Shipper | SHIP_NY_32K_100_16.3125-12.25-7.5625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | 6AP13472 | PDI Sani-Hands Shipper | Sani-Hands Shipper | SHIP_NY_32K_100_16.3125-12.25-7.5625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | 6AP13472 | PDI Sani-Hands Shipper | Sani-Hands Shipper | SHIP_NY_32K_100_16.3125-12.25-7.5625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | 6AP13472 | PDI Sani-Hands Shipper | Sani-Hands Shipper | SHIP_NY_32K_100_16.3125-12.25-7.5625 | NY-Analytical | 731 | Shipper Inspection | Height | 7.75 | in | FDA,4 | 7.375 | 7.75 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | 6AP13472 | PDI Sani-Hands Shipper | Sani-Hands Shipper | SHIP_NY_32K_100_16.3125-12.25-7.5625 | NY-Analytical | 731 | Shipper Inspection | Length | 16.25 | in | FDA,4 | 16.125 | 16.5 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | 6AP13472 | PDI Sani-Hands Shipper | Sani-Hands Shipper | SHIP_NY_32K_100_16.3125-12.25-7.5625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | 6AP13472 | PDI Sani-Hands Shipper | Sani-Hands Shipper | SHIP_NY_32K_100_16.3125-12.25-7.5625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | 6AP13472 | PDI Sani-Hands Shipper | Sani-Hands Shipper | SHIP_NY_32K_100_16.3125-12.25-7.5625 | NY-Analytical | 731 | Shipper Inspection | Width | 12.25 | in | FDA,4 | 12.0625 | 12.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | 6AP15984 | PDI Sani-Hands Shipper | Sani-Hands Shipper | SHIP_NY_44K_100_15.25-10.375-7.6875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | 6AP15984 | PDI Sani-Hands Shipper | Sani-Hands Shipper | SHIP_NY_44K_100_15.25-10.375-7.6875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | 6AP15984 | PDI Sani-Hands Shipper | Sani-Hands Shipper | SHIP_NY_44K_100_15.25-10.375-7.6875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | 6AP15984 | PDI Sani-Hands Shipper | Sani-Hands Shipper | SHIP_NY_44K_100_15.25-10.375-7.6875 | NY-Analytical | 731 | Shipper Inspection | Height | 7.75 | in | FDA,4 | 7.5 | 7.875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | 6AP15984 | PDI Sani-Hands Shipper | Sani-Hands Shipper | SHIP_NY_44K_100_15.25-10.375-7.6875 | NY-Analytical | 731 | Shipper Inspection | Length | 15.125 | in | FDA,4 | 15.0625 | 15.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | 6AP15984 | PDI Sani-Hands Shipper | Sani-Hands Shipper | SHIP_NY_44K_100_15.25-10.375-7.6875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | 6AP15984 | PDI Sani-Hands Shipper | Sani-Hands Shipper | SHIP_NY_44K_100_15.25-10.375-7.6875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | 6AP15984 | PDI Sani-Hands Shipper | Sani-Hands Shipper | SHIP_NY_44K_100_15.25-10.375-7.6875 | NY-Analytical | 731 | Shipper Inspection | Width | 10.5 | in | FDA,4 | 10.1875 | 10.5625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP22325/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | 6AP22325 | Simple Kind To Skin Facial Shipper | Simple Kind To Skin Facial Shipper | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Shipper Inspection |  | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | 6AP24284 | Sani-Cloth Prime Extra Large Wipe | Sani-Cloth Prime Extra Large Wipe | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | 6AP24284 | Sani-Cloth Prime Extra Large Wipe | Sani-Cloth Prime Extra Large Wipe | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 40 | Lbs./In. | FDA,0 | 40 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | 6AP24284 | Sani-Cloth Prime Extra Large Wipe | Sani-Cloth Prime Extra Large Wipe | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | 6AP24284 | Sani-Cloth Prime Extra Large Wipe | Sani-Cloth Prime Extra Large Wipe | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Height | 9.625 | in | FDA,4 | 9.4375 | 9.8125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | 6AP24284 | Sani-Cloth Prime Extra Large Wipe | Sani-Cloth Prime Extra Large Wipe | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Length | 15.375 | in | FDA,4 | 15.1875 | 15.5625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | 6AP24284 | Sani-Cloth Prime Extra Large Wipe | Sani-Cloth Prime Extra Large Wipe | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | 6AP24284 | Sani-Cloth Prime Extra Large Wipe | Sani-Cloth Prime Extra Large Wipe | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | 6AP24284 | Sani-Cloth Prime Extra Large Wipe | Sani-Cloth Prime Extra Large Wipe | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Width | 10.5625 | in | FDA,4 | 10.375 | 10.75 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP24372/SHIP_NY_44K_100_15.25-10.375-15.4375 | 6AP24372 | PDI Sani-Cloth Bleach Shipper | Sani-Cloth Bleach Shipper | SHIP_NY_44K_100_15.25-10.375-15.4375 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP24372/SHIP_NY_44K_100_15.25-10.375-15.4375 | 6AP24372 | PDI Sani-Cloth Bleach Shipper | Sani-Cloth Bleach Shipper | SHIP_NY_44K_100_15.25-10.375-15.4375 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP24372/SHIP_NY_44K_100_15.25-10.375-15.4375 | 6AP24372 | PDI Sani-Cloth Bleach Shipper | Sani-Cloth Bleach Shipper | SHIP_NY_44K_100_15.25-10.375-15.4375 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP24372/SHIP_NY_44K_100_15.25-10.375-15.4375 | 6AP24372 | PDI Sani-Cloth Bleach Shipper | Sani-Cloth Bleach Shipper | SHIP_NY_44K_100_15.25-10.375-15.4375 | NY-Analytical | 731 | Shipper Inspection | Height | 15.4375 | in | FDA,4 | 15.25 | 15.625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP24372/SHIP_NY_44K_100_15.25-10.375-15.4375 | 6AP24372 | PDI Sani-Cloth Bleach Shipper | Sani-Cloth Bleach Shipper | SHIP_NY_44K_100_15.25-10.375-15.4375 | NY-Analytical | 731 | Shipper Inspection | Length | 15.25 | in | FDA,4 | 15.0625 | 15.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP24372/SHIP_NY_44K_100_15.25-10.375-15.4375 | 6AP24372 | PDI Sani-Cloth Bleach Shipper | Sani-Cloth Bleach Shipper | SHIP_NY_44K_100_15.25-10.375-15.4375 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP24372/SHIP_NY_44K_100_15.25-10.375-15.4375 | 6AP24372 | PDI Sani-Cloth Bleach Shipper | Sani-Cloth Bleach Shipper | SHIP_NY_44K_100_15.25-10.375-15.4375 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP24372/SHIP_NY_44K_100_15.25-10.375-15.4375 | 6AP24372 | PDI Sani-Cloth Bleach Shipper | Sani-Cloth Bleach Shipper | SHIP_NY_44K_100_15.25-10.375-15.4375 | NY-Analytical | 731 | Shipper Inspection | Width | 10.375 | in | FDA,4 | 10.1875 | 10.5625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP24684/SHIP_NY_40K_100_15.375-10.5625-9.625 | 6AP24684 | PDI Sani-Cloth Bleach Extra large Shipper | Sani-Cloth Bleach Extra Large Shipper | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP27284/SHIP_NY_44K_100_15.375-10.5625-9.625 | 6AP27284 | Eco-Wipe Duo | Eco-Wipe Duo | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP27284/SHIP_NY_44K_100_15.375-10.5625-9.625 | 6AP27284 | Eco-Wipe Duo | Eco-Wipe Duo | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP27284/SHIP_NY_44K_100_15.375-10.5625-9.625 | 6AP27284 | Eco-Wipe Duo | Eco-Wipe Duo | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP27284/SHIP_NY_44K_100_15.375-10.5625-9.625 | 6AP27284 | Eco-Wipe Duo | Eco-Wipe Duo | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Height | 9.625 | in | FDA,4 | 9.4375 | 9.8125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP27284/SHIP_NY_44K_100_15.375-10.5625-9.625 | 6AP27284 | Eco-Wipe Duo | Eco-Wipe Duo | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Length | 15.25 | in | FDA,4 | 15.1875 | 15.5625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP27284/SHIP_NY_44K_100_15.375-10.5625-9.625 | 6AP27284 | Eco-Wipe Duo | Eco-Wipe Duo | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP27284/SHIP_NY_44K_100_15.375-10.5625-9.625 | 6AP27284 | Eco-Wipe Duo | Eco-Wipe Duo | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AP27284/SHIP_NY_44K_100_15.375-10.5625-9.625 | 6AP27284 | Eco-Wipe Duo | Eco-Wipe Duo | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Width | 10.437 | in | FDA,4 | 10.375 | 10.75 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
| 6AP71310 | Cetaphil 10ct Unprinted Shipper |  | Analytical - SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual | TM141 |  |  |  | left_only
| 6AP71310 | Cetaphil 10ct Unprinted Shipper |  | Analytical - SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 |  |  |  | left_only
| 6AP71310 | Cetaphil 10ct Unprinted Shipper |  | Analytical - SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | TM141 |  |  |  | left_only
| 6AP71310 | Cetaphil 10ct Unprinted Shipper |  | Analytical - SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 |  |  |  | left_only
| 6AP71310 | Cetaphil 10ct Unprinted Shipper |  | Analytical - SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | Height | 9 | in | FDA,4 | 8.8125 | 9.1875 | TM 141 / Visual | TM141 |  |  |  | left_only
| 6AP71310 | Cetaphil 10ct Unprinted Shipper |  | Analytical - SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | Length | 18.75 | in | FDA,4 | 18.6875 | 19.0625 | TM 141 / Visual | TM141 |  |  |  | left_only
| 6AP71310 | Cetaphil 10ct Unprinted Shipper |  | Analytical - SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 |  |  |  | left_only
| 6AP71310 | Cetaphil 10ct Unprinted Shipper |  | Analytical - SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 |  |  |  | left_only
| 6AP71310 | Cetaphil 10ct Unprinted Shipper |  | Analytical - SHIP_NY_44K_100_18.875-11.625-9 | NY-Analytical | 731 | Shipper Inspection | Width | 11.625 | in | FDA,4 | 11.4375 | 11.8125 | TM 141 / Visual | TM141 |  |  |  | left_only
6AS27350/SHIP_44_2O | 6AS27350 | COR CHLORASCRUB SWAB BULK | Chlorascrub Swab Bulk | SHIP_44_2O | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AS27350/SHIP_44_2O | 6AS27350 | COR CHLORASCRUB SWAB BULK | Chlorascrub Swab Bulk | SHIP_44_2O | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AS27350/SHIP_44_2O | 6AS27350 | COR CHLORASCRUB SWAB BULK | Chlorascrub Swab Bulk | SHIP_44_2O | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AS27350/SHIP_44_2O | 6AS27350 | COR CHLORASCRUB SWAB BULK | Chlorascrub Swab Bulk | SHIP_44_2O | NY-Analytical | 731 | Shipper Inspection | Height | 9 | in | FDA,4 | 8.8125 | 9.1875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AS27350/SHIP_44_2O | 6AS27350 | COR CHLORASCRUB SWAB BULK | Chlorascrub Swab Bulk | SHIP_44_2O | NY-Analytical | 731 | Shipper Inspection | Length | 15.375 | in | FDA,4 | 15.1875 | 15.5625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AS27350/SHIP_44_2O | 6AS27350 | COR CHLORASCRUB SWAB BULK | Chlorascrub Swab Bulk | SHIP_44_2O | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AS27350/SHIP_44_2O | 6AS27350 | COR CHLORASCRUB SWAB BULK | Chlorascrub Swab Bulk | SHIP_44_2O | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AS27350/SHIP_44_2O | 6AS27350 | COR CHLORASCRUB SWAB BULK | Chlorascrub Swab Bulk | SHIP_44_2O | NY-Analytical | 731 | Shipper Inspection | Width | 11.375 | in | FDA,4 | 11.1875 | 11.5625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AS40750/SHIP_32_1AP | 6AS40750 | COR CHLORASCRUB SWABSTICK | Chlorascrub Swabstick | SHIP_32_1AP | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AS40750/SHIP_32_1AP | 6AS40750 | COR CHLORASCRUB SWABSTICK | Chlorascrub Swabstick | SHIP_32_1AP | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AS40750/SHIP_32_1AP | 6AS40750 | COR CHLORASCRUB SWABSTICK | Chlorascrub Swabstick | SHIP_32_1AP | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AS40750/SHIP_32_1AP | 6AS40750 | COR CHLORASCRUB SWABSTICK | Chlorascrub Swabstick | SHIP_32_1AP | NY-Analytical | 731 | Shipper Inspection | Height | 13.4375 | in | FDA,4 | 13.25 | 13.625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AS40750/SHIP_32_1AP | 6AS40750 | COR CHLORASCRUB SWABSTICK | Chlorascrub Swabstick | SHIP_32_1AP | NY-Analytical | 731 | Shipper Inspection | Length | 15.625 | in | FDA,4 | 15.4375 | 15.8125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AS40750/SHIP_32_1AP | 6AS40750 | COR CHLORASCRUB SWABSTICK | Chlorascrub Swabstick | SHIP_32_1AP | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AS40750/SHIP_32_1AP | 6AS40750 | COR CHLORASCRUB SWABSTICK | Chlorascrub Swabstick | SHIP_32_1AP | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AS40750/SHIP_32_1AP | 6AS40750 | COR CHLORASCRUB SWABSTICK | Chlorascrub Swabstick | SHIP_32_1AP | NY-Analytical | 731 | Shipper Inspection | Width | 7.25 | in | FDA,4 | 7.0625 | 7.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AS41950/SHIP_40_1O | 6AS41950 | COR CHLORASCRUB MAXI SWABSTICK | Chlorascrub Maxi Swabstick | SHIP_40_1O | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AS41950/SHIP_40_1O | 6AS41950 | COR CHLORASCRUB MAXI SWABSTICK | Chlorascrub Maxi Swabstick | SHIP_40_1O | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 40 | Lbs./In. | FDA,0 | 40 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AS41950/SHIP_40_1O | 6AS41950 | COR CHLORASCRUB MAXI SWABSTICK | Chlorascrub Maxi Swabstick | SHIP_40_1O | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AS41950/SHIP_40_1O | 6AS41950 | COR CHLORASCRUB MAXI SWABSTICK | Chlorascrub Maxi Swabstick | SHIP_40_1O | NY-Analytical | 731 | Shipper Inspection | Height | 8.3125 | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AS41950/SHIP_40_1O | 6AS41950 | COR CHLORASCRUB MAXI SWABSTICK | Chlorascrub Maxi Swabstick | SHIP_40_1O | NY-Analytical | 731 | Shipper Inspection | Length | 21.125 | in | FDA,4 | 20.9375 | 21.3125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AS41950/SHIP_40_1O | 6AS41950 | COR CHLORASCRUB MAXI SWABSTICK | Chlorascrub Maxi Swabstick | SHIP_40_1O | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AS41950/SHIP_40_1O | 6AS41950 | COR CHLORASCRUB MAXI SWABSTICK | Chlorascrub Maxi Swabstick | SHIP_40_1O | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AS41950/SHIP_40_1O | 6AS41950 | COR CHLORASCRUB MAXI SWABSTICK | Chlorascrub Maxi Swabstick | SHIP_40_1O | NY-Analytical | 731 | Shipper Inspection | Width | 13.5625 | in | FDA,4 | 13.375 | 13.75 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU12310/SHIP_32_1U | 6AU12310 | COR PREP-H 10CT SINGLE SK17132 | Prep H 10Ct. Single Sk17132 | SHIP_32_1U | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU12310/SHIP_32_1U | 6AU12310 | COR PREP-H 10CT SINGLE SK17132 | Prep H 10Ct. Single Sk17132 | SHIP_32_1U | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU12310/SHIP_32_1U | 6AU12310 | COR PREP-H 10CT SINGLE SK17132 | Prep H 10Ct. Single Sk17132 | SHIP_32_1U | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU12310/SHIP_32_1U | 6AU12310 | COR PREP-H 10CT SINGLE SK17132 | Prep H 10Ct. Single Sk17132 | SHIP_32_1U | NY-Analytical | 731 | Shipper Inspection | Height | 8 | in | FDA,4 | 7.8125 | 8.1875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU12310/SHIP_32_1U | 6AU12310 | COR PREP-H 10CT SINGLE SK17132 | Prep H 10Ct. Single Sk17132 | SHIP_32_1U | NY-Analytical | 731 | Shipper Inspection | Length | 12.125 | in | FDA,4 | 11.8125 | 12.1875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU12310/SHIP_32_1U | 6AU12310 | COR PREP-H 10CT SINGLE SK17132 | Prep H 10Ct. Single Sk17132 | SHIP_32_1U | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU12310/SHIP_32_1U | 6AU12310 | COR PREP-H 10CT SINGLE SK17132 | Prep H 10Ct. Single Sk17132 | SHIP_32_1U | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU12310/SHIP_32_1U | 6AU12310 | COR PREP-H 10CT SINGLE SK17132 | Prep H 10Ct. Single Sk17132 | SHIP_32_1U | NY-Analytical | 731 | Shipper Inspection | Width | 6 | in | FDA,4 | 5.9375 | 6.3125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | 6AU130TY | Grime Boss Heavy Duty Wipes | Grime Boss Heavy Duty Wipes | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | 6AU130TY | Grime Boss Heavy Duty Wipes | Grime Boss Heavy Duty Wipes | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | 6AU130TY | Grime Boss Heavy Duty Wipes | Grime Boss Heavy Duty Wipes | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | 6AU130TY | Grime Boss Heavy Duty Wipes | Grime Boss Heavy Duty Wipes | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | 6AU130TY | Grime Boss Heavy Duty Wipes | Grime Boss Heavy Duty Wipes | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | Height | 4.5 | in | FDA,4 | 4.375 | 4.625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | 6AU130TY | Grime Boss Heavy Duty Wipes | Grime Boss Heavy Duty Wipes | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | Length | 6.625 | in | FDA,4 | 6.375 | 6.625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | 6AU130TY | Grime Boss Heavy Duty Wipes | Grime Boss Heavy Duty Wipes | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | White (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | 6AU130TY | Grime Boss Heavy Duty Wipes | Grime Boss Heavy Duty Wipes | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Fold Over Tray w/header (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | 6AU130TY | Grime Boss Heavy Duty Wipes | Grime Boss Heavy Duty Wipes | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | Width | 9.625 | in | FDA,4 | 9.5 | 9.75 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | 6AU25206 | L 'Oreal Paris Sublime Bronze Self Tan Twlts | L'Oreal Paris Sublime Bronze Self Tan Twlts | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | 6AU25206 | L 'Oreal Paris Sublime Bronze Self Tan Twlts | L'Oreal Paris Sublime Bronze Self Tan Twlts | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | No Go - Does not meet Specifications (OOS) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | 6AU25206 | L 'Oreal Paris Sublime Bronze Self Tan Twlts | L'Oreal Paris Sublime Bronze Self Tan Twlts | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | 6AU25206 | L 'Oreal Paris Sublime Bronze Self Tan Twlts | L'Oreal Paris Sublime Bronze Self Tan Twlts | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | 6AU25206 | L 'Oreal Paris Sublime Bronze Self Tan Twlts | L'Oreal Paris Sublime Bronze Self Tan Twlts | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | Height | 4.75 | in | FDA,4 | 4.375 | 4.75 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | 6AU25206 | L 'Oreal Paris Sublime Bronze Self Tan Twlts | L'Oreal Paris Sublime Bronze Self Tan Twlts | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | Length | 11.375 | in | FDA,4 | 11.0625 | 11.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | 6AU25206 | L 'Oreal Paris Sublime Bronze Self Tan Twlts | L'Oreal Paris Sublime Bronze Self Tan Twlts | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | 6AU25206 | L 'Oreal Paris Sublime Bronze Self Tan Twlts | L'Oreal Paris Sublime Bronze Self Tan Twlts | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | 6AU25206 | L 'Oreal Paris Sublime Bronze Self Tan Twlts | L'Oreal Paris Sublime Bronze Self Tan Twlts | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | Width | 8.5 | in | FDA,4 | 8.25 | 8.625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 6AU53601 | Cetaphil Gentle Skin Cleans. Cloths Shipper | Cetaphil Gentle Skin Cleans. Cloths Shipper | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 50 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 6AU53601 | Cetaphil Gentle Skin Cleans. Cloths Shipper | Cetaphil Gentle Skin Cleans. Cloths Shipper | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 6AU53601 | Cetaphil Gentle Skin Cleans. Cloths Shipper | Cetaphil Gentle Skin Cleans. Cloths Shipper | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 6AU53601 | Cetaphil Gentle Skin Cleans. Cloths Shipper | Cetaphil Gentle Skin Cleans. Cloths Shipper | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 6AU53601 | Cetaphil Gentle Skin Cleans. Cloths Shipper | Cetaphil Gentle Skin Cleans. Cloths Shipper | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | Height | 8.5 | in | FDA,4 | 8.3125 | 8.6875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 6AU53601 | Cetaphil Gentle Skin Cleans. Cloths Shipper | Cetaphil Gentle Skin Cleans. Cloths Shipper | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | Length | 11.875 | in | FDA,4 | 11.625 | 12 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 6AU53601 | Cetaphil Gentle Skin Cleans. Cloths Shipper | Cetaphil Gentle Skin Cleans. Cloths Shipper | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 6AU53601 | Cetaphil Gentle Skin Cleans. Cloths Shipper | Cetaphil Gentle Skin Cleans. Cloths Shipper | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | 6AU53601 | Cetaphil Gentle Skin Cleans. Cloths Shipper | Cetaphil Gentle Skin Cleans. Cloths Shipper | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | Width | 9.5 | in | FDA,4 | 9.1875 | 9.5625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 6AU56400 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 50 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 6AU56400 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 6AU56400 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 6AU56400 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 6AU56400 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | Height | 7.8125 | in | FDA,4 | 7.5005 | 7.8755 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 6AU56400 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | Length | 11.75 | in | FDA,4 | 11.5625 | 11.9375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 6AU56400 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 6AU56400 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | 6AU56400 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | Width | 9.625 | in | FDA,4 | 9.4375 | 9.8125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | 6AU57314 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | 6AU57314 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | 6AU57314 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | 6AU57314 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | 6AU57314 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | Height | 8.0625 | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | 6AU57314 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | Length | 7.5625 | in | FDA,4 | 7.5625 | 7.9375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | 6AU57314 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | 6AU57314 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | 6AU57314 | Summer's Eve Fresh Cycle Cleansing Cloths | Summer's Eve Fresh Cycle Cleansing Cloths | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | Width | 6.375 | in | FDA,4 | 6.0625 | 6.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 6AU71230 | Preparation H Rapid Relief Totables w/Lidocaine | Prep H Rapid Relief Totables W/Lidocaine | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 50 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 6AU71230 | Preparation H Rapid Relief Totables w/Lidocaine | Prep H Rapid Relief Totables W/Lidocaine | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 6AU71230 | Preparation H Rapid Relief Totables w/Lidocaine | Prep H Rapid Relief Totables W/Lidocaine | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 6AU71230 | Preparation H Rapid Relief Totables w/Lidocaine | Prep H Rapid Relief Totables W/Lidocaine | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 6AU71230 | Preparation H Rapid Relief Totables w/Lidocaine | Prep H Rapid Relief Totables W/Lidocaine | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | Height | 7.75 | in | FDA,4 | 7.5625 | 7.9375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 6AU71230 | Preparation H Rapid Relief Totables w/Lidocaine | Prep H Rapid Relief Totables W/Lidocaine | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | Length | 9.5625 | in | FDA,4 | 9.3755 | 9.7505 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 6AU71230 | Preparation H Rapid Relief Totables w/Lidocaine | Prep H Rapid Relief Totables W/Lidocaine | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 6AU71230 | Preparation H Rapid Relief Totables w/Lidocaine | Prep H Rapid Relief Totables W/Lidocaine | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | 6AU71230 | Preparation H Rapid Relief Totables w/Lidocaine | Prep H Rapid Relief Totables W/Lidocaine | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | Width | 9.188 | in | FDA,4 | 9.0005 | 9.3755 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 6BA709SWN | Preparation H for Women Shipper | Prep H For Women Shipper | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 6BA709SWN | Preparation H for Women Shipper | Prep H For Women Shipper | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 6BA709SWN | Preparation H for Women Shipper | Prep H For Women Shipper | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 40 | Lbs./In. | FDA,0 | 40 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 6BA709SWN | Preparation H for Women Shipper | Prep H For Women Shipper | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 6BA709SWN | Preparation H for Women Shipper | Prep H For Women Shipper | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | Height | 6.125 | in | FDA,4 | 5.9375 | 6.3125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 6BA709SWN | Preparation H for Women Shipper | Prep H For Women Shipper | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | Length | 14.75 | in | FDA,4 | 14.5625 | 14.9375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 6BA709SWN | Preparation H for Women Shipper | Prep H For Women Shipper | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 6BA709SWN | Preparation H for Women Shipper | Prep H For Women Shipper | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | 6BA709SWN | Preparation H for Women Shipper | Prep H For Women Shipper | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | Width | 9.375 | in | FDA,4 | 9.1875 | 9.5625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BB10800/SHIP_32_1S | 6BB10800 | COR CHLORASCRUB SWAB SHIPPER | Chlorascrub Swab Shipper | SHIP_32_1S | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BB10800/SHIP_32_1S | 6BB10800 | COR CHLORASCRUB SWAB SHIPPER | Chlorascrub Swab Shipper | SHIP_32_1S | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BB10800/SHIP_32_1S | 6BB10800 | COR CHLORASCRUB SWAB SHIPPER | Chlorascrub Swab Shipper | SHIP_32_1S | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BB10800/SHIP_32_1S | 6BB10800 | COR CHLORASCRUB SWAB SHIPPER | Chlorascrub Swab Shipper | SHIP_32_1S | NY-Analytical | 731 | Shipper Inspection | Height | 5.3125 | in | FDA,4 | 5.125 | 5.5 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BB10800/SHIP_32_1S | 6BB10800 | COR CHLORASCRUB SWAB SHIPPER | Chlorascrub Swab Shipper | SHIP_32_1S | NY-Analytical | 731 | Shipper Inspection | Length | 11.875 | in | FDA,4 | 11.6875 | 12.0625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BB10800/SHIP_32_1S | 6BB10800 | COR CHLORASCRUB SWAB SHIPPER | Chlorascrub Swab Shipper | SHIP_32_1S | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BB10800/SHIP_32_1S | 6BB10800 | COR CHLORASCRUB SWAB SHIPPER | Chlorascrub Swab Shipper | SHIP_32_1S | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BB10800/SHIP_32_1S | 6BB10800 | COR CHLORASCRUB SWAB SHIPPER | Chlorascrub Swab Shipper | SHIP_32_1S | NY-Analytical | 731 | Shipper Inspection | Width | 6.25 | in | FDA,4 | 6.0625 | 6.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 6BP25372 | PDI Sani Cloth prime Shipper | Sani-Cloth Prime Shipper | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 6BP25372 | PDI Sani Cloth prime Shipper | Sani-Cloth Prime Shipper | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | Certificate of Compliance | CofC meets specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 6BP25372 | PDI Sani Cloth prime Shipper | Sani-Cloth Prime Shipper | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 6BP25372 | PDI Sani Cloth prime Shipper | Sani-Cloth Prime Shipper | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 6BP25372 | PDI Sani Cloth prime Shipper | Sani-Cloth Prime Shipper | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | Height | 15.25 | in | FDA,4 | 15 | 15.375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 6BP25372 | PDI Sani Cloth prime Shipper | Sani-Cloth Prime Shipper | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | Length | 15.125 | in | FDA,4 | 15 | 15.375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 6BP25372 | PDI Sani Cloth prime Shipper | Sani-Cloth Prime Shipper | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 6BP25372 | PDI Sani Cloth prime Shipper | Sani-Cloth Prime Shipper | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 6BP25372 | PDI Sani Cloth prime Shipper | Sani-Cloth Prime Shipper | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | Width | 10.25 | in | FDA,4 | 10 | 10.375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BP54072/SHIP_44_2L | 6BP54072 | PDI Sani Cloth Bleach Shipper | Sani-Cloth Bleach Shipper | SHIP_44_2L | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BP54072/SHIP_44_2L | 6BP54072 | PDI Sani Cloth Bleach Shipper | Sani-Cloth Bleach Shipper | SHIP_44_2L | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BP54072/SHIP_44_2L | 6BP54072 | PDI Sani Cloth Bleach Shipper | Sani-Cloth Bleach Shipper | SHIP_44_2L | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BP54072/SHIP_44_2L | 6BP54072 | PDI Sani Cloth Bleach Shipper | Sani-Cloth Bleach Shipper | SHIP_44_2L | NY-Analytical | 731 | Shipper Inspection | Height | 15.4375 | in | FDA,4 | 15.25 | 15.625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BP54072/SHIP_44_2L | 6BP54072 | PDI Sani Cloth Bleach Shipper | Sani-Cloth Bleach Shipper | SHIP_44_2L | NY-Analytical | 731 | Shipper Inspection | Length | 15.25 | in | FDA,4 | 15.0625 | 15.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BP54072/SHIP_44_2L | 6BP54072 | PDI Sani Cloth Bleach Shipper | Sani-Cloth Bleach Shipper | SHIP_44_2L | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BP54072/SHIP_44_2L | 6BP54072 | PDI Sani Cloth Bleach Shipper | Sani-Cloth Bleach Shipper | SHIP_44_2L | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BP54072/SHIP_44_2L | 6BP54072 | PDI Sani Cloth Bleach Shipper | Sani-Cloth Bleach Shipper | SHIP_44_2L | NY-Analytical | 731 | Shipper Inspection | Width | 10.375 | in | FDA,4 | 10.1875 | 10.5625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | 6BU13195 | PDI Sani-Cloth Prime Shipper | Sani-Cloth Prime Shipper | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | 6BU13195 | PDI Sani-Cloth Prime Shipper | Sani-Cloth Prime Shipper | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | 6BU13195 | PDI Sani-Cloth Prime Shipper | Sani-Cloth Prime Shipper | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | 6BU13195 | PDI Sani-Cloth Prime Shipper | Sani-Cloth Prime Shipper | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | 6BU13195 | PDI Sani-Cloth Prime Shipper | Sani-Cloth Prime Shipper | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | Height | 12.25 | in | FDA,4 | 12.0625 | 12.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | 6BU13195 | PDI Sani-Cloth Prime Shipper | Sani-Cloth Prime Shipper | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | Length | 13.5 | in | FDA,4 | 13.4375 | 13.8125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | 6BU13195 | PDI Sani-Cloth Prime Shipper | Sani-Cloth Prime Shipper | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | 6BU13195 | PDI Sani-Cloth Prime Shipper | Sani-Cloth Prime Shipper | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | 6BU13195 | PDI Sani-Cloth Prime Shipper | Sani-Cloth Prime Shipper | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | Width | 4.375 | in | FDA,4 | 4.3125 | 4.6875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | 6BU25206 | L 'Oreal Paris Sublime Bronze Self Tan Twlts | L'Oreal Paris Sublime Bronze Self Tan Twlts | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 10 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | 6BU25206 | L 'Oreal Paris Sublime Bronze Self Tan Twlts | L'Oreal Paris Sublime Bronze Self Tan Twlts | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | 6BU25206 | L 'Oreal Paris Sublime Bronze Self Tan Twlts | L'Oreal Paris Sublime Bronze Self Tan Twlts | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | 6BU25206 | L 'Oreal Paris Sublime Bronze Self Tan Twlts | L'Oreal Paris Sublime Bronze Self Tan Twlts | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | 6BU25206 | L 'Oreal Paris Sublime Bronze Self Tan Twlts | L'Oreal Paris Sublime Bronze Self Tan Twlts | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | Height | 4.5 | in | FDA,4 | 4.4375 | 4.8125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | 6BU25206 | L 'Oreal Paris Sublime Bronze Self Tan Twlts | L'Oreal Paris Sublime Bronze Self Tan Twlts | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | Length | 11.3125 | in | FDA,4 | 11.125 | 11.5 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | 6BU25206 | L 'Oreal Paris Sublime Bronze Self Tan Twlts | L'Oreal Paris Sublime Bronze Self Tan Twlts | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | 6BU25206 | L 'Oreal Paris Sublime Bronze Self Tan Twlts | L'Oreal Paris Sublime Bronze Self Tan Twlts | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | 6BU25206 | L 'Oreal Paris Sublime Bronze Self Tan Twlts | L'Oreal Paris Sublime Bronze Self Tan Twlts | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | Width | 8.25 | in | FDA,4 | 8.25 | 8.625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | 6CA972CK | Sani Pro. - Table Turners Cleaning Multi-Surface | Sani Professional - Table Turners Cleaning Multi-Surface | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 32 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | 6CA972CK | Sani Pro. - Table Turners Cleaning Multi-Surface | Sani Professional - Table Turners Cleaning Multi-Surface | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | 6CA972CK | Sani Pro. - Table Turners Cleaning Multi-Surface | Sani Professional - Table Turners Cleaning Multi-Surface | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | 6CA972CK | Sani Pro. - Table Turners Cleaning Multi-Surface | Sani Professional - Table Turners Cleaning Multi-Surface | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | 6CA972CK | Sani Pro. - Table Turners Cleaning Multi-Surface | Sani Professional - Table Turners Cleaning Multi-Surface | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | Height | 10.25 | in | FDA,4 | 10.0625 | 10.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | 6CA972CK | Sani Pro. - Table Turners Cleaning Multi-Surface | Sani Professional - Table Turners Cleaning Multi-Surface | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | Length | 18 | in | FDA,4 | 17.8125 | 18.1875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | 6CA972CK | Sani Pro. - Table Turners Cleaning Multi-Surface | Sani Professional - Table Turners Cleaning Multi-Surface | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | 6CA972CK | Sani Pro. - Table Turners Cleaning Multi-Surface | Sani Professional - Table Turners Cleaning Multi-Surface | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | 6CA972CK | Sani Pro. - Table Turners Cleaning Multi-Surface | Sani Professional - Table Turners Cleaning Multi-Surface | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | Width | 8.25 | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | 6CJ21396 | Hygea Multi-Purpose Washcloths | Hygea Multi-Purpose Washcloths | SHIP_NY_44K_100_16.625-9.375-7 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | 6CJ21396 | Hygea Multi-Purpose Washcloths | Hygea Multi-Purpose Washcloths | SHIP_NY_44K_100_16.625-9.375-7 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | 6CJ21396 | Hygea Multi-Purpose Washcloths | Hygea Multi-Purpose Washcloths | SHIP_NY_44K_100_16.625-9.375-7 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | 6CJ21396 | Hygea Multi-Purpose Washcloths | Hygea Multi-Purpose Washcloths | SHIP_NY_44K_100_16.625-9.375-7 | NY-Analytical | 731 | Shipper Inspection | Height | 7.1875 | in | FDA,4 | 6.8125 | 7.1875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | 6CJ21396 | Hygea Multi-Purpose Washcloths | Hygea Multi-Purpose Washcloths | SHIP_NY_44K_100_16.625-9.375-7 | NY-Analytical | 731 | Shipper Inspection | Length | 16.75 | in | FDA,4 | 16.4375 | 16.8125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | 6CJ21396 | Hygea Multi-Purpose Washcloths | Hygea Multi-Purpose Washcloths | SHIP_NY_44K_100_16.625-9.375-7 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | 6CJ21396 | Hygea Multi-Purpose Washcloths | Hygea Multi-Purpose Washcloths | SHIP_NY_44K_100_16.625-9.375-7 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | 6CJ21396 | Hygea Multi-Purpose Washcloths | Hygea Multi-Purpose Washcloths | SHIP_NY_44K_100_16.625-9.375-7 | NY-Analytical | 731 | Shipper Inspection | Width | 9.5 | in | FDA,4 | 9.1875 | 9.5625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6CP22884/SHIP_NY_44K_100_15.375-10.5625-9.625 | 6CP22884 | Sani Professional Sani-Cloth Disinfecting Wipes | Sani Professional Sani-Cloth DisinfeCt. ing Wipes | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | [5 empty] | Shipper Inspection |  | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 6DC22370 | PDI Hygea Sterile Saline Wipe | PDI Hygea Sterile Saline Wipe | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 6DC22370 | PDI Hygea Sterile Saline Wipe | PDI Hygea Sterile Saline Wipe | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 6DC22370 | PDI Hygea Sterile Saline Wipe | PDI Hygea Sterile Saline Wipe | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 6DC22370 | PDI Hygea Sterile Saline Wipe | PDI Hygea Sterile Saline Wipe | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 6DC22370 | PDI Hygea Sterile Saline Wipe | PDI Hygea Sterile Saline Wipe | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | Height | 10.375 | in | FDA,4 | 10.1875 | 10.5625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 6DC22370 | PDI Hygea Sterile Saline Wipe | PDI Hygea Sterile Saline Wipe | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | Length | 11 | in | FDA,4 | 10.9375 | 11.3125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 6DC22370 | PDI Hygea Sterile Saline Wipe | PDI Hygea Sterile Saline Wipe | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 6DC22370 | PDI Hygea Sterile Saline Wipe | PDI Hygea Sterile Saline Wipe | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | 6DC22370 | PDI Hygea Sterile Saline Wipe | PDI Hygea Sterile Saline Wipe | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | Width | 6.25 | in | FDA,4 | 6.1875 | 6.5625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | 6DD33333 | Sani Professional Sani Hands Shipper | Sani Professional Sani-Hands Shipper | SHIP_NY_44K_100_21-12.5-10.5 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | 6DD33333 | Sani Professional Sani Hands Shipper | Sani Professional Sani-Hands Shipper | SHIP_NY_44K_100_21-12.5-10.5 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | 6DD33333 | Sani Professional Sani Hands Shipper | Sani Professional Sani-Hands Shipper | SHIP_NY_44K_100_21-12.5-10.5 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | 6DD33333 | Sani Professional Sani Hands Shipper | Sani Professional Sani-Hands Shipper | SHIP_NY_44K_100_21-12.5-10.5 | NY-Analytical | 731 | Shipper Inspection | Height | 10.5 | in | FDA,4 | 10.3125 | 10.6875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | 6DD33333 | Sani Professional Sani Hands Shipper | Sani Professional Sani-Hands Shipper | SHIP_NY_44K_100_21-12.5-10.5 | NY-Analytical | 731 | Shipper Inspection | Length | 21 | in | FDA,4 | 20.8125 | 21.1875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | 6DD33333 | Sani Professional Sani Hands Shipper | Sani Professional Sani-Hands Shipper | SHIP_NY_44K_100_21-12.5-10.5 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | 6DD33333 | Sani Professional Sani Hands Shipper | Sani Professional Sani-Hands Shipper | SHIP_NY_44K_100_21-12.5-10.5 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | 6DD33333 | Sani Professional Sani Hands Shipper | Sani Professional Sani-Hands Shipper | SHIP_NY_44K_100_21-12.5-10.5 | NY-Analytical | 731 | Shipper Inspection | Width | 12.625 | in | FDA,4 | 12.3125 | 12.6875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6DP63884/SHIP_NY_44K_100_15.375-10.5625-9.625 | 6DP63884 | PDI Sani Cloth AF3 Shipper | Sani-Cloth AF3 Shipper | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | 6EP44584 | Sani Professional - Sani Hands Antibac. Wipes. | Sani Professional - Sani-Hands Antibac. Wipes. | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 50 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | 6EP44584 | Sani Professional - Sani Hands Antibac. Wipes. | Sani Professional - Sani-Hands Antibac. Wipes. | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | 6EP44584 | Sani Professional - Sani Hands Antibac. Wipes. | Sani Professional - Sani-Hands Antibac. Wipes. | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | 6EP44584 | Sani Professional - Sani Hands Antibac. Wipes. | Sani Professional - Sani-Hands Antibac. Wipes. | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | 6EP44584 | Sani Professional - Sani Hands Antibac. Wipes. | Sani Professional - Sani-Hands Antibac. Wipes. | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | Height | 9.75 | in | FDA,4 | 9.4375 | 9.8125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | 6EP44584 | Sani Professional - Sani Hands Antibac. Wipes. | Sani Professional - Sani-Hands Antibac. Wipes. | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | Length | 15.25 | in | FDA,4 | 15 | 15.375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | 6EP44584 | Sani Professional - Sani Hands Antibac. Wipes. | Sani Professional - Sani-Hands Antibac. Wipes. | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | 6EP44584 | Sani Professional - Sani Hands Antibac. Wipes. | Sani Professional - Sani-Hands Antibac. Wipes. | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | 6EP44584 | Sani Professional - Sani Hands Antibac. Wipes. | Sani Professional - Sani-Hands Antibac. Wipes. | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | Width | 10.25 | in | FDA,4 | 10 | 10.375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 6FB71200 | PDI Nail Polish Remover Pad | PDI Nail Polish Remover Pad | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 50 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 6FB71200 | PDI Nail Polish Remover Pad | PDI Nail Polish Remover Pad | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 6FB71200 | PDI Nail Polish Remover Pad | PDI Nail Polish Remover Pad | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 6FB71200 | PDI Nail Polish Remover Pad | PDI Nail Polish Remover Pad | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 6FB71200 | PDI Nail Polish Remover Pad | PDI Nail Polish Remover Pad | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Height | 5.125 | in | FDA,4 | 4.8125 | 5.1875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 6FB71200 | PDI Nail Polish Remover Pad | PDI Nail Polish Remover Pad | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Length | 14.75 | in | FDA,4 | 14.4375 | 14.8125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 6FB71200 | PDI Nail Polish Remover Pad | PDI Nail Polish Remover Pad | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 6FB71200 | PDI Nail Polish Remover Pad | PDI Nail Polish Remover Pad | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC w/ shyed flaps (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | 6FB71200 | PDI Nail Polish Remover Pad | PDI Nail Polish Remover Pad | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Width | 6.344 | in | FDA,4 | 6.1875 | 6.5625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | 6FJ14108 | Hygea Multi Purpose Washcloth Shipper | Hygea Multi Purpose Washcloth Shipper | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | 6FJ14108 | Hygea Multi Purpose Washcloth Shipper | Hygea Multi Purpose Washcloth Shipper | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | 6FJ14108 | Hygea Multi Purpose Washcloth Shipper | Hygea Multi Purpose Washcloth Shipper | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | 6FJ14108 | Hygea Multi Purpose Washcloth Shipper | Hygea Multi Purpose Washcloth Shipper | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | 6FJ14108 | Hygea Multi Purpose Washcloth Shipper | Hygea Multi Purpose Washcloth Shipper | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | Height | 7 | in | FDA,4 | 6.8125 | 7.1875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | 6FJ14108 | Hygea Multi Purpose Washcloth Shipper | Hygea Multi Purpose Washcloth Shipper | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | Length | 16.625 | in | FDA,4 | 16.4375 | 16.8125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | 6FJ14108 | Hygea Multi Purpose Washcloth Shipper | Hygea Multi Purpose Washcloth Shipper | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | 6FJ14108 | Hygea Multi Purpose Washcloth Shipper | Hygea Multi Purpose Washcloth Shipper | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | 6FJ14108 | Hygea Multi Purpose Washcloth Shipper | Hygea Multi Purpose Washcloth Shipper | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | Width | 8.5 | in | FDA,4 | 8.4375 | 8.8125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6FJ14143/SHIP_NY_44K_100_16.5-10.875-5.5 | 6FJ14143 | Hygea Premium Multi-Purpose Washcloth | Hygea Premium Multi-Purpose Washcloth | SHIP_NY_44K_100_16.5-10.875-5.5 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6FJ14143/SHIP_NY_44K_100_16.5-10.875-5.5 | 6FJ14143 | Hygea Premium Multi-Purpose Washcloth | Hygea Premium Multi-Purpose Washcloth | SHIP_NY_44K_100_16.5-10.875-5.5 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6FJ14143/SHIP_NY_44K_100_16.5-10.875-5.5 | 6FJ14143 | Hygea Premium Multi-Purpose Washcloth | Hygea Premium Multi-Purpose Washcloth | SHIP_NY_44K_100_16.5-10.875-5.5 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6FJ14143/SHIP_NY_44K_100_16.5-10.875-5.5 | 6FJ14143 | Hygea Premium Multi-Purpose Washcloth | Hygea Premium Multi-Purpose Washcloth | SHIP_NY_44K_100_16.5-10.875-5.5 | NY-Analytical | 731 | Shipper Inspection | Height | 5.5 | in | FDA,4 | 5.3125 | 5.6875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6FJ14143/SHIP_NY_44K_100_16.5-10.875-5.5 | 6FJ14143 | Hygea Premium Multi-Purpose Washcloth | Hygea Premium Multi-Purpose Washcloth | SHIP_NY_44K_100_16.5-10.875-5.5 | NY-Analytical | 731 | Shipper Inspection | Length | 16.5 | in | FDA,4 | 16.3125 | 16.6875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6FJ14143/SHIP_NY_44K_100_16.5-10.875-5.5 | 6FJ14143 | Hygea Premium Multi-Purpose Washcloth | Hygea Premium Multi-Purpose Washcloth | SHIP_NY_44K_100_16.5-10.875-5.5 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6FJ14143/SHIP_NY_44K_100_16.5-10.875-5.5 | 6FJ14143 | Hygea Premium Multi-Purpose Washcloth | Hygea Premium Multi-Purpose Washcloth | SHIP_NY_44K_100_16.5-10.875-5.5 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6FJ14143/SHIP_NY_44K_100_16.5-10.875-5.5 | 6FJ14143 | Hygea Premium Multi-Purpose Washcloth | Hygea Premium Multi-Purpose Washcloth | SHIP_NY_44K_100_16.5-10.875-5.5 | NY-Analytical | 731 | Shipper Inspection | Width | 10.875 | in | FDA,4 | 10.6875 | 11.0625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | 6GB10300 | PDI CHG / Isopropyl Alcohol Shipper | PDI CHG / Isopropyl Alcohol Shipper | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | 6GB10300 | PDI CHG / Isopropyl Alcohol Shipper | PDI CHG / Isopropyl Alcohol Shipper | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | 6GB10300 | PDI CHG / Isopropyl Alcohol Shipper | PDI CHG / Isopropyl Alcohol Shipper | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | 6GB10300 | PDI CHG / Isopropyl Alcohol Shipper | PDI CHG / Isopropyl Alcohol Shipper | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | 6GB10300 | PDI CHG / Isopropyl Alcohol Shipper | PDI CHG / Isopropyl Alcohol Shipper | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Height | 5 | in | FDA,4 | 4.8125 | 5.1875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | 6GB10300 | PDI CHG / Isopropyl Alcohol Shipper | PDI CHG / Isopropyl Alcohol Shipper | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Length | 14.75 | in | FDA,4 | 14.4375 | 14.8125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | 6GB10300 | PDI CHG / Isopropyl Alcohol Shipper | PDI CHG / Isopropyl Alcohol Shipper | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | 6GB10300 | PDI CHG / Isopropyl Alcohol Shipper | PDI CHG / Isopropyl Alcohol Shipper | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC w/ shyed flaps (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | 6GB10300 | PDI CHG / Isopropyl Alcohol Shipper | PDI CHG / Isopropyl Alcohol Shipper | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Width | 6.25 | in | FDA,4 | 6.1875 | 6.5625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 6GD35185 | Hygea BZK Shipper | Hygea BZK Shipper | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 50 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 6GD35185 | Hygea BZK Shipper | Hygea BZK Shipper | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 6GD35185 | Hygea BZK Shipper | Hygea BZK Shipper | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 6GD35185 | Hygea BZK Shipper | Hygea BZK Shipper | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 6GD35185 | Hygea BZK Shipper | Hygea BZK Shipper | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | Height | 12.4375 | in | FDA,4 | 12.25 | 12.625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 6GD35185 | Hygea BZK Shipper | Hygea BZK Shipper | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | Length | 13.25 | in | FDA,4 | 13.0625 | 13.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 6GD35185 | Hygea BZK Shipper | Hygea BZK Shipper | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 6GD35185 | Hygea BZK Shipper | Hygea BZK Shipper | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | 6GD35185 | Hygea BZK Shipper | Hygea BZK Shipper | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | Width | 12 | in | FDA,4 | 11.8125 | 12.1875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | 6GS48450 | PDI CHG Swab shipper | PDI CHG Swab Shipper | SHIP_NY_29K_100_17.125-6.6875-11.75 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | 6GS48450 | PDI CHG Swab shipper | PDI CHG Swab Shipper | SHIP_NY_29K_100_17.125-6.6875-11.75 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 29 | Lbs./In. | FDA,0 | 29 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | 6GS48450 | PDI CHG Swab shipper | PDI CHG Swab Shipper | SHIP_NY_29K_100_17.125-6.6875-11.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | 6GS48450 | PDI CHG Swab shipper | PDI CHG Swab Shipper | SHIP_NY_29K_100_17.125-6.6875-11.75 | NY-Analytical | 731 | Shipper Inspection | Height | 11.875 | in | FDA,4 | 11.5625 | 11.9375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | 6GS48450 | PDI CHG Swab shipper | PDI CHG Swab Shipper | SHIP_NY_29K_100_17.125-6.6875-11.75 | NY-Analytical | 731 | Shipper Inspection | Length | 17.125 | in | FDA,4 | 16.9375 | 17.3125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | 6GS48450 | PDI CHG Swab shipper | PDI CHG Swab Shipper | SHIP_NY_29K_100_17.125-6.6875-11.75 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | 6GS48450 | PDI CHG Swab shipper | PDI CHG Swab Shipper | SHIP_NY_29K_100_17.125-6.6875-11.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | 6GS48450 | PDI CHG Swab shipper | PDI CHG Swab Shipper | SHIP_NY_29K_100_17.125-6.6875-11.75 | NY-Analytical | 731 | Shipper Inspection | Width | 6.75 | in | FDA,4 | 6.5 | 6.875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | 6HB60307 | PDI Alcohol Prep Pad Shipper | PDI Alcohol Prep Pad Shipper | SHIP_NY_40K_100_16.5-12.3125-6.1875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | 6HB60307 | PDI Alcohol Prep Pad Shipper | PDI Alcohol Prep Pad Shipper | SHIP_NY_40K_100_16.5-12.3125-6.1875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 40 | Lbs./In. | FDA,0 | 40 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | 6HB60307 | PDI Alcohol Prep Pad Shipper | PDI Alcohol Prep Pad Shipper | SHIP_NY_40K_100_16.5-12.3125-6.1875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | 6HB60307 | PDI Alcohol Prep Pad Shipper | PDI Alcohol Prep Pad Shipper | SHIP_NY_40K_100_16.5-12.3125-6.1875 | NY-Analytical | 731 | Shipper Inspection | Height | 6.312 | in | FDA,4 | 6 | 6.375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | 6HB60307 | PDI Alcohol Prep Pad Shipper | PDI Alcohol Prep Pad Shipper | SHIP_NY_40K_100_16.5-12.3125-6.1875 | NY-Analytical | 731 | Shipper Inspection | Length | 16.5 | in | FDA,4 | 16.3125 | 16.6875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | 6HB60307 | PDI Alcohol Prep Pad Shipper | PDI Alcohol Prep Pad Shipper | SHIP_NY_40K_100_16.5-12.3125-6.1875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | 6HB60307 | PDI Alcohol Prep Pad Shipper | PDI Alcohol Prep Pad Shipper | SHIP_NY_40K_100_16.5-12.3125-6.1875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | 6HB60307 | PDI Alcohol Prep Pad Shipper | PDI Alcohol Prep Pad Shipper | SHIP_NY_40K_100_16.5-12.3125-6.1875 | NY-Analytical | 731 | Shipper Inspection | Width | 12.25 | in | FDA,4 | 12.125 | 12.5 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 6HC69900 | PDI Alcohol Prep Pad Shipper | PDI Alcohol Prep Pad Shipper | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 50 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 6HC69900 | PDI Alcohol Prep Pad Shipper | PDI Alcohol Prep Pad Shipper | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 6HC69900 | PDI Alcohol Prep Pad Shipper | PDI Alcohol Prep Pad Shipper | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 6HC69900 | PDI Alcohol Prep Pad Shipper | PDI Alcohol Prep Pad Shipper | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 6HC69900 | PDI Alcohol Prep Pad Shipper | PDI Alcohol Prep Pad Shipper | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | Height | 5.25 | in | FDA,4 | 5 | 5.375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 6HC69900 | PDI Alcohol Prep Pad Shipper | PDI Alcohol Prep Pad Shipper | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | Length | 15.75 | in | FDA,4 | 15.5625 | 15.9375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 6HC69900 | PDI Alcohol Prep Pad Shipper | PDI Alcohol Prep Pad Shipper | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 6HC69900 | PDI Alcohol Prep Pad Shipper | PDI Alcohol Prep Pad Shipper | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | 6HC69900 | PDI Alcohol Prep Pad Shipper | PDI Alcohol Prep Pad Shipper | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | Width | 8.625 | in | FDA,4 | 8.4375 | 8.8125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6JA580FW/Shipper_NY_Skip lot testing | 6JA580FW | Sani Pro. - Table Turners Cleaning Shipper | Sani Professional - Table Turners Cleaning Shipper | Shipper_NY_Skip lot testing | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Shipper Inspection |  | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | 6JA580FW | Sani Pro. - Table Turners Cleaning Shipper | Sani Professional - Table Turners Cleaning Shipper | SHIP_NY_44K_100_16.750-8.125-10.750 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | 6JA580FW | Sani Pro. - Table Turners Cleaning Shipper | Sani Professional - Table Turners Cleaning Shipper | SHIP_NY_44K_100_16.750-8.125-10.750 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | 6JA580FW | Sani Pro. - Table Turners Cleaning Shipper | Sani Professional - Table Turners Cleaning Shipper | SHIP_NY_44K_100_16.750-8.125-10.750 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | 6JA580FW | Sani Pro. - Table Turners Cleaning Shipper | Sani Professional - Table Turners Cleaning Shipper | SHIP_NY_44K_100_16.750-8.125-10.750 | NY-Analytical | 731 | Shipper Inspection | Height | 10.875 | in | FDA,4 | 10.5625 | 10.9375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | 6JA580FW | Sani Pro. - Table Turners Cleaning Shipper | Sani Professional - Table Turners Cleaning Shipper | SHIP_NY_44K_100_16.750-8.125-10.750 | NY-Analytical | 731 | Shipper Inspection | Length | 16.9375 | in | FDA,4 | 16.5625 | 16.9375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | 6JA580FW | Sani Pro. - Table Turners Cleaning Shipper | Sani Professional - Table Turners Cleaning Shipper | SHIP_NY_44K_100_16.750-8.125-10.750 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | 6JA580FW | Sani Pro. - Table Turners Cleaning Shipper | Sani Professional - Table Turners Cleaning Shipper | SHIP_NY_44K_100_16.750-8.125-10.750 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | 6JA580FW | Sani Pro. - Table Turners Cleaning Shipper | Sani Professional - Table Turners Cleaning Shipper | SHIP_NY_44K_100_16.750-8.125-10.750 | NY-Analytical | 731 | Shipper Inspection | Width | 8.1875 | in | FDA,4 | 7.9375 | 8.3125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | 6LB40600 | Povidone Iodine Prep Pad Shipper | Povidone-Iodine Prep Pad Shipper | SHIP_NY_29K_100_12.375-5.5-4.625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | 6LB40600 | Povidone Iodine Prep Pad Shipper | Povidone-Iodine Prep Pad Shipper | SHIP_NY_29K_100_12.375-5.5-4.625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 29 | Lbs./In. | FDA,0 | 29 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | 6LB40600 | Povidone Iodine Prep Pad Shipper | Povidone-Iodine Prep Pad Shipper | SHIP_NY_29K_100_12.375-5.5-4.625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | 6LB40600 | Povidone Iodine Prep Pad Shipper | Povidone-Iodine Prep Pad Shipper | SHIP_NY_29K_100_12.375-5.5-4.625 | NY-Analytical | 731 | Shipper Inspection | Height | 4.75 | in | FDA,4 | 4.4375 | 4.8125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | 6LB40600 | Povidone Iodine Prep Pad Shipper | Povidone-Iodine Prep Pad Shipper | SHIP_NY_29K_100_12.375-5.5-4.625 | NY-Analytical | 731 | Shipper Inspection | Length | 12.437 | in | FDA,4 | 12.1875 | 12.5625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | 6LB40600 | Povidone Iodine Prep Pad Shipper | Povidone-Iodine Prep Pad Shipper | SHIP_NY_29K_100_12.375-5.5-4.625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | 6LB40600 | Povidone Iodine Prep Pad Shipper | Povidone-Iodine Prep Pad Shipper | SHIP_NY_29K_100_12.375-5.5-4.625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | 6LB40600 | Povidone Iodine Prep Pad Shipper | Povidone-Iodine Prep Pad Shipper | SHIP_NY_29K_100_12.375-5.5-4.625 | NY-Analytical | 731 | Shipper Inspection | Width | 5.5 | in | FDA,4 | 5.3125 | 5.6875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | 6OA404S48 | Prep H Shipper | Prep H Shipper | SHIP_NY_44K_100_15.625-12.5-6.1875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | 6OA404S48 | Prep H Shipper | Prep H Shipper | SHIP_NY_44K_100_15.625-12.5-6.1875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | 6OA404S48 | Prep H Shipper | Prep H Shipper | SHIP_NY_44K_100_15.625-12.5-6.1875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | 6OA404S48 | Prep H Shipper | Prep H Shipper | SHIP_NY_44K_100_15.625-12.5-6.1875 | NY-Analytical | 731 | Shipper Inspection | Height | 6 | in | FDA,4 | 6 | 6.375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | 6OA404S48 | Prep H Shipper | Prep H Shipper | SHIP_NY_44K_100_15.625-12.5-6.1875 | NY-Analytical | 731 | Shipper Inspection | Length | 15.625 | in | FDA,4 | 15.4375 | 15.8125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | 6OA404S48 | Prep H Shipper | Prep H Shipper | SHIP_NY_44K_100_15.625-12.5-6.1875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | 6OA404S48 | Prep H Shipper | Prep H Shipper | SHIP_NY_44K_100_15.625-12.5-6.1875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | 6OA404S48 | Prep H Shipper | Prep H Shipper | SHIP_NY_44K_100_15.625-12.5-6.1875 | NY-Analytical | 731 | Shipper Inspection | Width | 12.5625 | in | FDA,4 | 12.3125 | 12.6875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 6OA540S48 | Preparation H Medicated Wipes Shipper | Prep H Medicated Wipes Shipper | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 6OA540S48 | Preparation H Medicated Wipes Shipper | Prep H Medicated Wipes Shipper | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 6OA540S48 | Preparation H Medicated Wipes Shipper | Prep H Medicated Wipes Shipper | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 6OA540S48 | Preparation H Medicated Wipes Shipper | Prep H Medicated Wipes Shipper | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 6OA540S48 | Preparation H Medicated Wipes Shipper | Prep H Medicated Wipes Shipper | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | Height | 12.625 | in | FDA,4 | 12.4375 | 12.8125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 6OA540S48 | Preparation H Medicated Wipes Shipper | Prep H Medicated Wipes Shipper | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | Length | 11.875 | in | FDA,4 | 11.6875 | 12.0625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 6OA540S48 | Preparation H Medicated Wipes Shipper | Prep H Medicated Wipes Shipper | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 6OA540S48 | Preparation H Medicated Wipes Shipper | Prep H Medicated Wipes Shipper | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | 6OA540S48 | Preparation H Medicated Wipes Shipper | Prep H Medicated Wipes Shipper | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | Width | 7.75 | in | FDA,4 | 7.5625 | 7.9375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 6OA632S60 | Preparation H Shipper | Prep H Shipper | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 6OA632S60 | Preparation H Shipper | Prep H Shipper | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 6OA632S60 | Preparation H Shipper | Prep H Shipper | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 6OA632S60 | Preparation H Shipper | Prep H Shipper | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 6OA632S60 | Preparation H Shipper | Prep H Shipper | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | Height | 14.25 | in | FDA,4 | 13.9375 | 14.3125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 6OA632S60 | Preparation H Shipper | Prep H Shipper | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | Length | 19.5 | in | FDA,4 | 19.1875 | 19.5625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 6OA632S60 | Preparation H Shipper | Prep H Shipper | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 6OA632S60 | Preparation H Shipper | Prep H Shipper | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | HSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | 6OA632S60 | Preparation H Shipper | Prep H Shipper | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | Width | 8.25 | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 6OB123ST | Prevantics Device Swab Strip Case Shipper | Prevantics Device Swab Strip Case Shipper | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 6OB123ST | Prevantics Device Swab Strip Case Shipper | Prevantics Device Swab Strip Case Shipper | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 6OB123ST | Prevantics Device Swab Strip Case Shipper | Prevantics Device Swab Strip Case Shipper | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 6OB123ST | Prevantics Device Swab Strip Case Shipper | Prevantics Device Swab Strip Case Shipper | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 6OB123ST | Prevantics Device Swab Strip Case Shipper | Prevantics Device Swab Strip Case Shipper | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Height | 5.875 | in | FDA,4 | 5.5625 | 5.9375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 6OB123ST | Prevantics Device Swab Strip Case Shipper | Prevantics Device Swab Strip Case Shipper | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Length | 12.5 | in | FDA,4 | 12.1875 | 12.5625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 6OB123ST | Prevantics Device Swab Strip Case Shipper | Prevantics Device Swab Strip Case Shipper | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 6OB123ST | Prevantics Device Swab Strip Case Shipper | Prevantics Device Swab Strip Case Shipper | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | 6OB123ST | Prevantics Device Swab Strip Case Shipper | Prevantics Device Swab Strip Case Shipper | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Width | 10.75 | in | FDA,4 | 10.5625 | 10.9375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OB31400/SHIP_32_1AL | 6OB31400 | CRU 14.625X6.375X5 200# | CRU 14.625X6.375X5 200# | SHIP_32_1AL | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OB31400/SHIP_32_1AL | 6OB31400 | CRU 14.625X6.375X5 200# | CRU 14.625X6.375X5 200# | SHIP_32_1AL | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OB31400/SHIP_32_1AL | 6OB31400 | CRU 14.625X6.375X5 200# | CRU 14.625X6.375X5 200# | SHIP_32_1AL | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OB31400/SHIP_32_1AL | 6OB31400 | CRU 14.625X6.375X5 200# | CRU 14.625X6.375X5 200# | SHIP_32_1AL | NY-Analytical | 731 | Shipper Inspection | Height | 5 | in | FDA,4 | 4.8125 | 5.1875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OB31400/SHIP_32_1AL | 6OB31400 | CRU 14.625X6.375X5 200# | CRU 14.625X6.375X5 200# | SHIP_32_1AL | NY-Analytical | 731 | Shipper Inspection | Length | 14.625 | in | FDA,4 | 14.4375 | 14.8125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OB31400/SHIP_32_1AL | 6OB31400 | CRU 14.625X6.375X5 200# | CRU 14.625X6.375X5 200# | SHIP_32_1AL | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OB31400/SHIP_32_1AL | 6OB31400 | CRU 14.625X6.375X5 200# | CRU 14.625X6.375X5 200# | SHIP_32_1AL | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OB31400/SHIP_32_1AL | 6OB31400 | CRU 14.625X6.375X5 200# | CRU 14.625X6.375X5 200# | SHIP_32_1AL | NY-Analytical | 731 | Shipper Inspection | Width | 6.375 | in | FDA,4 | 6.1875 | 6.5625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OB389PD/Ship_NA_3A | 6OB389PD | PAD 14.5X6.25 ANY CHIP BRD | Pad 14.5X6.25 Any Chip Brd | Ship_NA_3A | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OB389PD/Ship_NA_3A | 6OB389PD | PAD 14.5X6.25 ANY CHIP BRD | Pad 14.5X6.25 Any Chip Brd | Ship_NA_3A | NY-Analytical | 731 | Shipper Inspection | Length | 14.5 | in | FDA,4 | 14.3125 | 14.6875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OB389PD/Ship_NA_3A | 6OB389PD | PAD 14.5X6.25 ANY CHIP BRD | Pad 14.5X6.25 Any Chip Brd | Ship_NA_3A | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Layer Pad (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OB389PD/Ship_NA_3A | 6OB389PD | PAD 14.5X6.25 ANY CHIP BRD | Pad 14.5X6.25 Any Chip Brd | Ship_NA_3A | NY-Analytical | 731 | Shipper Inspection | Width | 6.437 | in | FDA,4 | 6.0625 | 6.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OC22324/SHIP_32_1N | 6OC22324 | CRU 11.125X6.375X10.375 32ECT | CRU 11.125X6.375X10.375 32ECt. | SHIP_32_1N | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OC22324/SHIP_32_1N | 6OC22324 | CRU 11.125X6.375X10.375 32ECT | CRU 11.125X6.375X10.375 32ECt. | SHIP_32_1N | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OC22324/SHIP_32_1N | 6OC22324 | CRU 11.125X6.375X10.375 32ECT | CRU 11.125X6.375X10.375 32ECt. | SHIP_32_1N | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OC22324/SHIP_32_1N | 6OC22324 | CRU 11.125X6.375X10.375 32ECT | CRU 11.125X6.375X10.375 32ECt. | SHIP_32_1N | NY-Analytical | 731 | Shipper Inspection | Height | 10.5 | in | FDA,4 | 10.1875 | 10.5625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OC22324/SHIP_32_1N | 6OC22324 | CRU 11.125X6.375X10.375 32ECT | CRU 11.125X6.375X10.375 32ECt. | SHIP_32_1N | NY-Analytical | 731 | Shipper Inspection | Length | 11.125 | in | FDA,4 | 10.9375 | 11.3125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OC22324/SHIP_32_1N | 6OC22324 | CRU 11.125X6.375X10.375 32ECT | CRU 11.125X6.375X10.375 32ECt. | SHIP_32_1N | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OC22324/SHIP_32_1N | 6OC22324 | CRU 11.125X6.375X10.375 32ECT | CRU 11.125X6.375X10.375 32ECt. | SHIP_32_1N | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OC22324/SHIP_32_1N | 6OC22324 | CRU 11.125X6.375X10.375 32ECT | CRU 11.125X6.375X10.375 32ECt. | SHIP_32_1N | NY-Analytical | 731 | Shipper Inspection | Width | 6.375 | in | FDA,4 | 6.1875 | 6.5625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | 6OD01524 | Walmart Wet Nap Unprinted Shipper | Walmart Wet Nap Unprinted Shipper | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | 6OD01524 | Walmart Wet Nap Unprinted Shipper | Walmart Wet Nap Unprinted Shipper | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | 6OD01524 | Walmart Wet Nap Unprinted Shipper | Walmart Wet Nap Unprinted Shipper | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | 6OD01524 | Walmart Wet Nap Unprinted Shipper | Walmart Wet Nap Unprinted Shipper | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | 6OD01524 | Walmart Wet Nap Unprinted Shipper | Walmart Wet Nap Unprinted Shipper | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | Height | 4.75 | in | FDA,4 | 4.5 | 4.875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | 6OD01524 | Walmart Wet Nap Unprinted Shipper | Walmart Wet Nap Unprinted Shipper | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | Length | 8.25 | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | 6OD01524 | Walmart Wet Nap Unprinted Shipper | Walmart Wet Nap Unprinted Shipper | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | 6OD01524 | Walmart Wet Nap Unprinted Shipper | Walmart Wet Nap Unprinted Shipper | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | 6OD01524 | Walmart Wet Nap Unprinted Shipper | Walmart Wet Nap Unprinted Shipper | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | Width | 6.25 | in | FDA,4 | 6.0625 | 6.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OD13500A/SHIP_32_1AD | 6OD13500A | CRU 13.25X12.0X6.125 32 ECT/26 | CRU 13.25X12.0X6.125 32 ECt. /26 | SHIP_32_1AD | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OD13500A/SHIP_32_1AD | 6OD13500A | CRU 13.25X12.0X6.125 32 ECT/26 | CRU 13.25X12.0X6.125 32 ECt. /26 | SHIP_32_1AD | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OD13500A/SHIP_32_1AD | 6OD13500A | CRU 13.25X12.0X6.125 32 ECT/26 | CRU 13.25X12.0X6.125 32 ECt. /26 | SHIP_32_1AD | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OD13500A/SHIP_32_1AD | 6OD13500A | CRU 13.25X12.0X6.125 32 ECT/26 | CRU 13.25X12.0X6.125 32 ECt. /26 | SHIP_32_1AD | NY-Analytical | 731 | Shipper Inspection | Height | 6.25 | in | FDA,4 | 5.9375 | 6.3125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OD13500A/SHIP_32_1AD | 6OD13500A | CRU 13.25X12.0X6.125 32 ECT/26 | CRU 13.25X12.0X6.125 32 ECt. /26 | SHIP_32_1AD | NY-Analytical | 731 | Shipper Inspection | Length | 13.25 | in | FDA,4 | 13.0625 | 13.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OD13500A/SHIP_32_1AD | 6OD13500A | CRU 13.25X12.0X6.125 32 ECT/26 | CRU 13.25X12.0X6.125 32 ECt. /26 | SHIP_32_1AD | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OD13500A/SHIP_32_1AD | 6OD13500A | CRU 13.25X12.0X6.125 32 ECT/26 | CRU 13.25X12.0X6.125 32 ECt. /26 | SHIP_32_1AD | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OD13500A/SHIP_32_1AD | 6OD13500A | CRU 13.25X12.0X6.125 32 ECT/26 | CRU 13.25X12.0X6.125 32 ECt. /26 | SHIP_32_1AD | NY-Analytical | 731 | Shipper Inspection | Width | 12 | in | FDA,4 | 11.8125 | 12.1875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
| 6OD239DR | DIV COR 7.625 X 15.375 |  | Analytical - SHIP_44_1A | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Divider |  |  |  |  | Shipper Style | TM141 |  |  |  | left_only
| 6OD239DR | DIV COR 7.625 X 15.375 |  | Analytical - SHIP_44_1A | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 |  |  |  | left_only
| 6OD239DR | DIV COR 7.625 X 15.375 |  | Analytical - SHIP_44_1A | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | TM141 |  |  |  | left_only
| 6OD239DR | DIV COR 7.625 X 15.375 |  | Analytical - SHIP_44_1A | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 |  |  |  | left_only
| 6OD239DR | DIV COR 7.625 X 15.375 |  | Analytical - SHIP_44_1A | NY-Analytical | 731 | Shipper Inspection | Length | 7.75 | in | FDA,4 | 7.4375 | 7.8125 | TM 141 / Visual | TM141 |  |  |  | left_only
| 6OD239DR | DIV COR 7.625 X 15.375 |  | Analytical - SHIP_44_1A | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 |  |  |  | left_only
| 6OD239DR | DIV COR 7.625 X 15.375 |  | Analytical - SHIP_44_1A | NY-Analytical | 731 | Shipper Inspection | Width | 15.375 | in | FDA,4 | 15.1875 | 15.5625 | TM 141 / Visual | TM141 |  |  |  | left_only
6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | 6OD36055 | Unprinted Shipper | Unprinted Shipper | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | 6OD36055 | Unprinted Shipper | Unprinted Shipper | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | 6OD36055 | Unprinted Shipper | Unprinted Shipper | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | 6OD36055 | Unprinted Shipper | Unprinted Shipper | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | 6OD36055 | Unprinted Shipper | Unprinted Shipper | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Height | 11.375 | in | FDA,4 | 11.0625 | 11.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | 6OD36055 | Unprinted Shipper | Unprinted Shipper | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Length | 10.5 | in | FDA,4 | 10.4375 | 10.8125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | 6OD36055 | Unprinted Shipper | Unprinted Shipper | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Does not meet specifications (OOS) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | 6OD36055 | Unprinted Shipper | Unprinted Shipper | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC w/ shyed flaps (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | 6OD36055 | Unprinted Shipper | Unprinted Shipper | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Width | 6.375 | in | FDA,4 | 6.1875 | 6.5625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OH14324/SHIP_32_1Y | 6OH14324 | CRU SANI HANDS 24CT | CRU Sani-Hands 24Ct. | SHIP_32_1Y | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OH14324/SHIP_32_1Y | 6OH14324 | CRU SANI HANDS 24CT | CRU Sani-Hands 24Ct. | SHIP_32_1Y | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OH14324/SHIP_32_1Y | 6OH14324 | CRU SANI HANDS 24CT | CRU Sani-Hands 24Ct. | SHIP_32_1Y | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OH14324/SHIP_32_1Y | 6OH14324 | CRU SANI HANDS 24CT | CRU Sani-Hands 24Ct. | SHIP_32_1Y | NY-Analytical | 731 | Shipper Inspection | Height | 6.25 | in | FDA,4 | 6.0625 | 6.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OH14324/SHIP_32_1Y | 6OH14324 | CRU SANI HANDS 24CT | CRU Sani-Hands 24Ct. | SHIP_32_1Y | NY-Analytical | 731 | Shipper Inspection | Length | 12.4375 | in | FDA,4 | 12.25 | 12.625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OH14324/SHIP_32_1Y | 6OH14324 | CRU SANI HANDS 24CT | CRU Sani-Hands 24Ct. | SHIP_32_1Y | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OH14324/SHIP_32_1Y | 6OH14324 | CRU SANI HANDS 24CT | CRU Sani-Hands 24Ct. | SHIP_32_1Y | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OH14324/SHIP_32_1Y | 6OH14324 | CRU SANI HANDS 24CT | CRU Sani-Hands 24Ct. | SHIP_32_1Y | NY-Analytical | 731 | Shipper Inspection | Width | 9 | in | FDA,4 | 8.8125 | 9.1875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OH21112/SHIP_32_1F | 6OH21112 | CRU CVS ANTIBAC 9.75 x 7.625 x 6.75=32 ECT/26MED | CRU CVS Antibac 9.75 X 7.625 X 6.75=32 ECt. /26Med | SHIP_32_1F | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OH21112/SHIP_32_1F | 6OH21112 | CRU CVS ANTIBAC 9.75 x 7.625 x 6.75=32 ECT/26MED | CRU CVS Antibac 9.75 X 7.625 X 6.75=32 ECt. /26Med | SHIP_32_1F | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OH21112/SHIP_32_1F | 6OH21112 | CRU CVS ANTIBAC 9.75 x 7.625 x 6.75=32 ECT/26MED | CRU CVS Antibac 9.75 X 7.625 X 6.75=32 ECt. /26Med | SHIP_32_1F | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OH21112/SHIP_32_1F | 6OH21112 | CRU CVS ANTIBAC 9.75 x 7.625 x 6.75=32 ECT/26MED | CRU CVS Antibac 9.75 X 7.625 X 6.75=32 ECt. /26Med | SHIP_32_1F | NY-Analytical | 731 | Shipper Inspection | Height | 6.75 | in | FDA,4 | 6.5625 | 6.9375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OH21112/SHIP_32_1F | 6OH21112 | CRU CVS ANTIBAC 9.75 x 7.625 x 6.75=32 ECT/26MED | CRU CVS Antibac 9.75 X 7.625 X 6.75=32 ECt. /26Med | SHIP_32_1F | NY-Analytical | 731 | Shipper Inspection | Length | 9.75 | in | FDA,4 | 9.5625 | 9.9375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OH21112/SHIP_32_1F | 6OH21112 | CRU CVS ANTIBAC 9.75 x 7.625 x 6.75=32 ECT/26MED | CRU CVS Antibac 9.75 X 7.625 X 6.75=32 ECt. /26Med | SHIP_32_1F | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OH21112/SHIP_32_1F | 6OH21112 | CRU CVS ANTIBAC 9.75 x 7.625 x 6.75=32 ECT/26MED | CRU CVS Antibac 9.75 X 7.625 X 6.75=32 ECt. /26Med | SHIP_32_1F | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OH21112/SHIP_32_1F | 6OH21112 | CRU CVS ANTIBAC 9.75 x 7.625 x 6.75=32 ECT/26MED | CRU CVS Antibac 9.75 X 7.625 X 6.75=32 ECt. /26Med | SHIP_32_1F | NY-Analytical | 731 | Shipper Inspection | Width | 7.5 | in | FDA,4 | 7.4375 | 7.8125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | 6OH45016 | Individual Pack 16X Box 12x Case | Individual Pack 16X Box 12X Case | SHIP_NY_32K_100_8.25-6.875-7.375 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | 6OH45016 | Individual Pack 16X Box 12x Case | Individual Pack 16X Box 12X Case | SHIP_NY_32K_100_8.25-6.875-7.375 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | 6OH45016 | Individual Pack 16X Box 12x Case | Individual Pack 16X Box 12X Case | SHIP_NY_32K_100_8.25-6.875-7.375 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | 6OH45016 | Individual Pack 16X Box 12x Case | Individual Pack 16X Box 12X Case | SHIP_NY_32K_100_8.25-6.875-7.375 | NY-Analytical | 731 | Shipper Inspection | Height | 7.25 | in | FDA,4 | 7.1875 | 7.5625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | 6OH45016 | Individual Pack 16X Box 12x Case | Individual Pack 16X Box 12X Case | SHIP_NY_32K_100_8.25-6.875-7.375 | NY-Analytical | 731 | Shipper Inspection | Length | 8.1562 | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | 6OH45016 | Individual Pack 16X Box 12x Case | Individual Pack 16X Box 12X Case | SHIP_NY_32K_100_8.25-6.875-7.375 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | 6OH45016 | Individual Pack 16X Box 12x Case | Individual Pack 16X Box 12X Case | SHIP_NY_32K_100_8.25-6.875-7.375 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | 6OH45016 | Individual Pack 16X Box 12x Case | Individual Pack 16X Box 12X Case | SHIP_NY_32K_100_8.25-6.875-7.375 | NY-Analytical | 731 | Shipper Inspection | Width | 7 | in | FDA,4 | 6.6875 | 7.0625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OK75595/SHIP_32_1R | 6OK75595 | CRU 11.75X9.625X7.6875 32ECT | CRU 11.75X9.625X7.6875 32ECt. | SHIP_32_1R | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OK75595/SHIP_32_1R | 6OK75595 | CRU 11.75X9.625X7.6875 32ECT | CRU 11.75X9.625X7.6875 32ECt. | SHIP_32_1R | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OK75595/SHIP_32_1R | 6OK75595 | CRU 11.75X9.625X7.6875 32ECT | CRU 11.75X9.625X7.6875 32ECt. | SHIP_32_1R | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OK75595/SHIP_32_1R | 6OK75595 | CRU 11.75X9.625X7.6875 32ECT | CRU 11.75X9.625X7.6875 32ECt. | SHIP_32_1R | NY-Analytical | 731 | Shipper Inspection | Height | 7.75 | in | FDA,4 | 7.5 | 7.875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OK75595/SHIP_32_1R | 6OK75595 | CRU 11.75X9.625X7.6875 32ECT | CRU 11.75X9.625X7.6875 32ECt. | SHIP_32_1R | NY-Analytical | 731 | Shipper Inspection | Length | 11.75 | in | FDA,4 | 11.5625 | 11.9375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OK75595/SHIP_32_1R | 6OK75595 | CRU 11.75X9.625X7.6875 32ECT | CRU 11.75X9.625X7.6875 32ECt. | SHIP_32_1R | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OK75595/SHIP_32_1R | 6OK75595 | CRU 11.75X9.625X7.6875 32ECT | CRU 11.75X9.625X7.6875 32ECt. | SHIP_32_1R | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OK75595/SHIP_32_1R | 6OK75595 | CRU 11.75X9.625X7.6875 32ECT | CRU 11.75X9.625X7.6875 32ECt. | SHIP_32_1R | NY-Analytical | 731 | Shipper Inspection | Width | 9.5 | in | FDA,4 | 9.4375 | 9.8125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OM960CB/SHIP_NY_NA_200_2.5-2.5-33.75 | 6OM960CB | 2x60ct Grime Boss Costco Corner Board | Grime Boss Costco Corner Board 2X60Ct. | SHIP_NY_NA_200_2.5-2.5-33.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OM960CB/SHIP_NY_NA_200_2.5-2.5-33.75 | 6OM960CB | 2x60ct Grime Boss Costco Corner Board | Grime Boss Costco Corner Board 2X60Ct. | SHIP_NY_NA_200_2.5-2.5-33.75 | NY-Analytical | 731 | Shipper Inspection | Height | 33.625 | in | FDA,4 | 33.5625 | 33.9375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OM960CB/SHIP_NY_NA_200_2.5-2.5-33.75 | 6OM960CB | 2x60ct Grime Boss Costco Corner Board | Grime Boss Costco Corner Board 2X60Ct. | SHIP_NY_NA_200_2.5-2.5-33.75 | NY-Analytical | 731 | Shipper Inspection | Length | 2.5 | in | FDA,4 | 2.3125 | 2.6875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OM960CB/SHIP_NY_NA_200_2.5-2.5-33.75 | 6OM960CB | 2x60ct Grime Boss Costco Corner Board | Grime Boss Costco Corner Board 2X60Ct. | SHIP_NY_NA_200_2.5-2.5-33.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | ITW STACK MASTER 2600 (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OM960CB/SHIP_NY_NA_200_2.5-2.5-33.75 | 6OM960CB | 2x60ct Grime Boss Costco Corner Board | Grime Boss Costco Corner Board 2X60Ct. | SHIP_NY_NA_200_2.5-2.5-33.75 | NY-Analytical | 731 | Shipper Inspection | Width | 2.5 | in | FDA,4 | 2.3125 | 2.6875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OP081LP/SHIP_NY_32NA_250_17.250-11.938 | 6OP081LP | Layer Pad - Pond's 50ct WIP | Layer Pad - Pond'S 50Ct. Wip | SHIP_NY_32NA_250_17.250-11.938 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | SUB-ASSEMBLY - NY | RM_SHIPPER | both
6OP081LP/SHIP_NY_32NA_250_17.250-11.938 | 6OP081LP | Layer Pad - Pond's 50ct WIP | Layer Pad - Pond'S 50Ct. Wip | SHIP_NY_32NA_250_17.250-11.938 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | SUB-ASSEMBLY - NY | RM_SHIPPER | both
6OP081LP/SHIP_NY_32NA_250_17.250-11.938 | 6OP081LP | Layer Pad - Pond's 50ct WIP | Layer Pad - Pond'S 50Ct. Wip | SHIP_NY_32NA_250_17.250-11.938 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | SUB-ASSEMBLY - NY | RM_SHIPPER | both
6OP081LP/SHIP_NY_32NA_250_17.250-11.938 | 6OP081LP | Layer Pad - Pond's 50ct WIP | Layer Pad - Pond'S 50Ct. Wip | SHIP_NY_32NA_250_17.250-11.938 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | SUB-ASSEMBLY - NY | RM_SHIPPER | both
6OP081LP/SHIP_NY_32NA_250_17.250-11.938 | 6OP081LP | Layer Pad - Pond's 50ct WIP | Layer Pad - Pond'S 50Ct. Wip | SHIP_NY_32NA_250_17.250-11.938 | NY-Analytical | 731 | Shipper Inspection | Length | 17.125 | in | FDA,3 | 17.125 | 17.375 | TM 141 / Visual | TM141 | NY | SUB-ASSEMBLY - NY | RM_SHIPPER | both
6OP081LP/SHIP_NY_32NA_250_17.250-11.938 | 6OP081LP | Layer Pad - Pond's 50ct WIP | Layer Pad - Pond'S 50Ct. Wip | SHIP_NY_32NA_250_17.250-11.938 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Inserts (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | SUB-ASSEMBLY - NY | RM_SHIPPER | both
6OP081LP/SHIP_NY_32NA_250_17.250-11.938 | 6OP081LP | Layer Pad - Pond's 50ct WIP | Layer Pad - Pond'S 50Ct. Wip | SHIP_NY_32NA_250_17.250-11.938 | NY-Analytical | 731 | Shipper Inspection | Width | 11.875 | in | FDA,3 | 11.813 | 12.063 | TM 141 / Visual | TM141 | NY | SUB-ASSEMBLY - NY | RM_SHIPPER | both
6OP34524/SHIP_44_2G | 6OP34524 | CRU AVON UNPRINTED=44 ECT | CRU Avon Unprinted=44 ECt. | SHIP_44_2G | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OP34524/SHIP_44_2G | 6OP34524 | CRU AVON UNPRINTED=44 ECT | CRU Avon Unprinted=44 ECt. | SHIP_44_2G | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OP34524/SHIP_44_2G | 6OP34524 | CRU AVON UNPRINTED=44 ECT | CRU Avon Unprinted=44 ECt. | SHIP_44_2G | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OP34524/SHIP_44_2G | 6OP34524 | CRU AVON UNPRINTED=44 ECT | CRU Avon Unprinted=44 ECt. | SHIP_44_2G | NY-Analytical | 731 | Shipper Inspection | Height | 8.25 | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OP34524/SHIP_44_2G | 6OP34524 | CRU AVON UNPRINTED=44 ECT | CRU Avon Unprinted=44 ECt. | SHIP_44_2G | NY-Analytical | 731 | Shipper Inspection | Length | 13.5 | in | FDA,4 | 13.3125 | 13.6875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OP34524/SHIP_44_2G | 6OP34524 | CRU AVON UNPRINTED=44 ECT | CRU Avon Unprinted=44 ECt. | SHIP_44_2G | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OP34524/SHIP_44_2G | 6OP34524 | CRU AVON UNPRINTED=44 ECT | CRU Avon Unprinted=44 ECt. | SHIP_44_2G | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OP34524/SHIP_44_2G | 6OP34524 | CRU AVON UNPRINTED=44 ECT | CRU Avon Unprinted=44 ECt. | SHIP_44_2G | NY-Analytical | 731 | Shipper Inspection | Width | 11.75 | in | FDA,4 | 11.5625 | 11.9375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 6OP75510 | Cetaphil Unprinted Shipper | Cetaphil Unprinted Shipper | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 32 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 6OP75510 | Cetaphil Unprinted Shipper | Cetaphil Unprinted Shipper | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 6OP75510 | Cetaphil Unprinted Shipper | Cetaphil Unprinted Shipper | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 6OP75510 | Cetaphil Unprinted Shipper | Cetaphil Unprinted Shipper | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 6OP75510 | Cetaphil Unprinted Shipper | Cetaphil Unprinted Shipper | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | Height | 10.5625 | in | FDA,4 | 10.375 | 10.75 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 6OP75510 | Cetaphil Unprinted Shipper | Cetaphil Unprinted Shipper | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | Length | 15.375 | in | FDA,4 | 15.1875 | 15.5625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 6OP75510 | Cetaphil Unprinted Shipper | Cetaphil Unprinted Shipper | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 6OP75510 | Cetaphil Unprinted Shipper | Cetaphil Unprinted Shipper | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | 6OP75510 | Cetaphil Unprinted Shipper | Cetaphil Unprinted Shipper | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | Width | 11.5625 | in | FDA,4 | 11.375 | 11.75 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | 6OQ22272 | 12ct Large Canister shipper | Large Canister Shipper 12Ct. | SHIP_NY_44K_100_15.188-10.188-15.188 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | 6OQ22272 | 12ct Large Canister shipper | Large Canister Shipper 12Ct. | SHIP_NY_44K_100_15.188-10.188-15.188 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | 6OQ22272 | 12ct Large Canister shipper | Large Canister Shipper 12Ct. | SHIP_NY_44K_100_15.188-10.188-15.188 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | 6OQ22272 | 12ct Large Canister shipper | Large Canister Shipper 12Ct. | SHIP_NY_44K_100_15.188-10.188-15.188 | NY-Analytical | 731 | Shipper Inspection | Height | 15.188 | in | FDA,4 | 15.0005 | 15.3755 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | 6OQ22272 | 12ct Large Canister shipper | Large Canister Shipper 12Ct. | SHIP_NY_44K_100_15.188-10.188-15.188 | NY-Analytical | 731 | Shipper Inspection | Length | 15.188 | in | FDA,4 | 15.0005 | 15.3755 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | 6OQ22272 | 12ct Large Canister shipper | Large Canister Shipper 12Ct. | SHIP_NY_44K_100_15.188-10.188-15.188 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | 6OQ22272 | 12ct Large Canister shipper | Large Canister Shipper 12Ct. | SHIP_NY_44K_100_15.188-10.188-15.188 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | 6OQ22272 | 12ct Large Canister shipper | Large Canister Shipper 12Ct. | SHIP_NY_44K_100_15.188-10.188-15.188 | NY-Analytical | 731 | Shipper Inspection | Width | 10.188 | in | FDA,4 | 10.0005 | 10.3755 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 6OS12350 | Betadine Solution Swabstick Shipper | Betadine Solution Swabstick Shipper | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 6OS12350 | Betadine Solution Swabstick Shipper | Betadine Solution Swabstick Shipper | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 6OS12350 | Betadine Solution Swabstick Shipper | Betadine Solution Swabstick Shipper | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 6OS12350 | Betadine Solution Swabstick Shipper | Betadine Solution Swabstick Shipper | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 6OS12350 | Betadine Solution Swabstick Shipper | Betadine Solution Swabstick Shipper | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 6OS12350 | Betadine Solution Swabstick Shipper | Betadine Solution Swabstick Shipper | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Height | 5.938 | in | FDA,4 | 5.7505 | 6.1255 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 6OS12350 | Betadine Solution Swabstick Shipper | Betadine Solution Swabstick Shipper | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Length | 11.75 | in | FDA,4 | 11.4375 | 11.8125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 6OS12350 | Betadine Solution Swabstick Shipper | Betadine Solution Swabstick Shipper | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 6OS12350 | Betadine Solution Swabstick Shipper | Betadine Solution Swabstick Shipper | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | 6OS12350 | Betadine Solution Swabstick Shipper | Betadine Solution Swabstick Shipper | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Width | 6.375 | in | FDA,4 | 6.3125 | 6.6875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS42850/SHIP_44_2N | 6OS42850 | COR UN-PRINTED=44 ECT | Un-Printed=44 ECt. | SHIP_44_2N | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS42850/SHIP_44_2N | 6OS42850 | COR UN-PRINTED=44 ECT | Un-Printed=44 ECt. | SHIP_44_2N | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS42850/SHIP_44_2N | 6OS42850 | COR UN-PRINTED=44 ECT | Un-Printed=44 ECt. | SHIP_44_2N | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS42850/SHIP_44_2N | 6OS42850 | COR UN-PRINTED=44 ECT | Un-Printed=44 ECt. | SHIP_44_2N | NY-Analytical | 731 | Shipper Inspection | Height | 5.625 | in | FDA,4 | 5.3125 | 5.6875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS42850/SHIP_44_2N | 6OS42850 | COR UN-PRINTED=44 ECT | Un-Printed=44 ECt. | SHIP_44_2N | NY-Analytical | 731 | Shipper Inspection | Length | 15.375 | in | FDA,4 | 15.1875 | 15.5625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS42850/SHIP_44_2N | 6OS42850 | COR UN-PRINTED=44 ECT | Un-Printed=44 ECt. | SHIP_44_2N | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS42850/SHIP_44_2N | 6OS42850 | COR UN-PRINTED=44 ECT | Un-Printed=44 ECt. | SHIP_44_2N | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS42850/SHIP_44_2N | 6OS42850 | COR UN-PRINTED=44 ECT | Un-Printed=44 ECt. | SHIP_44_2N | NY-Analytical | 731 | Shipper Inspection | Width | 11.375 | in | FDA,4 | 11.1875 | 11.5625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 6OS45625 | Betadine Solution Swabstick Shipper | Betadine Solution Swabstick Shipper | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 6OS45625 | Betadine Solution Swabstick Shipper | Betadine Solution Swabstick Shipper | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 6OS45625 | Betadine Solution Swabstick Shipper | Betadine Solution Swabstick Shipper | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 6OS45625 | Betadine Solution Swabstick Shipper | Betadine Solution Swabstick Shipper | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 6OS45625 | Betadine Solution Swabstick Shipper | Betadine Solution Swabstick Shipper | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 6OS45625 | Betadine Solution Swabstick Shipper | Betadine Solution Swabstick Shipper | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Height | 6 | in | FDA,4 | 5.7505 | 6.1255 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 6OS45625 | Betadine Solution Swabstick Shipper | Betadine Solution Swabstick Shipper | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Length | 6.75 | in | FDA,4 | 6.6875 | 7.0625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 6OS45625 | Betadine Solution Swabstick Shipper | Betadine Solution Swabstick Shipper | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 6OS45625 | Betadine Solution Swabstick Shipper | Betadine Solution Swabstick Shipper | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | 6OS45625 | Betadine Solution Swabstick Shipper | Betadine Solution Swabstick Shipper | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Width | 6.5 | in | FDA,4 | 6.3125 | 6.6875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS46250/SHIP_29_1J | 6OS46250 | CRU 17.125X6.6875X11.75 | CRU 17.125X6.6875X11.75 | SHIP_29_1J | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS46250/SHIP_29_1J | 6OS46250 | CRU 17.125X6.6875X11.75 | CRU 17.125X6.6875X11.75 | SHIP_29_1J | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 29 | Lbs./In. | FDA,0 | 29 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS46250/SHIP_29_1J | 6OS46250 | CRU 17.125X6.6875X11.75 | CRU 17.125X6.6875X11.75 | SHIP_29_1J | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS46250/SHIP_29_1J | 6OS46250 | CRU 17.125X6.6875X11.75 | CRU 17.125X6.6875X11.75 | SHIP_29_1J | NY-Analytical | 731 | Shipper Inspection | Height | 11.75 | in | FDA,4 | 11.5625 | 11.9375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS46250/SHIP_29_1J | 6OS46250 | CRU 17.125X6.6875X11.75 | CRU 17.125X6.6875X11.75 | SHIP_29_1J | NY-Analytical | 731 | Shipper Inspection | Length | 17.125 | in | FDA,4 | 16.9375 | 17.3125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS46250/SHIP_29_1J | 6OS46250 | CRU 17.125X6.6875X11.75 | CRU 17.125X6.6875X11.75 | SHIP_29_1J | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS46250/SHIP_29_1J | 6OS46250 | CRU 17.125X6.6875X11.75 | CRU 17.125X6.6875X11.75 | SHIP_29_1J | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OS46250/SHIP_29_1J | 6OS46250 | CRU 17.125X6.6875X11.75 | CRU 17.125X6.6875X11.75 | SHIP_29_1J | NY-Analytical | 731 | Shipper Inspection | Width | 6.6875 | in | FDA,4 | 6.5 | 6.875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU12066/SHIP_40_1M | 6OU12066 | CRU 17-13/16X15-1/8X9 40ECT | CRU 17-13/16X15-1/8X9 40ECt. | SHIP_40_1M | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU12066/SHIP_40_1M | 6OU12066 | CRU 17-13/16X15-1/8X9 40ECT | CRU 17-13/16X15-1/8X9 40ECt. | SHIP_40_1M | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 40 | Lbs./In. | FDA,0 | 40 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU12066/SHIP_40_1M | 6OU12066 | CRU 17-13/16X15-1/8X9 40ECT | CRU 17-13/16X15-1/8X9 40ECt. | SHIP_40_1M | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU12066/SHIP_40_1M | 6OU12066 | CRU 17-13/16X15-1/8X9 40ECT | CRU 17-13/16X15-1/8X9 40ECt. | SHIP_40_1M | NY-Analytical | 731 | Shipper Inspection | Height | 9 | in | FDA,4 | 8.8125 | 9.1875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU12066/SHIP_40_1M | 6OU12066 | CRU 17-13/16X15-1/8X9 40ECT | CRU 17-13/16X15-1/8X9 40ECt. | SHIP_40_1M | NY-Analytical | 731 | Shipper Inspection | Length | 17.75 | in | FDA,4 | 17.625 | 18 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU12066/SHIP_40_1M | 6OU12066 | CRU 17-13/16X15-1/8X9 40ECT | CRU 17-13/16X15-1/8X9 40ECt. | SHIP_40_1M | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU12066/SHIP_40_1M | 6OU12066 | CRU 17-13/16X15-1/8X9 40ECT | CRU 17-13/16X15-1/8X9 40ECt. | SHIP_40_1M | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU12066/SHIP_40_1M | 6OU12066 | CRU 17-13/16X15-1/8X9 40ECT | CRU 17-13/16X15-1/8X9 40ECt. | SHIP_40_1M | NY-Analytical | 731 | Shipper Inspection | Width | 15.125 | in | FDA,4 | 14.9375 | 15.3125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU21470/SHIP_40_1J | 6OU21470 | COR KIRKLAND HH GENERIC SHIPPR | Kirkland Hh Generic Shippr | SHIP_40_1J | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU21470/SHIP_40_1J | 6OU21470 | COR KIRKLAND HH GENERIC SHIPPR | Kirkland Hh Generic Shippr | SHIP_40_1J | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 40 | Lbs./In. | FDA,0 | 40 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU21470/SHIP_40_1J | 6OU21470 | COR KIRKLAND HH GENERIC SHIPPR | Kirkland Hh Generic Shippr | SHIP_40_1J | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU21470/SHIP_40_1J | 6OU21470 | COR KIRKLAND HH GENERIC SHIPPR | Kirkland Hh Generic Shippr | SHIP_40_1J | NY-Analytical | 731 | Shipper Inspection | Height | 9.4375 | in | FDA,4 | 9.0625 | 9.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU21470/SHIP_40_1J | 6OU21470 | COR KIRKLAND HH GENERIC SHIPPR | Kirkland Hh Generic Shippr | SHIP_40_1J | NY-Analytical | 731 | Shipper Inspection | Length | 15.5 | in | FDA,4 | 15.1875 | 15.5625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU21470/SHIP_40_1J | 6OU21470 | COR KIRKLAND HH GENERIC SHIPPR | Kirkland Hh Generic Shippr | SHIP_40_1J | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU21470/SHIP_40_1J | 6OU21470 | COR KIRKLAND HH GENERIC SHIPPR | Kirkland Hh Generic Shippr | SHIP_40_1J | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU21470/SHIP_40_1J | 6OU21470 | COR KIRKLAND HH GENERIC SHIPPR | Kirkland Hh Generic Shippr | SHIP_40_1J | NY-Analytical | 731 | Shipper Inspection | Width | 11.25 | in | FDA,4 | 11.1875 | 11.5625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 6OU31450 | Prep H Totables 50ct Display | Prep H Totables 50Ct. Display | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 6OU31450 | Prep H Totables 50ct Display | Prep H Totables 50Ct. Display | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 6OU31450 | Prep H Totables 50ct Display | Prep H Totables 50Ct. Display | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 44 | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 6OU31450 | Prep H Totables 50ct Display | Prep H Totables 50Ct. Display | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 6OU31450 | Prep H Totables 50ct Display | Prep H Totables 50Ct. Display | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | Height | 6.125 | in | FDA,4 | 5.9375 | 6.3125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 6OU31450 | Prep H Totables 50ct Display | Prep H Totables 50Ct. Display | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | Length | 18.5 | in | FDA,4 | 18.3125 | 18.6875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 6OU31450 | Prep H Totables 50ct Display | Prep H Totables 50Ct. Display | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 6OU31450 | Prep H Totables 50ct Display | Prep H Totables 50Ct. Display | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | 6OU31450 | Prep H Totables 50ct Display | Prep H Totables 50Ct. Display | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | Width | 13.5 | in | FDA,4 | 13.3125 | 13.6875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 6OU35199 | Dermira Shipper | Dermira Shipper | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 6OU35199 | Dermira Shipper | Dermira Shipper | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 6OU35199 | Dermira Shipper | Dermira Shipper | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 6OU35199 | Dermira Shipper | Dermira Shipper | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 6OU35199 | Dermira Shipper | Dermira Shipper | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | Height | 11.5 | in | FDA,4 | 11.25 | 11.5 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 6OU35199 | Dermira Shipper | Dermira Shipper | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | Length | 19.75 | in | FDA,4 | 19.625 | 19.875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 6OU35199 | Dermira Shipper | Dermira Shipper | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 6OU35199 | Dermira Shipper | Dermira Shipper | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | 6OU35199 | Dermira Shipper | Dermira Shipper | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | Width | 12 | in | FDA,4 | 11.875 | 12.125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | 6OU351LP | 1440ct DRM4 Layer Pad | Drm4 Layer Pad 1440Ct. | SHIP_NY_32K_NA_19.500X11.750 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Pad |  |  |  |  | Shipper Style | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | 6OU351LP | 1440ct DRM4 Layer Pad | Drm4 Layer Pad 1440Ct. | SHIP_NY_32K_NA_19.500X11.750 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | 6OU351LP | 1440ct DRM4 Layer Pad | Drm4 Layer Pad 1440Ct. | SHIP_NY_32K_NA_19.500X11.750 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | 6OU351LP | 1440ct DRM4 Layer Pad | Drm4 Layer Pad 1440Ct. | SHIP_NY_32K_NA_19.500X11.750 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | 6OU351LP | 1440ct DRM4 Layer Pad | Drm4 Layer Pad 1440Ct. | SHIP_NY_32K_NA_19.500X11.750 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | 6OU351LP | 1440ct DRM4 Layer Pad | Drm4 Layer Pad 1440Ct. | SHIP_NY_32K_NA_19.500X11.750 | NY-Analytical | 731 | Shipper Inspection | Length | 19.5 | in | FDA,4 | 19.375 | 19.625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | 6OU351LP | 1440ct DRM4 Layer Pad | Drm4 Layer Pad 1440Ct. | SHIP_NY_32K_NA_19.500X11.750 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | 6OU351LP | 1440ct DRM4 Layer Pad | Drm4 Layer Pad 1440Ct. | SHIP_NY_32K_NA_19.500X11.750 | NY-Analytical | 731 | Shipper Inspection | Width | 11.75 | in | FDA,4 | 11.625 | 11.875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 6OU351TY | 120ct DRM4 Tray | Drm4 Tray 120Ct. | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | D/C Tray |  |  |  |  | Shipper Style | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 6OU351TY | 120ct DRM4 Tray | Drm4 Tray 120Ct. | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 6OU351TY | 120ct DRM4 Tray | Drm4 Tray 120Ct. | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 6OU351TY | 120ct DRM4 Tray | Drm4 Tray 120Ct. | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 6OU351TY | 120ct DRM4 Tray | Drm4 Tray 120Ct. | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 6OU351TY | 120ct DRM4 Tray | Drm4 Tray 120Ct. | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | Height | 3.5 | in | FDA,4 | 3.4375 | 3.5625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 6OU351TY | 120ct DRM4 Tray | Drm4 Tray 120Ct. | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | Length | 19.062 | in | FDA,4 | 19 | 19.125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 6OU351TY | 120ct DRM4 Tray | Drm4 Tray 120Ct. | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | 6OU351TY | 120ct DRM4 Tray | Drm4 Tray 120Ct. | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | Width | 2.812 | in | FDA,4 | 2.75 | 2.875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU38510/SHIP_32_1BG | 6OU38510 | CVS Medicated Wipes Unprinted Shipper | CVS Medicated Wipes Unprinted Shipper | SHIP_32_1BG | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU38510/SHIP_32_1BG | 6OU38510 | CVS Medicated Wipes Unprinted Shipper | CVS Medicated Wipes Unprinted Shipper | SHIP_32_1BG | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU38510/SHIP_32_1BG | 6OU38510 | CVS Medicated Wipes Unprinted Shipper | CVS Medicated Wipes Unprinted Shipper | SHIP_32_1BG | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU38510/SHIP_32_1BG | 6OU38510 | CVS Medicated Wipes Unprinted Shipper | CVS Medicated Wipes Unprinted Shipper | SHIP_32_1BG | NY-Analytical | 731 | Shipper Inspection | Height | 6 | in | FDA,4 | 5.8125 | 6.1875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU38510/SHIP_32_1BG | 6OU38510 | CVS Medicated Wipes Unprinted Shipper | CVS Medicated Wipes Unprinted Shipper | SHIP_32_1BG | NY-Analytical | 731 | Shipper Inspection | Length | 8 | in | FDA,4 | 7.8125 | 8.1875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU38510/SHIP_32_1BG | 6OU38510 | CVS Medicated Wipes Unprinted Shipper | CVS Medicated Wipes Unprinted Shipper | SHIP_32_1BG | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU38510/SHIP_32_1BG | 6OU38510 | CVS Medicated Wipes Unprinted Shipper | CVS Medicated Wipes Unprinted Shipper | SHIP_32_1BG | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU38510/SHIP_32_1BG | 6OU38510 | CVS Medicated Wipes Unprinted Shipper | CVS Medicated Wipes Unprinted Shipper | SHIP_32_1BG | NY-Analytical | 731 | Shipper Inspection | Width | 6.125 | in | FDA,4 | 5.9375 | 6.3125 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 6OU55510 | Preparation H Rapid Relief Totables w/Lidocaine | Prep H Rapid Relief Totables W/Lidocaine | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 80 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 6OU55510 | Preparation H Rapid Relief Totables w/Lidocaine | Prep H Rapid Relief Totables W/Lidocaine | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 6OU55510 | Preparation H Rapid Relief Totables w/Lidocaine | Prep H Rapid Relief Totables W/Lidocaine | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 6OU55510 | Preparation H Rapid Relief Totables w/Lidocaine | Prep H Rapid Relief Totables W/Lidocaine | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 6OU55510 | Preparation H Rapid Relief Totables w/Lidocaine | Prep H Rapid Relief Totables W/Lidocaine | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | Height | 8 | in | FDA,4 | 7.8125 | 8.1875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 6OU55510 | Preparation H Rapid Relief Totables w/Lidocaine | Prep H Rapid Relief Totables W/Lidocaine | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | Length | 11.75 | in | FDA,4 | 11.5625 | 11.9375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 6OU55510 | Preparation H Rapid Relief Totables w/Lidocaine | Prep H Rapid Relief Totables W/Lidocaine | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 6OU55510 | Preparation H Rapid Relief Totables w/Lidocaine | Prep H Rapid Relief Totables W/Lidocaine | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | 6OU55510 | Preparation H Rapid Relief Totables w/Lidocaine | Prep H Rapid Relief Totables W/Lidocaine | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | Width | 6.375 | in | FDA,4 | 6.0625 | 6.4375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU62620/SHIP_NY_32K_100_15-9.5-8 | 6OU62620 | Preparation H Rapid Relief Totables w/Lidocaine | Prep H Rapid Relief Totables W/Lidocaine | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU62620/SHIP_NY_32K_100_15-9.5-8 | 6OU62620 | Preparation H Rapid Relief Totables w/Lidocaine | Prep H Rapid Relief Totables W/Lidocaine | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU62620/SHIP_NY_32K_100_15-9.5-8 | 6OU62620 | Preparation H Rapid Relief Totables w/Lidocaine | Prep H Rapid Relief Totables W/Lidocaine | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU62620/SHIP_NY_32K_100_15-9.5-8 | 6OU62620 | Preparation H Rapid Relief Totables w/Lidocaine | Prep H Rapid Relief Totables W/Lidocaine | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU62620/SHIP_NY_32K_100_15-9.5-8 | 6OU62620 | Preparation H Rapid Relief Totables w/Lidocaine | Prep H Rapid Relief Totables W/Lidocaine | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | Height | 8 | in | FDA,4 | 7.8125 | 8.1875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU62620/SHIP_NY_32K_100_15-9.5-8 | 6OU62620 | Preparation H Rapid Relief Totables w/Lidocaine | Prep H Rapid Relief Totables W/Lidocaine | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | Length | 15.125 | in | FDA,4 | 14.8125 | 15.1875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU62620/SHIP_NY_32K_100_15-9.5-8 | 6OU62620 | Preparation H Rapid Relief Totables w/Lidocaine | Prep H Rapid Relief Totables W/Lidocaine | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU62620/SHIP_NY_32K_100_15-9.5-8 | 6OU62620 | Preparation H Rapid Relief Totables w/Lidocaine | Prep H Rapid Relief Totables W/Lidocaine | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | RSC (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OU62620/SHIP_NY_32K_100_15-9.5-8 | 6OU62620 | Preparation H Rapid Relief Totables w/Lidocaine | Prep H Rapid Relief Totables W/Lidocaine | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | Width | 9.375 | in | FDA,4 | 9.3125 | 9.6875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 6OX46500 | Shipper - Project X (5.1 mL, 10.5 mL & 26 mL) | Shipper - ProjeCt. X (5.1 Ml, 10.5 Ml & 26 Ml) | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | 125 |  | FDA,0 |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 6OX46500 | Shipper - Project X (5.1 mL, 10.5 mL & 26 mL) | Shipper - ProjeCt. X (5.1 Ml, 10.5 Ml & 26 Ml) | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 6OX46500 | Shipper - Project X (5.1 mL, 10.5 mL & 26 mL) | Shipper - ProjeCt. X (5.1 Ml, 10.5 Ml & 26 Ml) | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 6OX46500 | Shipper - Project X (5.1 mL, 10.5 mL & 26 mL) | Shipper - ProjeCt. X (5.1 Ml, 10.5 Ml & 26 Ml) | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 6OX46500 | Shipper - Project X (5.1 mL, 10.5 mL & 26 mL) | Shipper - ProjeCt. X (5.1 Ml, 10.5 Ml & 26 Ml) | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | Height | 7.5 | in | FDA,4 | 7.3125 | 7.6875 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 6OX46500 | Shipper - Project X (5.1 mL, 10.5 mL & 26 mL) | Shipper - ProjeCt. X (5.1 Ml, 10.5 Ml & 26 Ml) | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | Length | 23.75 | in | FDA,4 | 23.5005 | 23.8755 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 6OX46500 | Shipper - Project X (5.1 mL, 10.5 mL & 26 mL) | Shipper - ProjeCt. X (5.1 Ml, 10.5 Ml & 26 Ml) | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Kraft (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 6OX46500 | Shipper - Project X (5.1 mL, 10.5 mL & 26 mL) | Shipper - ProjeCt. X (5.1 Ml, 10.5 Ml & 26 Ml) | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | 6OX46500 | Shipper - Project X (5.1 mL, 10.5 mL & 26 mL) | Shipper - ProjeCt. X (5.1 Ml, 10.5 Ml & 26 Ml) | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | Width | 15.6 | in | FDA,4 | 15.5005 | 15.8755 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6RQ55172/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | 6RQ55172 | Super Sani Cloth Shipper | Super Sani-Cloth Shipper | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Shipper Inspection |  | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6ZS505PD/SHIP_32_14A | 6ZS505PD | Generic Pad | Generic Pad | SHIP_32_14A | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Go - Meets Specifications (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6ZS505PD/SHIP_32_14A | 6ZS505PD | Generic Pad | Generic Pad | SHIP_32_14A | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | 32 | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6ZS505PD/SHIP_32_14A | 6ZS505PD | Generic Pad | Generic Pad | SHIP_32_14A | NY-Analytical | 731 | Shipper Inspection | General Inspection | Accept (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6ZS505PD/SHIP_32_14A | 6ZS505PD | Generic Pad | Generic Pad | SHIP_32_14A | NY-Analytical | 731 | Shipper Inspection | Length | 11.4375 | in | FDA,4 | 11.1875 | 11.5625 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6ZS505PD/SHIP_32_14A | 6ZS505PD | Generic Pad | Generic Pad | SHIP_32_14A | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Cut Pad - Sheet (Done) |  |  |  |  | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
6ZS505PD/SHIP_32_14A | 6ZS505PD | Generic Pad | Generic Pad | SHIP_32_14A | NY-Analytical | 731 | Shipper Inspection | Width | 5.75 | in | FDA,4 | 5.5625 | 5.9375 | TM 141 / Visual | TM141 | NY | RAW MATERIAL - NY | RM_SHIPPER | both
7A4AA580/Chem_NY_Disodium Cocoamphodipropinate | 7A4AA580 | Disodium Cocoamphodipropinate | Disodium Cocoamphodipropinate | Chem_NY_Disodium Cocoamphodipropinate | IN-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA | COA | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4AA580/Chem_NY_Disodium Cocoamphodipropinate | 7A4AA580 | Disodium Cocoamphodipropinate | Disodium Cocoamphodipropinate | Chem_NY_Disodium Cocoamphodipropinate | IN-Analytical | 447 | pH as is | pH | 9.92 |  | FDA,1 | 9 | 10 | TM 200 | TM200 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4AA580/Chem_NY_Disodium Cocoamphodipropinate | 7A4AA580 | Disodium Cocoamphodipropinate | Disodium Cocoamphodipropinate | Chem_NY_Disodium Cocoamphodipropinate | IN-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4AA580/Chem_NY_Disodium Cocoamphodipropinate | 7A4AA580 | Disodium Cocoamphodipropinate | Disodium Cocoamphodipropinate | Chem_NY_Disodium Cocoamphodipropinate | IN-Analytical | 792 | Chemical Inspection | Appearance | Clear to Hazy liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4AA580/Chem_NY_Disodium Cocoamphodipropinate | 7A4AA580 | Disodium Cocoamphodipropinate | Disodium Cocoamphodipropinate | Chem_NY_Disodium Cocoamphodipropinate | IN-Analytical | 792 | Chemical Inspection | Color | Colorless to Yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4AA580/Chem_NY_Disodium Cocoamphodipropinate | 7A4AA580 | Disodium Cocoamphodipropinate | Disodium Cocoamphodipropinate | Chem_NY_Disodium Cocoamphodipropinate | IN-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4AA580/Chem_NY_Disodium Cocoamphodipropinate | 7A4AA580 | Disodium Cocoamphodipropinate | Disodium Cocoamphodipropinate | Chem_NY_Disodium Cocoamphodipropinate | IN-Analytical | 792 | Chemical Inspection | Odor | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4AL930/Chem_08A | 7A4AL930 | Potassium Hydroxide Pellets USP/NF | Potassium Hydroxide Pellets USP/Nf | Chem_08A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4AL930/Chem_08A | 7A4AL930 | Potassium Hydroxide Pellets USP/NF | Potassium Hydroxide Pellets USP/Nf | Chem_08A | NY-Analytical | 792 | Chemical Inspection | Appearance | Pellets (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4AL930/Chem_08A | 7A4AL930 | Potassium Hydroxide Pellets USP/NF | Potassium Hydroxide Pellets USP/Nf | Chem_08A | NY-Analytical | 792 | Chemical Inspection | Color | White/Grey (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4AL930/Chem_08A | 7A4AL930 | Potassium Hydroxide Pellets USP/NF | Potassium Hydroxide Pellets USP/Nf | Chem_08A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4AL930/Chem_08A | 7A4AL930 | Potassium Hydroxide Pellets USP/NF | Potassium Hydroxide Pellets USP/Nf | Chem_08A | NY-Analytical | 792 | Chemical Inspection | Odor | Odorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4AL930/Chem_08A | 7A4AL930 | Potassium Hydroxide Pellets USP/NF | Potassium Hydroxide Pellets USP/Nf | Chem_08A | NY-Analytical | 390 | Identification | Identification for Potassium | Pass (Done) |  |  |  |  | TM400/408b/421/422/USP | TM400_TM408b_TM421_TM422_USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4AS162/Chem_09A | 7A4AS162 | Glycerin USP | Glycerin USP | Chem_09A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25�C |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4AS162/Chem_09A | 7A4AS162 | Glycerin USP | Glycerin USP | Chem_09A | NY-Analytical | 402 | Specific Gravity at 25�C | Specific Gravity at 25°C | 1.2846 |  | FDA,3 | 1.249 |  | TM 202 | TM202 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4AS162/Chem_09A | 7A4AS162 | Glycerin USP | Glycerin USP | Chem_09A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4AS162/Chem_09A | 7A4AS162 | Glycerin USP | Glycerin USP | Chem_09A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear, syrupy liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4AS162/Chem_09A | 7A4AS162 | Glycerin USP | Glycerin USP | Chem_09A | NY-Analytical | 792 | Chemical Inspection | Color | Meet USP requirements (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4AS162/Chem_09A | 7A4AS162 | Glycerin USP | Glycerin USP | Chem_09A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4AS162/Chem_09A | 7A4AS162 | Glycerin USP | Glycerin USP | Chem_09A | NY-Analytical | 792 | Chemical Inspection | Odor | Odorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4AS162/Chem_09A | 7A4AS162 | Glycerin USP | Glycerin USP | Chem_09A | NY-Analytical | 816 | Related Compounds | Individual Impurity | 0.01 | % | FDA,3 |  | 0.1 | USP <621> | USP_621 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4AS162/Chem_09A | 7A4AS162 | Glycerin USP | Glycerin USP | Chem_09A | NY-Analytical | 816 | Related Compounds | Total Impurities | 0.02 | % | FDA,3 |  | 1 | USP <621> | USP_621 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4BL625/Chem_11A | 7A4BL625 | Oliec Acid FCC | Oliec Acid FCC | Chem_11A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4BL625/Chem_11A | 7A4BL625 | Oliec Acid FCC | Oliec Acid FCC | Chem_11A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4BL625/Chem_11A | 7A4BL625 | Oliec Acid FCC | Oliec Acid FCC | Chem_11A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4BL625/Chem_11A | 7A4BL625 | Oliec Acid FCC | Oliec Acid FCC | Chem_11A | NY-Analytical | 792 | Chemical Inspection | Color | Light Yellow to Yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4BL625/Chem_11A | 7A4BL625 | Oliec Acid FCC | Oliec Acid FCC | Chem_11A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4BL625/Chem_11A | 7A4BL625 | Oliec Acid FCC | Oliec Acid FCC | Chem_11A | NY-Analytical | 792 | Chemical Inspection | Odor | Lard-Like (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4CD419/Chem_13A | 7A4CD419 | Parachlorometaxylenol USP grade | Parachlorometaxylenol USP Grade | Chem_13A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4CD419/Chem_13A | 7A4CD419 | Parachlorometaxylenol USP grade | Parachlorometaxylenol USP Grade | Chem_13A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4CD419/Chem_13A | 7A4CD419 | Parachlorometaxylenol USP grade | Parachlorometaxylenol USP Grade | Chem_13A | NY-Analytical | 792 | Chemical Inspection | Appearance | Crystalline Powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4CD419/Chem_13A | 7A4CD419 | Parachlorometaxylenol USP grade | Parachlorometaxylenol USP Grade | Chem_13A | NY-Analytical | 792 | Chemical Inspection | Color | White to off White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4CD419/Chem_13A | 7A4CD419 | Parachlorometaxylenol USP grade | Parachlorometaxylenol USP Grade | Chem_13A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4CD419/Chem_13A | 7A4CD419 | Parachlorometaxylenol USP grade | Parachlorometaxylenol USP Grade | Chem_13A | NY-Analytical | 792 | Chemical Inspection | Odor | Faint phenolic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4FD110/Chem_Sodium Chloride USP/NF | 7A4FD110 | Sodium Chloride USP/NF | Sodium Chloride USP/NF | Chem_Sodium Chloride USP/NF | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4FD110/Chem_Sodium Chloride USP/NF | 7A4FD110 | Sodium Chloride USP/NF | Sodium Chloride USP/NF | Chem_Sodium Chloride USP/NF | NY-Analytical | 792 | Chemical Inspection | Appearance | Crystals/crystalline powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4FD110/Chem_Sodium Chloride USP/NF | 7A4FD110 | Sodium Chloride USP/NF | Sodium Chloride USP/NF | Chem_Sodium Chloride USP/NF | NY-Analytical | 792 | Chemical Inspection | Color | Colorless to white (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4FD110/Chem_Sodium Chloride USP/NF | 7A4FD110 | Sodium Chloride USP/NF | Sodium Chloride USP/NF | Chem_Sodium Chloride USP/NF | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4FD110/Chem_Sodium Chloride USP/NF | 7A4FD110 | Sodium Chloride USP/NF | Sodium Chloride USP/NF | Chem_Sodium Chloride USP/NF | NY-Analytical | 792 | Chemical Inspection | Odor | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4FD110/Chem_Sodium Chloride USP/NF | 7A4FD110 | Sodium Chloride USP/NF | Sodium Chloride USP/NF | Chem_Sodium Chloride USP/NF | NY-Analytical | 390 | Identification | Identification for Chloride | Meets USP (Done) |  |  |  |  | TM400/408b/421/422/USP | TM400_TM408b_TM421_TM422_USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4FD110/Chem_Sodium Chloride USP/NF | 7A4FD110 | Sodium Chloride USP/NF | Sodium Chloride USP/NF | Chem_Sodium Chloride USP/NF | NY-Analytical | 390 | Identification | Identification for Sodium | Meets USP (Done) |  |  |  |  | TM400/408b/421/422/USP | TM400_TM408b_TM421_TM422_USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4FD110/Chem_Sodium Chloride USP/NF | 7A4FD110 | Sodium Chloride USP/NF | Sodium Chloride USP/NF | Chem_Sodium Chloride USP/NF | NY-Analytical | 802 | Loss on Drying | Loss on Drying |  | % | FDA,1 |  | 0.5 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4FD110/Chem_Sodium Chloride USP/NF | 7A4FD110 | Sodium Chloride USP/NF | Sodium Chloride USP/NF | Chem_Sodium Chloride USP/NF | NY-Analytical | 802 | Loss on Drying | Loss on Drying Final |  | % | FDA,1 |  | 0.5 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4FD110/Chem_Sodium Chloride USP/NF | 7A4FD110 | Sodium Chloride USP/NF | Sodium Chloride USP/NF | Chem_Sodium Chloride USP/NF | NY-Analytical | 796 | Assay | Assay on Dried Basis | 99.7 | % | FDA,1 | 99 | 100.5 | USP <541> | USP_541 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4FD110/Chem_Sodium Chloride USP/NF | 7A4FD110 | Sodium Chloride USP/NF | Sodium Chloride USP/NF | Chem_Sodium Chloride USP/NF | NY-Analytical | 796 | Assay | Assay on Dried Basis Final | 99.7 | % | FDA,1 | 99 | 100.5 | USP <541> | USP_541 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OA347/Chem_NY_Sodium-L(+) Lactate 60% | 7A4OA347 | Sodium-L (+)-Lactate 60% | Sodium-L (+)-LaCt. ate 60% | Chem_NY_Sodium-L(+) Lactate 60% | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OA347/Chem_NY_Sodium-L(+) Lactate 60% | 7A4OA347 | Sodium-L (+)-Lactate 60% | Sodium-L (+)-LaCt. ate 60% | Chem_NY_Sodium-L(+) Lactate 60% | NY-Analytical | 398 | pH | pH (1 in 20 solution) | 6.86 |  | FDA,1 | 6.5 | 7.5 | TM 200 | TM200 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OA347/Chem_NY_Sodium-L(+) Lactate 60% | 7A4OA347 | Sodium-L (+)-Lactate 60% | Sodium-L (+)-LaCt. ate 60% | Chem_NY_Sodium-L(+) Lactate 60% | NY-Analytical | 390 | Identification | IR Identification | Comparison of scans of sample vs standard (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OA347/Chem_NY_Sodium-L(+) Lactate 60% | 7A4OA347 | Sodium-L (+)-Lactate 60% | Sodium-L (+)-LaCt. ate 60% | Chem_NY_Sodium-L(+) Lactate 60% | NY-Analytical | 792 | Chemical Inspection | Appearance | Transparent liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OA347/Chem_NY_Sodium-L(+) Lactate 60% | 7A4OA347 | Sodium-L (+)-Lactate 60% | Sodium-L (+)-LaCt. ate 60% | Chem_NY_Sodium-L(+) Lactate 60% | NY-Analytical | 792 | Chemical Inspection | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OA347/Chem_NY_Sodium-L(+) Lactate 60% | 7A4OA347 | Sodium-L (+)-Lactate 60% | Sodium-L (+)-LaCt. ate 60% | Chem_NY_Sodium-L(+) Lactate 60% | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OA347/Chem_NY_Sodium-L(+) Lactate 60% | 7A4OA347 | Sodium-L (+)-Lactate 60% | Sodium-L (+)-LaCt. ate 60% | Chem_NY_Sodium-L(+) Lactate 60% | NY-Analytical | 792 | Chemical Inspection | Odor | Odorless or with a slight Characteristic Odor (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OA709/Chem_NY_Chamomile Extract # 11-19-2 | 7A4OA709 | Chamomile Extract # 11-19-2 | Chamomile ExtraCt. # 11-19-2 | Chem_NY_Chamomile Extract # 11-19-2 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OA709/Chem_NY_Chamomile Extract # 11-19-2 | 7A4OA709 | Chamomile Extract # 11-19-2 | Chamomile ExtraCt. # 11-19-2 | Chem_NY_Chamomile Extract # 11-19-2 | NY-Analytical | 398 | pH | pH | 4.92 |  | FDA,1 | 4 | 6.5 | TM 200 | TM200 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OA709/Chem_NY_Chamomile Extract # 11-19-2 | 7A4OA709 | Chamomile Extract # 11-19-2 | Chamomile ExtraCt. # 11-19-2 | Chem_NY_Chamomile Extract # 11-19-2 | NY-Analytical | 402 | Specific Gravity at 25�C | Specific Gravity at 25°C | 1.0482 |  | FDA,3 | 1.043 | 1.053 | TM 202 | TM202 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OA709/Chem_NY_Chamomile Extract # 11-19-2 | 7A4OA709 | Chamomile Extract # 11-19-2 | Chamomile ExtraCt. # 11-19-2 | Chem_NY_Chamomile Extract # 11-19-2 | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | 1.412 |  | FDA,3 | 1.409 | 1.419 | TM 204 | TM204 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OA709/Chem_NY_Chamomile Extract # 11-19-2 | 7A4OA709 | Chamomile Extract # 11-19-2 | Chamomile ExtraCt. # 11-19-2 | Chem_NY_Chamomile Extract # 11-19-2 | NY-Analytical | 792 | Chemical Inspection | Contamination | None- free of foreign particles (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OA709/Chem_NY_Chamomile Extract # 11-19-2 | 7A4OA709 | Chamomile Extract # 11-19-2 | Chamomile ExtraCt. # 11-19-2 | Chem_NY_Chamomile Extract # 11-19-2 | NY-Analytical | 792 | Chemical Inspection | Description | Clear Amber to brown liquid having a characteristic odor (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | 7A4OA911 | Fragrance Hugs & Kisses Sheer 1 | Fragrance Hugs & Kisses Sheer 1 | Chem_NY_Fragrance Hugs & Kisses Sheer1 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Refractive Index at 25�C |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | 7A4OA911 | Fragrance Hugs & Kisses Sheer 1 | Fragrance Hugs & Kisses Sheer 1 | Chem_NY_Fragrance Hugs & Kisses Sheer1 | NY-Analytical | 451 | Refractive Index at 25�C | Refractive Index | 1.4537 |  | FDA,4 | 1.4536 | 1.4636 | TM 204 | TM204 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | 7A4OA911 | Fragrance Hugs & Kisses Sheer 1 | Fragrance Hugs & Kisses Sheer 1 | Chem_NY_Fragrance Hugs & Kisses Sheer1 | NY-Analytical | 792 | Chemical Inspection | Appearance | Liquid (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | 7A4OA911 | Fragrance Hugs & Kisses Sheer 1 | Fragrance Hugs & Kisses Sheer 1 | Chem_NY_Fragrance Hugs & Kisses Sheer1 | NY-Analytical | 792 | Chemical Inspection | Color | Colorless to Slightly Yellow (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | 7A4OA911 | Fragrance Hugs & Kisses Sheer 1 | Fragrance Hugs & Kisses Sheer 1 | Chem_NY_Fragrance Hugs & Kisses Sheer1 | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | 7A4OA911 | Fragrance Hugs & Kisses Sheer 1 | Fragrance Hugs & Kisses Sheer 1 | Chem_NY_Fragrance Hugs & Kisses Sheer1 | NY-Analytical | 792 | Chemical Inspection | Odor | Match standard (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | 7A4OA911 | Fragrance Hugs & Kisses Sheer 1 | Fragrance Hugs & Kisses Sheer 1 | Chem_NY_Fragrance Hugs & Kisses Sheer1 | NY-Analytical | 390 | Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
| 7A4OB223 | Sodium Thiosulfate |  | Analytical - Chem_189A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  |  |  |  | left_only
| 7A4OB223 | Sodium Thiosulfate |  | Analytical - Chem_189A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 |  |  |  | left_only
| 7A4OB223 | Sodium Thiosulfate |  | Analytical - Chem_189A | NY-Analytical | 792 | Chemical Inspection | Appearance | Translucent crystals or crystalline powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7A4OB223 | Sodium Thiosulfate |  | Analytical - Chem_189A | NY-Analytical | 792 | Chemical Inspection | Color | Colorless to white (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7A4OB223 | Sodium Thiosulfate |  | Analytical - Chem_189A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7A4OB223 | Sodium Thiosulfate |  | Analytical - Chem_189A | NY-Analytical | 792 | Chemical Inspection | Odor | Odorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis | COA |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_125A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | [5 empty] | Certificate of Analysis | COA |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH |  |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_125A | NY-Analytical | 398 | pH | pH (10% Solution) |  |  | FDA,1 | 6 | 8 | TM 200 | TM200 |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 402 | Specific Gravity at 25�C | Specific Gravity at 25°C | 0.7837 |  | FDA,3 | 0.783 | 0.787 | TM 202 | TM202 |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 402 | Specific Gravity at 25�C | Specific Gravity at 25°C | 0.7839 |  | FDA,3 | 0.783 | 0.787 | TM 202 | TM202 |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C |  |  | FDA,3 | 1.376 | 1.378 | TM 204 | TM204 |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear, mobile liquid (Done) |  |  |  |  | TM 208 | TM208 |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 792 | Chemical Inspection | Color | Colorless (Done) |  |  |  |  | TM 208 | TM208 |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 792 | Chemical Inspection | Odor | IPA - Characteristic (Done) |  |  |  |  | TM 208 | TM208 |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 797 | Nonvolatile Residue | End Weight |  | g | FDA,4 |  |  | TM 226 | TM226 |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 797 | Nonvolatile Residue | Nonvolatile Residue |  | % | FDA,3 |  | 0.005 | TM 226 | TM226 |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 797 | Nonvolatile Residue | Nonvolatile Residue Final |  | % | FDA,3 |  | 0.005 | TM 226 | TM226 |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 797 | Nonvolatile Residue | Start Weight |  | g | FDA,4 |  |  | TM 226 | TM226 |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 797 | Nonvolatile Residue | Weight of Residue |  | g | FDA,4 |  |  | TM 226 | TM226 |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_125A | NY-Analytical | 796 | Assay | Assay |  | % | FDA,1 | 76 | 84 | TM 301/434 | TM301_TM434 |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_125A | NY-Analytical | 796 | Assay | Assay Final |  | % | FDA,1 | 76 | 84 | TM 301/434 | TM301_TM434 |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 795 | Acidity | Acidity |  | mL | FDA,2 |  | 0.7 | TM 375 | TM375 |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 795 | Acidity | Acidity Final (Calc) |  | mL | FDA,2 |  | 0.7 | TM 375 | TM375 |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 795 | Acidity | Titrant End Volume |  | mL | FDA,2 |  |  | TM 375 | TM375 |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 795 | Acidity | Titrant Start Volume |  | mL | FDA,2 |  |  | TM 375 | TM375 |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 796 | Assay | Assay |  | % | FDA,1 | 99 |  | TM 376 | TM376 |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 796 | Assay | Assay | 99.96 | % | FDA,1 | 99 |  | TM 376 | TM376 |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 796 | Assay | Assay Final |  | % | FDA,1 | 99 |  | TM 376 | TM376 |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 796 | Assay | Assay Final |  | % | FDA,1 | 99 |  | TM 376 | TM376 |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_125A | NY-Analytical | 390 | Identification | IR Identification | [5 empty] | TM 411 | TM411 |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 390 | Identification | IR Identification | Match standard, per USP (Done) |  |  |  |  | TM 411 | TM411 |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 390 | Identification | IR Identification | [5 empty] | TM 411 | TM411 |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_125A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear viscous liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 792 | Chemical Inspection | Appearance | [5 empty] | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_125A | NY-Analytical | 792 | Chemical Inspection | Color | [5 empty] | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 792 | Chemical Inspection | Color | [5 empty] | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_125A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 792 | Chemical Inspection | Contamination | [5 empty] | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_125A | NY-Analytical | 792 | Chemical Inspection | Odor | [5 empty] | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7A4OB314 | Isopropyl alcohol 99% USP |  | Analytical - Chem_NY_Isopropyl alcohol 99% USP_Reduced Test | NY-Analytical | 792 | Chemical Inspection | Odor | [5 empty] | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB339 | Isopropyl alcohol 99% USP-Inline | Isopropyl Alcohol 99% USP-Inline | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | 1.376 |  | FDA,3 | 1.376 | 1.378 | TM 204 | TM204 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB339 | Isopropyl alcohol 99% USP-Inline | Isopropyl Alcohol 99% USP-Inline | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 797 | Nonvolatile Residue | End Weight | 92.0885 | g | FDA,4 |  |  | TM 226 | TM226 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB339 | Isopropyl alcohol 99% USP-Inline | Isopropyl Alcohol 99% USP-Inline | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 797 | Nonvolatile Residue | Nonvolatile Residue | 0 | % | FDA,3 |  | 0.005 | TM 226 | TM226 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB339 | Isopropyl alcohol 99% USP-Inline | Isopropyl Alcohol 99% USP-Inline | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 797 | Nonvolatile Residue | Nonvolatile Residue Final | 0 | % | FDA,3 |  | 0.005 | TM 226 | TM226 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB339 | Isopropyl alcohol 99% USP-Inline | Isopropyl Alcohol 99% USP-Inline | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 797 | Nonvolatile Residue | Start Weight | 92.0885 | g | FDA,4 |  |  | TM 226 | TM226 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB339 | Isopropyl alcohol 99% USP-Inline | Isopropyl Alcohol 99% USP-Inline | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 797 | Nonvolatile Residue | Weight of Residue | 0 | g | FDA,4 |  |  | TM 226 | TM226 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB339 | Isopropyl alcohol 99% USP-Inline | Isopropyl Alcohol 99% USP-Inline | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 795 | Acidity | Acidity | 0.4 | mL | FDA,2 |  | 0.7 | TM 375 | TM375 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB339 | Isopropyl alcohol 99% USP-Inline | Isopropyl Alcohol 99% USP-Inline | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 795 | Acidity | Acidity Final (Calc) | 0.425 | mL | FDA,2 |  | 0.7 | TM 375 | TM375 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB339 | Isopropyl alcohol 99% USP-Inline | Isopropyl Alcohol 99% USP-Inline | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 795 | Acidity | Titrant End Volume | 0.5 | mL | FDA,2 |  |  | TM 375 | TM375 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | 7A4OB339 | Isopropyl alcohol 99% USP-Inline | Isopropyl Alcohol 99% USP-Inline | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 795 | Acidity | Titrant Start Volume | 0 | mL | FDA,2 |  |  | TM 375 | TM375 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
| 7A4OB397 | Benzalkonium Chloride 50% |  | Analytical - Chem_04A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Assay |  |  |  |  | left_only
| 7A4OB397 | Benzalkonium Chloride 50% |  | Analytical - Chem_04A | NY-Analytical | 796 | Assay | Assay | 55.85 | % | FDA,1 | 47.5 | 52.5 | TM 301/434 | TM301_TM434 |  |  |  | left_only
| 7A4OB397 | Benzalkonium Chloride 50% |  | Analytical - Chem_04A | NY-Analytical | 796 | Assay | Assay Final | 55.77 | % | FDA,1 | 47.5 | 52.5 | TM 301/434 | TM301_TM434 |  |  |  | left_only
| 7A4OB397 | Benzalkonium Chloride 50% |  | Analytical - Chem_04A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear viscous liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7A4OB397 | Benzalkonium Chloride 50% |  | Analytical - Chem_04A | NY-Analytical | 792 | Chemical Inspection | Color | Colorless to light yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7A4OB397 | Benzalkonium Chloride 50% |  | Analytical - Chem_04A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7A4OB397 | Benzalkonium Chloride 50% |  | Analytical - Chem_04A | NY-Analytical | 792 | Chemical Inspection | Odor | Mild almond (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7A4OB397 | Benzalkonium Chloride 50% |  | Analytical - Chem_04A | NY-Analytical | 390 | Identification | Identification | Formation of precipitate per test (Done) |  |  |  |  | TM400/408b/421/422/USP | TM400_TM408b_TM421_TM422_USP |  |  |  | left_only
7A4OB690/Chem_01A | 7A4OB690 | Isopar K | Isopar K | Chem_01A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OB690/Chem_01A | 7A4OB690 | Isopar K | Isopar K | Chem_01A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OB690/Chem_01A | 7A4OB690 | Isopar K | Isopar K | Chem_01A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear, mobile liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OB690/Chem_01A | 7A4OB690 | Isopar K | Isopar K | Chem_01A | NY-Analytical | 792 | Chemical Inspection | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OB690/Chem_01A | 7A4OB690 | Isopar K | Isopar K | Chem_01A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OB690/Chem_01A | 7A4OB690 | Isopar K | Isopar K | Chem_01A | NY-Analytical | 792 | Chemical Inspection | Odor | Solvent (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OF602/Chem_19A | 7A4OF602 | Avon Dry Facial Fragrance/Watery Floral | Avon Dry Facial Fragrance/Watery Floral | Chem_19A | NY-Analytical | 402 | Specific Gravity at 25�C | Specific Gravity at 25°C | 0.968 |  | FDA,3 | 0.964 | 0.968 | TM 202 | TM202 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7A4OF602/Chem_19A | 7A4OF602 | Avon Dry Facial Fragrance/Watery Floral | Avon Dry Facial Fragrance/Watery Floral | Chem_19A | NY-Analytical | 451 | Refractive Index at 25�C | Refractive Index | 1.458 |  | FDA,3 | 1.455 | 1.461 | TM 204 | TM204 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7A4OF602/Chem_19A | 7A4OF602 | Avon Dry Facial Fragrance/Watery Floral | Avon Dry Facial Fragrance/Watery Floral | Chem_19A | NY-Analytical | 390 | Identification | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411 | TM411 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7A4OF602/Chem_19A | 7A4OF602 | Avon Dry Facial Fragrance/Watery Floral | Avon Dry Facial Fragrance/Watery Floral | Chem_19A | NY-Analytical | 792 | Chemical Inspection | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7A4OF602/Chem_19A | 7A4OF602 | Avon Dry Facial Fragrance/Watery Floral | Avon Dry Facial Fragrance/Watery Floral | Chem_19A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7A4OG169/Chem_20A | 7A4OG169 | Dowfax C10L | Dowfax C10L | Chem_20A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OG169/Chem_20A | 7A4OG169 | Dowfax C10L | Dowfax C10L | Chem_20A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OG169/Chem_20A | 7A4OG169 | Dowfax C10L | Dowfax C10L | Chem_20A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OG169/Chem_20A | 7A4OG169 | Dowfax C10L | Dowfax C10L | Chem_20A | NY-Analytical | 792 | Chemical Inspection | Color | Amber (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OG169/Chem_20A | 7A4OG169 | Dowfax C10L | Dowfax C10L | Chem_20A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OH247/Chem_21A | 7A4OH247 | Trisodium Phosphate Crystalline (TSPc) | Trisodium pHosphate Crystalline (Tspc) | Chem_21A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OH247/Chem_21A | 7A4OH247 | Trisodium Phosphate Crystalline (TSPc) | Trisodium pHosphate Crystalline (Tspc) | Chem_21A | NY-Analytical | 792 | Chemical Inspection | Appearance | White granular material (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OH247/Chem_21A | 7A4OH247 | Trisodium Phosphate Crystalline (TSPc) | Trisodium pHosphate Crystalline (Tspc) | Chem_21A | NY-Analytical | 792 | Chemical Inspection | Color | Match reference (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OH247/Chem_21A | 7A4OH247 | Trisodium Phosphate Crystalline (TSPc) | Trisodium pHosphate Crystalline (Tspc) | Chem_21A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OH247/Chem_21A | 7A4OH247 | Trisodium Phosphate Crystalline (TSPc) | Trisodium pHosphate Crystalline (Tspc) | Chem_21A | NY-Analytical | 792 | Chemical Inspection | Odor | Match reference (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | CHE SALICYLIC ACID USP | CHE Salicylic Acid USP | Chem_NY_Salicylic Acid USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | CHE SALICYLIC ACID USP | CHE Salicylic Acid USP | Chem_NY_Salicylic Acid USP | NY-Analytical | 792 | Chemical Inspection | Appearance | White crystalline powder (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | CHE SALICYLIC ACID USP | CHE Salicylic Acid USP | Chem_NY_Salicylic Acid USP | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | CHE SALICYLIC ACID USP | CHE Salicylic Acid USP | Chem_NY_Salicylic Acid USP | NY-Analytical | 832 | Chloride | Chloride | <0.014 | % | FDA,3 |  | 0.014 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | CHE SALICYLIC ACID USP | CHE Salicylic Acid USP | Chem_NY_Salicylic Acid USP | NY-Analytical | 833 | Heavy Metals | Heavy Metals | <20 | ug/g | FDA,0 |  | <20 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | CHE SALICYLIC ACID USP | CHE Salicylic Acid USP | Chem_NY_Salicylic Acid USP | NY-Analytical | 816 | Related Compounds | Impurity | 0 | % | FDA,4 |  | 0.05 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | CHE SALICYLIC ACID USP | CHE Salicylic Acid USP | Chem_NY_Salicylic Acid USP | NY-Analytical | 802 | Loss on Drying | Loss on Drying | 0.03 | % | FDA,3 |  | 0.5 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | CHE SALICYLIC ACID USP | CHE Salicylic Acid USP | Chem_NY_Salicylic Acid USP | NY-Analytical | 816 | Related Compounds | Phenol | 0 | % | FDA,4 |  | 0.01 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | CHE SALICYLIC ACID USP | CHE Salicylic Acid USP | Chem_NY_Salicylic Acid USP | NY-Analytical | 816 | Related Compounds | Related Compound A | 0.043 | % | FDA,4 |  | 0.1 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | CHE SALICYLIC ACID USP | CHE Salicylic Acid USP | Chem_NY_Salicylic Acid USP | NY-Analytical | 816 | Related Compounds | Related Compound B | 0.037 | % | FDA,4 |  | 0.05 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | CHE SALICYLIC ACID USP | CHE Salicylic Acid USP | Chem_NY_Salicylic Acid USP | NY-Analytical | 834 | Residue on Ignition | Residue on Ignition | 0.02 | % | FDA,2 |  | 0.05 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | CHE SALICYLIC ACID USP | CHE Salicylic Acid USP | Chem_NY_Salicylic Acid USP | NY-Analytical | 835 | Sulfate Content | Sulfate Content | <0.020 | % | FDA,3 |  | 0.02 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | CHE SALICYLIC ACID USP | CHE Salicylic Acid USP | Chem_NY_Salicylic Acid USP | NY-Analytical | 816 | Related Compounds | Total Impurities | 0.081 | % | FDA,4 |  | 0.2 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | CHE SALICYLIC ACID USP | CHE Salicylic Acid USP | Chem_NY_Salicylic Acid USP | NY-Analytical | 796 | Assay | Assay | 99.7 | % | FDA,1 | 98 | 102 | USP/NF | USP_NF | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | CHE SALICYLIC ACID USP | CHE Salicylic Acid USP | Chem_NY_Salicylic Acid USP | NY-Analytical | 455 | Identification A | Identification A | Passes Test (Done) |  |  |  |  | USP/NF | USP_NF | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OK459/Chem_NY_Salicylic Acid USP | 7A4OK459 | CHE SALICYLIC ACID USP | CHE Salicylic Acid USP | Chem_NY_Salicylic Acid USP | NY-Analytical | 456 | Identification B | Identification B | Passes Test (Done) |  |  |  |  | USP/NF | USP_NF | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OM002/Chem_NY_ Dissolvine GL-47-S | 7A4OM002 | Dissolvine GL-47-S | Dissolvine GL-47-S | Chem_NY_ Dissolvine GL-47-S | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH |  | JB | RAW MATERIAL - ARKANSAS | RM_CHEMICAL | both
7A4OM002/Chem_NY_ Dissolvine GL-47-S | 7A4OM002 | Dissolvine GL-47-S | Dissolvine GL-47-S | Chem_NY_ Dissolvine GL-47-S | NY-Analytical | 398 | pH | pH (1% w/v aqueous solution) | 11.14 |  | FDA,1 | 11 | 11.8 | TM 200 | TM200 | JB | RAW MATERIAL - ARKANSAS | RM_CHEMICAL | both
7A4OM002/Chem_NY_ Dissolvine GL-47-S | 7A4OM002 | Dissolvine GL-47-S | Dissolvine GL-47-S | Chem_NY_ Dissolvine GL-47-S | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear Liquid (Done) |  |  |  |  | TM 208 | TM208 | JB | RAW MATERIAL - ARKANSAS | RM_CHEMICAL | both
7A4OM002/Chem_NY_ Dissolvine GL-47-S | 7A4OM002 | Dissolvine GL-47-S | Dissolvine GL-47-S | Chem_NY_ Dissolvine GL-47-S | NY-Analytical | 792 | Chemical Inspection | Color | Characteristic (Done) |  |  |  |  | TM 208 | TM208 | JB | RAW MATERIAL - ARKANSAS | RM_CHEMICAL | both
7A4OM002/Chem_NY_ Dissolvine GL-47-S | 7A4OM002 | Dissolvine GL-47-S | Dissolvine GL-47-S | Chem_NY_ Dissolvine GL-47-S | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | JB | RAW MATERIAL - ARKANSAS | RM_CHEMICAL | both
7A4OM002/Chem_NY_ Dissolvine GL-47-S | 7A4OM002 | Dissolvine GL-47-S | Dissolvine GL-47-S | Chem_NY_ Dissolvine GL-47-S | NY-Analytical | 792 | Chemical Inspection | Odor | Characteristic (Done) |  |  |  |  | TM 208 | TM208 | JB | RAW MATERIAL - ARKANSAS | RM_CHEMICAL | both
7A4OM002/Chem_NY_ Dissolvine GL-47-S | 7A4OM002 | Dissolvine GL-47-S | Dissolvine GL-47-S | Chem_NY_ Dissolvine GL-47-S | NY-Analytical | 390 | Identification | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411 | TM411 | JB | RAW MATERIAL - ARKANSAS | RM_CHEMICAL | both
7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP027 | Ceraphyl 41 | Ceraphyl 41 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP027 | Ceraphyl 41 | Ceraphyl 41 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 792 | Chemical Inspection | Appearance | A clear colorless to pale straw colored liquid. (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP027 | Ceraphyl 41 | Ceraphyl 41 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 792 | Chemical Inspection | Odor | Passes Test (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP027 | Ceraphyl 41 | Ceraphyl 41 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP027 | Ceraphyl 41 | Ceraphyl 41 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 390 | Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP027 | Ceraphyl 41 | Ceraphyl 41 | MASTER - Microbiology - J&J Chemical Testing) | NY-Microbiology | 524 | J&J Fungal | Fungal | 0 | CFU/g | FDA,0 | 0 | 200 | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP027 | Ceraphyl 41 | Ceraphyl 41 | MASTER - Microbiology - J&J Chemical Testing) | NY-Microbiology | 700 | J&J SDA Controls | Negative Control | [5 empty] | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP027 | Ceraphyl 41 | Ceraphyl 41 | MASTER - Microbiology - J&J Chemical Testing) | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | No Growth TSB |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP027 | Ceraphyl 41 | Ceraphyl 41 | MASTER - Microbiology - J&J Chemical Testing) | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | 0 | cfu/mL | FDA,0 | 0 | 500 | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP065/Chem_NY_ Fragrance Plaisir Gourmand 23 | 7A4OP065 | Fragrance Plaisir Gourmand 23 | Fragrance Plaisir Gourmand 23 | Chem_NY_ Fragrance Plaisir Gourmand 23 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP065/Chem_NY_ Fragrance Plaisir Gourmand 23 | 7A4OP065 | Fragrance Plaisir Gourmand 23 | Fragrance Plaisir Gourmand 23 | Chem_NY_ Fragrance Plaisir Gourmand 23 | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear Liquid (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP065/Chem_NY_ Fragrance Plaisir Gourmand 23 | 7A4OP065 | Fragrance Plaisir Gourmand 23 | Fragrance Plaisir Gourmand 23 | Chem_NY_ Fragrance Plaisir Gourmand 23 | NY-Analytical | 792 | Chemical Inspection | Color | Pale Yellow to Yellow (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP065/Chem_NY_ Fragrance Plaisir Gourmand 23 | 7A4OP065 | Fragrance Plaisir Gourmand 23 | Fragrance Plaisir Gourmand 23 | Chem_NY_ Fragrance Plaisir Gourmand 23 | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP065/Chem_NY_ Fragrance Plaisir Gourmand 23 | 7A4OP065 | Fragrance Plaisir Gourmand 23 | Fragrance Plaisir Gourmand 23 | Chem_NY_ Fragrance Plaisir Gourmand 23 | NY-Analytical | 792 | Chemical Inspection | Odor | Match standard (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP065/Chem_NY_ Fragrance Plaisir Gourmand 23 | 7A4OP065 | Fragrance Plaisir Gourmand 23 | Fragrance Plaisir Gourmand 23 | Chem_NY_ Fragrance Plaisir Gourmand 23 | NY-Analytical | 390 | Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP073/Chem_NY_Fragrance Bali Leia 5 | 7A4OP073 | Fragrance Bali Leia 5 | Fragrance Bali Leia 5 | Chem_NY_Fragrance Bali Leia 5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP073/Chem_NY_Fragrance Bali Leia 5 | 7A4OP073 | Fragrance Bali Leia 5 | Fragrance Bali Leia 5 | Chem_NY_Fragrance Bali Leia 5 | NY-Analytical | 792 | Chemical Inspection | Appearance | Liquid (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP073/Chem_NY_Fragrance Bali Leia 5 | 7A4OP073 | Fragrance Bali Leia 5 | Fragrance Bali Leia 5 | Chem_NY_Fragrance Bali Leia 5 | NY-Analytical | 792 | Chemical Inspection | Color | Dark yellow to lemon yellow (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP073/Chem_NY_Fragrance Bali Leia 5 | 7A4OP073 | Fragrance Bali Leia 5 | Fragrance Bali Leia 5 | Chem_NY_Fragrance Bali Leia 5 | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP073/Chem_NY_Fragrance Bali Leia 5 | 7A4OP073 | Fragrance Bali Leia 5 | Fragrance Bali Leia 5 | Chem_NY_Fragrance Bali Leia 5 | NY-Analytical | 792 | Chemical Inspection | Odor | Match standard (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP073/Chem_NY_Fragrance Bali Leia 5 | 7A4OP073 | Fragrance Bali Leia 5 | Fragrance Bali Leia 5 | Chem_NY_Fragrance Bali Leia 5 | NY-Analytical | 390 | Identification | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP081/Chem_NY_Fragrance SC 2705 LF | 7A4OP081 | Fragrance SC 2705 LF | Fragrance Sc 2705 Lf | Chem_NY_Fragrance SC 2705 LF | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP081/Chem_NY_Fragrance SC 2705 LF | 7A4OP081 | Fragrance SC 2705 LF | Fragrance Sc 2705 Lf | Chem_NY_Fragrance SC 2705 LF | NY-Analytical | 792 | Chemical Inspection | Appearance | Liquid (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP081/Chem_NY_Fragrance SC 2705 LF | 7A4OP081 | Fragrance SC 2705 LF | Fragrance Sc 2705 Lf | Chem_NY_Fragrance SC 2705 LF | NY-Analytical | 792 | Chemical Inspection | Color | Colorless to very light yellow (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP081/Chem_NY_Fragrance SC 2705 LF | 7A4OP081 | Fragrance SC 2705 LF | Fragrance Sc 2705 Lf | Chem_NY_Fragrance SC 2705 LF | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP081/Chem_NY_Fragrance SC 2705 LF | 7A4OP081 | Fragrance SC 2705 LF | Fragrance Sc 2705 Lf | Chem_NY_Fragrance SC 2705 LF | NY-Analytical | 792 | Chemical Inspection | Odor | Match standard (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP081/Chem_NY_Fragrance SC 2705 LF | 7A4OP081 | Fragrance SC 2705 LF | Fragrance Sc 2705 Lf | Chem_NY_Fragrance SC 2705 LF | NY-Analytical | 390 | Identification | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP093/Chem_NY_Essential Oil 27A Fragrance | 7A4OP093 | Essential Oil 27A Fragrance | Essential Oil 27A Fragrance | Chem_NY_Essential Oil 27A Fragrance | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP093/Chem_NY_Essential Oil 27A Fragrance | 7A4OP093 | Essential Oil 27A Fragrance | Essential Oil 27A Fragrance | Chem_NY_Essential Oil 27A Fragrance | NY-Analytical | 792 | Chemical Inspection | Appearance | Liquid (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP093/Chem_NY_Essential Oil 27A Fragrance | 7A4OP093 | Essential Oil 27A Fragrance | Essential Oil 27A Fragrance | Chem_NY_Essential Oil 27A Fragrance | NY-Analytical | 792 | Chemical Inspection | Color | Amber (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP093/Chem_NY_Essential Oil 27A Fragrance | 7A4OP093 | Essential Oil 27A Fragrance | Essential Oil 27A Fragrance | Chem_NY_Essential Oil 27A Fragrance | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP093/Chem_NY_Essential Oil 27A Fragrance | 7A4OP093 | Essential Oil 27A Fragrance | Essential Oil 27A Fragrance | Chem_NY_Essential Oil 27A Fragrance | NY-Analytical | 792 | Chemical Inspection | Odor | Match standard (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP093/Chem_NY_Essential Oil 27A Fragrance | 7A4OP093 | Essential Oil 27A Fragrance | Essential Oil 27A Fragrance | Chem_NY_Essential Oil 27A Fragrance | NY-Analytical | 390 | Identification | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP113/Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | 7A4OP113 | Vitis Vinifera (Grape) Fruit Water | Vitis Vinifera (Grape) Fruit Water | Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP113/Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | 7A4OP113 | Vitis Vinifera (Grape) Fruit Water | Vitis Vinifera (Grape) Fruit Water | Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | NY-Analytical | 398 | pH | pH | 4.11 |  | FDA,1 | 2.5 | 6 | TM 200 | TM200 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP113/Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | 7A4OP113 | Vitis Vinifera (Grape) Fruit Water | Vitis Vinifera (Grape) Fruit Water | Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear to slightly cloudy liquid (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP113/Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | 7A4OP113 | Vitis Vinifera (Grape) Fruit Water | Vitis Vinifera (Grape) Fruit Water | Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | NY-Analytical | 792 | Chemical Inspection | Color | Colorless (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP113/Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | 7A4OP113 | Vitis Vinifera (Grape) Fruit Water | Vitis Vinifera (Grape) Fruit Water | Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP113/Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | 7A4OP113 | Vitis Vinifera (Grape) Fruit Water | Vitis Vinifera (Grape) Fruit Water | Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | NY-Analytical | 792 | Chemical Inspection | Odor | Characteristic (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP197/Chem_NY_D-Panthenol 75W_01 | 7A4OP197 | D-Panthenol 75W | D-Panthenol 75W | Chem_NY_D-Panthenol 75W_01 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP197/Chem_NY_D-Panthenol 75W_01 | 7A4OP197 | D-Panthenol 75W | D-Panthenol 75W | Chem_NY_D-Panthenol 75W_01 | NY-Analytical | 398 | pH | pH | 6.96 |  | FDA,1 | 6 | 7.5 | TM 200 | TM200 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP197/Chem_NY_D-Panthenol 75W_01 | 7A4OP197 | D-Panthenol 75W | D-Panthenol 75W | Chem_NY_D-Panthenol 75W_01 | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | 1.466 |  | FDA,3 | 1.466 | 1.475 | TM 204 | TM204 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP197/Chem_NY_D-Panthenol 75W_01 | 7A4OP197 | D-Panthenol 75W | D-Panthenol 75W | Chem_NY_D-Panthenol 75W_01 | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear Liquid (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP197/Chem_NY_D-Panthenol 75W_01 | 7A4OP197 | D-Panthenol 75W | D-Panthenol 75W | Chem_NY_D-Panthenol 75W_01 | NY-Analytical | 792 | Chemical Inspection | Color | Slightly Yellowish or Colorless (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP197/Chem_NY_D-Panthenol 75W_01 | 7A4OP197 | D-Panthenol 75W | D-Panthenol 75W | Chem_NY_D-Panthenol 75W_01 | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP197/Chem_NY_D-Panthenol 75W_01 | 7A4OP197 | D-Panthenol 75W | D-Panthenol 75W | Chem_NY_D-Panthenol 75W_01 | NY-Analytical | 792 | Chemical Inspection | Odor | Match standard (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP197/Chem_NY_D-Panthenol 75W_01 | 7A4OP197 | D-Panthenol 75W | D-Panthenol 75W | Chem_NY_D-Panthenol 75W_01 | NY-Analytical | 390 | Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | 7A4OP205 | Cucumis Sativus (Cucumber) Fruit Rxtract | Cucumis Sativus (Cucumber) Fruit RxtraCt. | Chem_NY_Cucumis Sativus Fruit Extract | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | 7A4OP205 | Cucumis Sativus (Cucumber) Fruit Rxtract | Cucumis Sativus (Cucumber) Fruit RxtraCt. | Chem_NY_Cucumis Sativus Fruit Extract | NY-Analytical | 792 | Chemical Inspection | Appearance | Transparent Liquid to slightly Turbid (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | 7A4OP205 | Cucumis Sativus (Cucumber) Fruit Rxtract | Cucumis Sativus (Cucumber) Fruit RxtraCt. | Chem_NY_Cucumis Sativus Fruit Extract | NY-Analytical | 792 | Chemical Inspection | Color | Colorless to pale yellow (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | 7A4OP205 | Cucumis Sativus (Cucumber) Fruit Rxtract | Cucumis Sativus (Cucumber) Fruit RxtraCt. | Chem_NY_Cucumis Sativus Fruit Extract | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | 7A4OP205 | Cucumis Sativus (Cucumber) Fruit Rxtract | Cucumis Sativus (Cucumber) Fruit RxtraCt. | Chem_NY_Cucumis Sativus Fruit Extract | NY-Analytical | 792 | Chemical Inspection | Odor | Match standard (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | 7A4OP205 | Cucumis Sativus (Cucumber) Fruit Rxtract | Cucumis Sativus (Cucumber) Fruit RxtraCt. | Chem_NY_Cucumis Sativus Fruit Extract | NY-Analytical | 390 | Identification | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | 7A4OP205 | Cucumis Sativus (Cucumber) Fruit Rxtract | Cucumis Sativus (Cucumber) Fruit RxtraCt. | Chem_NY_Cucumis Sativus Fruit Extract | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | 7A4OP205 | Cucumis Sativus (Cucumber) Fruit Rxtract | Cucumis Sativus (Cucumber) Fruit RxtraCt. | Chem_NY_Cucumis Sativus Fruit Extract | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | 7A4OP205 | Cucumis Sativus (Cucumber) Fruit Rxtract | Cucumis Sativus (Cucumber) Fruit RxtraCt. | Chem_NY_Cucumis Sativus Fruit Extract | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | 7A4OP205 | Cucumis Sativus (Cucumber) Fruit Rxtract | Cucumis Sativus (Cucumber) Fruit RxtraCt. | Chem_NY_Cucumis Sativus Fruit Extract | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP345/Chem_195A | 7A4OP345 | CHE ACB BIO-CHELATE | CHE Acb Bio-CHElate | Chem_195A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH as is |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP345/Chem_195A | 7A4OP345 | CHE ACB BIO-CHELATE | CHE Acb Bio-CHElate | Chem_195A | NY-Analytical | 447 | pH as is | pH | 3.92 |  | FDA,1 | 3.5 | 4.5 | TM 200 | TM200 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP345/Chem_195A | 7A4OP345 | CHE ACB BIO-CHELATE | CHE Acb Bio-CHElate | Chem_195A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP345/Chem_195A | 7A4OP345 | CHE ACB BIO-CHELATE | CHE Acb Bio-CHElate | Chem_195A | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram) | 0 | CFU/g | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP345/Chem_195A | 7A4OP345 | CHE ACB BIO-CHELATE | CHE Acb Bio-CHElate | Chem_195A | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP345/Chem_195A | 7A4OP345 | CHE ACB BIO-CHELATE | CHE Acb Bio-CHElate | Chem_195A | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram) | 0 | CFU/g | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP345/Chem_195A | 7A4OP345 | CHE ACB BIO-CHELATE | CHE Acb Bio-CHElate | Chem_195A | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP345/Chem_195A | 7A4OP345 | CHE ACB BIO-CHELATE | CHE Acb Bio-CHElate | Chem_195A | NY-Analytical | 792 | Chemical Inspection | Appearance | Hazy liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP345/Chem_195A | 7A4OP345 | CHE ACB BIO-CHELATE | CHE Acb Bio-CHElate | Chem_195A | NY-Analytical | 792 | Chemical Inspection | Color | Light green/amber (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP345/Chem_195A | 7A4OP345 | CHE ACB BIO-CHELATE | CHE Acb Bio-CHElate | Chem_195A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP345/Chem_195A | 7A4OP345 | CHE ACB BIO-CHELATE | CHE Acb Bio-CHElate | Chem_195A | NY-Analytical | 792 | Chemical Inspection | Odor | Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP448/Chem_208A | 7A4OP448 | Fresh Scent Type #2 Fragrance | Fresh Scent Type #2 Fragrance | Chem_208A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis | COA | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP448/Chem_197A | 7A4OP448 | Fresh Scent Type #2 Fragrance | Fresh Scent Type #2 Fragrance | Chem_197A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP448/Chem_197A | 7A4OP448 | Fresh Scent Type #2 Fragrance | Fresh Scent Type #2 Fragrance | Chem_197A | NY-Analytical | 398 | pH | pH (10% in water) | 5.9 |  | FDA,1 | 9 | 11 | TM 200 | TM200 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP448/Chem_208A | 7A4OP448 | Fresh Scent Type #2 Fragrance | Fresh Scent Type #2 Fragrance | Chem_208A | NY-Analytical | 402 | Specific Gravity at 25�C | Specific Gravity at 25°C | 0.9801 |  | FDA,4 | 0.9694 | 0.9894 | TM 202 | TM202 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP448/Chem_208A | 7A4OP448 | Fresh Scent Type #2 Fragrance | Fresh Scent Type #2 Fragrance | Chem_208A | NY-Analytical | 451 | Refractive Index at 25�C | Refractive Index | 1.469 |  | FDA,4 | 1.4655 | 1.4855 | TM 204 | TM204 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP448/Chem_197A | 7A4OP448 | Fresh Scent Type #2 Fragrance | Fresh Scent Type #2 Fragrance | Chem_197A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP448/Chem_208A | 7A4OP448 | Fresh Scent Type #2 Fragrance | Fresh Scent Type #2 Fragrance | Chem_208A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP448/Chem_197A | 7A4OP448 | Fresh Scent Type #2 Fragrance | Fresh Scent Type #2 Fragrance | Chem_197A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP448/Chem_208A | 7A4OP448 | Fresh Scent Type #2 Fragrance | Fresh Scent Type #2 Fragrance | Chem_208A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear to pale yellow liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP448/Chem_197A | 7A4OP448 | Fresh Scent Type #2 Fragrance | Fresh Scent Type #2 Fragrance | Chem_197A | NY-Analytical | 792 | Chemical Inspection | Color | Yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP448/Chem_197A | 7A4OP448 | Fresh Scent Type #2 Fragrance | Fresh Scent Type #2 Fragrance | Chem_197A | NY-Analytical | 792 | Chemical Inspection | Color |  |  | FDA,1 | 0 | 3 | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP448/Chem_197A | 7A4OP448 | Fresh Scent Type #2 Fragrance | Fresh Scent Type #2 Fragrance | Chem_197A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP448/Chem_208A | 7A4OP448 | Fresh Scent Type #2 Fragrance | Fresh Scent Type #2 Fragrance | Chem_208A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP448/Chem_197A | 7A4OP448 | Fresh Scent Type #2 Fragrance | Fresh Scent Type #2 Fragrance | Chem_197A | NY-Analytical | 792 | Chemical Inspection | Odor | Fatty odor (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP448/Chem_208A | 7A4OP448 | Fresh Scent Type #2 Fragrance | Fresh Scent Type #2 Fragrance | Chem_208A | NY-Analytical | 792 | Chemical Inspection | Odor | Fresh Bouquet (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP567/Chem_NY_Surfonic L12-6_Surfactant | 7A4OP567 | Surfonic L12-6 Surfactant | Surfonic L12-6 SurfaCt. ant | Chem_NY_Surfonic L12-6_Surfactant | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA | COA | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP567/Chem_NY_Surfonic L12-6_Surfactant | 7A4OP567 | Surfonic L12-6 Surfactant | Surfonic L12-6 SurfaCt. ant | Chem_NY_Surfonic L12-6_Surfactant | NY-Analytical | 398 | pH | pH (1%) in 10:6 IPA: H2O | 6.9 |  | FDA,1 | 6.5 | 8 | TM 200 | TM200 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP567/Chem_NY_Surfonic L12-6_Surfactant | 7A4OP567 | Surfonic L12-6 Surfactant | Surfonic L12-6 SurfaCt. ant | Chem_NY_Surfonic L12-6_Surfactant | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear to slightly turbid, free of suspended matter (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP567/Chem_NY_Surfonic L12-6_Surfactant | 7A4OP567 | Surfonic L12-6 Surfactant | Surfonic L12-6 SurfaCt. ant | Chem_NY_Surfonic L12-6_Surfactant | NY-Analytical | 792 | Chemical Inspection | Color | Colorless to Slightly Yellow (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP567/Chem_NY_Surfonic L12-6_Surfactant | 7A4OP567 | Surfonic L12-6 Surfactant | Surfonic L12-6 SurfaCt. ant | Chem_NY_Surfonic L12-6_Surfactant | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP567/Chem_NY_Surfonic L12-6_Surfactant | 7A4OP567 | Surfonic L12-6 Surfactant | Surfonic L12-6 SurfaCt. ant | Chem_NY_Surfonic L12-6_Surfactant | NY-Analytical | 792 | Chemical Inspection | Odor | To match standard (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | 7A4OP601 | Tegosoft LSE 65 K Soft | Tegosoft Lse 65 K Soft | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Iodine Value |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | 7A4OP601 | Tegosoft LSE 65 K Soft | Tegosoft Lse 65 K Soft | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 828 | Iodine Value | Iodine Value | 2.2 | g I/100g | FDA,2 |  | 5 | J&J TM-003029 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | 7A4OP601 | Tegosoft LSE 65 K Soft | Tegosoft Lse 65 K Soft | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 828 | Iodine Value | Iodine Value Final | 2.2 | g I/100g | FDA,2 |  | 5 | J&J TM-003029 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | 7A4OP601 | Tegosoft LSE 65 K Soft | Tegosoft Lse 65 K Soft | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 792 | Chemical Inspection | Appearance | Light yellow pasty fluid (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | 7A4OP601 | Tegosoft LSE 65 K Soft | Tegosoft Lse 65 K Soft | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | 7A4OP601 | Tegosoft LSE 65 K Soft | Tegosoft Lse 65 K Soft | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 390 | Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | 7A4OP601 | Tegosoft LSE 65 K Soft | Tegosoft Lse 65 K Soft | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 398 | pH | pH (5% Solution in DI Water:Methanol 1:1) | 7 |  | FDA,1 | 6.5 | 7.5 | J&J TMWW78014 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | 7A4OP601 | Tegosoft LSE 65 K Soft | Tegosoft Lse 65 K Soft | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 810 | Saponification Value | Saponification Value | 54.2 | KOH/g | FDA,1 | 47 | 57 | J&J TMWW78016 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | 7A4OP601 | Tegosoft LSE 65 K Soft | Tegosoft Lse 65 K Soft | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 810 | Saponification Value | Saponification Value Final | 53.35 | KOH/g | FDA,1 | 47 | 57 | J&J TMWW78016 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | 7A4OP601 | Tegosoft LSE 65 K Soft | Tegosoft Lse 65 K Soft | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 801 | % Water | Water Content | 35.15 | % | FDA,2 | 33 | 37 | J&J TMWW78017 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP806/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP806 | Actipone Lamina | ACt. ipone Lamina | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP806/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP806 | Actipone Lamina | ACt. ipone Lamina | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 432 | Specific Gravity at 20�C | Specific Gravity at 20°C | 1.1155 |  | FDA,4 | 1.11 | 1.12 | J&J TMWW78012 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP806/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP806 | Actipone Lamina | ACt. ipone Lamina | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | 1.391 |  | FDA,4 | 1.386 | 1.396 | J&J TMWW78015 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP806/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP806 | Actipone Lamina | ACt. ipone Lamina | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 825 | Limit of Ethylene Glycol and Diethylene Glycol | Diethylene Glycol | 0 | % | FDA,2 |  | 0.1 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OP806/MASTER - Microbiology - J&J Chemical Testing) | 7A4OP806 | Actipone Lamina | ACt. ipone Lamina | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 825 | Limit of Ethylene Glycol and Diethylene Glycol | Ethylene Glycol | 0 | % | FDA,2 |  | 0.1 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ028/Chem_NY_Emulgin VL-75 | 7A4OQ028 | Emulgin VL-75 | Emulgin VL-75 | Chem_NY_Emulgin VL-75 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH as is |  | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7A4OQ028/Chem_NY_Emulgin VL-75 | 7A4OQ028 | Emulgin VL-75 | Emulgin VL-75 | Chem_NY_Emulgin VL-75 | NY-Analytical | 447 | pH as is | pH | 5.76 |  | FDA,1 | 5 | 6 | TM 200 | TM200 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7A4OQ028/Chem_NY_Emulgin VL-75 | 7A4OQ028 | Emulgin VL-75 | Emulgin VL-75 | Chem_NY_Emulgin VL-75 | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7A4OQ028/Chem_NY_Emulgin VL-75 | 7A4OQ028 | Emulgin VL-75 | Emulgin VL-75 | Chem_NY_Emulgin VL-75 | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear to slightly Turbid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7A4OQ028/Chem_NY_Emulgin VL-75 | 7A4OQ028 | Emulgin VL-75 | Emulgin VL-75 | Chem_NY_Emulgin VL-75 | NY-Analytical | 792 | Chemical Inspection | Color | Yellowish Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7A4OQ028/Chem_NY_Emulgin VL-75 | 7A4OQ028 | Emulgin VL-75 | Emulgin VL-75 | Chem_NY_Emulgin VL-75 | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7A4OQ028/Chem_NY_Emulgin VL-75 | 7A4OQ028 | Emulgin VL-75 | Emulgin VL-75 | Chem_NY_Emulgin VL-75 | NY-Analytical | 792 | Chemical Inspection | Odor | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7A4OQ084/Chem_23A | 7A4OQ084 | Versene 100 | Versene 100 | Chem_23A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ084/Chem_23A | 7A4OQ084 | Versene 100 | Versene 100 | Chem_23A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ084/Chem_23A | 7A4OQ084 | Versene 100 | Versene 100 | Chem_23A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ084/Chem_23A | 7A4OQ084 | Versene 100 | Versene 100 | Chem_23A | NY-Analytical | 792 | Chemical Inspection | Color | None or Pale Straw (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ084/Chem_23A | 7A4OQ084 | Versene 100 | Versene 100 | Chem_23A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ084/Chem_23A | 7A4OQ084 | Versene 100 | Versene 100 | Chem_23A | NY-Analytical | 792 | Chemical Inspection | Odor | Little or no distinguisable odor (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
| 7A4OQ132 | Triton X-100 |  | Analytical - Chem_25A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  |  |  |  | left_only
| 7A4OQ132 | Triton X-100 |  | Analytical - Chem_25A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 |  |  |  | left_only
| 7A4OQ132 | Triton X-100 |  | Analytical - Chem_25A | NY-Analytical | 792 | Chemical Inspection | Appearance | Transparent liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7A4OQ132 | Triton X-100 |  | Analytical - Chem_25A | NY-Analytical | 792 | Chemical Inspection | Color | Colorless to pale yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7A4OQ132 | Triton X-100 |  | Analytical - Chem_25A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7A4OQ132 | Triton X-100 |  | Analytical - Chem_25A | NY-Analytical | 792 | Chemical Inspection | Odor | Mild (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
7A4OQ203/MASTER - Microbiology - J&J Chemical Testing) | 7A4OQ203 | Carbopol Ultrez-10 | Carbopol Ultrez-10 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 802 | Loss on Drying | Loss on Drying | 0.75 | % | FDA,2 |  | 2 | J&J TM7221 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ203/MASTER - Microbiology - J&J Chemical Testing) | 7A4OQ203 | Carbopol Ultrez-10 | Carbopol Ultrez-10 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 802 | Loss on Drying | Loss on Drying Final | 0.75 | % | FDA,2 |  | 2 | J&J TM7221 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ203/MASTER - Microbiology - J&J Chemical Testing) | 7A4OQ203 | Carbopol Ultrez-10 | Carbopol Ultrez-10 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 747 | Viscosity | Viscosity (Spindle 7@20RPM@30 sec 0.5% solution) | 60667 | cps | FDA,0 | 45000 | 65000 | J&J TM7423 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ229/Chem_Trisodium Citrate Dihydrate USP | 7A4OQ229 | Trisodium Citrate Dihydrate, USP | Trisodium Citrate Dihydrate, USP | Chem_Trisodium Citrate Dihydrate USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA | COA | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ229/Chem_NY_Witch Hazel | 7A4OQ229 | Trisodium Citrate Dihydrate, USP | Trisodium Citrate Dihydrate, USP | Chem_NY_Witch Hazel | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25�C |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ229/Chem_NY_Witch Hazel | 7A4OQ229 | Trisodium Citrate Dihydrate, USP | Trisodium Citrate Dihydrate, USP | Chem_NY_Witch Hazel | NY-Analytical | 402 | Specific Gravity at 25�C | Specific Gravity at 25°C | 0.981 |  | FDA,3 | 0.979 | 0.983 | TM 202 | TM202 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ229/Chem_Trisodium Citrate Dihydrate USP | 7A4OQ229 | Trisodium Citrate Dihydrate, USP | Trisodium Citrate Dihydrate, USP | Chem_Trisodium Citrate Dihydrate USP | NY-Analytical | 792 | Chemical Inspection | Appearance | Fine granular (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ229/Chem_Trisodium Citrate Dihydrate USP | 7A4OQ229 | Trisodium Citrate Dihydrate, USP | Trisodium Citrate Dihydrate, USP | Chem_Trisodium Citrate Dihydrate USP | NY-Analytical | 792 | Chemical Inspection | Color | White (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ229/Chem_Trisodium Citrate Dihydrate USP | 7A4OQ229 | Trisodium Citrate Dihydrate, USP | Trisodium Citrate Dihydrate, USP | Chem_Trisodium Citrate Dihydrate USP | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ229/Chem_Trisodium Citrate Dihydrate USP | 7A4OQ229 | Trisodium Citrate Dihydrate, USP | Trisodium Citrate Dihydrate, USP | Chem_Trisodium Citrate Dihydrate USP | NY-Analytical | 792 | Chemical Inspection | Odor | Odorless (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ229/Chem_Trisodium Citrate Dihydrate USP | 7A4OQ229 | Trisodium Citrate Dihydrate, USP | Trisodium Citrate Dihydrate, USP | Chem_Trisodium Citrate Dihydrate USP | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ229/Chem_NY_Witch Hazel | 7A4OQ229 | Trisodium Citrate Dihydrate, USP | Trisodium Citrate Dihydrate, USP | Chem_NY_Witch Hazel | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ229/Chem_NY_Witch Hazel | 7A4OQ229 | Trisodium Citrate Dihydrate, USP | Trisodium Citrate Dihydrate, USP | Chem_NY_Witch Hazel | NY-Analytical | 792 | Chemical Inspection | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ229/Chem_NY_Witch Hazel | 7A4OQ229 | Trisodium Citrate Dihydrate, USP | Trisodium Citrate Dihydrate, USP | Chem_NY_Witch Hazel | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ229/Chem_NY_Witch Hazel | 7A4OQ229 | Trisodium Citrate Dihydrate, USP | Trisodium Citrate Dihydrate, USP | Chem_NY_Witch Hazel | NY-Analytical | 792 | Chemical Inspection | Odor | Characteristic witch hazel, free from acetous odor (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ229/Chem_NY_Witch Hazel | 7A4OQ229 | Trisodium Citrate Dihydrate, USP | Trisodium Citrate Dihydrate, USP | Chem_NY_Witch Hazel | NY-Analytical | 797 | Nonvolatile Residue | End Weight | 240.3127 | g | FDA,0 |  |  | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ229/Chem_NY_Witch Hazel | 7A4OQ229 | Trisodium Citrate Dihydrate, USP | Trisodium Citrate Dihydrate, USP | Chem_NY_Witch Hazel | NY-Analytical | 797 | Nonvolatile Residue | Start Weight | 240.3107 | g | FDA,0 |  |  | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ229/Chem_NY_Witch Hazel | 7A4OQ229 | Trisodium Citrate Dihydrate, USP | Trisodium Citrate Dihydrate, USP | Chem_NY_Witch Hazel | NY-Analytical | 797 | Nonvolatile Residue | Total non-volatile reside | 2 | mg/100 mL | FDA,0 |  | 25 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ229/Chem_NY_Witch Hazel | 7A4OQ229 | Trisodium Citrate Dihydrate, USP | Trisodium Citrate Dihydrate, USP | Chem_NY_Witch Hazel | NY-Analytical | 447 | pH as is | pH | 3.64 |  | FDA,1 | 3 | 5 | USP <791> | USP_791 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ260/Chem_29A | 7A4OQ260 | Adult Wipe Fragrance | Adult Wipe Fragrance | Chem_29A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25�C |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ260/Chem_29A | 7A4OQ260 | Adult Wipe Fragrance | Adult Wipe Fragrance | Chem_29A | NY-Analytical | 402 | Specific Gravity at 25�C | Specific Gravity at 25°C | 0.9659 |  | FDA,3 | 0.954 | 0.974 | TM 202 | TM202 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ260/Chem_29A | 7A4OQ260 | Adult Wipe Fragrance | Adult Wipe Fragrance | Chem_29A | NY-Analytical | 451 | Refractive Index at 25�C | Refractive Index | 1.48 |  | FDA,3 | 1.464 | 1.484 | TM 204 | TM204 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ260/Chem_29A | 7A4OQ260 | Adult Wipe Fragrance | Adult Wipe Fragrance | Chem_29A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ260/Chem_29A | 7A4OQ260 | Adult Wipe Fragrance | Adult Wipe Fragrance | Chem_29A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ260/Chem_29A | 7A4OQ260 | Adult Wipe Fragrance | Adult Wipe Fragrance | Chem_29A | NY-Analytical | 792 | Chemical Inspection | Color | Colorless to pale yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ260/Chem_29A | 7A4OQ260 | Adult Wipe Fragrance | Adult Wipe Fragrance | Chem_29A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ260/Chem_29A | 7A4OQ260 | Adult Wipe Fragrance | Adult Wipe Fragrance | Chem_29A | NY-Analytical | 792 | Chemical Inspection | Odor | Powder Rose (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ518/Chem_31A | 7A4OQ518 | BTC 2125M-50% Maquat MQ2525M-50% Barquat 4250Z | BTC 2125M-50% Maquat Mq2525M-50% Barquat 4250Z | Chem_31A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ518/Chem_31A | 7A4OQ518 | BTC 2125M-50% Maquat MQ2525M-50% Barquat 4250Z | BTC 2125M-50% Maquat Mq2525M-50% Barquat 4250Z | Chem_31A | NY-Analytical | 398 | pH | pH (10% Solution) | 7.57 |  | FDA,1 | 6 | 9 | TM 200 | TM200 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ518/Chem_31A | 7A4OQ518 | BTC 2125M-50% Maquat MQ2525M-50% Barquat 4250Z | BTC 2125M-50% Maquat Mq2525M-50% Barquat 4250Z | Chem_31A | NY-Analytical | 796 | Assay | Assay | 51.97 | % | FDA,1 | 50 | 54 | TM 301/434 | TM301_TM434 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ518/Chem_31A | 7A4OQ518 | BTC 2125M-50% Maquat MQ2525M-50% Barquat 4250Z | BTC 2125M-50% Maquat Mq2525M-50% Barquat 4250Z | Chem_31A | NY-Analytical | 796 | Assay | Assay Final | 51.98 | % | FDA,1 | 50 | 54 | TM 301/434 | TM301_TM434 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ518/Chem_31A | 7A4OQ518 | BTC 2125M-50% Maquat MQ2525M-50% Barquat 4250Z | BTC 2125M-50% Maquat Mq2525M-50% Barquat 4250Z | Chem_31A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear viscous liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ518/Chem_31A | 7A4OQ518 | BTC 2125M-50% Maquat MQ2525M-50% Barquat 4250Z | BTC 2125M-50% Maquat Mq2525M-50% Barquat 4250Z | Chem_31A | NY-Analytical | 792 | Chemical Inspection | Color | Colorless to light straw (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ518/Chem_31A | 7A4OQ518 | BTC 2125M-50% Maquat MQ2525M-50% Barquat 4250Z | BTC 2125M-50% Maquat Mq2525M-50% Barquat 4250Z | Chem_31A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ518/Chem_31A | 7A4OQ518 | BTC 2125M-50% Maquat MQ2525M-50% Barquat 4250Z | BTC 2125M-50% Maquat Mq2525M-50% Barquat 4250Z | Chem_31A | NY-Analytical | 792 | Chemical Inspection | Odor | Pleasant almond (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ890/Chem_34A | 7A4OQ890 | Sodium Metasilicate Pentahydrate | Sodium Metasilicate Pentahydrate | Chem_34A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ890/Chem_34A | 7A4OQ890 | Sodium Metasilicate Pentahydrate | Sodium Metasilicate Pentahydrate | Chem_34A | NY-Analytical | 792 | Chemical Inspection | Appearance | Granular Powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ890/Chem_34A | 7A4OQ890 | Sodium Metasilicate Pentahydrate | Sodium Metasilicate Pentahydrate | Chem_34A | NY-Analytical | 792 | Chemical Inspection | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ890/Chem_34A | 7A4OQ890 | Sodium Metasilicate Pentahydrate | Sodium Metasilicate Pentahydrate | Chem_34A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OQ890/Chem_34A | 7A4OQ890 | Sodium Metasilicate Pentahydrate | Sodium Metasilicate Pentahydrate | Chem_34A | NY-Analytical | 792 | Chemical Inspection | Odor | Odorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Solumed Isopropyl Alcohol 99% USP | Solumed Isopropyl Alcohol 99% USP | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Solumed Isopropyl Alcohol 99% USP | Solumed Isopropyl Alcohol 99% USP | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 792 | Chemical Inspection | Description | Transparent, colorless, mobile, volatile liquid (Done) |  |  |  |  | ORGANOLEPTIC | TM201 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Solumed Isopropyl Alcohol 99% USP | Solumed Isopropyl Alcohol 99% USP | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 795 | Acidity | Acidity | 0.4 | mL | FDA,2 |  | 0.7 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Solumed Isopropyl Alcohol 99% USP | Solumed Isopropyl Alcohol 99% USP | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 795 | Acidity | Acidity Final (Calc) | 0.45 | mL | FDA,2 |  | 0.7 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Solumed Isopropyl Alcohol 99% USP | Solumed Isopropyl Alcohol 99% USP | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 796 | Assay | Assay | 100 | % | FDA,1 | 99 |  | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Solumed Isopropyl Alcohol 99% USP | Solumed Isopropyl Alcohol 99% USP | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 796 | Assay | Assay Final | 99.975 | % | FDA,1 | 99 |  | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Solumed Isopropyl Alcohol 99% USP | Solumed Isopropyl Alcohol 99% USP | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 797 | Nonvolatile Residue | End Weight | 90.0835 | g | FDA,4 |  |  | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Solumed Isopropyl Alcohol 99% USP | Solumed Isopropyl Alcohol 99% USP | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 797 | Nonvolatile Residue | Nonvolatile Residue | 0 | % | FDA,3 |  | 0.005 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Solumed Isopropyl Alcohol 99% USP | Solumed Isopropyl Alcohol 99% USP | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 797 | Nonvolatile Residue | Nonvolatile Residue Final | 0.0008 | % | FDA,3 |  | 0.005 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Solumed Isopropyl Alcohol 99% USP | Solumed Isopropyl Alcohol 99% USP | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | 1.376 |  | FDA,3 | 1.376 | 1.378 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Solumed Isopropyl Alcohol 99% USP | Solumed Isopropyl Alcohol 99% USP | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 797 | Nonvolatile Residue | Start Weight | 90.7647 | g | FDA,4 |  |  | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Solumed Isopropyl Alcohol 99% USP | Solumed Isopropyl Alcohol 99% USP | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 795 | Acidity | Titrant End Volume | 0.9 | mL | FDA,2 |  |  | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Solumed Isopropyl Alcohol 99% USP | Solumed Isopropyl Alcohol 99% USP | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 795 | Acidity | Titrant Start Volume | 0.5 | mL | FDA,2 |  |  | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Solumed Isopropyl Alcohol 99% USP | Solumed Isopropyl Alcohol 99% USP | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 797 | Nonvolatile Residue | Weight of Residue | 0.0004 | g | FDA,4 |  |  | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Solumed Isopropyl Alcohol 99% USP | Solumed Isopropyl Alcohol 99% USP | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 390 | Identification | IR Identification | IR conform to standard (Done) |  |  |  |  | USP <197F> | USP_197F | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | 7A4OS407 | Solumed Isopropyl Alcohol 99% USP | Solumed Isopropyl Alcohol 99% USP | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 402 | Specific Gravity at 25�C | Specific Gravity at 25°C | 0.7832 |  | FDA,3 | 0.783 | 0.787 | USP/NF <841> | USP_NF_841 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
| 7A4OS411 | Povidone Iodine Powder USP/NF |  | Analytical - Chem_36A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  |  |  |  | left_only
| 7A4OS411 | Povidone Iodine Powder USP/NF |  | Analytical - Chem_36A | NY-Analytical | 792 | Chemical Inspection | Appearance | Free Flowing Powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7A4OS411 | Povidone Iodine Powder USP/NF |  | Analytical - Chem_36A | NY-Analytical | 792 | Chemical Inspection | Color | Brownish-Red (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7A4OS411 | Povidone Iodine Powder USP/NF |  | Analytical - Chem_36A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7A4OS411 | Povidone Iodine Powder USP/NF |  | Analytical - Chem_36A | NY-Analytical | 792 | Chemical Inspection | Odor | Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7A4OS411 | Povidone Iodine Powder USP/NF |  | Analytical - Chem_36A | NY-Analytical | 796 | Assay | Assay as is (Without Drying) | 10.17 | % | FDA,1 |  |  | USP | USP |  |  |  | left_only
| 7A4OS411 | Povidone Iodine Powder USP/NF |  | Analytical - Chem_36A | NY-Analytical | 796 | Assay | Assay as is Final (Without Drying) | 10.15 | % | FDA,1 |  |  | USP | USP |  |  |  | left_only
| 7A4OS411 | Povidone Iodine Powder USP/NF |  | Analytical - Chem_36A | NY-Analytical | 796 | Assay | Assay on Dried Basis | 10.59 | % | FDA,1 | 9 | 12 | USP | USP |  |  |  | left_only
| 7A4OS411 | Povidone Iodine Powder USP/NF |  | Analytical - Chem_36A | NY-Analytical | 796 | Assay | Assay on Dried Basis Final | 10.54 | % | FDA,1 | 9 | 12 | USP | USP |  |  |  | left_only
| 7A4OS411 | Povidone Iodine Powder USP/NF |  | Analytical - Chem_36A | NY-Analytical | 802 | Loss on Drying | Loss on Drying | 5.55 | % | FDA,1 |  | 8 | USP | USP |  |  |  | left_only
| 7A4OS411 | Povidone Iodine Powder USP/NF |  | Analytical - Chem_36A | NY-Analytical | 802 | Loss on Drying | Loss on Drying Final | 5.55 | % | FDA,1 |  | 8 | USP | USP |  |  |  | left_only
7A4OU351/Chem_NY_Dehydrated Alcohol USP | 7A4OU351 | Dehydrated Alcohol USP | Dehydrated Alcohol USP | Chem_NY_Dehydrated Alcohol USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA | COA | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OU351/Chem_NY_Dehydrated Alcohol USP | 7A4OU351 | Dehydrated Alcohol USP | Dehydrated Alcohol USP | Chem_NY_Dehydrated Alcohol USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | COA Meets Specification (Done) |  |  |  |  | COA | COA | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OU351/Chem_NY_Dehydrated Alcohol USP | 7A4OU351 | Dehydrated Alcohol USP | Dehydrated Alcohol USP | Chem_NY_Dehydrated Alcohol USP | NY-Analytical | 822 | Specific Gravity @ 15.56°C/ 60° F | Specific Gravity at 15.56°C | 0.794 |  | FDA,4 |  | 0.7962 | TM 202 | TM202 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OU351/Chem_NY_Dehydrated Alcohol USP | 7A4OU351 | Dehydrated Alcohol USP | Dehydrated Alcohol USP | Chem_NY_Dehydrated Alcohol USP | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear Liquid (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OU351/Chem_NY_Dehydrated Alcohol USP | 7A4OU351 | Dehydrated Alcohol USP | Dehydrated Alcohol USP | Chem_NY_Dehydrated Alcohol USP | NY-Analytical | 792 | Chemical Inspection | Color | Clear to Colorless and Match standard (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OU351/Chem_NY_Dehydrated Alcohol USP | 7A4OU351 | Dehydrated Alcohol USP | Dehydrated Alcohol USP | Chem_NY_Dehydrated Alcohol USP | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OU351/Chem_NY_Dehydrated Alcohol USP | 7A4OU351 | Dehydrated Alcohol USP | Dehydrated Alcohol USP | Chem_NY_Dehydrated Alcohol USP | NY-Analytical | 792 | Chemical Inspection | Odor | Alcohol (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OU351/Chem_NY_Dehydrated Alcohol USP | 7A4OU351 | Dehydrated Alcohol USP | Dehydrated Alcohol USP | Chem_NY_Dehydrated Alcohol USP | NY-Analytical | 390 | Identification | IR Identification | Compare to reference Standard (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OU555/Chem_NY_Phenylephrine Reduced | 7A4OU555 | Phenylephrine HCL, USP | Phenylephrine Hcl, USP | Chem_NY_Phenylephrine Reduced | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis | COA | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Phenylephrine HCL, USP | Phenylephrine Hcl, USP | Chem_NY_Phenylephrine HCL | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OU555/Chem_NY_Phenylephrine Reduced | 7A4OU555 | Phenylephrine HCL, USP | Phenylephrine Hcl, USP | Chem_NY_Phenylephrine Reduced | NY-Analytical | 792 | Chemical Inspection | Appearance | Crystalline Powder or Crystals (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Phenylephrine HCL, USP | Phenylephrine Hcl, USP | Chem_NY_Phenylephrine HCL | NY-Analytical | 792 | Chemical Inspection | Appearance | Crystalline Powder or Crystals (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OU555/Chem_NY_Phenylephrine Reduced | 7A4OU555 | Phenylephrine HCL, USP | Phenylephrine Hcl, USP | Chem_NY_Phenylephrine Reduced | NY-Analytical | 792 | Chemical Inspection | Color | White to almost white (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Phenylephrine HCL, USP | Phenylephrine Hcl, USP | Chem_NY_Phenylephrine HCL | NY-Analytical | 792 | Chemical Inspection | Color | White to almost white (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OU555/Chem_NY_Phenylephrine Reduced | 7A4OU555 | Phenylephrine HCL, USP | Phenylephrine Hcl, USP | Chem_NY_Phenylephrine Reduced | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Phenylephrine HCL, USP | Phenylephrine Hcl, USP | Chem_NY_Phenylephrine HCL | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OU555/Chem_NY_Phenylephrine Reduced | 7A4OU555 | Phenylephrine HCL, USP | Phenylephrine Hcl, USP | Chem_NY_Phenylephrine Reduced | NY-Analytical | 792 | Chemical Inspection | Odor | Match standard (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Phenylephrine HCL, USP | Phenylephrine Hcl, USP | Chem_NY_Phenylephrine HCL | NY-Analytical | 792 | Chemical Inspection | Odor | Match standard (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Phenylephrine HCL, USP | Phenylephrine Hcl, USP | Chem_NY_Phenylephrine HCL | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OU555/Chem_NY_Phenylephrine Reduced | 7A4OU555 | Phenylephrine HCL, USP | Phenylephrine Hcl, USP | Chem_NY_Phenylephrine Reduced | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Phenylephrine HCL, USP | Phenylephrine Hcl, USP | Chem_NY_Phenylephrine HCL | NY-Analytical | 812 | Organic Impurites | Any individual unspecified impuritv |  | % | FDA,2 |  | 0.1 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Phenylephrine HCL, USP | Phenylephrine Hcl, USP | Chem_NY_Phenylephrine HCL | NY-Analytical | 796 | Assay | Assay | 99.4 | % | FDA,1 | 98 | 102 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OU555/Chem_NY_Phenylephrine Reduced | 7A4OU555 | Phenylephrine HCL, USP | Phenylephrine Hcl, USP | Chem_NY_Phenylephrine Reduced | NY-Analytical | 796 | Assay | Assay | 99.9 | % | FDA,1 | 98 | 102 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Phenylephrine HCL, USP | Phenylephrine Hcl, USP | Chem_NY_Phenylephrine HCL | NY-Analytical | 802 | Loss on Drying | Loss on Drying |  | % | FDA,1 |  | 1 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Phenylephrine HCL, USP | Phenylephrine Hcl, USP | Chem_NY_Phenylephrine HCL | NY-Analytical | 802 | Loss on Drying | Loss on Drying Final |  | % | FDA,1 |  | 1 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Phenylephrine HCL, USP | Phenylephrine Hcl, USP | Chem_NY_Phenylephrine HCL | NY-Analytical | 812 | Organic Impurites | Norphenylephrine |  | % | FDA,2 |  | 0.1 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Phenylephrine HCL, USP | Phenylephrine Hcl, USP | Chem_NY_Phenylephrine HCL | NY-Analytical | 1043 | Optical Rotation | Optical Rotation |  | ° | FDA,0 | -47 | -43 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Phenylephrine HCL, USP | Phenylephrine Hcl, USP | Chem_NY_Phenylephrine HCL | NY-Analytical | 812 | Organic Impurites | Phenylephrine Related Compound C |  | % | FDA,2 |  | 0.1 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Phenylephrine HCL, USP | Phenylephrine Hcl, USP | Chem_NY_Phenylephrine HCL | NY-Analytical | 812 | Organic Impurites | Phenylephrine Related Compound D |  | % | FDA,2 |  | 0.1 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Phenylephrine HCL, USP | Phenylephrine Hcl, USP | Chem_NY_Phenylephrine HCL | NY-Analytical | 812 | Organic Impurites | Phenylephrine Related Compound E |  | % | FDA,2 |  | 0.1 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Phenylephrine HCL, USP | Phenylephrine Hcl, USP | Chem_NY_Phenylephrine HCL | NY-Analytical | 834 | Residue on Ignition | Residue on Ignition |  | % | FDA,1 |  | 0.2 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Phenylephrine HCL, USP | Phenylephrine Hcl, USP | Chem_NY_Phenylephrine HCL | NY-Analytical | 835 | Sulfate Content | Sulfate |  | % | FDA,2 |  | 0.2 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OU555/Chem_NY_Phenylephrine HCL | 7A4OU555 | Phenylephrine HCL, USP | Phenylephrine Hcl, USP | Chem_NY_Phenylephrine HCL | NY-Analytical | 812 | Organic Impurites | Total Impurities |  | % | FDA,2 |  | 0.2 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
| 7A4OX110 | Polysorbate 20 |  | Analytical - Chem_39A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  |  |  |  | left_only
| 7A4OX110 | Polysorbate 20 |  | Analytical - Chem_39A | NY-Analytical | 390 | Identification | IR Identification | [5 empty] | TM 411 | TM411 |  |  |  | left_only
| 7A4OX110 | Polysorbate 20 |  | Analytical - Chem_39A | NY-Analytical | 792 | Chemical Inspection | Appearance | [5 empty] | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7A4OX110 | Polysorbate 20 |  | Analytical - Chem_39A | NY-Analytical | 792 | Chemical Inspection | Color | [5 empty] | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7A4OX110 | Polysorbate 20 |  | Analytical - Chem_39A | NY-Analytical | 792 | Chemical Inspection | Contamination | [5 empty] | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7A4OX110 | Polysorbate 20 |  | Analytical - Chem_39A | NY-Analytical | 792 | Chemical Inspection | Odor | [5 empty] | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
7A4OX231/Chem_NY_D&C Yellow 10 | 7A4OX231 | D&C Yellow 10 (D&C Grade) | D&C Yellow 10 (D&C Grade) | Chem_NY_D&C Yellow 10 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Appearance |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OX231/Chem_NY_D&C Yellow 10 | 7A4OX231 | D&C Yellow 10 (D&C Grade) | D&C Yellow 10 (D&C Grade) | Chem_NY_D&C Yellow 10 | NY-Analytical | 397 | Appearance | Physical Properties | Yellow Powder (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OX231/Chem_NY_D&C Yellow 10 | 7A4OX231 | D&C Yellow 10 (D&C Grade) | D&C Yellow 10 (D&C Grade) | Chem_NY_D&C Yellow 10 | NY-Analytical | 390 | Identification | Spectrophotometric | 410.88 | λ | FDA,0 | 408 | 416 | TM 3014 | TM3014 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OX231/Chem_NY_D&C Yellow 10 | 7A4OX231 | D&C Yellow 10 (D&C Grade) | D&C Yellow 10 (D&C Grade) | Chem_NY_D&C Yellow 10 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram) | 0 | CFU/g | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OX231/Chem_NY_D&C Yellow 10 | 7A4OX231 | D&C Yellow 10 (D&C Grade) | D&C Yellow 10 (D&C Grade) | Chem_NY_D&C Yellow 10 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OX231/Chem_NY_D&C Yellow 10 | 7A4OX231 | D&C Yellow 10 (D&C Grade) | D&C Yellow 10 (D&C Grade) | Chem_NY_D&C Yellow 10 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram) | 0 | CFU/g | FDA,1 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7A4OX231/Chem_NY_D&C Yellow 10 | 7A4OX231 | D&C Yellow 10 (D&C Grade) | D&C Yellow 10 (D&C Grade) | Chem_NY_D&C Yellow 10 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7ACHE61P/Chem_NY_Sodium Benzoate | 7ACHE61P | Sodium Benzoate EDF NF/FCC | Sodium Benzoate EDF NF/FCC | Chem_NY_Sodium Benzoate | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7ACHE61P/Chem_NY_Sodium Benzoate | 7ACHE61P | Sodium Benzoate EDF NF/FCC | Sodium Benzoate EDF NF/FCC | Chem_NY_Sodium Benzoate | NY-Analytical | 792 | Chemical Inspection | Appearance | Extruded Pellets (Done) |  |  |  |  | TM 208 | TM208 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7ACHE61P/Chem_NY_Sodium Benzoate | 7ACHE61P | Sodium Benzoate EDF NF/FCC | Sodium Benzoate EDF NF/FCC | Chem_NY_Sodium Benzoate | NY-Analytical | 792 | Chemical Inspection | Color | White (Done) |  |  |  |  | TM 208 | TM208 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7ACHE61P/Chem_NY_Sodium Benzoate | 7ACHE61P | Sodium Benzoate EDF NF/FCC | Sodium Benzoate EDF NF/FCC | Chem_NY_Sodium Benzoate | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7ACHE61P/Chem_NY_Sodium Benzoate | 7ACHE61P | Sodium Benzoate EDF NF/FCC | Sodium Benzoate EDF NF/FCC | Chem_NY_Sodium Benzoate | NY-Analytical | 792 | Chemical Inspection | Odor | Odorless (Done) |  |  |  |  | TM 208 | TM208 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7ACHE61P/Chem_NY_Sodium Benzoate | 7ACHE61P | Sodium Benzoate EDF NF/FCC | Sodium Benzoate EDF NF/FCC | Chem_NY_Sodium Benzoate | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7B4AA580/Chem_NY_Acetic Acid - USP | 7B4AA580 | Acetic Acid - USP | Acetic Acid - USP | Chem_NY_Acetic Acid - USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4AA580/Chem_NY_Acetic Acid - USP | 7B4AA580 | Acetic Acid - USP | Acetic Acid - USP | Chem_NY_Acetic Acid - USP | NY-Analytical | 792 | Chemical Inspection | Appearance | Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4AA580/Chem_NY_Acetic Acid - USP | 7B4AA580 | Acetic Acid - USP | Acetic Acid - USP | Chem_NY_Acetic Acid - USP | NY-Analytical | 792 | Chemical Inspection | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4AA580/Chem_NY_Acetic Acid - USP | 7B4AA580 | Acetic Acid - USP | Acetic Acid - USP | Chem_NY_Acetic Acid - USP | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4AA580/Chem_NY_Acetic Acid - USP | 7B4AA580 | Acetic Acid - USP | Acetic Acid - USP | Chem_NY_Acetic Acid - USP | NY-Analytical | 792 | Chemical Inspection | Odor | Pungent (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4AL930/Chem_42A | 7B4AL930 | Gluconic Acid 50% Technical | Gluconic Acid 50% Technical | Chem_42A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH as is |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4AL930/Chem_42A | 7B4AL930 | Gluconic Acid 50% Technical | Gluconic Acid 50% Technical | Chem_42A | NY-Analytical | 447 | pH as is | pH |  |  | FDA,1 | 1.1 | 1.5 | TM 200 | TM200 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4AL930/Chem_42A | 7B4AL930 | Gluconic Acid 50% Technical | Gluconic Acid 50% Technical | Chem_42A | NY-Analytical | 402 | Specific Gravity at 25�C | Specific Gravity at 25°C | 1.2444 |  | FDA,2 | 1.21 | 1.27 | TM 202 | TM202 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4AL930/Chem_42A | 7B4AL930 | Gluconic Acid 50% Technical | Gluconic Acid 50% Technical | Chem_42A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4AL930/Chem_42A | 7B4AL930 | Gluconic Acid 50% Technical | Gluconic Acid 50% Technical | Chem_42A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4AL930/Chem_42A | 7B4AL930 | Gluconic Acid 50% Technical | Gluconic Acid 50% Technical | Chem_42A | NY-Analytical | 792 | Chemical Inspection | Color | Light Yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4AL930/Chem_42A | 7B4AL930 | Gluconic Acid 50% Technical | Gluconic Acid 50% Technical | Chem_42A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4AL930/Chem_42A | 7B4AL930 | Gluconic Acid 50% Technical | Gluconic Acid 50% Technical | Chem_42A | NY-Analytical | 792 | Chemical Inspection | Odor | Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4FD110/Chem_44A | 7B4FD110 | Boric Acid USP/NF | Boric Acid USP/Nf | Chem_44A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4FD110/Chem_44A | 7B4FD110 | Boric Acid USP/NF | Boric Acid USP/Nf | Chem_44A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4FD110/Chem_44A | 7B4FD110 | Boric Acid USP/NF | Boric Acid USP/Nf | Chem_44A | NY-Analytical | 792 | Chemical Inspection | Appearance | White granular powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4FD110/Chem_44A | 7B4FD110 | Boric Acid USP/NF | Boric Acid USP/Nf | Chem_44A | NY-Analytical | 792 | Chemical Inspection | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4FD110/Chem_44A | 7B4FD110 | Boric Acid USP/NF | Boric Acid USP/Nf | Chem_44A | NY-Analytical | 792 | Chemical Inspection | Odor | Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4GB598/Chem_45A | 7B4GB598 | Pumice | Pumice | Chem_45A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4GB598/Chem_45A | 7B4GB598 | Pumice | Pumice | Chem_45A | NY-Analytical | 792 | Chemical Inspection | Appearance | Grain Powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4GB598/Chem_45A | 7B4GB598 | Pumice | Pumice | Chem_45A | NY-Analytical | 792 | Chemical Inspection | Color | Off-White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4GB598/Chem_45A | 7B4GB598 | Pumice | Pumice | Chem_45A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4GB598/Chem_45A | 7B4GB598 | Pumice | Pumice | Chem_45A | NY-Analytical | 792 | Chemical Inspection | Odor | Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OA242/Chem_48A | 7B4OA242 | Citric Acid USP Anhydrous Granular | Citric Acid USP Anhydrous Granular | Chem_48A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7B4OA242/Chem_48A | 7B4OA242 | Citric Acid USP Anhydrous Granular | Citric Acid USP Anhydrous Granular | Chem_48A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7B4OA242/Chem_48A | 7B4OA242 | Citric Acid USP Anhydrous Granular | Citric Acid USP Anhydrous Granular | Chem_48A | NY-Analytical | 792 | Chemical Inspection | Appearance | Crystals/crystalline powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7B4OA242/Chem_48A | 7B4OA242 | Citric Acid USP Anhydrous Granular | Citric Acid USP Anhydrous Granular | Chem_48A | NY-Analytical | 792 | Chemical Inspection | Color | Colorless to white (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7B4OA242/Chem_48A | 7B4OA242 | Citric Acid USP Anhydrous Granular | Citric Acid USP Anhydrous Granular | Chem_48A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7B4OA242/Chem_48A | 7B4OA242 | Citric Acid USP Anhydrous Granular | Citric Acid USP Anhydrous Granular | Chem_48A | NY-Analytical | 792 | Chemical Inspection | Odor | Odorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | 7B4OA347 | Calming Waters INCI w/Symfresh | Calming Waters Inci W/Symfresh | Chem_NY_Calming Waters INCI w/symfresh | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25�C |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | 7B4OA347 | Calming Waters INCI w/Symfresh | Calming Waters Inci W/Symfresh | Chem_NY_Calming Waters INCI w/symfresh | NY-Analytical | 402 | Specific Gravity at 25�C | Specific Gravity at 25°C | 0.964 |  | FDA,3 | 0.959 | 0.969 | TM 202 | TM202 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | 7B4OA347 | Calming Waters INCI w/Symfresh | Calming Waters Inci W/Symfresh | Chem_NY_Calming Waters INCI w/symfresh | NY-Analytical | 451 | Refractive Index at 25�C | Refractive Index | 1.472 |  | FDA,3 | 1.465 | 1.475 | TM 204 | TM204 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | 7B4OA347 | Calming Waters INCI w/Symfresh | Calming Waters Inci W/Symfresh | Chem_NY_Calming Waters INCI w/symfresh | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear Liquid (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | 7B4OA347 | Calming Waters INCI w/Symfresh | Calming Waters Inci W/Symfresh | Chem_NY_Calming Waters INCI w/symfresh | NY-Analytical | 792 | Chemical Inspection | Color | Colorless to pale yellow (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | 7B4OA347 | Calming Waters INCI w/Symfresh | Calming Waters Inci W/Symfresh | Chem_NY_Calming Waters INCI w/symfresh | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | 7B4OA347 | Calming Waters INCI w/Symfresh | Calming Waters Inci W/Symfresh | Chem_NY_Calming Waters INCI w/symfresh | NY-Analytical | 792 | Chemical Inspection | Odor | Match standard (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | 7B4OA347 | Calming Waters INCI w/Symfresh | Calming Waters Inci W/Symfresh | Chem_NY_Calming Waters INCI w/symfresh | NY-Analytical | 390 | Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OA709/Chem_NY_Bell fragrance | 7B4OA709 | Bell Fragrance | Bell Fragrance | Chem_NY_Bell fragrance | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25�C |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OA709/Chem_NY_Bell fragrance | 7B4OA709 | Bell Fragrance | Bell Fragrance | Chem_NY_Bell fragrance | NY-Analytical | 402 | Specific Gravity at 25�C | Specific Gravity at 25°C | 1.0639 |  | FDA,3 | 1.06 | 1.07 | TM 202 | TM202 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OA709/Chem_NY_Bell fragrance | 7B4OA709 | Bell Fragrance | Bell Fragrance | Chem_NY_Bell fragrance | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | 1.468 |  | FDA,3 | 1.467 | 1.477 | TM 204 | TM204 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OA709/Chem_NY_Bell fragrance | 7B4OA709 | Bell Fragrance | Bell Fragrance | Chem_NY_Bell fragrance | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear Liquid (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OA709/Chem_NY_Bell fragrance | 7B4OA709 | Bell Fragrance | Bell Fragrance | Chem_NY_Bell fragrance | NY-Analytical | 792 | Chemical Inspection | Color | Pale to light yellow (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OA709/Chem_NY_Bell fragrance | 7B4OA709 | Bell Fragrance | Bell Fragrance | Chem_NY_Bell fragrance | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OA709/Chem_NY_Bell fragrance | 7B4OA709 | Bell Fragrance | Bell Fragrance | Chem_NY_Bell fragrance | NY-Analytical | 792 | Chemical Inspection | Odor | Conforms to Standard (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OB690/Chem_50A | 7B4OB690 | VM&P Naphtha HT | Vm&P Naphtha Ht | Chem_50A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OB690/Chem_50A | 7B4OB690 | VM&P Naphtha HT | Vm&P Naphtha Ht | Chem_50A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear, mobile liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OB690/Chem_50A | 7B4OB690 | VM&P Naphtha HT | Vm&P Naphtha Ht | Chem_50A | NY-Analytical | 792 | Chemical Inspection | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OB690/Chem_50A | 7B4OB690 | VM&P Naphtha HT | Vm&P Naphtha Ht | Chem_50A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OB690/Chem_50A | 7B4OB690 | VM&P Naphtha HT | Vm&P Naphtha Ht | Chem_50A | NY-Analytical | 792 | Chemical Inspection | Odor | Solvent (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | Antifoam AF-Emulsion | Antifoam Af-Emulsion | CHE ANTIFOAM AF EMULSION | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | Antifoam AF-Emulsion | Antifoam Af-Emulsion | CHE ANTIFOAM AF EMULSION | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | Antifoam AF-Emulsion | Antifoam Af-Emulsion | CHE ANTIFOAM AF EMULSION | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per 10mL) | 0 | cfu/10mL | FDA,1 |  | 0 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | Antifoam AF-Emulsion | Antifoam Af-Emulsion | CHE ANTIFOAM AF EMULSION | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per mL - Undiluted) | 0 | cfu/mL |  |  | 0 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | Antifoam AF-Emulsion | Antifoam Af-Emulsion | CHE ANTIFOAM AF EMULSION | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu | FDA,0 |  | 0 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | Antifoam AF-Emulsion | Antifoam Af-Emulsion | CHE ANTIFOAM AF EMULSION | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per 10mL) | 0 | cfu/10mL | FDA,1 |  | 0 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | Antifoam AF-Emulsion | Antifoam Af-Emulsion | CHE ANTIFOAM AF EMULSION | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per mL - Undiluted) | 0 | cfu/mL |  |  | 0 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | Antifoam AF-Emulsion | Antifoam Af-Emulsion | CHE ANTIFOAM AF EMULSION | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | 0 | cfu | FDA,0 |  | 0 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | Antifoam AF-Emulsion | Antifoam Af-Emulsion | CHE ANTIFOAM AF EMULSION | NY-Analytical | 792 | Chemical Inspection | Appearance | Homogeneous Liquid, Free of Dark Particles (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | Antifoam AF-Emulsion | Antifoam Af-Emulsion | CHE ANTIFOAM AF EMULSION | NY-Analytical | 792 | Chemical Inspection | Color | White to off White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | Antifoam AF-Emulsion | Antifoam Af-Emulsion | CHE ANTIFOAM AF EMULSION | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | Antifoam AF-Emulsion | Antifoam Af-Emulsion | CHE ANTIFOAM AF EMULSION | NY-Analytical | 792 | Chemical Inspection | Odor | Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OC109/CHE ANTIFOAM AF EMULSION | 7B4OC109 | Antifoam AF-Emulsion | Antifoam Af-Emulsion | CHE ANTIFOAM AF EMULSION | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | [5 empty] | WI QAS 076 | WI_QAS_076 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OD232/Chem_52A | 7B4OD232 | Propylene Glycol USP | Propylene Glycol USP | Chem_52A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25�C |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OD232/Chem_52A | 7B4OD232 | Propylene Glycol USP | Propylene Glycol USP | Chem_52A | NY-Analytical | 402 | Specific Gravity at 25�C | Specific Gravity at 25°C | 1.0361 |  | FDA,3 | 1.035 | 1.037 | TM 202 | TM202 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OD232/Chem_52A | 7B4OD232 | Propylene Glycol USP | Propylene Glycol USP | Chem_52A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OD232/Chem_52A | 7B4OD232 | Propylene Glycol USP | Propylene Glycol USP | Chem_52A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear viscous liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OD232/Chem_52A | 7B4OD232 | Propylene Glycol USP | Propylene Glycol USP | Chem_52A | NY-Analytical | 792 | Chemical Inspection | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OD232/Chem_52A | 7B4OD232 | Propylene Glycol USP | Propylene Glycol USP | Chem_52A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OD232/Chem_52A | 7B4OD232 | Propylene Glycol USP | Propylene Glycol USP | Chem_52A | NY-Analytical | 792 | Chemical Inspection | Odor | Odorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | 7B4OM799 | Cucumber Extract in Glycerin NP | Cucumber Extract in Glycerin NP | Chem_NY_Lipotec_Cucumber Extract in Glycerin | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25�C |  | JB | RAW MATERIAL - ARKANSAS | RM_CHEMICAL | both
7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | 7B4OM799 | Cucumber Extract in Glycerin NP | Cucumber Extract in Glycerin NP | Chem_NY_Lipotec_Cucumber Extract in Glycerin | NY-Analytical | 402 | Specific Gravity at 25�C | Specific Gravity at 25°C | 1.2627 |  | FDA,3 | 1.25 | 1.29 | TM 202 | TM202 | JB | RAW MATERIAL - ARKANSAS | RM_CHEMICAL | both
7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | 7B4OM799 | Cucumber Extract in Glycerin NP | Cucumber Extract in Glycerin NP | Chem_NY_Lipotec_Cucumber Extract in Glycerin | NY-Analytical | 451 | Refractive Index at 25�C | Refractive Index | 1.473 |  | FDA,3 | 1.45 | 1.49 | TM 204 | TM204 | JB | RAW MATERIAL - ARKANSAS | RM_CHEMICAL | both
7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | 7B4OM799 | Cucumber Extract in Glycerin NP | Cucumber Extract in Glycerin NP | Chem_NY_Lipotec_Cucumber Extract in Glycerin | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear viscous liquid (Done) |  |  |  |  | TM 208 | TM208 | JB | RAW MATERIAL - ARKANSAS | RM_CHEMICAL | both
7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | 7B4OM799 | Cucumber Extract in Glycerin NP | Cucumber Extract in Glycerin NP | Chem_NY_Lipotec_Cucumber Extract in Glycerin | NY-Analytical | 792 | Chemical Inspection | Color | Light to Medium Yellow (Done) |  |  |  |  | TM 208 | TM208 | JB | RAW MATERIAL - ARKANSAS | RM_CHEMICAL | both
7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | 7B4OM799 | Cucumber Extract in Glycerin NP | Cucumber Extract in Glycerin NP | Chem_NY_Lipotec_Cucumber Extract in Glycerin | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | JB | RAW MATERIAL - ARKANSAS | RM_CHEMICAL | both
7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | 7B4OM799 | Cucumber Extract in Glycerin NP | Cucumber Extract in Glycerin NP | Chem_NY_Lipotec_Cucumber Extract in Glycerin | NY-Analytical | 792 | Chemical Inspection | Odor | Characteristic (Done) |  |  |  |  | TM 208 | TM208 | JB | RAW MATERIAL - ARKANSAS | RM_CHEMICAL | both
7B4OM799/Chem_NY_Lipotec_Cucumber Extract in Glycerin | 7B4OM799 | Cucumber Extract in Glycerin NP | Cucumber Extract in Glycerin NP | Chem_NY_Lipotec_Cucumber Extract in Glycerin | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | JB | RAW MATERIAL - ARKANSAS | RM_CHEMICAL | both
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Abil B88183 PH | Abil B88183 pH | Chem_NY_Abil B88183 PH | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Abil B88183 PH | Abil B88183 pH | Chem_NY_Abil B88183 PH | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear to slightly turbid.Colorless to pale yellow,slightly viscous liquid (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Abil B88183 PH | Abil B88183 pH | Chem_NY_Abil B88183 PH | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Abil B88183 PH | Abil B88183 pH | Chem_NY_Abil B88183 PH | NY-Analytical | 390 | Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Abil B88183 PH | Abil B88183 pH | Chem_NY_Abil B88183 PH | NY-Microbiology | 509 | J&J_Pathogens | Biggy or SDA | No Growth_SDA |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Abil B88183 PH | Abil B88183 pH | Chem_NY_Abil B88183 PH | NY-Microbiology | 509 | J&J_Pathogens | Cetrimide | No Growth |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Abil B88183 PH | Abil B88183 pH | Chem_NY_Abil B88183 PH | NY-Microbiology | 524 | J&J Fungal | Fungal | 0 | CFU/g | FDA,0 | 0 | 200 | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Abil B88183 PH | Abil B88183 pH | Chem_NY_Abil B88183 PH | NY-Microbiology | 509 | J&J_Pathogens | MacConkey | No Growth |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Abil B88183 PH | Abil B88183 pH | Chem_NY_Abil B88183 PH | NY-Microbiology | 509 | J&J_Pathogens | MSA or VJ | No Growth_MSA |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Abil B88183 PH | Abil B88183 pH | Chem_NY_Abil B88183 PH | NY-Microbiology | 700 | J&J SDA Controls | Negative Control | No Growth SDA |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Abil B88183 PH | Abil B88183 pH | Chem_NY_Abil B88183 PH | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | No Growth TSB |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Abil B88183 PH | Abil B88183 pH | Chem_NY_Abil B88183 PH | NY-Microbiology | 509 | J&J_Pathogens | TSA | No Growth |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Abil B88183 PH | Abil B88183 pH | Chem_NY_Abil B88183 PH | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | 0 | cfu/mL | FDA,0 | 0 | 200 | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP027/Chem_NY_Abil B88183 PH | 7B4OP027 | Abil B88183 PH | Abil B88183 pH | Chem_NY_Abil B88183 PH | NY-Microbiology | 517 | J&J_Enrichment | TSB | Not Turbid |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP065/Chem_NY_Chamomile Flower Extract | 7B4OP065 | Chamomile Flower Extract | Chamomile Flower ExtraCt. | Chem_NY_Chamomile Flower Extract | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP065/Chem_NY_Chamomile Flower Extract | 7B4OP065 | Chamomile Flower Extract | Chamomile Flower ExtraCt. | Chem_NY_Chamomile Flower Extract | NY-Analytical | 792 | Chemical Inspection | Appearance | Liquid (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP065/Chem_NY_Chamomile Flower Extract | 7B4OP065 | Chamomile Flower Extract | Chamomile Flower ExtraCt. | Chem_NY_Chamomile Flower Extract | NY-Analytical | 792 | Chemical Inspection | Color | Greenish Yellow to Yellow Orange (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP065/Chem_NY_Chamomile Flower Extract | 7B4OP065 | Chamomile Flower Extract | Chamomile Flower ExtraCt. | Chem_NY_Chamomile Flower Extract | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP065/Chem_NY_Chamomile Flower Extract | 7B4OP065 | Chamomile Flower Extract | Chamomile Flower ExtraCt. | Chem_NY_Chamomile Flower Extract | NY-Analytical | 792 | Chemical Inspection | Odor | Match standard (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP065/Chem_NY_Chamomile Flower Extract | 7B4OP065 | Chamomile Flower Extract | Chamomile Flower ExtraCt. | Chem_NY_Chamomile Flower Extract | NY-Analytical | 390 | Identification | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP065/Chem_NY_Chamomile Flower Extract | 7B4OP065 | Chamomile Flower Extract | Chamomile Flower ExtraCt. | Chem_NY_Chamomile Flower Extract | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP065/Chem_NY_Chamomile Flower Extract | 7B4OP065 | Chamomile Flower Extract | Chamomile Flower ExtraCt. | Chem_NY_Chamomile Flower Extract | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP065/Chem_NY_Chamomile Flower Extract | 7B4OP065 | Chamomile Flower Extract | Chamomile Flower ExtraCt. | Chem_NY_Chamomile Flower Extract | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram) |  | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP065/Chem_NY_Chamomile Flower Extract | 7B4OP065 | Chamomile Flower Extract | Chamomile Flower ExtraCt. | Chem_NY_Chamomile Flower Extract | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) |  | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP065/Chem_NY_Chamomile Flower Extract | 7B4OP065 | Chamomile Flower Extract | Chamomile Flower ExtraCt. | Chem_NY_Chamomile Flower Extract | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | Non-Subculturable Growth (Done) |  |  |  |  | WI QAS 076 | WI_QAS_076 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP093/Chem_NY_PEG-4 | 7B4OP093 | PEG-4 | PEG-4 | Chem_NY_PEG-4 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP093/Chem_NY_PEG-4 | 7B4OP093 | PEG-4 | PEG-4 | Chem_NY_PEG-4 | NY-Analytical | 398 | pH | pH (5% aqueous @25 degree C) | 6.01 |  | FDA,1 | 4.5 | 7.5 | TM 200 | TM200 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP093/Chem_NY_PEG-4 | 7B4OP093 | PEG-4 | PEG-4 | Chem_NY_PEG-4 | NY-Analytical | 792 | Chemical Inspection | Appearance | Not More Than Slightly Hazy (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP093/Chem_NY_PEG-4 | 7B4OP093 | PEG-4 | PEG-4 | Chem_NY_PEG-4 | NY-Analytical | 792 | Chemical Inspection | Color | Colorless (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP093/Chem_NY_PEG-4 | 7B4OP093 | PEG-4 | PEG-4 | Chem_NY_PEG-4 | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP093/Chem_NY_PEG-4 | 7B4OP093 | PEG-4 | PEG-4 | Chem_NY_PEG-4 | NY-Analytical | 792 | Chemical Inspection | Odor | Match standard (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP093/Chem_NY_PEG-4 | 7B4OP093 | PEG-4 | PEG-4 | Chem_NY_PEG-4 | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP113/Chem_NY_Mentha Piperita Extract_01 | 7B4OP113 | Mentha Piperita Extract | Mentha Piperita ExtraCt. | Chem_NY_Mentha Piperita Extract_01 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP113/Chem_NY_Mentha Piperita Extract_01 | 7B4OP113 | Mentha Piperita Extract | Mentha Piperita ExtraCt. | Chem_NY_Mentha Piperita Extract_01 | NY-Analytical | 398 | pH | pH | 4.24 |  | FDA,1 | 4 | 5 | TM 200 | TM200 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP113/Chem_NY_Mentha Piperita Extract_01 | 7B4OP113 | Mentha Piperita Extract | Mentha Piperita ExtraCt. | Chem_NY_Mentha Piperita Extract_01 | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear aqueous solution (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP113/Chem_NY_Mentha Piperita Extract_01 | 7B4OP113 | Mentha Piperita Extract | Mentha Piperita ExtraCt. | Chem_NY_Mentha Piperita Extract_01 | NY-Analytical | 792 | Chemical Inspection | Color | Amber (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP113/Chem_NY_Mentha Piperita Extract_01 | 7B4OP113 | Mentha Piperita Extract | Mentha Piperita ExtraCt. | Chem_NY_Mentha Piperita Extract_01 | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP113/Chem_NY_Mentha Piperita Extract_01 | 7B4OP113 | Mentha Piperita Extract | Mentha Piperita ExtraCt. | Chem_NY_Mentha Piperita Extract_01 | NY-Analytical | 792 | Chemical Inspection | Odor | Characteristic (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP203/Chem_NY_Diazolidinyl Urea – RI-1157 | 7B4OP203 | Diazolidinyl Urea – RI-1157 | Diazolidinyl Urea – Ri-1157 | Chem_NY_Diazolidinyl Urea – RI-1157 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP203/Chem_NY_Diazolidinyl Urea – RI-1157 | 7B4OP203 | Diazolidinyl Urea – RI-1157 | Diazolidinyl Urea – Ri-1157 | Chem_NY_Diazolidinyl Urea – RI-1157 | NY-Analytical | 398 | pH | pH (1% w/v aqueous solution) | 7.02 |  | FDA,1 | 6.5 | 8 | TM 200 | TM200 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP203/Chem_NY_Diazolidinyl Urea – RI-1157 | 7B4OP203 | Diazolidinyl Urea – RI-1157 | Diazolidinyl Urea – Ri-1157 | Chem_NY_Diazolidinyl Urea – RI-1157 | NY-Analytical | 891 | Solutions and Solubility | Solubility (1% w/v aqueous solution) | Clear and Colorless (Done) |  |  |  |  | TM 222 | TM222 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP203/Chem_NY_Diazolidinyl Urea – RI-1157 | 7B4OP203 | Diazolidinyl Urea – RI-1157 | Diazolidinyl Urea – Ri-1157 | Chem_NY_Diazolidinyl Urea – RI-1157 | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP203/Chem_NY_Diazolidinyl Urea – RI-1157 | 7B4OP203 | Diazolidinyl Urea – RI-1157 | Diazolidinyl Urea – Ri-1157 | Chem_NY_Diazolidinyl Urea – RI-1157 | NY-Analytical | 792 | Chemical Inspection | Appearance | Free flowing powder, free of foreign material (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP203/Chem_NY_Diazolidinyl Urea – RI-1157 | 7B4OP203 | Diazolidinyl Urea – RI-1157 | Diazolidinyl Urea – Ri-1157 | Chem_NY_Diazolidinyl Urea – RI-1157 | NY-Analytical | 792 | Chemical Inspection | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP203/Chem_NY_Diazolidinyl Urea – RI-1157 | 7B4OP203 | Diazolidinyl Urea – RI-1157 | Diazolidinyl Urea – Ri-1157 | Chem_NY_Diazolidinyl Urea – RI-1157 | NY-Analytical | 792 | Chemical Inspection | Odor | Similar in character and intensity to standard, slight pungent odor (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP209/MASTER - Microbiology - J&J Chemical Testing) | 7B4OP209 | Tinogard HS | Tinogard Hs | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 808 | Absorption | Absorption (1%/1cm) | 404.16 |  | FDA,1 | 401.8 |  | J&J TM-002747 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP209/MASTER - Microbiology - J&J Chemical Testing) | 7B4OP209 | Tinogard HS | Tinogard Hs | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 398 | pH | pH (1% Solution) | 6.2 |  | FDA,1 | 4.5 | 7 | J&J TMWW78014 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP230/Chem_NY_Cosvat (Elestab CPN Ultra Pure) | 7B4OP230 | Cosvat (Elestab CPN Ultra Pure) | Cosvat (Elestab Cpn Ultra Pure) | Chem_NY_Cosvat (Elestab CPN Ultra Pure) | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP230/Chem_NY_Cosvat (Elestab CPN Ultra Pure) | 7B4OP230 | Cosvat (Elestab CPN Ultra Pure) | Cosvat (Elestab Cpn Ultra Pure) | Chem_NY_Cosvat (Elestab CPN Ultra Pure) | NY-Analytical | 792 | Chemical Inspection | Appearance | White to off - white powder or crystalline aggregates. (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP230/Chem_NY_Cosvat (Elestab CPN Ultra Pure) | 7B4OP230 | Cosvat (Elestab CPN Ultra Pure) | Cosvat (Elestab Cpn Ultra Pure) | Chem_NY_Cosvat (Elestab CPN Ultra Pure) | NY-Analytical | 390 | Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP230/Chem_NY_Cosvat (Elestab CPN Ultra Pure) | 7B4OP230 | Cosvat (Elestab CPN Ultra Pure) | Cosvat (Elestab Cpn Ultra Pure) | Chem_NY_Cosvat (Elestab CPN Ultra Pure) | NY-Analytical | 801 | % Water | Water Content | 0.2 | % | FDA,1 |  | 1 | J&J TMWW78017 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP230/Chem_NY_Cosvat (Elestab CPN Ultra Pure) | 7B4OP230 | Cosvat (Elestab CPN Ultra Pure) | Cosvat (Elestab Cpn Ultra Pure) | Chem_NY_Cosvat (Elestab CPN Ultra Pure) | NY-Analytical | 796 | Assay | Assay | 100.7 | % | FDA,1 | 99 |  | TM-002291 | TM_002291 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP230/Chem_NY_Cosvat (Elestab CPN Ultra Pure) | 7B4OP230 | Cosvat (Elestab CPN Ultra Pure) | Cosvat (Elestab Cpn Ultra Pure) | Chem_NY_Cosvat (Elestab CPN Ultra Pure) | NY-Analytical | 796 | Assay | Assay Final | 100.7 | % | FDA,1 | 99 |  | TM-002291 | TM_002291 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP284/Chem_NY_Sodium Citrate USP | 7B4OP284 | Sodium Citrate USP | Sodium Citrate USP | Chem_NY_Sodium Citrate USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Odor |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP284/Chem_NY_Sodium Citrate USP | 7B4OP284 | Sodium Citrate USP | Sodium Citrate USP | Chem_NY_Sodium Citrate USP | NY-Analytical | 396 | Odor | Odor | Passes Test (Done) |  |  |  |  | J&J TM-003188 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP284/Chem_NY_Sodium Citrate USP | 7B4OP284 | Sodium Citrate USP | Sodium Citrate USP | Chem_NY_Sodium Citrate USP | NY-Analytical | 792 | Chemical Inspection | Appearance | Colorless crystals, or as a white crystalline powder (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP284/Chem_NY_Sodium Citrate USP | 7B4OP284 | Sodium Citrate USP | Sodium Citrate USP | Chem_NY_Sodium Citrate USP | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP284/Chem_NY_Sodium Citrate USP | 7B4OP284 | Sodium Citrate USP | Sodium Citrate USP | Chem_NY_Sodium Citrate USP | NY-Analytical | 813 | Identification - USP | Identification A - Sodium | Passes Test (Done) |  |  |  |  | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP284/Chem_NY_Sodium Citrate USP | 7B4OP284 | Sodium Citrate USP | Sodium Citrate USP | Chem_NY_Sodium Citrate USP | NY-Analytical | 813 | Identification - USP | Identification B - Citrate | Passes Test (Done) |  |  |  |  | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP284/Chem_NY_Sodium Citrate USP | 7B4OP284 | Sodium Citrate USP | Sodium Citrate USP | Chem_NY_Sodium Citrate USP | NY-Analytical | 813 | Identification - USP | Identification C | Passes Test (Done) |  |  |  |  | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP661/Chem_NY_Pemulen TR-2 | 7B4OP661 | Pemulen TR-2 | Pemulen Tr-2 | Chem_NY_Pemulen TR-2 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA | COA | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP661/Chem_NY_Pemulen TR-2 | 7B4OP661 | Pemulen TR-2 | Pemulen Tr-2 | Chem_NY_Pemulen TR-2 | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP661/Chem_NY_Pemulen TR-2 | 7B4OP661 | Pemulen TR-2 | Pemulen Tr-2 | Chem_NY_Pemulen TR-2 | NY-Analytical | 792 | Chemical Inspection | Appearance | Powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP661/Chem_NY_Pemulen TR-2 | 7B4OP661 | Pemulen TR-2 | Pemulen Tr-2 | Chem_NY_Pemulen TR-2 | NY-Analytical | 792 | Chemical Inspection | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP661/Chem_NY_Pemulen TR-2 | 7B4OP661 | Pemulen TR-2 | Pemulen Tr-2 | Chem_NY_Pemulen TR-2 | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OP661/Chem_NY_Pemulen TR-2 | 7B4OP661 | Pemulen TR-2 | Pemulen Tr-2 | Chem_NY_Pemulen TR-2 | NY-Analytical | 792 | Chemical Inspection | Odor | Slightly acetic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ028/Chem_NY_Sorbitan Monolaurate | 7B4OQ028 | Sorbitan Monolaurate | Sorbitan Monolaurate | Chem_NY_Sorbitan Monolaurate | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7B4OQ028/Chem_NY_Sorbitan Monolaurate | 7B4OQ028 | Sorbitan Monolaurate | Sorbitan Monolaurate | Chem_NY_Sorbitan Monolaurate | NY-Analytical | 390 | Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411 | TM411 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7B4OQ028/Chem_NY_Sorbitan Monolaurate | 7B4OQ028 | Sorbitan Monolaurate | Sorbitan Monolaurate | Chem_NY_Sorbitan Monolaurate | NY-Analytical | 792 | Chemical Inspection | Appearance | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7B4OQ028/Chem_NY_Sorbitan Monolaurate | 7B4OQ028 | Sorbitan Monolaurate | Sorbitan Monolaurate | Chem_NY_Sorbitan Monolaurate | NY-Analytical | 792 | Chemical Inspection | Color | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7B4OQ028/Chem_NY_Sorbitan Monolaurate | 7B4OQ028 | Sorbitan Monolaurate | Sorbitan Monolaurate | Chem_NY_Sorbitan Monolaurate | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7B4OQ028/Chem_NY_Sorbitan Monolaurate | 7B4OQ028 | Sorbitan Monolaurate | Sorbitan Monolaurate | Chem_NY_Sorbitan Monolaurate | NY-Analytical | 792 | Chemical Inspection | Odor | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7B4OQ084/Chem_60A | 7B4OQ084 | Tergitol 15-S-12 | Tergitol 15-S-12 | Chem_60A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ084/Chem_60A | 7B4OQ084 | Tergitol 15-S-12 | Tergitol 15-S-12 | Chem_60A | NY-Analytical | 792 | Chemical Inspection | Appearance | Slightly Opaque Viscous Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ084/Chem_60A | 7B4OQ084 | Tergitol 15-S-12 | Tergitol 15-S-12 | Chem_60A | NY-Analytical | 792 | Chemical Inspection | Color | Slightly White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ084/Chem_60A | 7B4OQ084 | Tergitol 15-S-12 | Tergitol 15-S-12 | Chem_60A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ084/Chem_60A | 7B4OQ084 | Tergitol 15-S-12 | Tergitol 15-S-12 | Chem_60A | NY-Analytical | 792 | Chemical Inspection | Odor | Mild (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ209/Chem_61A | 7B4OQ209 | Disodium Phosphate Food Grade | Disodium pHosphate Food Grade | Chem_61A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ209/Chem_61A | 7B4OQ209 | Disodium Phosphate Food Grade | Disodium pHosphate Food Grade | Chem_61A | NY-Analytical | 390 | Identification | IR Identification | [5 empty] | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ209/Chem_61A | 7B4OQ209 | Disodium Phosphate Food Grade | Disodium pHosphate Food Grade | Chem_61A | NY-Analytical | 792 | Chemical Inspection | Appearance | Free flowing granules or powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ209/Chem_61A | 7B4OQ209 | Disodium Phosphate Food Grade | Disodium pHosphate Food Grade | Chem_61A | NY-Analytical | 792 | Chemical Inspection | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ209/Chem_61A | 7B4OQ209 | Disodium Phosphate Food Grade | Disodium pHosphate Food Grade | Chem_61A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ209/Chem_61A | 7B4OQ209 | Disodium Phosphate Food Grade | Disodium pHosphate Food Grade | Chem_61A | NY-Analytical | 792 | Chemical Inspection | Odor | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ214/Chem_62A | 7B4OQ214 | Sorbic Acid/FCC | Sorbic Acid/FCC | Chem_62A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ214/Chem_62A | 7B4OQ214 | Sorbic Acid/FCC | Sorbic Acid/FCC | Chem_62A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ214/Chem_62A | 7B4OQ214 | Sorbic Acid/FCC | Sorbic Acid/FCC | Chem_62A | NY-Analytical | 792 | Chemical Inspection | Appearance | Free flowing crystalline powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ214/Chem_62A | 7B4OQ214 | Sorbic Acid/FCC | Sorbic Acid/FCC | Chem_62A | NY-Analytical | 792 | Chemical Inspection | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ214/Chem_62A | 7B4OQ214 | Sorbic Acid/FCC | Sorbic Acid/FCC | Chem_62A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ214/Chem_62A | 7B4OQ214 | Sorbic Acid/FCC | Sorbic Acid/FCC | Chem_62A | NY-Analytical | 792 | Chemical Inspection | Odor | Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ595/Chem_63A | 7B4OQ595 | Bronopol | Bronopol | Chem_63A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ595/Chem_63A | 7B4OQ595 | Bronopol | Bronopol | Chem_63A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ595/Chem_63A | 7B4OQ595 | Bronopol | Bronopol | Chem_63A | NY-Analytical | 792 | Chemical Inspection | Appearance | Crystalline Powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ595/Chem_63A | 7B4OQ595 | Bronopol | Bronopol | Chem_63A | NY-Analytical | 792 | Chemical Inspection | Color | White to almost white (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ595/Chem_63A | 7B4OQ595 | Bronopol | Bronopol | Chem_63A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ595/Chem_63A | 7B4OQ595 | Bronopol | Bronopol | Chem_63A | NY-Analytical | 792 | Chemical Inspection | Odor | Faint (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Disodium EDTA-USP | Disodium Edta-USP | Chem_NY_Disodium EDTA-USP | NY-Analytical | 841 | Residual Solvents | Residual Solvents | Supplier Attestation Letter |  |  |  |  | Residual Solvents | 0 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Disodium EDTA-USP | Disodium Edta-USP | Chem_NY_Disodium EDTA-USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Disodium EDTA-USP | Disodium Edta-USP | Chem_NY_Disodium EDTA-USP | NY-Analytical | 792 | Chemical Inspection | Appearance | White crystalline powder (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Disodium EDTA-USP | Disodium Edta-USP | Chem_NY_Disodium EDTA-USP | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Disodium EDTA-USP | Disodium Edta-USP | Chem_NY_Disodium EDTA-USP | NY-Analytical | 398 | pH | pH (1 in 20 solution) | 4.9 |  | FDA,1 | 4 | 6 | J&J TMWW78014 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Disodium EDTA-USP | Disodium Edta-USP | Chem_NY_Disodium EDTA-USP | NY-Analytical | 796 | Assay | Assay | 99.8 | % | FDA,1 | 99 | 101 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Disodium EDTA-USP | Disodium Edta-USP | Chem_NY_Disodium EDTA-USP | NY-Analytical | 796 | Assay | Assay Final | 99.8 | % | FDA,1 | 99 | 101 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Disodium EDTA-USP | Disodium Edta-USP | Chem_NY_Disodium EDTA-USP | NY-Analytical | 843 | Calcium | Calcium | Passes Test (Done) |  |  |  |  | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Disodium EDTA-USP | Disodium Edta-USP | Chem_NY_Disodium EDTA-USP | NY-Analytical | 833 | Heavy Metals | Heavy Metals | <0.005 | % | FDA,3 |  | 0.005 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Disodium EDTA-USP | Disodium Edta-USP | Chem_NY_Disodium EDTA-USP | NY-Analytical | 813 | Identification - USP | Identification A | Passes Test (Done) |  |  |  |  | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Disodium EDTA-USP | Disodium Edta-USP | Chem_NY_Disodium EDTA-USP | NY-Analytical | 813 | Identification - USP | Identification B | Passes Test (Done) |  |  |  |  | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Disodium EDTA-USP | Disodium Edta-USP | Chem_NY_Disodium EDTA-USP | NY-Analytical | 813 | Identification - USP | Identification C | Passes Test (Done) |  |  |  |  | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Disodium EDTA-USP | Disodium Edta-USP | Chem_NY_Disodium EDTA-USP | NY-Analytical | 802 | Loss on Drying | Loss on Drying | 9.6 | % | FDA,1 | 8.7 | 11.4 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Disodium EDTA-USP | Disodium Edta-USP | Chem_NY_Disodium EDTA-USP | NY-Analytical | 802 | Loss on Drying | Loss on Drying Final | 9.6 | % | FDA,1 | 8.7 | 11.4 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OQ794/Chem_NY_Disodium EDTA-USP | 7B4OQ794 | Disodium EDTA-USP | Disodium Edta-USP | Chem_NY_Disodium EDTA-USP | NY-Analytical | 842 | Nitrilotriacetic Acid | Nitrilotriacetic Acid | 0 | % | FDA,1 |  | 0.1 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OX232/Chem_40A | 7B4OX232 | Che alcohol SDA40-B 190 Proof | CHE Alcohol SDa40-B 190 Proof | Chem_40A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Apparent proof |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OX232/Chem_40A | 7B4OX232 | Che alcohol SDA40-B 190 Proof | CHE Alcohol SDa40-B 190 Proof | Chem_40A | NY-Analytical | 936 | Apparent proof | Apparent proof | 189.7 |  | FDA,1 | 189.7 | 190.2 | STM 101 | STM101 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OX232/Chem_40A | 7B4OX232 | Che alcohol SDA40-B 190 Proof | CHE Alcohol SDa40-B 190 Proof | Chem_40A | NY-Analytical | 402 | Specific Gravity at 25�C | Specific Gravity at 25°C | 0.8091 |  | FDA,4 | 0.8086 | 0.8098 | TM 202 | TM202 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OX232/Chem_40A | 7B4OX232 | Che alcohol SDA40-B 190 Proof | CHE Alcohol SDa40-B 190 Proof | Chem_40A | NY-Analytical | 390 | Identification | IR Identification | Meets USP requirement (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OX232/Chem_40A | 7B4OX232 | Che alcohol SDA40-B 190 Proof | CHE Alcohol SDa40-B 190 Proof | Chem_40A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear, mobile liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OX232/Chem_40A | 7B4OX232 | Che alcohol SDA40-B 190 Proof | CHE Alcohol SDa40-B 190 Proof | Chem_40A | NY-Analytical | 792 | Chemical Inspection | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OX232/Chem_40A | 7B4OX232 | Che alcohol SDA40-B 190 Proof | CHE Alcohol SDa40-B 190 Proof | Chem_40A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7B4OX232/Chem_40A | 7B4OX232 | Che alcohol SDA40-B 190 Proof | CHE Alcohol SDa40-B 190 Proof | Chem_40A | NY-Analytical | 792 | Chemical Inspection | Odor | Alcohol (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | 7C4AA580 | Fragrance - Grapefruit Blossom & Orange | Fragrance - Grapefruit Blossom & Orange | Chem_NY_Fragrance-Grapefruit Blossom &Orange | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25�C |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | 7C4AA580 | Fragrance - Grapefruit Blossom & Orange | Fragrance - Grapefruit Blossom & Orange | Chem_NY_Fragrance-Grapefruit Blossom &Orange | NY-Analytical | 402 | Specific Gravity at 25�C | Specific Gravity at 25°C | 1.0662 |  | FDA,4 | 1.0561 | 1.0761 | TM 202 | TM202 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | 7C4AA580 | Fragrance - Grapefruit Blossom & Orange | Fragrance - Grapefruit Blossom & Orange | Chem_NY_Fragrance-Grapefruit Blossom &Orange | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | 7C4AA580 | Fragrance - Grapefruit Blossom & Orange | Fragrance - Grapefruit Blossom & Orange | Chem_NY_Fragrance-Grapefruit Blossom &Orange | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | 7C4AA580 | Fragrance - Grapefruit Blossom & Orange | Fragrance - Grapefruit Blossom & Orange | Chem_NY_Fragrance-Grapefruit Blossom &Orange | NY-Analytical | 792 | Chemical Inspection | Color | Pale Yellow – Light Yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | 7C4AA580 | Fragrance - Grapefruit Blossom & Orange | Fragrance - Grapefruit Blossom & Orange | Chem_NY_Fragrance-Grapefruit Blossom &Orange | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | 7C4AA580 | Fragrance - Grapefruit Blossom & Orange | Fragrance - Grapefruit Blossom & Orange | Chem_NY_Fragrance-Grapefruit Blossom &Orange | NY-Analytical | 792 | Chemical Inspection | Odor | Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4FD110/Chem_69A | 7C4FD110 | Sodium Hydroxide Pellets NF/FCC | Sodium Hydroxide Pellets NF/FCC | Chem_69A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4FD110/Chem_69A | 7C4FD110 | Sodium Hydroxide Pellets NF/FCC | Sodium Hydroxide Pellets NF/FCC | Chem_69A | NY-Analytical | 792 | Chemical Inspection | Appearance | White or nearly white, hard, brittle material (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4FD110/Chem_69A | 7C4FD110 | Sodium Hydroxide Pellets NF/FCC | Sodium Hydroxide Pellets NF/FCC | Chem_69A | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4FD110/Chem_69A | 7C4FD110 | Sodium Hydroxide Pellets NF/FCC | Sodium Hydroxide Pellets NF/FCC | Chem_69A | NY-Analytical | 390 | Identification | Identification B | Pass (Done) |  |  |  |  | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4FD110/Chem_69A | 7C4FD110 | Sodium Hydroxide Pellets NF/FCC | Sodium Hydroxide Pellets NF/FCC | Chem_69A | NY-Analytical | 390 | Identification | Identification for Sodium | Positive response for Sodium (Done) |  |  |  |  | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 7C4OA709 | PEG-75 Shea Butter Glycerides (Lipex 102E-75 100%) | PEG-75 Shea Butter Glycerides (Lipex 102E-75 100%) | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 961 | Peroxide Value | Peroxide Value |  |  | FDA,1 |  | 1 | AOCS Cd 8-53 | AOCS_Cd_8_53 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 7C4OA709 | PEG-75 Shea Butter Glycerides (Lipex 102E-75 100%) | PEG-75 Shea Butter Glycerides (Lipex 102E-75 100%) | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 944 | color (Gardner) | color (Gardner) |  |  | FDA,0 |  | 3 | AOCS Td 1a-64) | AOCS_Td_1a_64 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 7C4OA709 | PEG-75 Shea Butter Glycerides (Lipex 102E-75 100%) | PEG-75 Shea Butter Glycerides (Lipex 102E-75 100%) | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA | COA | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 7C4OA709 | PEG-75 Shea Butter Glycerides (Lipex 102E-75 100%) | PEG-75 Shea Butter Glycerides (Lipex 102E-75 100%) | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 826 | Acid Value | Acid Value | 1.7 | mg KOH/g | FDA,2 |  | 2 | IUPAC 2.201(M) | IUPAC_2_201M | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 7C4OA709 | PEG-75 Shea Butter Glycerides (Lipex 102E-75 100%) | PEG-75 Shea Butter Glycerides (Lipex 102E-75 100%) | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 826 | Acid Value | Acid Value Final | 1.7 | mg KOH/g | FDA,2 |  | 2 | IUPAC 2.201(M) | IUPAC_2_201M | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 7C4OA709 | PEG-75 Shea Butter Glycerides (Lipex 102E-75 100%) | PEG-75 Shea Butter Glycerides (Lipex 102E-75 100%) | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 810 | Saponification Value | Saponification Value | 22.1 | mg/g | FDA,1 | 17 | 25 | IUPAC 2.202 | IUPAC_2_202 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 7C4OA709 | PEG-75 Shea Butter Glycerides (Lipex 102E-75 100%) | PEG-75 Shea Butter Glycerides (Lipex 102E-75 100%) | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 810 | Saponification Value | Saponification Value Final | 22.1 | mg/g | FDA,1 | 17 | 25 | IUPAC 2.202 | IUPAC_2_202 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 7C4OA709 | PEG-75 Shea Butter Glycerides (Lipex 102E-75 100%) | PEG-75 Shea Butter Glycerides (Lipex 102E-75 100%) | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 1000 | Cloudpoint | Cloudpoint | 67.4 | C | FDA,1 | 66 | 71 | TC276/WG2N30 930809C | TC276_WG2N30_930809C | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OA709/Chem_NY_PEG-75 Shea Butter Glycerides | 7C4OA709 | PEG-75 Shea Butter Glycerides (Lipex 102E-75 100%) | PEG-75 Shea Butter Glycerides (Lipex 102E-75 100%) | Chem_NY_PEG-75 Shea Butter Glycerides | NY-Analytical | 959 | Water Content | Water Content |  | % | FDA,2 |  | 1.5 | Vatten LB | Vatten_LB | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
| 7C4OD232 | Methylparaben NF/USP |  | Analytical - Chem_NY_Methylparaben NF/USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis | COA |  |  |  | left_only
| 7C4OD232 | Methylparaben NF/USP |  | Analytical - Chem_NY_Propylparaben USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  |  |  |  | left_only
| 7C4OD232 | Methylparaben NF/USP |  | Analytical - Chem_NY_Methylparaben NF/USP | NY-Analytical | 390 | Identification | IR Identification | [5 empty] | TM 411 | TM411 |  |  |  | left_only
| 7C4OD232 | Methylparaben NF/USP |  | Analytical - Chem_NY_Methylparaben NF/USP | NY-Analytical | 792 | Chemical Inspection | Appearance | Colorless crystals or white crystalline powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7C4OD232 | Methylparaben NF/USP |  | Analytical - Chem_NY_Propylparaben USP | NY-Analytical | 792 | Chemical Inspection | Appearance | Small, colorless crystals or white crystalline powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7C4OD232 | Methylparaben NF/USP |  | Analytical - Chem_NY_Methylparaben NF/USP | NY-Analytical | 792 | Chemical Inspection | Contamination | [5 empty] | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7C4OD232 | Methylparaben NF/USP |  | Analytical - Chem_NY_Propylparaben USP | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7C4OD232 | Methylparaben NF/USP |  | Analytical - Chem_NY_Methylparaben NF/USP | NY-Analytical | 792 | Chemical Inspection | Odor | [5 empty] | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7C4OD232 | Methylparaben NF/USP |  | Analytical - Chem_NY_Propylparaben USP | NY-Analytical | 792 | Chemical Inspection | Odor | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7C4OD232 | Methylparaben NF/USP |  | Analytical - Chem_NY_Propylparaben USP | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | USP <197M> | USP_197M |  |  |  | left_only
| 7C4OD232 | Methylparaben NF/USP |  | Analytical - Chem_NY_Propylparaben USP | NY-Analytical | 796 | Assay | Assay | 99.4 | % | FDA,1 | 98 | 102 | USP/NF | USP_NF |  |  |  | left_only
| 7C4OD232 | Methylparaben NF/USP |  | Analytical - Chem_NY_Methylparaben NF/USP | NY-Analytical | 796 | Assay | Assay |  | % | FDA,1 | 98 | 102 | USP/NF | USP_NF |  |  |  | left_only
| 7C4OD232 | Methylparaben NF/USP |  | Analytical - Chem_NY_Propylparaben USP | NY-Analytical | 796 | Assay | Assay Final | 99.4 | % | FDA,1 | 98 | 102 | USP/NF | USP_NF |  |  |  | left_only
| 7C4OD232 | Methylparaben NF/USP |  | Analytical - Chem_NY_Methylparaben NF/USP | NY-Analytical | 796 | Assay | Assay Final |  | % | FDA,1 | 98 | 102 | USP/NF | USP_NF |  |  |  | left_only
7C4OG169/Chem_75A | 7C4OG169 | Ammonium Chloride | Ammonium Chloride | Chem_75A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OG169/Chem_75A | 7C4OG169 | Ammonium Chloride | Ammonium Chloride | Chem_75A | NY-Analytical | 792 | Chemical Inspection | Appearance | Crystals (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OG169/Chem_75A | 7C4OG169 | Ammonium Chloride | Ammonium Chloride | Chem_75A | NY-Analytical | 792 | Chemical Inspection | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OG169/Chem_75A | 7C4OG169 | Ammonium Chloride | Ammonium Chloride | Chem_75A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OG169/Chem_75A | 7C4OG169 | Ammonium Chloride | Ammonium Chloride | Chem_75A | NY-Analytical | 390 | Identification | Identification for Ammonium | Pass (Done) |  |  |  |  | TM400/408b/421/422/USP | TM400_TM408b_TM421_TM422_USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OG169/Chem_75A | 7C4OG169 | Ammonium Chloride | Ammonium Chloride | Chem_75A | NY-Analytical | 390 | Identification | Identification for Chloride | Pass (Done) |  |  |  |  | TM400/408b/421/422/USP | TM400_TM408b_TM421_TM422_USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Camellia Oleifera Leaf Extract | Camellia Oleifera Leaf ExtraCt. | Chem_NY_Camellia Oleifera Leaf Extract | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Camellia Oleifera Leaf Extract | Camellia Oleifera Leaf ExtraCt. | Chem_NY_Camellia Oleifera Leaf Extract | NY-Analytical | 792 | Chemical Inspection | Appearance | Light yellow to amber liquid (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Camellia Oleifera Leaf Extract | Camellia Oleifera Leaf ExtraCt. | Chem_NY_Camellia Oleifera Leaf Extract | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Camellia Oleifera Leaf Extract | Camellia Oleifera Leaf ExtraCt. | Chem_NY_Camellia Oleifera Leaf Extract | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Camellia Oleifera Leaf Extract | Camellia Oleifera Leaf ExtraCt. | Chem_NY_Camellia Oleifera Leaf Extract | NY-Analytical | 390 | Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Camellia Oleifera Leaf Extract | Camellia Oleifera Leaf ExtraCt. | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 509 | J&J_Pathogens | Biggy or SDA | No Growth_SDA |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Camellia Oleifera Leaf Extract | Camellia Oleifera Leaf ExtraCt. | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 509 | J&J_Pathogens | Cetrimide | No Growth |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Camellia Oleifera Leaf Extract | Camellia Oleifera Leaf ExtraCt. | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 524 | J&J Fungal | Fungal | 0 | CFU/g | FDA,0 | 0 | 200 | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Camellia Oleifera Leaf Extract | Camellia Oleifera Leaf ExtraCt. | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 509 | J&J_Pathogens | MacConkey | No Growth |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Camellia Oleifera Leaf Extract | Camellia Oleifera Leaf ExtraCt. | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 509 | J&J_Pathogens | MSA or VJ | No Growth_MSA |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Camellia Oleifera Leaf Extract | Camellia Oleifera Leaf ExtraCt. | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 723 | J&J TSB Controls - TSA | Negative Control | No Growth TSB |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Camellia Oleifera Leaf Extract | Camellia Oleifera Leaf ExtraCt. | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | No Growth TSB |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Camellia Oleifera Leaf Extract | Camellia Oleifera Leaf ExtraCt. | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 509 | J&J_Pathogens | TSA | No Growth |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Camellia Oleifera Leaf Extract | Camellia Oleifera Leaf ExtraCt. | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | 0 | cfu/mL | FDA,0 | 0 | 500 | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | 7C4OP027 | Camellia Oleifera Leaf Extract | Camellia Oleifera Leaf ExtraCt. | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 517 | J&J_Enrichment | TSB | Not Turbid |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | 7C4OP065 | Sodium Lauroamphoacetate 30% | Sodium Lauroamphoacetate 30% | Chem_NY_ Sodium Lauroamphoacetate 30% | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | 7C4OP065 | Sodium Lauroamphoacetate 30% | Sodium Lauroamphoacetate 30% | Chem_NY_ Sodium Lauroamphoacetate 30% | NY-Analytical | 398 | pH | pH (1 in 20 solution) | 9.5 |  | FDA,1 | 9 | 9.5 | TM 200 | TM200 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | 7C4OP065 | Sodium Lauroamphoacetate 30% | Sodium Lauroamphoacetate 30% | Chem_NY_ Sodium Lauroamphoacetate 30% | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear Clean Liquid (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | 7C4OP065 | Sodium Lauroamphoacetate 30% | Sodium Lauroamphoacetate 30% | Chem_NY_ Sodium Lauroamphoacetate 30% | NY-Analytical | 792 | Chemical Inspection | Color | Colorless to light yellow (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | 7C4OP065 | Sodium Lauroamphoacetate 30% | Sodium Lauroamphoacetate 30% | Chem_NY_ Sodium Lauroamphoacetate 30% | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | 7C4OP065 | Sodium Lauroamphoacetate 30% | Sodium Lauroamphoacetate 30% | Chem_NY_ Sodium Lauroamphoacetate 30% | NY-Analytical | 792 | Chemical Inspection | Odor | Match standard (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | 7C4OP065 | Sodium Lauroamphoacetate 30% | Sodium Lauroamphoacetate 30% | Chem_NY_ Sodium Lauroamphoacetate 30% | NY-Analytical | 390 | Identification | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OP230/MASTER - Microbiology - J&J Chemical Testing) | 7C4OP230 | Cetiol CC | Cetiol Cc | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 826 | Acid Value | Acid Value | 0.01 |  | FDA,2 | 0 | 0.2 | J&J TMWW78011 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OP230/MASTER - Microbiology - J&J Chemical Testing) | 7C4OP230 | Cetiol CC | Cetiol Cc | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 831 | Density @ 20° C | Density | 0.8916 |  | FDA,4 | 0.89 | 0.893 | J&J TMWW78012 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OP284/Chem_NY_Grape Fruit Extract | 7C4OP284 | Grape Fruit Extract | Grape Fruit ExtraCt. | Chem_NY_Grape Fruit Extract | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OP284/Chem_NY_Grape Fruit Extract | 7C4OP284 | Grape Fruit Extract | Grape Fruit ExtraCt. | Chem_NY_Grape Fruit Extract | NY-Analytical | 390 | Identification | Identification (FTIR) | Matches that of the Reference Standard (Done) |  |  |  |  | J&J TMWW78013 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OP284/Chem_NY_Grape Fruit Extract | 7C4OP284 | Grape Fruit Extract | Grape Fruit ExtraCt. | Chem_NY_Grape Fruit Extract | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear to Pale Yellow Liquid (Done) |  |  |  |  | JJTM7252 | JJTM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OP284/Chem_NY_Grape Fruit Extract | 7C4OP284 | Grape Fruit Extract | Grape Fruit ExtraCt. | Chem_NY_Grape Fruit Extract | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | JJTM7252 | JJTM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OP601/MASTER - Microbiology - J&J Chemical Testing) | 7C4OP601 | Glycerox 767 | Glycerox 767 | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 748 | Visual Inspection | Appearance of Packaging | Meets specifications (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7C4OQ028/Chem_NY_Isohexadecane | 7C4OQ028 | Isohexadecane | Isohexadecane | Chem_NY_Isohexadecane | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Refractive Index at 20°C |  | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7C4OQ028/Chem_NY_Isohexadecane | 7C4OQ028 | Isohexadecane | Isohexadecane | Chem_NY_Isohexadecane | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | 1.439 |  | FDA,1 | 1.4 | 1.5 | TM 204 | TM204 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7C4OQ028/Chem_NY_Isohexadecane | 7C4OQ028 | Isohexadecane | Isohexadecane | Chem_NY_Isohexadecane | NY-Analytical | 390 | Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411 | TM411 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7C4OQ028/Chem_NY_Isohexadecane | 7C4OQ028 | Isohexadecane | Isohexadecane | Chem_NY_Isohexadecane | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7C4OQ028/Chem_NY_Isohexadecane | 7C4OQ028 | Isohexadecane | Isohexadecane | Chem_NY_Isohexadecane | NY-Analytical | 792 | Chemical Inspection | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7C4OQ028/Chem_NY_Isohexadecane | 7C4OQ028 | Isohexadecane | Isohexadecane | Chem_NY_Isohexadecane | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7C4OQ028/Chem_NY_Isohexadecane | 7C4OQ028 | Isohexadecane | Isohexadecane | Chem_NY_Isohexadecane | NY-Analytical | 792 | Chemical Inspection | Odor | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
| 7C4OQ595 | Potassium Sorbate USP/NF/FCC |  | Analytical - Chem_81A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  |  |  |  | left_only
| 7C4OQ595 | Potassium Sorbate USP/NF/FCC |  | Analytical - Chem_81A | NY-Analytical | 792 | Chemical Inspection | Appearance | Free flowing granules or powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7C4OQ595 | Potassium Sorbate USP/NF/FCC |  | Analytical - Chem_81A | NY-Analytical | 792 | Chemical Inspection | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7C4OQ595 | Potassium Sorbate USP/NF/FCC |  | Analytical - Chem_81A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7C4OQ595 | Potassium Sorbate USP/NF/FCC |  | Analytical - Chem_81A | NY-Analytical | 792 | Chemical Inspection | Odor | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7C4OQ595 | Potassium Sorbate USP/NF/FCC |  | Analytical - Chem_81A | NY-Analytical | 390 | Identification | Identification B | Pass (Done) |  |  |  |  | TM400/408b/421/422/USP | TM400_TM408b_TM421_TM422_USP |  |  |  | left_only
| 7C4OQ595 | Potassium Sorbate USP/NF/FCC |  | Analytical - Chem_81A | NY-Analytical | 390 | Identification | Identification for Potassium | Pass (Done) |  |  |  |  | TM400/408b/421/422/USP | TM400_TM408b_TM421_TM422_USP |  |  |  | left_only
| 7C4OX110 | IPP/IPM/IPS Blend |  | Analytical - Chem_84A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25�C |  |  |  |  | left_only
| 7C4OX110 | IPP/IPM/IPS Blend |  | Analytical - Chem_84A | NY-Analytical | 402 | Specific Gravity at 25�C | Specific Gravity at 25°C | 0.8529 |  | FDA,3 | 0.845 | 0.858 | TM 202 | TM202 |  |  |  | left_only
| 7C4OX110 | IPP/IPM/IPS Blend |  | Analytical - Chem_84A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7C4OX110 | IPP/IPM/IPS Blend |  | Analytical - Chem_84A | NY-Analytical | 792 | Chemical Inspection | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7C4OX110 | IPP/IPM/IPS Blend |  | Analytical - Chem_84A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7C4OX110 | IPP/IPM/IPS Blend |  | Analytical - Chem_84A | NY-Analytical | 792 | Chemical Inspection | Odor | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
7D4AA580/Chem_NY_Monoethanolamine | 7D4AA580 | Monoethanolamine | Monoethanolamine | Chem_NY_Monoethanolamine | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4AA580/Chem_NY_Monoethanolamine | 7D4AA580 | Monoethanolamine | Monoethanolamine | Chem_NY_Monoethanolamine | NY-Analytical | 390 | Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4AA580/Chem_NY_Monoethanolamine | 7D4AA580 | Monoethanolamine | Monoethanolamine | Chem_NY_Monoethanolamine | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear, Colorless, Viscous Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4AA580/Chem_NY_Monoethanolamine | 7D4AA580 | Monoethanolamine | Monoethanolamine | Chem_NY_Monoethanolamine | NY-Analytical | 792 | Chemical Inspection | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4AA580/Chem_NY_Monoethanolamine | 7D4AA580 | Monoethanolamine | Monoethanolamine | Chem_NY_Monoethanolamine | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4AA580/Chem_NY_Monoethanolamine | 7D4AA580 | Monoethanolamine | Monoethanolamine | Chem_NY_Monoethanolamine | NY-Analytical | 792 | Chemical Inspection | Odor | Ammoniacal (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OG169/Chem_88A | 7D4OG169 | Glass Wipe Fragrance | Glass Wipe Fragrance | Chem_88A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25�C |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OG169/Chem_88A | 7D4OG169 | Glass Wipe Fragrance | Glass Wipe Fragrance | Chem_88A | NY-Analytical | 402 | Specific Gravity at 25�C | Specific Gravity at 25°C | 1.068 |  | FDA,4 | 1.0594 | 1.0794 | TM 202 | TM202 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OG169/Chem_88A | 7D4OG169 | Glass Wipe Fragrance | Glass Wipe Fragrance | Chem_88A | NY-Analytical | 451 | Refractive Index at 25�C | Refractive Index | 1.4755 |  | FDA,4 | 1.4695 | 1.4895 | TM 204 | TM204 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OG169/Chem_88A | 7D4OG169 | Glass Wipe Fragrance | Glass Wipe Fragrance | Chem_88A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OG169/Chem_88A | 7D4OG169 | Glass Wipe Fragrance | Glass Wipe Fragrance | Chem_88A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OG169/Chem_88A | 7D4OG169 | Glass Wipe Fragrance | Glass Wipe Fragrance | Chem_88A | NY-Analytical | 792 | Chemical Inspection | Color | Pale straw (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OG169/Chem_88A | 7D4OG169 | Glass Wipe Fragrance | Glass Wipe Fragrance | Chem_88A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OP065/Chem_NY_ Glydant Plus Liquid | 7D4OP065 | Glydant Plus Liquid | Glydant Plus Liquid | Chem_NY_ Glydant Plus Liquid | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OP065/Chem_NY_ Glydant Plus Liquid | 7D4OP065 | Glydant Plus Liquid | Glydant Plus Liquid | Chem_NY_ Glydant Plus Liquid | NY-Analytical | 398 | pH | pH | 7.66 |  | FDA,1 | 6 | 8 | TM 200 | TM200 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OP065/Chem_NY_ Glydant Plus Liquid | 7D4OP065 | Glydant Plus Liquid | Glydant Plus Liquid | Chem_NY_ Glydant Plus Liquid | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear Liquid (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OP065/Chem_NY_ Glydant Plus Liquid | 7D4OP065 | Glydant Plus Liquid | Glydant Plus Liquid | Chem_NY_ Glydant Plus Liquid | NY-Analytical | 792 | Chemical Inspection | Color | Water-White to light yellow (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OP065/Chem_NY_ Glydant Plus Liquid | 7D4OP065 | Glydant Plus Liquid | Glydant Plus Liquid | Chem_NY_ Glydant Plus Liquid | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OP065/Chem_NY_ Glydant Plus Liquid | 7D4OP065 | Glydant Plus Liquid | Glydant Plus Liquid | Chem_NY_ Glydant Plus Liquid | NY-Analytical | 792 | Chemical Inspection | Odor | Characteristic (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OP065/Chem_NY_ Glydant Plus Liquid | 7D4OP065 | Glydant Plus Liquid | Glydant Plus Liquid | Chem_NY_ Glydant Plus Liquid | NY-Analytical | 390 | Identification | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OP230/Chem_NY_Zemea_01 | 7D4OP230 | Propanediol, USP/NF | Propanediol, USP/Nf | Chem_NY_Zemea_01 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Appearance |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OP230/Chem_NY_Zemea_01 | 7D4OP230 | Propanediol, USP/NF | Propanediol, USP/Nf | Chem_NY_Zemea_01 | NY-Analytical | 397 | Appearance | Physical Properties | A Clear, Colorless Liquid (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OP230/Chem_NY_Zemea_01 | 7D4OP230 | Propanediol, USP/NF | Propanediol, USP/Nf | Chem_NY_Zemea_01 | NY-Analytical | 390 | Identification | Identification A | Passes test. (Done) |  |  |  |  | USP/NF | USP_NF | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OP230/Chem_NY_Zemea_01 | 7D4OP230 | Propanediol, USP/NF | Propanediol, USP/Nf | Chem_NY_Zemea_01 | NY-Analytical | 390 | Identification | Identification B | Passes test. (Done) |  |  |  |  | USP/NF | USP_NF | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OP601/Chem_91A | 7D4OP601 | Euxyl K 702 | Euxyl K 702 | Chem_91A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OP601/Chem_91A | 7D4OP601 | Euxyl K 702 | Euxyl K 702 | Chem_91A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear, yellow liquid. (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OP601/Chem_91A | 7D4OP601 | Euxyl K 702 | Euxyl K 702 | Chem_91A | NY-Analytical | 792 | Chemical Inspection | Odor | Meets Odor Specification or of Characteristic Odor. [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OP601/Chem_91A | 7D4OP601 | Euxyl K 702 | Euxyl K 702 | Chem_91A | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OP601/Chem_91A | 7D4OP601 | Euxyl K 702 | Euxyl K 702 | Chem_91A | NY-Analytical | 831 | Density @ 20° C | Density | 1.1193 |  | FDA,3 | 1.111 | 1.131 | J&J TMWW78012 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OP601/Chem_91A | 7D4OP601 | Euxyl K 702 | Euxyl K 702 | Chem_91A | NY-Analytical | 390 | Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OP601/Chem_91A | 7D4OP601 | Euxyl K 702 | Euxyl K 702 | Chem_91A | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | 1.53 |  | FDA,3 | 1.526 | 1.538 | J&J TMWW78015 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OP601/Chem_91A | 7D4OP601 | Euxyl K 702 | Euxyl K 702 | Chem_91A | NY-Analytical | 844 | % Benzoic Acid Phenoxyethanol & Dehydroacetic Acid | Benzoic Acid | 12.13 | % | FDA,1 | 11 | 13 | TM 882 | TM882 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OP601/Chem_91A | 7D4OP601 | Euxyl K 702 | Euxyl K 702 | Chem_91A | NY-Analytical | 844 | % Benzoic Acid Phenoxyethanol & Dehydroacetic Acid | Dehydroacetic Acid | 6.91 | % | FDA,1 | 6.5 | 7.5 | TM 882 | TM882 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OP601/Chem_91A | 7D4OP601 | Euxyl K 702 | Euxyl K 702 | Chem_91A | NY-Analytical | 844 | % Benzoic Acid Phenoxyethanol & Dehydroacetic Acid | Phenoxyethanol | 76.51 | % | FDA,1 | 71 | 77 | TM 882 | TM882 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OQ028/Chem_NY_Dicaprylyl Ether | 7D4OQ028 | Dicaprylyl Ether | Dicaprylyl Ether | Chem_NY_Dicaprylyl Ether | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Refractive Index at 20°C |  | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7D4OQ028/Chem_NY_Dicaprylyl Ether | 7D4OQ028 | Dicaprylyl Ether | Dicaprylyl Ether | Chem_NY_Dicaprylyl Ether | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | 1.431 |  | FDA,3 | 1.431 | 1.435 | TM 204 | TM204 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7D4OQ028/Chem_NY_Dicaprylyl Ether | 7D4OQ028 | Dicaprylyl Ether | Dicaprylyl Ether | Chem_NY_Dicaprylyl Ether | NY-Analytical | 390 | Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411 | TM411 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7D4OQ028/Chem_NY_Dicaprylyl Ether | 7D4OQ028 | Dicaprylyl Ether | Dicaprylyl Ether | Chem_NY_Dicaprylyl Ether | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7D4OQ028/Chem_NY_Dicaprylyl Ether | 7D4OQ028 | Dicaprylyl Ether | Dicaprylyl Ether | Chem_NY_Dicaprylyl Ether | NY-Analytical | 792 | Chemical Inspection | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7D4OQ028/Chem_NY_Dicaprylyl Ether | 7D4OQ028 | Dicaprylyl Ether | Dicaprylyl Ether | Chem_NY_Dicaprylyl Ether | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7D4OQ028/Chem_NY_Dicaprylyl Ether | 7D4OQ028 | Dicaprylyl Ether | Dicaprylyl Ether | Chem_NY_Dicaprylyl Ether | NY-Analytical | 792 | Chemical Inspection | Odor | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
| 7D4OQ209 | NP Baby Wipe Fragrance |  | Analytical - Chem_92A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25�C |  |  |  |  | left_only
| 7D4OQ209 | NP Baby Wipe Fragrance |  | Analytical - Chem_92A | NY-Analytical | 402 | Specific Gravity at 25�C | Specific Gravity at 25°C | 1.0512 |  | FDA,4 | 1.0399 | 1.0599 | TM 202 | TM202 |  |  |  | left_only
| 7D4OQ209 | NP Baby Wipe Fragrance |  | Analytical - Chem_92A | NY-Analytical | 451 | Refractive Index at 25�C | Refractive Index | 1.511 |  | FDA,4 | 1.507 | 1.5272 | TM 204 | TM204 |  |  |  | left_only
| 7D4OQ209 | NP Baby Wipe Fragrance |  | Analytical - Chem_92A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 |  |  |  | left_only
| 7D4OQ209 | NP Baby Wipe Fragrance |  | Analytical - Chem_92A | NY-Analytical | 792 | Chemical Inspection | Appearance | Pale Yellow Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7D4OQ209 | NP Baby Wipe Fragrance |  | Analytical - Chem_92A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7D4OQ209 | NP Baby Wipe Fragrance |  | Analytical - Chem_92A | NY-Analytical | 792 | Chemical Inspection | Odor | Baby Powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
7D4OQ214/Chem_93A | 7D4OQ214 | Disodium EDTA | Disodium Edta | Chem_93A | NY-Analytical | 792 | Chemical Inspection | Color | White to Off-White |  |  |  |  | Color (TM 208) | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OQ214/Chem_93A | 7D4OQ214 | Disodium EDTA | Disodium Edta | Chem_93A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OQ214/Chem_93A | 7D4OQ214 | Disodium EDTA | Disodium Edta | Chem_93A | NY-Analytical | 398 | pH | pH (5% Solution) | 4.72 |  | FDA,1 | 4 | 6 | TM 200 | TM200 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OQ214/Chem_93A | 7D4OQ214 | Disodium EDTA | Disodium Edta | Chem_93A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OQ214/Chem_93A | 7D4OQ214 | Disodium EDTA | Disodium Edta | Chem_93A | NY-Analytical | 792 | Chemical Inspection | Appearance | Powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OQ214/Chem_93A | 7D4OQ214 | Disodium EDTA | Disodium Edta | Chem_93A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OQ214/Chem_93A | 7D4OQ214 | Disodium EDTA | Disodium Edta | Chem_93A | NY-Analytical | 792 | Chemical Inspection | Odor | Odorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OQ967/Chem_94A | 7D4OQ967 | PEG-75 Lanolin | PEG-75 Lanolin | Chem_94A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OQ967/Chem_94A | 7D4OQ967 | PEG-75 Lanolin | PEG-75 Lanolin | Chem_94A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OQ967/Chem_94A | 7D4OQ967 | PEG-75 Lanolin | PEG-75 Lanolin | Chem_94A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear viscous liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OQ967/Chem_94A | 7D4OQ967 | PEG-75 Lanolin | PEG-75 Lanolin | Chem_94A | NY-Analytical | 792 | Chemical Inspection | Color | Yellow to amber (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7D4OQ967/Chem_94A | 7D4OQ967 | PEG-75 Lanolin | PEG-75 Lanolin | Chem_94A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
| 7E4OD232 | Sodium Carbonate Anhydrous |  | Analytical - Chem_97A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  |  |  |  | left_only
| 7E4OD232 | Sodium Carbonate Anhydrous |  | Analytical - Chem_97A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 |  |  |  | left_only
| 7E4OD232 | Sodium Carbonate Anhydrous |  | Analytical - Chem_97A | NY-Analytical | 792 | Chemical Inspection | Appearance | Granular Solid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7E4OD232 | Sodium Carbonate Anhydrous |  | Analytical - Chem_97A | NY-Analytical | 792 | Chemical Inspection | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7E4OD232 | Sodium Carbonate Anhydrous |  | Analytical - Chem_97A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7E4OD232 | Sodium Carbonate Anhydrous |  | Analytical - Chem_97A | NY-Analytical | 792 | Chemical Inspection | Odor | Odorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
7E4OP065/Chem_NY_Benzyl Alcohol | 7E4OP065 | Benzyl Alcohol | Benzyl Alcohol | Chem_NY_Benzyl Alcohol | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP065/Chem_NY_Benzyl Alcohol | 7E4OP065 | Benzyl Alcohol | Benzyl Alcohol | Chem_NY_Benzyl Alcohol | NY-Analytical | 792 | Chemical Inspection | Appearance | Colorless Liquid (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP065/Chem_NY_Benzyl Alcohol | 7E4OP065 | Benzyl Alcohol | Benzyl Alcohol | Chem_NY_Benzyl Alcohol | NY-Analytical | 792 | Chemical Inspection | Color | Colorless (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP065/Chem_NY_Benzyl Alcohol | 7E4OP065 | Benzyl Alcohol | Benzyl Alcohol | Chem_NY_Benzyl Alcohol | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP065/Chem_NY_Benzyl Alcohol | 7E4OP065 | Benzyl Alcohol | Benzyl Alcohol | Chem_NY_Benzyl Alcohol | NY-Analytical | 792 | Chemical Inspection | Odor | Match standard (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP065/Chem_NY_Benzyl Alcohol | 7E4OP065 | Benzyl Alcohol | Benzyl Alcohol | Chem_NY_Benzyl Alcohol | NY-Analytical | 390 | Identification | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | 7E4OP203 | Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | Cocamidopropyl Dimonium Chloride pHosphate Ri-1059 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | 7E4OP203 | Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | Cocamidopropyl Dimonium Chloride pHosphate Ri-1059 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Analytical | 398 | pH | pH (10% w/v aqueous solution) | 6.93 |  | FDA,1 | 6.5 | 7.5 | TM 200 | TM200 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | 7E4OP203 | Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | Cocamidopropyl Dimonium Chloride pHosphate Ri-1059 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | 7E4OP203 | Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | Cocamidopropyl Dimonium Chloride pHosphate Ri-1059 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Microbiology | 867 | Celsis_Rapid | Bacteria Count (Obeserved) | 0 | cfu | FDA,0 |  | 100 | TM 603 | TM603 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | 7E4OP203 | Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | Cocamidopropyl Dimonium Chloride pHosphate Ri-1059 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Microbiology | 867 | Celsis_Rapid | Bacteria Count (Reported) | 0 | CFU/g | FDA,0 |  | 100 | TM 603 | TM603 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | 7E4OP203 | Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | Cocamidopropyl Dimonium Chloride pHosphate Ri-1059 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Microbiology | 867 | Celsis_Rapid | Celsis Result | Negative ATP (Done) |  |  |  |  | TM 603 | TM603 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | 7E4OP203 | Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | Cocamidopropyl Dimonium Chloride pHosphate Ri-1059 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Microbiology | 867 | Celsis_Rapid | Mold Count (Observed) | 0 | cfu | FDA,0 |  |  | TM 603 | TM603 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | 7E4OP203 | Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | Cocamidopropyl Dimonium Chloride pHosphate Ri-1059 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Microbiology | 867 | Celsis_Rapid | Mold Count (Reported) | 0 | CFU/g | FDA,0 |  |  | TM 603 | TM603 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | 7E4OP203 | Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | Cocamidopropyl Dimonium Chloride pHosphate Ri-1059 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear viscous liquid, free of foreign matter. (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | 7E4OP203 | Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | Cocamidopropyl Dimonium Chloride pHosphate Ri-1059 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Analytical | 792 | Chemical Inspection | Color | 2 |  | FDA,0 |  | 6 | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | 7E4OP203 | Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | Cocamidopropyl Dimonium Chloride pHosphate Ri-1059 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Analytical | 792 | Chemical Inspection | Odor | Similar in character and intensity to standard. Herbal Odor (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Lexfeel Natural | Lexfeel Natural | Chem_NY_Lexfeel Natural | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Lexfeel Natural | Lexfeel Natural | Chem_NY_Lexfeel Natural | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear colorless to pale yellow liquid. (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Lexfeel Natural | Lexfeel Natural | Chem_NY_Lexfeel Natural | NY-Analytical | 792 | Chemical Inspection | Colour | 2 |  | FDA,0 |  | 50 | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Lexfeel Natural | Lexfeel Natural | Chem_NY_Lexfeel Natural | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Lexfeel Natural | Lexfeel Natural | Chem_NY_Lexfeel Natural | NY-Analytical | 826 | Acid Value | Acid Value | 0.03 | mg KOH/g | FDA,1 |  | 0.1 | J&J TMWW78011 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Lexfeel Natural | Lexfeel Natural | Chem_NY_Lexfeel Natural | NY-Analytical | 826 | Acid Value | Acid Value Final | 0.03 | mg KOH/g | FDA,1 |  | 0.1 | J&J TMWW78011 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Lexfeel Natural | Lexfeel Natural | Chem_NY_Lexfeel Natural | NY-Analytical | 390 | Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Lexfeel Natural | Lexfeel Natural | Chem_NY_Lexfeel Natural | NY-Analytical | 801 | % Water | Moisture Content | 0.04 | % | FDA,1 |  | 0.1 | J&J TMWW78017 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Lexfeel Natural | Lexfeel Natural | Chem_NY_Lexfeel Natural | NY-Microbiology | 524 | J&J Fungal | Fungal | 0 | CFU/g | FDA,0 | 0 | 200 | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Lexfeel Natural | Lexfeel Natural | Chem_NY_Lexfeel Natural | NY-Microbiology | 701 | J&J TSA Controls | Negative Control | No Growth TSA |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Lexfeel Natural | Lexfeel Natural | Chem_NY_Lexfeel Natural | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | No Growth TSB |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP230/Chem_NY_Lexfeel Natural | 7E4OP230 | Lexfeel Natural | Lexfeel Natural | Chem_NY_Lexfeel Natural | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | 0 | cfu/mL | FDA,0 | 0 | 2000 | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP601/Chem_102A | 7E4OP601 | Nipaguard IPF | Nipaguard Ipf | Chem_102A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Assay |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP601/Chem_102A | 7E4OP601 | Nipaguard IPF | Nipaguard Ipf | Chem_102A | NY-Analytical | 796 | Assay | Assay | 10.16 | % | FDA,2 | 9 | 11 | J&J TM-002528 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP601/Chem_102A | 7E4OP601 | Nipaguard IPF | Nipaguard Ipf | Chem_102A | NY-Analytical | 796 | Assay | Assay Final | 10.16 | % | FDA,1 | 9 | 11 | J&J TM-002528 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP601/Chem_102A | 7E4OP601 | Nipaguard IPF | Nipaguard Ipf | Chem_102A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear, yellow liquid. (Done) |  |  |  |  | J&J TMWW78010 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP601/Chem_102A | 7E4OP601 | Nipaguard IPF | Nipaguard Ipf | Chem_102A | NY-Analytical | 390 | Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP601/Chem_102A | 7E4OP601 | Nipaguard IPF | Nipaguard Ipf | Chem_102A | NY-Analytical | 801 | % Water | Water Content | 0.4 | % | FDA,2 |  | 1 | J&J TMWW78017 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP806/Chem_58B | 7E4OP806 | Hexylene Glycol | Hexylene Glycol | Chem_58B | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP806/Chem_58B | 7E4OP806 | Hexylene Glycol | Hexylene Glycol | Chem_58B | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear colorless liquid free from suspended materials (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP806/Chem_58B | 7E4OP806 | Hexylene Glycol | Hexylene Glycol | Chem_58B | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP806/Chem_58B | 7E4OP806 | Hexylene Glycol | Hexylene Glycol | Chem_58B | NY-Analytical | 831 | Density @ 20° C | Density | 0.9221 |  | FDA,3 | 0.92 | 0.923 | J&J TMWW78012 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP806/Chem_58B | 7E4OP806 | Hexylene Glycol | Hexylene Glycol | Chem_58B | NY-Analytical | 390 | Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OP806/Chem_58B | 7E4OP806 | Hexylene Glycol | Hexylene Glycol | Chem_58B | NY-Analytical | 801 | % Water | Moisture Content (10 grams of sample) | 0.06 | % | FDA,3 |  | 0.1 | J&J TMWW78017 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | 7E4OQ028 | Ethylhexyl Hydroxystearate | Ethylhexyl Hydroxystearate | Chem_NY_Ethylhexyl Hydroxystearate | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25�C |  | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | 7E4OQ028 | Ethylhexyl Hydroxystearate | Ethylhexyl Hydroxystearate | Chem_NY_Ethylhexyl Hydroxystearate | NY-Analytical | 402 | Specific Gravity at 25�C | Specific Gravity at 25°C | 0.8964 |  | FDA,4 | 0.892 | 0.898 | TM 202 | TM202 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | 7E4OQ028 | Ethylhexyl Hydroxystearate | Ethylhexyl Hydroxystearate | Chem_NY_Ethylhexyl Hydroxystearate | NY-Analytical | 451 | Refractive Index at 25�C | Refractive Index | 1.4555 |  | FDA,4 | 1.4555 | 1.458 | TM 204 | TM204 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | 7E4OQ028 | Ethylhexyl Hydroxystearate | Ethylhexyl Hydroxystearate | Chem_NY_Ethylhexyl Hydroxystearate | NY-Analytical | 390 | Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411 | TM411 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | 7E4OQ028 | Ethylhexyl Hydroxystearate | Ethylhexyl Hydroxystearate | Chem_NY_Ethylhexyl Hydroxystearate | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear to slightly hazy liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | 7E4OQ028 | Ethylhexyl Hydroxystearate | Ethylhexyl Hydroxystearate | Chem_NY_Ethylhexyl Hydroxystearate | NY-Analytical | 792 | Chemical Inspection | Color | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | 7E4OQ028 | Ethylhexyl Hydroxystearate | Ethylhexyl Hydroxystearate | Chem_NY_Ethylhexyl Hydroxystearate | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | 7E4OQ028 | Ethylhexyl Hydroxystearate | Ethylhexyl Hydroxystearate | Chem_NY_Ethylhexyl Hydroxystearate | NY-Analytical | 792 | Chemical Inspection | Odor | Slightly Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7F4BP869/Chem_NY_Citric Acid Anhydrous - USP | 7F4BP869 | Citric Acid Anhydrous - USP | Citric Acid Anhydrous - USP | Chem_NY_Citric Acid Anhydrous - USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA | COA | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4BP869/Chem_NY_Citric Acid Anhydrous - USP | 7F4BP869 | Citric Acid Anhydrous - USP | Citric Acid Anhydrous - USP | Chem_NY_Citric Acid Anhydrous - USP | NY-Analytical | 792 | Chemical Inspection | Appearance | Crystals/crystalline powder (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4BP869/Chem_NY_Citric Acid Anhydrous - USP | 7F4BP869 | Citric Acid Anhydrous - USP | Citric Acid Anhydrous - USP | Chem_NY_Citric Acid Anhydrous - USP | NY-Analytical | 792 | Chemical Inspection | Color | Colorless to white (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4BP869/Chem_NY_Citric Acid Anhydrous - USP | 7F4BP869 | Citric Acid Anhydrous - USP | Citric Acid Anhydrous - USP | Chem_NY_Citric Acid Anhydrous - USP | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4BP869/Chem_NY_Citric Acid Anhydrous - USP | 7F4BP869 | Citric Acid Anhydrous - USP | Citric Acid Anhydrous - USP | Chem_NY_Citric Acid Anhydrous - USP | NY-Analytical | 792 | Chemical Inspection | Odor | Odorless (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4BP869/Chem_NY_Citric Acid Anhydrous - USP | 7F4BP869 | Citric Acid Anhydrous - USP | Citric Acid Anhydrous - USP | Chem_NY_Citric Acid Anhydrous - USP | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | 7F4OA160 | PEG-40 Hydrogenated Castor Oil | PEG-40 Hydrogenated Castor Oil | Chem_NY_PEG-40 Hydrogenated Castor Oil | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | 7F4OA160 | PEG-40 Hydrogenated Castor Oil | PEG-40 Hydrogenated Castor Oil | Chem_NY_PEG-40 Hydrogenated Castor Oil | NY-Analytical | 398 | pH | pH (10% in water) | 6.99 |  | FDA,1 | 6 | 7 | TM 200 | TM200 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | 7F4OA160 | PEG-40 Hydrogenated Castor Oil | PEG-40 Hydrogenated Castor Oil | Chem_NY_PEG-40 Hydrogenated Castor Oil | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | 7F4OA160 | PEG-40 Hydrogenated Castor Oil | PEG-40 Hydrogenated Castor Oil | Chem_NY_PEG-40 Hydrogenated Castor Oil | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear viscous liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | 7F4OA160 | PEG-40 Hydrogenated Castor Oil | PEG-40 Hydrogenated Castor Oil | Chem_NY_PEG-40 Hydrogenated Castor Oil | NY-Analytical | 792 | Chemical Inspection | Color / Appearance | White to yellowish soft or flowing paste. Close Match to Standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | 7F4OA160 | PEG-40 Hydrogenated Castor Oil | PEG-40 Hydrogenated Castor Oil | Chem_NY_PEG-40 Hydrogenated Castor Oil | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | 7F4OA160 | PEG-40 Hydrogenated Castor Oil | PEG-40 Hydrogenated Castor Oil | Chem_NY_PEG-40 Hydrogenated Castor Oil | NY-Analytical | 792 | Chemical Inspection | Odor | Close match to standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
| 7F4OD110 | Tetrasodium Edetate |  | Analytical - Chem_106A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH |  |  |  |  | left_only
| 7F4OD110 | Tetrasodium Edetate |  | Analytical - Chem_106A | NY-Analytical | 398 | pH | pH (1% Solution) | 10.6 |  | FDA,1 | 10.5 | 11.5 | TM 200 | TM200 |  |  |  | left_only
| 7F4OD110 | Tetrasodium Edetate |  | Analytical - Chem_106A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 |  |  |  | left_only
| 7F4OD110 | Tetrasodium Edetate |  | Analytical - Chem_106A | NY-Analytical | 792 | Chemical Inspection | Appearance | Crystalline Powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7F4OD110 | Tetrasodium Edetate |  | Analytical - Chem_106A | NY-Analytical | 792 | Chemical Inspection | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7F4OD110 | Tetrasodium Edetate |  | Analytical - Chem_106A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7F4OD110 | Tetrasodium Edetate |  | Analytical - Chem_106A | NY-Analytical | 792 | Chemical Inspection | Odor | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | CHE ARLASILK PHOSPHOLIPID PTC | CHE Arlasilk pHospholipid Ptc | Chem_NY_Arlasilk Phospholipid PTC | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | color (Gardner) |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | CHE ARLASILK PHOSPHOLIPID PTC | CHE Arlasilk pHospholipid Ptc | Chem_NY_Arlasilk Phospholipid PTC | NY-Analytical | 944 | color (Gardner) | color (Gardner) | 3 |  | FDA,0 |  | 4 | J&J TM701 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | CHE ARLASILK PHOSPHOLIPID PTC | CHE Arlasilk pHospholipid Ptc | Chem_NY_Arlasilk Phospholipid PTC | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear amber liquid. (Done) |  |  |  |  | J&J TMWW78010 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | CHE ARLASILK PHOSPHOLIPID PTC | CHE Arlasilk pHospholipid Ptc | Chem_NY_Arlasilk Phospholipid PTC | NY-Analytical | 792 | Chemical Inspection | Odor | Passes Test (Done) |  |  |  |  | J&J TMWW78010 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | CHE ARLASILK PHOSPHOLIPID PTC | CHE Arlasilk pHospholipid Ptc | Chem_NY_Arlasilk Phospholipid PTC | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | CHE ARLASILK PHOSPHOLIPID PTC | CHE Arlasilk pHospholipid Ptc | Chem_NY_Arlasilk Phospholipid PTC | NY-Analytical | 390 | Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | CHE ARLASILK PHOSPHOLIPID PTC | CHE Arlasilk pHospholipid Ptc | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 509 | J&J_Pathogens | Biggy or SDA | No Growth_SDA |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | CHE ARLASILK PHOSPHOLIPID PTC | CHE Arlasilk pHospholipid Ptc | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 509 | J&J_Pathogens | Cetrimide | No Growth |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | CHE ARLASILK PHOSPHOLIPID PTC | CHE Arlasilk pHospholipid Ptc | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 524 | J&J Fungal | Fungal | 0 | CFU/g | FDA,0 | 0 | 200 | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | CHE ARLASILK PHOSPHOLIPID PTC | CHE Arlasilk pHospholipid Ptc | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 509 | J&J_Pathogens | MacConkey | No Growth |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | CHE ARLASILK PHOSPHOLIPID PTC | CHE Arlasilk pHospholipid Ptc | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 509 | J&J_Pathogens | MSA or VJ | No Growth_MSA |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | CHE ARLASILK PHOSPHOLIPID PTC | CHE Arlasilk pHospholipid Ptc | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 700 | J&J SDA Controls | Negative Control | No Growth SDA |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | CHE ARLASILK PHOSPHOLIPID PTC | CHE Arlasilk pHospholipid Ptc | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | No Growth TSB |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | CHE ARLASILK PHOSPHOLIPID PTC | CHE Arlasilk pHospholipid Ptc | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 509 | J&J_Pathogens | TSA | No Growth |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | CHE ARLASILK PHOSPHOLIPID PTC | CHE Arlasilk pHospholipid Ptc | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | 0 | cfu/mL | FDA,0 | 0 | 2000 | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | 7F4OP027 | CHE ARLASILK PHOSPHOLIPID PTC | CHE Arlasilk pHospholipid Ptc | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 517 | J&J_Enrichment | TSB | Turbid |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OP601/Chem_109A | 7F4OP601 | CHE ISONONYL ISONONANOATE | CHE Isononyl Isononanoate | Chem_109A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Odor |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OP601/MASTER - Microbiology - J&J Chemical Testing) | 7F4OP601 | CHE ISONONYL ISONONANOATE | CHE Isononyl Isononanoate | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 396 | Odor | Odor | Passes Test (Done) |  |  |  |  | J&J TM-003188 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OP601/Chem_109A | 7F4OP601 | CHE ISONONYL ISONONANOATE | CHE Isononyl Isononanoate | Chem_109A | NY-Analytical | 396 | Odor | Odor | Passes Test (Done) |  |  |  |  | J&J TM-003188 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OP601/MASTER - Microbiology - J&J Chemical Testing) | 7F4OP601 | CHE ISONONYL ISONONANOATE | CHE Isononyl Isononanoate | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear colorless liquid with oily consistency (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OP601/Chem_109A | 7F4OP601 | CHE ISONONYL ISONONANOATE | CHE Isononyl Isononanoate | Chem_109A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear colorless liquid with oily consistency (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OP601/Chem_109A | 7F4OP601 | CHE ISONONYL ISONONANOATE | CHE Isononyl Isononanoate | Chem_109A | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OP601/Chem_109A | 7F4OP601 | CHE ISONONYL ISONONANOATE | CHE Isononyl Isononanoate | Chem_109A | NY-Analytical | 390 | Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OP806/Chem_NY_Phenoxyethanol_7F4OP806 | 7F4OP806 | Phenoxyethanol | Phenoxyethanol | Chem_NY_Phenoxyethanol_7F4OP806 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OP806/Chem_NY_Phenoxyethanol_7F4OP806 | 7F4OP806 | Phenoxyethanol | Phenoxyethanol | Chem_NY_Phenoxyethanol_7F4OP806 | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear, thin colorless liquid (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OP806/Chem_NY_Phenoxyethanol_7F4OP806 | 7F4OP806 | Phenoxyethanol | Phenoxyethanol | Chem_NY_Phenoxyethanol_7F4OP806 | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OP806/Chem_NY_Phenoxyethanol_7F4OP806 | 7F4OP806 | Phenoxyethanol | Phenoxyethanol | Chem_NY_Phenoxyethanol_7F4OP806 | NY-Analytical | 390 | Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OP806/Chem_NY_Phenoxyethanol_7F4OP806 | 7F4OP806 | Phenoxyethanol | Phenoxyethanol | Chem_NY_Phenoxyethanol_7F4OP806 | NY-Analytical | 905 | Phenol | Phenol | 21.5 | ppm | FDA,0 |  | 25 | USP/NF | USP_NF | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OP806/Chem_NY_Phenoxyethanol_7F4OP806 | 7F4OP806 | Phenoxyethanol | Phenoxyethanol | Chem_NY_Phenoxyethanol_7F4OP806 | NY-Analytical | 986 | Phenoxyethanol Content | Phenoxyethanol Assay | 100.9 | % | FDA,1 | 98 | 102 | USP/NF | USP_NF | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OP806/Chem_NY_Phenoxyethanol_7F4OP806 | 7F4OP806 | Phenoxyethanol | Phenoxyethanol | Chem_NY_Phenoxyethanol_7F4OP806 | NY-Analytical | 432 | Specific Gravity at 20�C | Specific Gravity at 20°C | 1.1095 |  | FDA,3 | 1.105 | 1.11 | USP/NF | USP_NF | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | 7F4OQ028 | Octyldodecyl Stearoyl Stearate | Octyldodecyl Stearoyl Stearate | Chem_NY_Octyldodecyl Stearoyl Stearate | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25�C |  | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | 7F4OQ028 | Octyldodecyl Stearoyl Stearate | Octyldodecyl Stearoyl Stearate | Chem_NY_Octyldodecyl Stearoyl Stearate | NY-Analytical | 402 | Specific Gravity at 25�C | Specific Gravity at 25°C | 0.8747 |  | FDA,4 | 0.865 | 0.885 | TM 202 | TM202 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | 7F4OQ028 | Octyldodecyl Stearoyl Stearate | Octyldodecyl Stearoyl Stearate | Chem_NY_Octyldodecyl Stearoyl Stearate | NY-Analytical | 451 | Refractive Index at 25�C | Refractive Index | 1.4585 |  | FDA,4 | 1.455 | 1.46 | TM 204 | TM204 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | 7F4OQ028 | Octyldodecyl Stearoyl Stearate | Octyldodecyl Stearoyl Stearate | Chem_NY_Octyldodecyl Stearoyl Stearate | NY-Analytical | 390 | Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411 | TM411 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | 7F4OQ028 | Octyldodecyl Stearoyl Stearate | Octyldodecyl Stearoyl Stearate | Chem_NY_Octyldodecyl Stearoyl Stearate | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | 7F4OQ028 | Octyldodecyl Stearoyl Stearate | Octyldodecyl Stearoyl Stearate | Chem_NY_Octyldodecyl Stearoyl Stearate | NY-Analytical | 792 | Chemical Inspection | Color | Light Yellow Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | 7F4OQ028 | Octyldodecyl Stearoyl Stearate | Octyldodecyl Stearoyl Stearate | Chem_NY_Octyldodecyl Stearoyl Stearate | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | 7F4OQ028 | Octyldodecyl Stearoyl Stearate | Octyldodecyl Stearoyl Stearate | Chem_NY_Octyldodecyl Stearoyl Stearate | NY-Analytical | 792 | Chemical Inspection | Odor | Mild Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | 7G4OP027 | SD Alcohol 40-B 200 Proof | SD Alcohol 40-B 200 Proof | Chem_NY_SD Alcohol 40-B 200 Proof_ADM | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis | COA | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | 7G4OP027 | SD Alcohol 40-B 200 Proof | SD Alcohol 40-B 200 Proof | Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | [5 empty] | Identity of Denatonium Benzoate |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | 7G4OP027 | SD Alcohol 40-B 200 Proof | SD Alcohol 40-B 200 Proof | Chem_NY_SD Alcohol 40-B 200 Proof_ADM | NY-Analytical | 824 | Identity of Denatonium Benzoate | Identification | Passes test with 5cm Quartz (Done) |  |  |  |  | J&J TM4028 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | 7G4OP027 | SD Alcohol 40-B 200 Proof | SD Alcohol 40-B 200 Proof | Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | NY-Analytical | 824 | Identity of Denatonium Benzoate | Identification | [5 empty] | J&J TM4028 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | 7G4OP027 | SD Alcohol 40-B 200 Proof | SD Alcohol 40-B 200 Proof | Chem_NY_SD Alcohol 40-B 200 Proof_ADM | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear colorless liquid. (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | 7G4OP027 | SD Alcohol 40-B 200 Proof | SD Alcohol 40-B 200 Proof | Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | NY-Analytical | 792 | Chemical Inspection | Appearance | [5 empty] | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | 7G4OP027 | SD Alcohol 40-B 200 Proof | SD Alcohol 40-B 200 Proof | Chem_NY_SD Alcohol 40-B 200 Proof_ADM | NY-Analytical | 792 | Chemical Inspection | Odor | Passes Test (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | 7G4OP027 | SD Alcohol 40-B 200 Proof | SD Alcohol 40-B 200 Proof | Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | NY-Analytical | 792 | Chemical Inspection | Odor | [5 empty] | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | 7G4OP027 | SD Alcohol 40-B 200 Proof | SD Alcohol 40-B 200 Proof | Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | [5 empty] | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | 7G4OP027 | SD Alcohol 40-B 200 Proof | SD Alcohol 40-B 200 Proof | Chem_NY_SD Alcohol 40-B 200 Proof_ADM | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | 7G4OP027 | SD Alcohol 40-B 200 Proof | SD Alcohol 40-B 200 Proof | Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | NY-Analytical | 432 | Specific Gravity at 20�C | Specific Gravity at 20°C |  |  | FDA,4 | 0.7905 | 0.792 | J&J TMWW78012 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | 7G4OP027 | SD Alcohol 40-B 200 Proof | SD Alcohol 40-B 200 Proof | Chem_NY_SD Alcohol 40-B 200 Proof_ADM | NY-Analytical | 432 | Specific Gravity at 20�C | Specific Gravity at 20°C |  |  | FDA,4 | 0.7905 | 0.792 | J&J TMWW78012 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | 7G4OP027 | SD Alcohol 40-B 200 Proof | SD Alcohol 40-B 200 Proof | Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | NY-Analytical | 822 | Specific Gravity @ 15.56°C/ 60° F | Specific Gravity | 0.7942 |  | FDA,4 |  | 0.7952 | J&J TMWW78012 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | 7G4OP027 | SD Alcohol 40-B 200 Proof | SD Alcohol 40-B 200 Proof | Chem_NY_SD Alcohol 40-B 200 Proof_ADM | NY-Analytical | 822 | Specific Gravity @ 15.56°C/ 60° F | Specific Gravity | 0.7941 |  | FDA,4 |  | 0.7952 | J&J TMWW78012 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | 7G4OP027 | SD Alcohol 40-B 200 Proof | SD Alcohol 40-B 200 Proof | Chem_NY_SD Alcohol 40-B 200 Proof_ADM | NY-Analytical | 390 | Identification | IR Identification | [5 empty] | J&J TMWW78013 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | 7G4OP027 | SD Alcohol 40-B 200 Proof | SD Alcohol 40-B 200 Proof | Chem_NY_SD Alcohol 40-B 200 Proof_Equistar_ MGP_ Grain | NY-Analytical | 390 | Identification | IR Identification | [5 empty] | J&J TMWW78013 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7G4OP308/Chem_111A | 7G4OP308 | Perfume Meadow Mist RQ-1222 | Perfume Meadow Mist Rq-1222 | Chem_111A | NY-Analytical | 792 | Chemical Inspection | Odor | Matches |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7G4OP308/Chem_111A | 7G4OP308 | Perfume Meadow Mist RQ-1222 | Perfume Meadow Mist Rq-1222 | Chem_111A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7G4OP308/Chem_111A | 7G4OP308 | Perfume Meadow Mist RQ-1222 | Perfume Meadow Mist Rq-1222 | Chem_111A | NY-Analytical | 792 | Chemical Inspection | Appearance | Colorless to Light Yellow Liquid, Free of Particulate Matter (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7G4OP308/Chem_111A | 7G4OP308 | Perfume Meadow Mist RQ-1222 | Perfume Meadow Mist Rq-1222 | Chem_111A | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7G4OP308/Chem_111A | 7G4OP308 | Perfume Meadow Mist RQ-1222 | Perfume Meadow Mist Rq-1222 | Chem_111A | NY-Analytical | 402 | Specific Gravity at 25�C | Specific Gravity at 25°C | 0.9464 |  | FDA,3 | 0.942 | 0.952 | J&J TMWW78012 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7G4OP308/Chem_111A | 7G4OP308 | Perfume Meadow Mist RQ-1222 | Perfume Meadow Mist Rq-1222 | Chem_111A | NY-Analytical | 390 | Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7G4OP308/Chem_111A | 7G4OP308 | Perfume Meadow Mist RQ-1222 | Perfume Meadow Mist Rq-1222 | Chem_111A | NY-Analytical | 451 | Refractive Index at 25�C | Refractive Index | 1.4632 |  | FDA,3 | 1.463 | 1.473 | J&J TMWW78015 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7G4OP601/MASTER - Microbiology - J&J Chemical Testing) | 7G4OP601 | CHE DERMOL ISP J & J | Dermol ISP | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 828 | Iodine Value | Iodine Value | 0.6 | g I2/100g | FDA,1 |  | 2 | J&J TM-003029 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7G4OP601/MASTER - Microbiology - J&J Chemical Testing) | 7G4OP601 | CHE DERMOL ISP J & J | Dermol ISP | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 828 | Iodine Value | Iodine Value Final | 0.6 | g I2/100g | FDA,1 |  | 2 | J&J TM-003029 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7G4OP601/MASTER - Microbiology - J&J Chemical Testing) | 7G4OP601 | CHE DERMOL ISP J & J | Dermol ISP | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 792 | Chemical Inspection | Color | 1 |  | FDA,1 |  | 1 | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7G4OP601/MASTER - Microbiology - J&J Chemical Testing) | 7G4OP601 | CHE DERMOL ISP J & J | Dermol ISP | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 826 | Acid Value | Acid Value Final | 0.1 | mg KOH/g | FDA,1 |  | 2 | J&J TMWW78011 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7G4OP601/MASTER - Microbiology - J&J Chemical Testing) | 7G4OP601 | CHE DERMOL ISP J & J | Dermol ISP | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 451 | Refractive Index at 25�C | Refractive Index | 1.451 |  | FDA,4 | 1.451 | 1.453 | J&J TMWW78015 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7G4OP601/MASTER - Microbiology - J&J Chemical Testing) | 7G4OP601 | CHE DERMOL ISP J & J | Dermol ISP | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 810 | Saponification Value | Saponification Value | 108.1 | mg KOH/g | FDA,1 | 100 | 115 | J&J TMWW78016 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7G4OP601/MASTER - Microbiology - J&J Chemical Testing) | 7G4OP601 | CHE DERMOL ISP J & J | Dermol ISP | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 810 | Saponification Value | Saponification Value Final | 108.1 | mg KOH/g | FDA,1 | 100 | 115 | J&J TMWW78016 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7G4OQ028/Chem_Aminomethylpropanol | 7G4OQ028 | Aminomethylpropanol | Aminomethylpropanol | Chem_Aminomethylpropanol | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7G4OQ028/Chem_Aminomethylpropanol | 7G4OQ028 | Aminomethylpropanol | Aminomethylpropanol | Chem_Aminomethylpropanol | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7G4OQ028/Chem_Aminomethylpropanol | 7G4OQ028 | Aminomethylpropanol | Aminomethylpropanol | Chem_Aminomethylpropanol | NY-Analytical | 792 | Chemical Inspection | Appearance | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7G4OQ028/Chem_Aminomethylpropanol | 7G4OQ028 | Aminomethylpropanol | Aminomethylpropanol | Chem_Aminomethylpropanol | NY-Analytical | 792 | Chemical Inspection | Color | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7G4OQ028/Chem_Aminomethylpropanol | 7G4OQ028 | Aminomethylpropanol | Aminomethylpropanol | Chem_Aminomethylpropanol | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7G4OQ028/Chem_Aminomethylpropanol | 7G4OQ028 | Aminomethylpropanol | Aminomethylpropanol | Chem_Aminomethylpropanol | NY-Analytical | 792 | Chemical Inspection | Odor | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Feverfew CMP-02A | Feverfew Cmp-02A | Chem_NY_Feverfew CMP-02A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Feverfew CMP-02A | Feverfew Cmp-02A | Chem_NY_Feverfew CMP-02A | NY-Analytical | 792 | Chemical Inspection | Appearance | Red Brown Liquid (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Feverfew CMP-02A | Feverfew Cmp-02A | Chem_NY_Feverfew CMP-02A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Feverfew CMP-02A | Feverfew Cmp-02A | Chem_NY_Feverfew CMP-02A | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Feverfew CMP-02A | Feverfew Cmp-02A | Chem_NY_Feverfew CMP-02A | NY-Analytical | 390 | Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Feverfew CMP-02A | Feverfew Cmp-02A | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 509 | J&J_Pathogens | Biggy or SDA | No Growth_SDA |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Feverfew CMP-02A | Feverfew Cmp-02A | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 509 | J&J_Pathogens | Cetrimide | No Growth |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Feverfew CMP-02A | Feverfew Cmp-02A | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 524 | J&J Fungal | Fungal | 0 | CFU/g | FDA,0 | 0 | 200 | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Feverfew CMP-02A | Feverfew Cmp-02A | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 509 | J&J_Pathogens | MacConkey | No Growth |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Feverfew CMP-02A | Feverfew Cmp-02A | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 509 | J&J_Pathogens | MSA or VJ | No Growth_MSA |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Feverfew CMP-02A | Feverfew Cmp-02A | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 700 | J&J SDA Controls | Negative Control | No Growth SDA |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Feverfew CMP-02A | Feverfew Cmp-02A | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | No Growth TSB |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Feverfew CMP-02A | Feverfew Cmp-02A | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 509 | J&J_Pathogens | TSA | No Growth |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Feverfew CMP-02A | Feverfew Cmp-02A | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | 0 | cfu/mL | FDA,0 | 0 | 200 | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP308/Chem_NY_Feverfew CMP-02A | 7H4OP308 | Feverfew CMP-02A | Feverfew Cmp-02A | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 517 | J&J_Enrichment | TSB | Not Turbid |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP601/Chem_NY_DUB PTO | 7H4OP601 | CHE DUB PTO | CHE Dub Pto | Chem_NY_DUB PTO | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP601/Chem_NY_DUB PTO | 7H4OP601 | CHE DUB PTO | CHE Dub Pto | Chem_NY_DUB PTO | NY-Analytical | 792 | Chemical Inspection | Appearance | Colorless to pale yellow liquid (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP601/Chem_NY_DUB PTO | 7H4OP601 | CHE DUB PTO | CHE Dub Pto | Chem_NY_DUB PTO | NY-Analytical | 792 | Chemical Inspection | Colour | 3 |  | FDA,0 |  | 60 | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP601/Chem_NY_DUB PTO | 7H4OP601 | CHE DUB PTO | CHE Dub Pto | Chem_NY_DUB PTO | NY-Analytical | 792 | Chemical Inspection | Odor | Passes Test (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP601/Chem_NY_DUB PTO | 7H4OP601 | CHE DUB PTO | CHE Dub Pto | Chem_NY_DUB PTO | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP601/Chem_NY_DUB PTO | 7H4OP601 | CHE DUB PTO | CHE Dub Pto | Chem_NY_DUB PTO | NY-Analytical | 390 | Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP601/Chem_NY_DUB PTO | 7H4OP601 | CHE DUB PTO | CHE Dub Pto | Chem_NY_DUB PTO | NY-Microbiology | 524 | J&J Fungal | Fungal |  | CFU/g | FDA,0 | 0 | 200 | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP601/J&J Chemical CHE DUB PTO | 7H4OP601 | CHE DUB PTO | CHE Dub Pto | J&J Chemical CHE DUB PTO | NY-Microbiology | 524 | J&J Fungal | Fungal |  | CFU/g | FDA,0 | 0 | 200 | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP601/Chem_NY_DUB PTO | 7H4OP601 | CHE DUB PTO | CHE Dub Pto | Chem_NY_DUB PTO | NY-Microbiology | 701 | J&J TSA Controls | Negative Control | [5 empty] | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP601/J&J Chemical CHE DUB PTO | 7H4OP601 | CHE DUB PTO | CHE Dub Pto | J&J Chemical CHE DUB PTO | NY-Microbiology | 701 | J&J TSA Controls | Negative Control | [5 empty] | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP601/Chem_NY_DUB PTO | 7H4OP601 | CHE DUB PTO | CHE Dub Pto | Chem_NY_DUB PTO | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | No Growth TSB |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP601/J&J Chemical CHE DUB PTO | 7H4OP601 | CHE DUB PTO | CHE Dub Pto | J&J Chemical CHE DUB PTO | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | [5 empty] | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP601/Chem_NY_DUB PTO | 7H4OP601 | CHE DUB PTO | CHE Dub Pto | Chem_NY_DUB PTO | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | 0 | cfu/mL | FDA,0 | 0 | 2000 | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP601/J&J Chemical CHE DUB PTO | 7H4OP601 | CHE DUB PTO | CHE Dub Pto | J&J Chemical CHE DUB PTO | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ |  | cfu/mL | FDA,0 | 0 | 2000 | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP806/Chem_NY_Kelp & Sea Salt Fragrance | 7H4OP806 | Kelp & Sea Salt Fragrance | Kelp & Sea Salt Fragrance | Chem_NY_Kelp & Sea Salt Fragrance | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | GC Profile |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP806/Chem_NY_Kelp & Sea Salt Fragrance | 7H4OP806 | Kelp & Sea Salt Fragrance | Kelp & Sea Salt Fragrance | Chem_NY_Kelp & Sea Salt Fragrance | NY-Analytical | 839 | GC Profile | GC Profile | To match standard. Not required when performing IR test. (Done) |  |  |  |  | J&J TM-003030 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP806/Chem_NY_Kelp & Sea Salt Fragrance | 7H4OP806 | Kelp & Sea Salt Fragrance | Kelp & Sea Salt Fragrance | Chem_NY_Kelp & Sea Salt Fragrance | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear yellow to dark yellow liquid (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP806/Chem_NY_Kelp & Sea Salt Fragrance | 7H4OP806 | Kelp & Sea Salt Fragrance | Kelp & Sea Salt Fragrance | Chem_NY_Kelp & Sea Salt Fragrance | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP806/Chem_NY_Kelp & Sea Salt Fragrance | 7H4OP806 | Kelp & Sea Salt Fragrance | Kelp & Sea Salt Fragrance | Chem_NY_Kelp & Sea Salt Fragrance | NY-Analytical | 432 | Specific Gravity at 20�C | Specific Gravity at 20°C | 0.9621 |  | FDA,3 | 0.957 | 0.967 | J&J TMWW78012 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP806/Chem_NY_Kelp & Sea Salt Fragrance | 7H4OP806 | Kelp & Sea Salt Fragrance | Kelp & Sea Salt Fragrance | Chem_NY_Kelp & Sea Salt Fragrance | NY-Analytical | 390 | Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7H4OP806/Chem_NY_Kelp & Sea Salt Fragrance | 7H4OP806 | Kelp & Sea Salt Fragrance | Kelp & Sea Salt Fragrance | Chem_NY_Kelp & Sea Salt Fragrance | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | 1.458 |  | FDA,4 | 1.456 | 1.466 | J&J TMWW78015 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7J4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7J4OP027 | Actiphyte of Cucumber GL | ACt. iphyte Of Cucumber Gl | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 792 | Chemical Inspection | Solubility | Passes Test Soluble in any proportion of water (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7J4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7J4OP027 | Actiphyte of Cucumber GL | ACt. iphyte Of Cucumber Gl | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 447 | pH as is | pH | 5.59 |  | FDA,1 | 4 | 6.5 | TM WW78014 | TMWW78014 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7K4OP027/Chem_116A | 7K4OP027 | Fragrance 5831-AT 01 GN MOD2 IF | Fragrance 5831-At 01 Gn Mod2 If | Chem_116A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7K4OP027/Chem_116A | 7K4OP027 | Fragrance 5831-AT 01 GN MOD2 IF | Fragrance 5831-At 01 Gn Mod2 If | Chem_116A | NY-Analytical | 792 | Chemical Inspection | Appearance | Pale yellow to yellow liquid (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7K4OP027/Chem_116A | 7K4OP027 | Fragrance 5831-AT 01 GN MOD2 IF | Fragrance 5831-At 01 Gn Mod2 If | Chem_116A | NY-Analytical | 792 | Chemical Inspection | Odor (CoA) | Certified on CoA as Meets Odor Specification or of Characteristic Odor. (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7K4OP027/Chem_116A | 7K4OP027 | Fragrance 5831-AT 01 GN MOD2 IF | Fragrance 5831-At 01 Gn Mod2 If | Chem_116A | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7K4OP027/Chem_116A | 7K4OP027 | Fragrance 5831-AT 01 GN MOD2 IF | Fragrance 5831-At 01 Gn Mod2 If | Chem_116A | NY-Analytical | 390 | Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Menthol USP | Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Menthol USP | Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 792 | Chemical Inspection | Appearance | Colorless to white needle crystal or granule (Done) |  |  |  |  | JJTMWW78010 / JJTM7252 | JJTMWW78010__JJTM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Menthol USP | Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | JJTMWW78010 / JJTM7252 | JJTMWW78010__JJTM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Menthol USP | Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 809 | Melting Range | Melting Range - End | 42.5 | C | FDA,0 | 41 | 44 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Menthol USP | Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 809 | Melting Range | Melting Range - Start | 41.2 | C | FDA,0 | 41 | 44 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Menthol USP | Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 797 | Nonvolatile Residue | Nonvolatile Residue | 0.02 | % w/v | FDA,2 |  | 0.05 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Menthol USP | Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 863 | Specific Rotation | Specific Rotation | -49.9 | ° | FDA,0 | -51 | -45 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Menthol USP | Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 796 | Assay | Assay | 100.4 | % | FDA,1 | 98 | 102 | USP/NF | USP_NF | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Menthol USP | Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 796 | Assay | Assay Final | 100.4 | % | FDA,1 | 98 | 102 | USP/NF | USP_NF | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Menthol USP | Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 455 | Identification A | Identification A | Passes Test (Done) |  |  |  |  | USP/NF | USP_NF | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Menthol USP | Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 456 | Identification B | Identification B | Passes Test (Done) |  |  |  |  | USP/NF | USP_NF | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Menthol USP | Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 964 | Individual Impurity | Iso-Menthone | 0.3 | % | FDA,1 |  | 0.3 | USP/NF | USP_NF | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Menthol USP | Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 964 | Individual Impurity | Iso-Pulegol | 0.14 | % | FDA,1 |  | 0.3 | USP/NF | USP_NF | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Menthol USP | Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 964 | Individual Impurity | Menthone | 0.04 | % | FDA,1 |  | 0.3 | USP/NF | USP_NF | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Menthol USP | Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 964 | Individual Impurity | Neo-Menthol | 0.18 | % | FDA,1 |  | 0.3 | USP/NF | USP_NF | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Menthol USP | Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 964 | Individual Impurity | Pulegone | 0.08 | % | FDA,1 |  | 0.3 | USP/NF | USP_NF | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Menthol USP | Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 965 | Total Impurities | Total Impurities | 0.53 | % | FDA,1 |  | 2 | USP/NF | USP_NF | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP027/Chem_NY_Menthol USP | 7L4OP027 | Menthol USP | Menthol USP | Chem_NY_Menthol USP | NY-Analytical | 964 | Individual Impurity | Unknown impurities | 0.05 | % | FDA,1 |  | 0.1 | USP/NF | USP_NF | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Glycerin - RI-0420 | Glycerin - Ri-0420 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 801 | % Water | Water | 0.09 | % | FDA,1 |  | 5 | APAM 6.018 | APAM_6_018 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Glycerin - RI-0420 | Glycerin - Ri-0420 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Glycerin - RI-0420 | Glycerin - Ri-0420 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear, syrupy liquid, free of foreign matter. (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Glycerin - RI-0420 | Glycerin - Ri-0420 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 792 | Chemical Inspection | Color | White (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Glycerin - RI-0420 | Glycerin - Ri-0420 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 792 | Chemical Inspection | Odor | Similar in character and intensity to standard, practically odorless (Done) |  |  |  |  | See Analyte for Method # | See_Analyte_for_Method | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Glycerin - RI-0420 | Glycerin - Ri-0420 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 402 | Specific Gravity at 25�C | Specific Gravity at 25°C | 1.273 |  | FDA,3 | 1.249 |  | TM 202 | TM202 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Glycerin - RI-0420 | Glycerin - Ri-0420 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Glycerin - RI-0420 | Glycerin - Ri-0420 | Chem_NY_Glycerin - RI-0420 | NY-Microbiology | 867 | Celsis_Rapid | Bacteria Count (Obeserved) | 0 | cfu | FDA,0 |  | 100 | TM 603 | TM603 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Glycerin - RI-0420 | Glycerin - Ri-0420 | Chem_NY_Glycerin - RI-0420 | NY-Microbiology | 867 | Celsis_Rapid | Bacteria Count (Reported) | 0 | CFU/g | FDA,0 |  | 100 | TM 603 | TM603 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Glycerin - RI-0420 | Glycerin - Ri-0420 | Chem_NY_Glycerin - RI-0420 | NY-Microbiology | 867 | Celsis_Rapid | Celsis Result | Negative ATP (Done) |  |  |  |  | TM 603 | TM603 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Glycerin - RI-0420 | Glycerin - Ri-0420 | Chem_NY_Glycerin - RI-0420 | NY-Microbiology | 867 | Celsis_Rapid | Mold Count (Observed) | 0 | cfu | FDA,0 |  |  | TM 603 | TM603 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Glycerin - RI-0420 | Glycerin - Ri-0420 | Chem_NY_Glycerin - RI-0420 | NY-Microbiology | 867 | Celsis_Rapid | Mold Count (Reported) | 0 | CFU/g | FDA,0 |  |  | TM 603 | TM603 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Glycerin - RI-0420 | Glycerin - Ri-0420 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 825 | Limit of Ethylene Glycol and Diethylene Glycol | Diethylene Glycol | 0 | % | FDA,1 |  | 0.1 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Glycerin - RI-0420 | Glycerin - Ri-0420 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 825 | Limit of Ethylene Glycol and Diethylene Glycol | Ethylene Glycol | 0 | % | FDA,1 |  | 0.1 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Glycerin - RI-0420 | Glycerin - Ri-0420 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 825 | Limit of Ethylene Glycol and Diethylene Glycol | Individual Impurity | 0.01 | % | FDA,1 |  | 0.1 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP203/Chem_NY_Glycerin - RI-0420 | 7L4OP203 | Glycerin - RI-0420 | Glycerin - Ri-0420 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 825 | Limit of Ethylene Glycol and Diethylene Glycol | Total Impurities | 0.02 | % | FDA,0 |  | 1 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | 7L4OP485 | Polysorbate 20 NF | Polysorbate 20 Nf | Chem_NY_Polysorbate 20_7L4OP485 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Odor |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | 7L4OP485 | Polysorbate 20 NF | Polysorbate 20 Nf | Chem_NY_Polysorbate 20_7L4OP485 | NY-Analytical | 396 | Odor | Odor | Passes Test (Done) |  |  |  |  | J&J TM-003188 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | 7L4OP485 | Polysorbate 20 NF | Polysorbate 20 Nf | Chem_NY_Polysorbate 20_7L4OP485 | NY-Analytical | 792 | Chemical Inspection | Appearance | A pale yellow to an amber colored liquid, free of foreign matter. (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | 7L4OP485 | Polysorbate 20 NF | Polysorbate 20 Nf | Chem_NY_Polysorbate 20_7L4OP485 | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | 7L4OP485 | Polysorbate 20 NF | Polysorbate 20 Nf | Chem_NY_Polysorbate 20_7L4OP485 | NY-Analytical | 390 | Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | 7L4OP485 | Polysorbate 20 NF | Polysorbate 20 Nf | Chem_NY_Polysorbate 20_7L4OP485 | NY-Microbiology | 524 | J&J Fungal | Fungal | 0 | CFU/g | FDA,0 | 0 | 200 | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | 7L4OP485 | Polysorbate 20 NF | Polysorbate 20 Nf | Chem_NY_Polysorbate 20_7L4OP485 | NY-Microbiology | 723 | J&J TSB Controls - TSA | Negative Control | No Growth TSB |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | 7L4OP485 | Polysorbate 20 NF | Polysorbate 20 Nf | Chem_NY_Polysorbate 20_7L4OP485 | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | No Growth TSB |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | 7L4OP485 | Polysorbate 20 NF | Polysorbate 20 Nf | Chem_NY_Polysorbate 20_7L4OP485 | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | 0 | CFU/g | FDA,0 | 0 | 2000 | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Benzalkonium Chloride 50% - USP | Benzalkonium Chloride 50% - USP | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Benzalkonium Chloride 50% - USP | Benzalkonium Chloride 50% - USP | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear colorless liquid. (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Benzalkonium Chloride 50% - USP | Benzalkonium Chloride 50% - USP | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Benzalkonium Chloride 50% - USP | Benzalkonium Chloride 50% - USP | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 796 | Assay | Assay | 48.2 | % | FDA,1 | 47.5 | 52.5 | USP/NF | USP_NF | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Benzalkonium Chloride 50% - USP | Benzalkonium Chloride 50% - USP | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 796 | Assay | Assay Final | 48.2 | % | FDA,1 | 47.5 | 52.5 | USP/NF | USP_NF | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Benzalkonium Chloride 50% - USP | Benzalkonium Chloride 50% - USP | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 846 | Ratio of Alkyl Compounds | C12 + C14 Content of Total | 95.7 | % | FDA,1 | 70 |  | USP/NF | USP_NF | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Benzalkonium Chloride 50% - USP | Benzalkonium Chloride 50% - USP | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 846 | Ratio of Alkyl Compounds | C12 Content of Total | 56.6 | % | FDA,1 | 40 |  | USP/NF | USP_NF | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Benzalkonium Chloride 50% - USP | Benzalkonium Chloride 50% - USP | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 846 | Ratio of Alkyl Compounds | C14 Content of Total | 39.1 | % | FDA,1 | 20 |  | USP/NF | USP_NF | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Benzalkonium Chloride 50% - USP | Benzalkonium Chloride 50% - USP | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 390 | Identification | Identification A | Passes test. (Done) |  |  |  |  | USP/NF | USP_NF | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Benzalkonium Chloride 50% - USP | Benzalkonium Chloride 50% - USP | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 390 | Identification | Identification B | Passes test. (Done) |  |  |  |  | USP/NF | USP_NF | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Benzalkonium Chloride 50% - USP | Benzalkonium Chloride 50% - USP | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 390 | Identification | Identification C | Passes test. (Done) |  |  |  |  | USP/NF | USP_NF | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | 7M4OP027 | Benzalkonium Chloride 50% - USP | Benzalkonium Chloride 50% - USP | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 390 | Identification | Identification D | Passes test. (Done) |  |  |  |  | USP/NF | USP_NF | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4AA580/Chem_NY_Ethylene Glycol Monohexyl Ether | 7O4AA580 | Ethylene Glycol Monohexyl ether | Ethylene Glycol Monohexyl Ether | Chem_NY_Ethylene Glycol Monohexyl Ether | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4AA580/Chem_NY_Ethylene Glycol Monohexyl Ether | 7O4AA580 | Ethylene Glycol Monohexyl ether | Ethylene Glycol Monohexyl Ether | Chem_NY_Ethylene Glycol Monohexyl Ether | NY-Analytical | 390 | Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4AA580/Chem_NY_Ethylene Glycol Monohexyl Ether | 7O4AA580 | Ethylene Glycol Monohexyl ether | Ethylene Glycol Monohexyl Ether | Chem_NY_Ethylene Glycol Monohexyl Ether | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear, Colorless, Viscous Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4AA580/Chem_NY_Ethylene Glycol Monohexyl Ether | 7O4AA580 | Ethylene Glycol Monohexyl ether | Ethylene Glycol Monohexyl Ether | Chem_NY_Ethylene Glycol Monohexyl Ether | NY-Analytical | 792 | Chemical Inspection | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4AA580/Chem_NY_Ethylene Glycol Monohexyl Ether | 7O4AA580 | Ethylene Glycol Monohexyl ether | Ethylene Glycol Monohexyl Ether | Chem_NY_Ethylene Glycol Monohexyl Ether | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4AA580/Chem_NY_Ethylene Glycol Monohexyl Ether | 7O4AA580 | Ethylene Glycol Monohexyl ether | Ethylene Glycol Monohexyl Ether | Chem_NY_Ethylene Glycol Monohexyl Ether | NY-Analytical | 792 | Chemical Inspection | Odor | Pungent (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4AQ084/Chem_162A | 7O4AQ084 | Integra 44 | Integra 44 | Chem_162A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4AQ084/Chem_162A | 7O4AQ084 | Integra 44 | Integra 44 | Chem_162A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4AQ084/Chem_162A | 7O4AQ084 | Integra 44 | Integra 44 | Chem_162A | NY-Analytical | 792 | Chemical Inspection | Color | Clear/Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4AQ084/Chem_162A | 7O4AQ084 | Integra 44 | Integra 44 | Chem_162A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4AQ084/Chem_162A | 7O4AQ084 | Integra 44 | Integra 44 | Chem_162A | NY-Analytical | 792 | Chemical Inspection | Odor | Light Amine (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4BJ131/Chem_NY_Acticide CBM-2 | 7O4BJ131 | Acticide CBM-2 | ACt. icide Cbm-2 | Chem_NY_Acticide CBM-2 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4BJ131/Chem_NY_Acticide CBM-2 | 7O4BJ131 | Acticide CBM-2 | ACt. icide Cbm-2 | Chem_NY_Acticide CBM-2 | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4BJ131/Chem_NY_Acticide CBM-2 | 7O4BJ131 | Acticide CBM-2 | ACt. icide Cbm-2 | Chem_NY_Acticide CBM-2 | NY-Analytical | 792 | Chemical Inspection | Appearance | Off-white - greenish dispersion (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4BJ131/Chem_NY_Acticide CBM-2 | 7O4BJ131 | Acticide CBM-2 | ACt. icide Cbm-2 | Chem_NY_Acticide CBM-2 | NY-Analytical | 792 | Chemical Inspection | Color | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4BJ131/Chem_NY_Acticide CBM-2 | 7O4BJ131 | Acticide CBM-2 | ACt. icide Cbm-2 | Chem_NY_Acticide CBM-2 | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4BJ131/Chem_NY_Acticide CBM-2 | 7O4BJ131 | Acticide CBM-2 | ACt. icide Cbm-2 | Chem_NY_Acticide CBM-2 | NY-Analytical | 792 | Chemical Inspection | Odor | Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4BM301/Chem_126A | 7O4BM301 | IPBC | Glycacil-NP | Chem_126A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4BM301/Chem_126A | 7O4BM301 | IPBC | Glycacil-NP | Chem_126A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4BM301/Chem_126A | 7O4BM301 | IPBC | Glycacil-NP | Chem_126A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4BM301/Chem_126A | 7O4BM301 | IPBC | Glycacil-NP | Chem_126A | NY-Analytical | 792 | Chemical Inspection | Color | Colorless to Amber (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4BM301/Chem_126A | 7O4BM301 | IPBC | Glycacil-NP | Chem_126A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4BM301/Chem_126A | 7O4BM301 | IPBC | Glycacil-NP | Chem_126A | NY-Analytical | 792 | Chemical Inspection | Odor | Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4DH247/Chem_201A | 7O4DH247 | Dequest 7000 | Dequest 7000 | Chem_201A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 20�C |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4DH247/Chem_201A | 7O4DH247 | Dequest 7000 | Dequest 7000 | Chem_201A | NY-Analytical | 432 | Specific Gravity at 20�C | Specific Gravity at 20°C | 1.2883 |  | FDA,3 | 1.27 | 1.31 | TM 202 | TM202 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4DH247/Chem_201A | 7O4DH247 | Dequest 7000 | Dequest 7000 | Chem_201A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear solution, free from solids (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4DH247/Chem_201A | 7O4DH247 | Dequest 7000 | Dequest 7000 | Chem_201A | NY-Analytical | 792 | Chemical Inspection | Color | 2 |  | FDA,1 |  | 2 | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4DH247/Chem_201A | 7O4DH247 | Dequest 7000 | Dequest 7000 | Chem_201A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4DH247/Chem_201A | 7O4DH247 | Dequest 7000 | Dequest 7000 | Chem_201A | NY-Analytical | 792 | Chemical Inspection | Odor | Bleach/chlorine (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OA234/Chem_131A | 7O4OA234 | Emulgade CM | Emulgade CM | Chem_131A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH as is |  | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4OA234/Chem_131A | 7O4OA234 | Emulgade CM | Emulgade CM | Chem_131A | NY-Analytical | 447 | pH as is | pH | 3.22 |  | FDA,2 | 3 | 3.3 | TM 200 | TM200 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4OA234/Chem_131A | 7O4OA234 | Emulgade CM | Emulgade CM | Chem_131A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4OA234/Chem_131A | 7O4OA234 | Emulgade CM | Emulgade CM | Chem_131A | NY-Analytical | 792 | Chemical Inspection | Appearance | Opaque, Thin Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4OA234/Chem_131A | 7O4OA234 | Emulgade CM | Emulgade CM | Chem_131A | NY-Analytical | 792 | Chemical Inspection | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4OA234/Chem_131A | 7O4OA234 | Emulgade CM | Emulgade CM | Chem_131A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4OA234/Chem_131A | 7O4OA234 | Emulgade CM | Emulgade CM | Chem_131A | NY-Analytical | 792 | Chemical Inspection | Odor | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4OA347/Chem_NY_Lauryl Glucoside | 7O4OA347 | Lauryl Glucoside | Lauryl Glucoside | Chem_NY_Lauryl Glucoside | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OA347/Chem_NY_Lauryl Glucoside | 7O4OA347 | Lauryl Glucoside | Lauryl Glucoside | Chem_NY_Lauryl Glucoside | NY-Analytical | 792 | Chemical Inspection | Appearance | Match standard (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OA347/Chem_NY_Lauryl Glucoside | 7O4OA347 | Lauryl Glucoside | Lauryl Glucoside | Chem_NY_Lauryl Glucoside | NY-Analytical | 792 | Chemical Inspection | Color | Match standard (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OA347/Chem_NY_Lauryl Glucoside | 7O4OA347 | Lauryl Glucoside | Lauryl Glucoside | Chem_NY_Lauryl Glucoside | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OA347/Chem_NY_Lauryl Glucoside | 7O4OA347 | Lauryl Glucoside | Lauryl Glucoside | Chem_NY_Lauryl Glucoside | NY-Analytical | 792 | Chemical Inspection | Odor | Match standard (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OA347/Chem_NY_Lauryl Glucoside | 7O4OA347 | Lauryl Glucoside | Lauryl Glucoside | Chem_NY_Lauryl Glucoside | NY-Analytical | 390 | Identification | IR Identification | Comparison of scans of sample vs standard (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OA401/Chem_132A | 7O4OA401 | Feminine Hygiene Wipe Fragrance | Feminine Hygiene Wipe Fragrance | Chem_132A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25�C |  | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4OA401/Chem_132A | 7O4OA401 | Feminine Hygiene Wipe Fragrance | Feminine Hygiene Wipe Fragrance | Chem_132A | NY-Analytical | 402 | Specific Gravity at 25�C | Specific Gravity at 25°C | 1.0225 |  | FDA,4 | 1.0115 | 1.0315 | TM 202 | TM202 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4OA401/Chem_132A | 7O4OA401 | Feminine Hygiene Wipe Fragrance | Feminine Hygiene Wipe Fragrance | Chem_132A | NY-Analytical | 451 | Refractive Index at 25�C | Refractive Index | 1.485 |  | FDA,3 | 1.4794 | 1.4994 | TM 204 | TM204 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4OA401/Chem_132A | 7O4OA401 | Feminine Hygiene Wipe Fragrance | Feminine Hygiene Wipe Fragrance | Chem_132A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4OA401/Chem_132A | 7O4OA401 | Feminine Hygiene Wipe Fragrance | Feminine Hygiene Wipe Fragrance | Chem_132A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4OA401/Chem_132A | 7O4OA401 | Feminine Hygiene Wipe Fragrance | Feminine Hygiene Wipe Fragrance | Chem_132A | NY-Analytical | 792 | Chemical Inspection | Color | Colorless to light yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4OA401/Chem_132A | 7O4OA401 | Feminine Hygiene Wipe Fragrance | Feminine Hygiene Wipe Fragrance | Chem_132A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4OA401/Chem_132A | 7O4OA401 | Feminine Hygiene Wipe Fragrance | Feminine Hygiene Wipe Fragrance | Chem_132A | NY-Analytical | 792 | Chemical Inspection | Odor | Fresh Floral (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4OA490/Chem_133A | 7O4OA490 | Dowanol DPnP | Dowanol Dpnp | Chem_133A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25�C |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OA490/Chem_133A | 7O4OA490 | Dowanol DPnP | Dowanol Dpnp | Chem_133A | NY-Analytical | 402 | Specific Gravity at 25�C | Specific Gravity at 25°C | 0.9182 |  | FDA,3 | 0.909 | 0.929 | TM 202 | TM202 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OA490/Chem_133A | 7O4OA490 | Dowanol DPnP | Dowanol Dpnp | Chem_133A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OA490/Chem_133A | 7O4OA490 | Dowanol DPnP | Dowanol Dpnp | Chem_133A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OA490/Chem_133A | 7O4OA490 | Dowanol DPnP | Dowanol Dpnp | Chem_133A | NY-Analytical | 792 | Chemical Inspection | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OA490/Chem_133A | 7O4OA490 | Dowanol DPnP | Dowanol Dpnp | Chem_133A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OA490/Chem_133A | 7O4OA490 | Dowanol DPnP | Dowanol Dpnp | Chem_133A | NY-Analytical | 792 | Chemical Inspection | Odor | Ether (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OA580/Chem_NY_Glycacil SG | 7O4OA580 | Glycacil SG | Glycacil Sg | Chem_NY_Glycacil SG | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OA580/Chem_NY_Glycacil SG | 7O4OA580 | Glycacil SG | Glycacil Sg | Chem_NY_Glycacil SG | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OA580/Chem_NY_Glycacil SG | 7O4OA580 | Glycacil SG | Glycacil Sg | Chem_NY_Glycacil SG | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OA580/Chem_NY_Glycacil SG | 7O4OA580 | Glycacil SG | Glycacil Sg | Chem_NY_Glycacil SG | NY-Analytical | 792 | Chemical Inspection | Color | Colorless to Amber (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OA580/Chem_NY_Glycacil SG | 7O4OA580 | Glycacil SG | Glycacil Sg | Chem_NY_Glycacil SG | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OA580/Chem_NY_Glycacil SG | 7O4OA580 | Glycacil SG | Glycacil Sg | Chem_NY_Glycacil SG | NY-Analytical | 792 | Chemical Inspection | Odor | Characteristic Match Standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | 7O4OA634 | Water Soluble Flowering Herbs Fragrance | Water Soluble Flowering Herbs Fragrance | Chem_NY_Water Soluble Flowering Herbs Fragrance | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25�C |  | JB | RAW MATERIAL - ARKANSAS | RM_CHEMICAL | both
7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | 7O4OA634 | Water Soluble Flowering Herbs Fragrance | Water Soluble Flowering Herbs Fragrance | Chem_NY_Water Soluble Flowering Herbs Fragrance | NY-Analytical | 402 | Specific Gravity at 25�C | Specific Gravity at 25°C | 1.0629 |  | FDA,4 | 1.0531 | 1.0831 | TM 202 | TM202 | JB | RAW MATERIAL - ARKANSAS | RM_CHEMICAL | both
7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | 7O4OA634 | Water Soluble Flowering Herbs Fragrance | Water Soluble Flowering Herbs Fragrance | Chem_NY_Water Soluble Flowering Herbs Fragrance | NY-Analytical | 451 | Refractive Index at 25�C | Refractive Index | 1.47 |  | FDA,4 | 1.4655 | 1.4855 | TM 204 | TM204 | JB | RAW MATERIAL - ARKANSAS | RM_CHEMICAL | both
7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | 7O4OA634 | Water Soluble Flowering Herbs Fragrance | Water Soluble Flowering Herbs Fragrance | Chem_NY_Water Soluble Flowering Herbs Fragrance | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | JB | RAW MATERIAL - ARKANSAS | RM_CHEMICAL | both
7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | 7O4OA634 | Water Soluble Flowering Herbs Fragrance | Water Soluble Flowering Herbs Fragrance | Chem_NY_Water Soluble Flowering Herbs Fragrance | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | JB | RAW MATERIAL - ARKANSAS | RM_CHEMICAL | both
7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | 7O4OA634 | Water Soluble Flowering Herbs Fragrance | Water Soluble Flowering Herbs Fragrance | Chem_NY_Water Soluble Flowering Herbs Fragrance | NY-Analytical | 792 | Chemical Inspection | Color | Pale Yellow to Straw (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | JB | RAW MATERIAL - ARKANSAS | RM_CHEMICAL | both
7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | 7O4OA634 | Water Soluble Flowering Herbs Fragrance | Water Soluble Flowering Herbs Fragrance | Chem_NY_Water Soluble Flowering Herbs Fragrance | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | JB | RAW MATERIAL - ARKANSAS | RM_CHEMICAL | both
7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | 7O4OA634 | Water Soluble Flowering Herbs Fragrance | Water Soluble Flowering Herbs Fragrance | Chem_NY_Water Soluble Flowering Herbs Fragrance | NY-Analytical | 792 | Chemical Inspection | Odor | Floral (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | JB | RAW MATERIAL - ARKANSAS | RM_CHEMICAL | both
7O4OA709/Chem_NY_Cucumber Extract # 11-19-3 | 7O4OA709 | Cucumber Extract # 11-19-3 | Cucumber ExtraCt. # 11-19-3 | Chem_NY_Cucumber Extract # 11-19-3 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | [5 empty] | Specific Gravity at 25�C |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OA709/Chem_NY_Cucumber Extract # 11-19-3 | 7O4OA709 | Cucumber Extract # 11-19-3 | Cucumber ExtraCt. # 11-19-3 | Chem_NY_Cucumber Extract # 11-19-3 | NY-Analytical | 402 | Specific Gravity at 25�C | Specific Gravity at 25°C |  |  | FDA,3 | 1.034 | 1.044 | TM 202 | TM202 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OA709/Chem_NY_Cucumber Extract # 11-19-3 | 7O4OA709 | Cucumber Extract # 11-19-3 | Cucumber ExtraCt. # 11-19-3 | Chem_NY_Cucumber Extract # 11-19-3 | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | 1.425 |  | FDA,3 | 1.42 | 1.43 | TM 204 | TM204 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OA709/Chem_NY_Cucumber Extract # 11-19-3 | 7O4OA709 | Cucumber Extract # 11-19-3 | Cucumber ExtraCt. # 11-19-3 | Chem_NY_Cucumber Extract # 11-19-3 | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OA709/Chem_NY_Cucumber Extract # 11-19-3 | 7O4OA709 | Cucumber Extract # 11-19-3 | Cucumber ExtraCt. # 11-19-3 | Chem_NY_Cucumber Extract # 11-19-3 | NY-Analytical | 792 | Chemical Inspection | Description | Clear colorless liquid having a characteristic odor (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OA911/Chem_NY_Zinc Sulfate Heptahydrate | 7O4OA911 | Zinc Sulfate Heptahydrate | Zinc Sulfate Heptahydrate | Chem_NY_Zinc Sulfate Heptahydrate | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OA911/Chem_NY_Zinc Sulfate Heptahydrate | 7O4OA911 | Zinc Sulfate Heptahydrate | Zinc Sulfate Heptahydrate | Chem_NY_Zinc Sulfate Heptahydrate | NY-Analytical | 792 | Chemical Inspection | Appearance | Crystals/crystalline powder (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OA911/Chem_NY_Zinc Sulfate Heptahydrate | 7O4OA911 | Zinc Sulfate Heptahydrate | Zinc Sulfate Heptahydrate | Chem_NY_Zinc Sulfate Heptahydrate | NY-Analytical | 792 | Chemical Inspection | Color | Colorless to white (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OA911/Chem_NY_Zinc Sulfate Heptahydrate | 7O4OA911 | Zinc Sulfate Heptahydrate | Zinc Sulfate Heptahydrate | Chem_NY_Zinc Sulfate Heptahydrate | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OA911/Chem_NY_Zinc Sulfate Heptahydrate | 7O4OA911 | Zinc Sulfate Heptahydrate | Zinc Sulfate Heptahydrate | Chem_NY_Zinc Sulfate Heptahydrate | NY-Analytical | 792 | Chemical Inspection | Odor | Match standard (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OA911/Chem_NY_Zinc Sulfate Heptahydrate | 7O4OA911 | Zinc Sulfate Heptahydrate | Zinc Sulfate Heptahydrate | Chem_NY_Zinc Sulfate Heptahydrate | NY-Analytical | 390 | Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OD351/Chem_NY_Benzalkonium Chloride 50% | 7O4OD351 | Benzalkonium Chloride 50% | Benzalkonium Chloride 50% | Chem_NY_Benzalkonium Chloride 50% | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | [5 empty] | Appearance |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OD351/Chem_NY_Benzalkonium Chloride 50% | 7O4OD351 | Benzalkonium Chloride 50% | Benzalkonium Chloride 50% | Chem_NY_Benzalkonium Chloride 50% | NY-Analytical | 397 | Appearance | Appearance | Clear liquid, free from foreign material (FFFM) (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OD351/Chem_NY_Benzalkonium Chloride 50% | 7O4OD351 | Benzalkonium Chloride 50% | Benzalkonium Chloride 50% | Chem_NY_Benzalkonium Chloride 50% | NY-Analytical | 938 | Ethanol | Ethanol |  | % v/v | FDA,2 | 5.9 | 6.5 | TM 323 | TM323 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OD351/Chem_NY_Benzalkonium Chloride 50% | 7O4OD351 | Benzalkonium Chloride 50% | Benzalkonium Chloride 50% | Chem_NY_Benzalkonium Chloride 50% | NY-Analytical | 938 | Ethanol | Ethanol Final |  | % v/v | FDA,2 | 5.9 | 6.5 | TM 323 | TM323 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OD351/Chem_NY_Benzalkonium Chloride 50% | 7O4OD351 | Benzalkonium Chloride 50% | Benzalkonium Chloride 50% | Chem_NY_Benzalkonium Chloride 50% | NY-Analytical | 455 | Identification A | Identification A - Chloride | [5 empty] | TM 4001 | TM4001 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OD351/Chem_NY_Benzalkonium Chloride 50% | 7O4OD351 | Benzalkonium Chloride 50% | Benzalkonium Chloride 50% | Chem_NY_Benzalkonium Chloride 50% | NY-Analytical | 456 | Identification B | Identification B - Benzalkonium | [5 empty] | TM 4001 | TM4001 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OD351/Chem_NY_Benzalkonium Chloride 50% | 7O4OD351 | Benzalkonium Chloride 50% | Benzalkonium Chloride 50% | Chem_NY_Benzalkonium Chloride 50% | NY-Analytical | 1064 | Total % ADBAC | % BZK and Dual-Quat |  | % | FDA,1 | 47.5 | 52.5 | TM 434 | TM434 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OD363/Chem_137A | 7O4OD363 | Tego Betaine 810 | Tego Betaine 810 | Chem_137A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH as is |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OD363/Chem_137A | 7O4OD363 | Tego Betaine 810 | Tego Betaine 810 | Chem_137A | NY-Analytical | 447 | pH as is | pH | 5.06 |  | FDA,1 | 4.5 | 5.5 | TM 200 | TM200 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OD363/Chem_137A | 7O4OD363 | Tego Betaine 810 | Tego Betaine 810 | Chem_137A | NY-Analytical | 806 | Total Solids | Total Solids (2 g for 3 hr @ 105°C) | 43.5 | % | FDA,1 | 41 | 45 | TM 206 | TM206 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OD363/Chem_137A | 7O4OD363 | Tego Betaine 810 | Tego Betaine 810 | Chem_137A | NY-Analytical | 806 | Total Solids | Total Solids Final (2 g for 3 hr @ 105°C) | 43.45 | % | FDA,1 | 41 | 45 | TM 206 | TM206 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OD363/Chem_137A | 7O4OD363 | Tego Betaine 810 | Tego Betaine 810 | Chem_137A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OD363/Chem_137A | 7O4OD363 | Tego Betaine 810 | Tego Betaine 810 | Chem_137A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OD363/Chem_137A | 7O4OD363 | Tego Betaine 810 | Tego Betaine 810 | Chem_137A | NY-Analytical | 792 | Chemical Inspection | Color | Colorless to pale yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
| 7O4OD419 | Coconut Fatty Acid |  | Analytical - Chem_138A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA does not meet specifications (OOS) |  |  |  |  | Identification |  |  |  |  | left_only
| 7O4OD419 | Coconut Fatty Acid |  | Analytical - Chem_138A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 |  |  |  | left_only
| 7O4OD419 | Coconut Fatty Acid |  | Analytical - Chem_138A | NY-Analytical | 792 | Chemical Inspection | Appearance | Semi Solid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7O4OD419 | Coconut Fatty Acid |  | Analytical - Chem_138A | NY-Analytical | 792 | Chemical Inspection | Color | White to light yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7O4OD419 | Coconut Fatty Acid |  | Analytical - Chem_138A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
| 7O4OD419 | Coconut Fatty Acid |  | Analytical - Chem_138A | NY-Analytical | 792 | Chemical Inspection | Odor | Fatty (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM |  |  |  | left_only
7O4OD748/Chem_139A | 7O4OD748 | Thymol USP/NF | Thymol USP/Nf | Chem_139A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OD748/Chem_139A | 7O4OD748 | Thymol USP/NF | Thymol USP/Nf | Chem_139A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OD748/Chem_139A | 7O4OD748 | Thymol USP/NF | Thymol USP/Nf | Chem_139A | NY-Analytical | 792 | Chemical Inspection | Appearance | White solid crystals (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OD748/Chem_139A | 7O4OD748 | Thymol USP/NF | Thymol USP/Nf | Chem_139A | NY-Analytical | 792 | Chemical Inspection | Color | White Crystals (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OD748/Chem_139A | 7O4OD748 | Thymol USP/NF | Thymol USP/Nf | Chem_139A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OD748/Chem_139A | 7O4OD748 | Thymol USP/NF | Thymol USP/Nf | Chem_139A | NY-Analytical | 792 | Chemical Inspection | Odor | To match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OF575/Che_NY_Plantaren 2000N - UP | 7O4OF575 | Plantaren 2000N-UP | Plantaren 2000N-Up | Che_NY_Plantaren 2000N - UP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OF575/Che_NY_Plantaren 2000N - UP | 7O4OF575 | Plantaren 2000N-UP | Plantaren 2000N-Up | Che_NY_Plantaren 2000N - UP | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OF575/Che_NY_Plantaren 2000N - UP | 7O4OF575 | Plantaren 2000N-UP | Plantaren 2000N-Up | Che_NY_Plantaren 2000N - UP | NY-Analytical | 792 | Chemical Inspection | Appearance | Hazy, Pale Yellow Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OF575/Che_NY_Plantaren 2000N - UP | 7O4OF575 | Plantaren 2000N-UP | Plantaren 2000N-Up | Che_NY_Plantaren 2000N - UP | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OF602/Chem_141A | 7O4OF602 | Phenoxyethanol | Phenoxyethanol | Chem_141A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4OF602/Chem_141A | 7O4OF602 | Phenoxyethanol | Phenoxyethanol | Chem_141A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4OF602/Chem_141A | 7O4OF602 | Phenoxyethanol | Phenoxyethanol | Chem_141A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4OG169/Chem_142A | 7O4OG169 | Bronidox-L | Bronidox-L | Chem_142A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 20�C |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OG169/Chem_142A | 7O4OG169 | Bronidox-L | Bronidox-L | Chem_142A | NY-Analytical | 432 | Specific Gravity at 20�C | Specific Gravity at 20°C | 1.08 |  | FDA,3 | 1.08 | 1.09 | TM 202 | TM202 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OG169/Chem_142A | 7O4OG169 | Bronidox-L | Bronidox-L | Chem_142A | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | 1.435 |  | FDA,3 | 1.435 | 1.437 | TM 204 | TM204 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OG169/Chem_142A | 7O4OG169 | Bronidox-L | Bronidox-L | Chem_142A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OG169/Chem_142A | 7O4OG169 | Bronidox-L | Bronidox-L | Chem_142A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OG169/Chem_142A | 7O4OG169 | Bronidox-L | Bronidox-L | Chem_142A | NY-Analytical | 792 | Chemical Inspection | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OG169/Chem_142A | 7O4OG169 | Bronidox-L | Bronidox-L | Chem_142A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OH247/Chem_143A | 7O4OH247 | Multitrope 1214 | Multitrope 1214 | Chem_143A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OH247/Chem_143A | 7O4OH247 | Multitrope 1214 | Multitrope 1214 | Chem_143A | NY-Analytical | 398 | pH | pH (1% w/v aqueous solution) | 2 |  | FDA,1 |  | 2 | TM 200 | TM200 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OH247/Chem_143A | 7O4OH247 | Multitrope 1214 | Multitrope 1214 | Chem_143A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OH247/Chem_143A | 7O4OH247 | Multitrope 1214 | Multitrope 1214 | Chem_143A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear viscous liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OH247/Chem_143A | 7O4OH247 | Multitrope 1214 | Multitrope 1214 | Chem_143A | NY-Analytical | 792 | Chemical Inspection | Color | Colorless to dark amber (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OH247/Chem_143A | 7O4OH247 | Multitrope 1214 | Multitrope 1214 | Chem_143A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OH460/MASTER - Microbiology - J&J Chemical Testing) | 7O4OH460 | CHE SOOTHEX | CHE Soothex | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 801 | % Water | Moisture Content | 0.6 | % | FDA,1 |  | 1.5 | J&J TMWW78017 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OJ123/Chem_NY_NET-KOB | 7O4OJ123 | NET-KOB | Net-Kob | Chem_NY_NET-KOB | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OJ123/Chem_NY_NET-KOB | 7O4OJ123 | NET-KOB | Net-Kob | Chem_NY_NET-KOB | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OJ123/Chem_NY_NET-KOB | 7O4OJ123 | NET-KOB | Net-Kob | Chem_NY_NET-KOB | NY-Analytical | 792 | Chemical Inspection | Appearance | Viscous Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OJ123/Chem_NY_NET-KOB | 7O4OJ123 | NET-KOB | Net-Kob | Chem_NY_NET-KOB | NY-Analytical | 792 | Chemical Inspection | Color | Colorless to pale yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OJ123/Chem_NY_NET-KOB | 7O4OJ123 | NET-KOB | Net-Kob | Chem_NY_NET-KOB | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OJ123/Chem_NY_NET-KOB | 7O4OJ123 | NET-KOB | Net-Kob | Chem_NY_NET-KOB | NY-Analytical | 792 | Chemical Inspection | Odor | Characteristic (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OM165/Chem_146A | 7O4OM165 | Tocopheryl Acetate | Tocopheryl Acetate | Chem_146A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OM165/Chem_146A | 7O4OM165 | Tocopheryl Acetate | Tocopheryl Acetate | Chem_146A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OM165/Chem_146A | 7O4OM165 | Tocopheryl Acetate | Tocopheryl Acetate | Chem_146A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear, slightly oily viscous liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OM165/Chem_146A | 7O4OM165 | Tocopheryl Acetate | Tocopheryl Acetate | Chem_146A | NY-Analytical | 792 | Chemical Inspection | Color | Light Yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OM165/Chem_146A | 7O4OM165 | Tocopheryl Acetate | Tocopheryl Acetate | Chem_146A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OM165/Chem_146A | 7O4OM165 | Tocopheryl Acetate | Tocopheryl Acetate | Chem_146A | NY-Analytical | 792 | Chemical Inspection | Odor | To match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OM666/Chem_NY_Sodium Benzoate EDF NF/FCC | 7O4OM666 | Sodium Benzoate EDF NF/FCC | Sodium Benzoate EDF NF/FCC | Chem_NY_Sodium Benzoate EDF NF/FCC | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | WI | RAW MATERIAL - GREEN BAY | RM_CHEMICAL | both
7O4OM666/Chem_NY_Sodium Benzoate EDF NF/FCC | 7O4OM666 | Sodium Benzoate EDF NF/FCC | Sodium Benzoate EDF NF/FCC | Chem_NY_Sodium Benzoate EDF NF/FCC | NY-Analytical | 792 | Chemical Inspection | Appearance | Extruded Pellets (Done) |  |  |  |  | TM 208 | TM208 | WI | RAW MATERIAL - GREEN BAY | RM_CHEMICAL | both
7O4OM666/Chem_NY_Sodium Benzoate EDF NF/FCC | 7O4OM666 | Sodium Benzoate EDF NF/FCC | Sodium Benzoate EDF NF/FCC | Chem_NY_Sodium Benzoate EDF NF/FCC | NY-Analytical | 792 | Chemical Inspection | Color | White (Done) |  |  |  |  | TM 208 | TM208 | WI | RAW MATERIAL - GREEN BAY | RM_CHEMICAL | both
7O4OM666/Chem_NY_Sodium Benzoate EDF NF/FCC | 7O4OM666 | Sodium Benzoate EDF NF/FCC | Sodium Benzoate EDF NF/FCC | Chem_NY_Sodium Benzoate EDF NF/FCC | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | WI | RAW MATERIAL - GREEN BAY | RM_CHEMICAL | both
7O4OM666/Chem_NY_Sodium Benzoate EDF NF/FCC | 7O4OM666 | Sodium Benzoate EDF NF/FCC | Sodium Benzoate EDF NF/FCC | Chem_NY_Sodium Benzoate EDF NF/FCC | NY-Analytical | 792 | Chemical Inspection | Odor | Odorless (Done) |  |  |  |  | TM 208 | TM208 | WI | RAW MATERIAL - GREEN BAY | RM_CHEMICAL | both
7O4OM666/Chem_NY_Sodium Benzoate EDF NF/FCC | 7O4OM666 | Sodium Benzoate EDF NF/FCC | Sodium Benzoate EDF NF/FCC | Chem_NY_Sodium Benzoate EDF NF/FCC | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | WI | RAW MATERIAL - GREEN BAY | RM_CHEMICAL | both
7O4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7O4OP027 | CHE PROCETYL AWS | CHE Procetyl Aws | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 827 | Hydroxyl Value | Hydroxyl Value | 38.5 | mg KOH/g | FDA,1 | 35 | 50 | J&J TMWW7671 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP027/MASTER - Microbiology - J&J Chemical Testing) | 7O4OP027 | CHE PROCETYL AWS | CHE Procetyl Aws | MASTER - Microbiology - J&J Chemical Testing) | NY-Analytical | 827 | Hydroxyl Value | Hydroxyl Value Final | 38.5 | mg KOH/g | FDA,1 | 35 | 50 | J&J TMWW7671 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Actipone White Tea | ACt. ipone White Tea | Chem_NY_Actipone White Tea | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Density @ 20° C |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Actipone White Tea | ACt. ipone White Tea | Chem_NY_Actipone White Tea | NY-Analytical | 831 | Density @ 20° C | Density | 1.1173 |  | FDA,3 | 1.112 | 1.122 | TM 202 | TM202 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Actipone White Tea | ACt. ipone White Tea | Chem_NY_Actipone White Tea | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | 1.392 |  | FDA,3 | 1.388 | 1.398 | TM 204 | TM204 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Actipone White Tea | ACt. ipone White Tea | Chem_NY_Actipone White Tea | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear to slightly cloudy liquid (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Actipone White Tea | ACt. ipone White Tea | Chem_NY_Actipone White Tea | NY-Analytical | 792 | Chemical Inspection | Color | Yellow to Dark Brown (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Actipone White Tea | ACt. ipone White Tea | Chem_NY_Actipone White Tea | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Actipone White Tea | ACt. ipone White Tea | Chem_NY_Actipone White Tea | NY-Analytical | 792 | Chemical Inspection | Odor | Match standard (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Actipone White Tea | ACt. ipone White Tea | Chem_NY_Actipone White Tea | NY-Analytical | 390 | Identification | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Actipone White Tea | ACt. ipone White Tea | Chem_NY_Actipone White Tea | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram) | 0 | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Actipone White Tea | ACt. ipone White Tea | Chem_NY_Actipone White Tea | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | 0 | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Actipone White Tea | ACt. ipone White Tea | Chem_NY_Actipone White Tea | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram) |  | CFU/g | FDA,1 |  | 100 | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Actipone White Tea | ACt. ipone White Tea | Chem_NY_Actipone White Tea | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) |  | cfu | FDA,0 |  |  | TM 610 | TM610 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP065/Chem_NY_Actipone White Tea | 7O4OP065 | Actipone White Tea | ACt. ipone White Tea | Chem_NY_Actipone White Tea | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | Gram Negative Rods (OOS) |  |  |  |  | WI QAS 076 | WI_QAS_076 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP075/Chem_NY_Magique Lift 02LF | 7O4OP075 | Magique Lift 02 LF | Magique Lift 02 Lf | Chem_NY_Magique Lift 02LF | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA does not meet specifications (OOS) |  |  |  |  | Specific Gravity at 25�C |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP075/Chem_NY_Magique Lift 02LF | 7O4OP075 | Magique Lift 02 LF | Magique Lift 02 Lf | Chem_NY_Magique Lift 02LF | NY-Analytical | 402 | Specific Gravity at 25�C | Specific Gravity at 25°C | 0.9516 |  | FDA,3 | 0.938 | 0.948 | TM 202 | TM202 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP075/Chem_NY_Magique Lift 02LF | 7O4OP075 | Magique Lift 02 LF | Magique Lift 02 Lf | Chem_NY_Magique Lift 02LF | NY-Analytical | 451 | Refractive Index at 25�C | Refractive Index | 1.48 |  | FDA,3 | 1.474 | 1.484 | TM 204 | TM204 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP075/Chem_NY_Magique Lift 02LF | 7O4OP075 | Magique Lift 02 LF | Magique Lift 02 Lf | Chem_NY_Magique Lift 02LF | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP075/Chem_NY_Magique Lift 02LF | 7O4OP075 | Magique Lift 02 LF | Magique Lift 02 Lf | Chem_NY_Magique Lift 02LF | NY-Analytical | 792 | Chemical Inspection | Appearance | Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP075/Chem_NY_Magique Lift 02LF | 7O4OP075 | Magique Lift 02 LF | Magique Lift 02 Lf | Chem_NY_Magique Lift 02LF | NY-Analytical | 792 | Chemical Inspection | Color | Colorless to very light yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP075/Chem_NY_Magique Lift 02LF | 7O4OP075 | Magique Lift 02 LF | Magique Lift 02 Lf | Chem_NY_Magique Lift 02LF | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP075/Chem_NY_Magique Lift 02LF | 7O4OP075 | Magique Lift 02 LF | Magique Lift 02 Lf | Chem_NY_Magique Lift 02LF | NY-Analytical | 792 | Chemical Inspection | Odor | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP093/Chem_NY_PEG-11 | 7O4OP093 | PEG-11Stearyl Ether | PEG-11Stearyl Ether | Chem_NY_PEG-11 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA | COA | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP093/Chem_NY_PEG-11 | 7O4OP093 | PEG-11Stearyl Ether | PEG-11Stearyl Ether | Chem_NY_PEG-11 | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear, Light liquid (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP093/Chem_NY_PEG-11 | 7O4OP093 | PEG-11Stearyl Ether | PEG-11Stearyl Ether | Chem_NY_PEG-11 | NY-Analytical | 792 | Chemical Inspection | Color | Match standard (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP093/Chem_NY_PEG-11 | 7O4OP093 | PEG-11Stearyl Ether | PEG-11Stearyl Ether | Chem_NY_PEG-11 | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP093/Chem_NY_PEG-11 | 7O4OP093 | PEG-11Stearyl Ether | PEG-11Stearyl Ether | Chem_NY_PEG-11 | NY-Analytical | 792 | Chemical Inspection | Odor | Match standard (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP093/Chem_NY_PEG-11 | 7O4OP093 | PEG-11Stearyl Ether | PEG-11Stearyl Ether | Chem_NY_PEG-11 | NY-Analytical | 390 | Identification | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP113/Chem_NY_Robertet The Matcha I | 7O4OP113 | Fragrance Robertet The Matcha I | Fragrance Robertet The Matcha I | Chem_NY_Robertet The Matcha I | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 20�C |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP113/Chem_NY_Robertet The Matcha I | 7O4OP113 | Fragrance Robertet The Matcha I | Fragrance Robertet The Matcha I | Chem_NY_Robertet The Matcha I | NY-Analytical | 432 | Specific Gravity at 20�C | Specific Gravity at 20°C | 0.9652 |  | FDA,3 | 0.96 | 0.97 | TM 202 | TM202 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP113/Chem_NY_Robertet The Matcha I | 7O4OP113 | Fragrance Robertet The Matcha I | Fragrance Robertet The Matcha I | Chem_NY_Robertet The Matcha I | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | 1.47 |  | FDA,3 | 1.469 | 1.479 | TM 204 | TM204 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP113/Chem_NY_Robertet The Matcha I | 7O4OP113 | Fragrance Robertet The Matcha I | Fragrance Robertet The Matcha I | Chem_NY_Robertet The Matcha I | NY-Analytical | 792 | Chemical Inspection | Appearance | Oily liquid (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP113/Chem_NY_Robertet The Matcha I | 7O4OP113 | Fragrance Robertet The Matcha I | Fragrance Robertet The Matcha I | Chem_NY_Robertet The Matcha I | NY-Analytical | 792 | Chemical Inspection | Color | Pale Yellow – Light Yellow (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP113/Chem_NY_Robertet The Matcha I | 7O4OP113 | Fragrance Robertet The Matcha I | Fragrance Robertet The Matcha I | Chem_NY_Robertet The Matcha I | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP113/Chem_NY_Robertet The Matcha I | 7O4OP113 | Fragrance Robertet The Matcha I | Fragrance Robertet The Matcha I | Chem_NY_Robertet The Matcha I | NY-Analytical | 792 | Chemical Inspection | Odor | Match standard (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP113/Chem_NY_Robertet The Matcha I | 7O4OP113 | Fragrance Robertet The Matcha I | Fragrance Robertet The Matcha I | Chem_NY_Robertet The Matcha I | NY-Analytical | 390 | Identification | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP197/Chem_NY_Emulgade Smart/MB_01 | 7O4OP197 | Emulgade Smart/MB | Emulgade Smart/Mb | Chem_NY_Emulgade Smart/MB_01 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP197/Chem_NY_Emulgade Smart/MB_01 | 7O4OP197 | Emulgade Smart/MB | Emulgade Smart/Mb | Chem_NY_Emulgade Smart/MB_01 | NY-Analytical | 398 | pH | pH | 3 |  | FDA,1 | 3 | 3.5 | TM 200 | TM200 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP197/Chem_NY_Emulgade Smart/MB_01 | 7O4OP197 | Emulgade Smart/MB | Emulgade Smart/Mb | Chem_NY_Emulgade Smart/MB_01 | NY-Analytical | 792 | Chemical Inspection | Appearance | Match standard (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP197/Chem_NY_Emulgade Smart/MB_01 | 7O4OP197 | Emulgade Smart/MB | Emulgade Smart/Mb | Chem_NY_Emulgade Smart/MB_01 | NY-Analytical | 792 | Chemical Inspection | Color | White Liquid (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP197/Chem_NY_Emulgade Smart/MB_01 | 7O4OP197 | Emulgade Smart/MB | Emulgade Smart/Mb | Chem_NY_Emulgade Smart/MB_01 | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP197/Chem_NY_Emulgade Smart/MB_01 | 7O4OP197 | Emulgade Smart/MB | Emulgade Smart/Mb | Chem_NY_Emulgade Smart/MB_01 | NY-Analytical | 792 | Chemical Inspection | Odor | Match standard (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP197/Chem_NY_Emulgade Smart/MB_01 | 7O4OP197 | Emulgade Smart/MB | Emulgade Smart/Mb | Chem_NY_Emulgade Smart/MB_01 | NY-Analytical | 390 | Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Hydromanil H. GL. MS | Hydromanil H. GL MS | Chem_NY_Hydromanil | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Hydromanil H. GL. MS | Hydromanil H. GL MS | Chem_NY_Hydromanil | NY-Analytical | 792 | Chemical Inspection | Appearance | An opalescent, pale yellow to pale brown viscous liquid. (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Hydromanil H. GL. MS | Hydromanil H. GL MS | Chem_NY_Hydromanil | NY-Analytical | 792 | Chemical Inspection | Odor | Passes Test (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Hydromanil H. GL. MS | Hydromanil H. GL MS | Chem_NY_Hydromanil | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Hydromanil H. GL. MS | Hydromanil H. GL MS | Chem_NY_Hydromanil | NY-Analytical | 432 | Specific Gravity at 20�C | Specific Gravity at 20°C | 1.1158 |  | FDA,3 | 1.1 | 1.14 | J&J TMWW78012 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Hydromanil H. GL. MS | Hydromanil H. GL MS | Chem_NY_Hydromanil | NY-Analytical | 390 | Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Hydromanil H. GL. MS | Hydromanil H. GL MS | Chem_NY_Hydromanil | NY-Microbiology | 509 | J&J_Pathogens | Biggy or SDA | No Growth_SDA |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Hydromanil H. GL. MS | Hydromanil H. GL MS | Chem_NY_Hydromanil | NY-Microbiology | 509 | J&J_Pathogens | Cetrimide | No Growth |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Hydromanil H. GL. MS | Hydromanil H. GL MS | Chem_NY_Hydromanil | NY-Microbiology | 524 | J&J Fungal | Fungal | 0 | CFU/g | FDA,0 | 0 | 200 | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Hydromanil H. GL. MS | Hydromanil H. GL MS | Chem_NY_Hydromanil | NY-Microbiology | 509 | J&J_Pathogens | MacConkey | No Growth |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Hydromanil H. GL. MS | Hydromanil H. GL MS | Chem_NY_Hydromanil | NY-Microbiology | 509 | J&J_Pathogens | MSA or VJ | No Growth_MSA |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Hydromanil H. GL. MS | Hydromanil H. GL MS | Chem_NY_Hydromanil | NY-Microbiology | 700 | J&J SDA Controls | Negative Control | No Growth SDA |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Hydromanil H. GL. MS | Hydromanil H. GL MS | Chem_NY_Hydromanil | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | No Growth TSB |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Hydromanil H. GL. MS | Hydromanil H. GL MS | Chem_NY_Hydromanil | NY-Microbiology | 509 | J&J_Pathogens | TSA | No Growth |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Hydromanil H. GL. MS | Hydromanil H. GL MS | Chem_NY_Hydromanil | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | 0 | cfu/mL | FDA,0 | 0 | 2000 | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Hydromanil H. GL. MS | Hydromanil H. GL MS | Chem_NY_Hydromanil | NY-Microbiology | 517 | J&J_Enrichment | TSB | Not Turbid |  |  |  |  | TM 869 | TM869 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Hydromanil H. GL. MS | Hydromanil H. GL MS | Chem_NY_Hydromanil | NY-Analytical | 447 | pH as is | pH | 4.87 |  | FDA,1 | 4.5 | 5.5 | TM WW78014 | TMWW78014 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Hydromanil H. GL. MS | Hydromanil H. GL MS | Chem_NY_Hydromanil | NY-Analytical | 917 | Total Sugar | Total Sugar | 8.2 | % | FDA,1 | 6.5 | 9 | TM-002568 | TM_002568 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP230/Chem_NY_Hydromanil | 7O4OP230 | Hydromanil H. GL. MS | Hydromanil H. GL MS | Chem_NY_Hydromanil | NY-Analytical | 916 | Reducing Sugars | Reducing Sugars | 3.3 | % | FDA,1 | 2.5 | 3.6 | TM-002569 | TM_002569 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP253/Chem_NY_BTC 1010 | 7O4OP253 | BTC 1010 | BTC 1010 | Chem_NY_BTC 1010 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP253/Chem_NY_BTC 1010 | 7O4OP253 | BTC 1010 | BTC 1010 | Chem_NY_BTC 1010 | NY-Analytical | 398 | pH | pH (10% in water) | 7.19 |  | FDA,1 | 6 | 9 | TM 200 | TM200 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP253/Chem_NY_BTC 1010 | 7O4OP253 | BTC 1010 | BTC 1010 | Chem_NY_BTC 1010 | NY-Analytical | 796 | Assay | Assay | 50.8 | % | FDA,0 | 50 | 55 | TM 301/434 | TM301_TM434 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP253/Chem_NY_BTC 1010 | 7O4OP253 | BTC 1010 | BTC 1010 | Chem_NY_BTC 1010 | NY-Analytical | 796 | Assay | Assay Final | 50.85 | % | FDA,0 | 50 | 55 | TM 301/434 | TM301_TM434 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP284/Chem_155A | 7O4OP284 | Grape Fruit Bliss A 297424 | Grape Fruit Bliss A 297424 | Chem_155A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | GC Profile |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP284/Chem_155A | 7O4OP284 | Grape Fruit Bliss A 297424 | Grape Fruit Bliss A 297424 | Chem_155A | NY-Analytical | 839 | GC Profile | GC Profile | To match standard. Not required when performing IR test. (Done) |  |  |  |  | J&J TM-001862 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP284/Chem_155A | 7O4OP284 | Grape Fruit Bliss A 297424 | Grape Fruit Bliss A 297424 | Chem_155A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear, yellow liquid. (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP284/Chem_155A | 7O4OP284 | Grape Fruit Bliss A 297424 | Grape Fruit Bliss A 297424 | Chem_155A | NY-Analytical | 792 | Chemical Inspection | Odor | Supplier certifies on C of A that odor matches standard (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP284/Chem_155A | 7O4OP284 | Grape Fruit Bliss A 297424 | Grape Fruit Bliss A 297424 | Chem_155A | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP284/Chem_155A | 7O4OP284 | Grape Fruit Bliss A 297424 | Grape Fruit Bliss A 297424 | Chem_155A | NY-Analytical | 432 | Specific Gravity at 20�C | Specific Gravity at 20°C | 0.9722 |  | FDA,3 | 0.963 | 0.973 | J&J TMWW78012 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP284/Chem_155A | 7O4OP284 | Grape Fruit Bliss A 297424 | Grape Fruit Bliss A 297424 | Chem_155A | NY-Analytical | 390 | Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP284/Chem_155A | 7O4OP284 | Grape Fruit Bliss A 297424 | Grape Fruit Bliss A 297424 | Chem_155A | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | 1.442 |  | FDA,4 | 1.442 | 1.452 | J&J TMWW78015 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP345/Chem_157A | 7O4OP345 | Euxyl K 701 | Euxyl K 701 | Chem_157A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP345/Chem_157A | 7O4OP345 | Euxyl K 701 | Euxyl K 701 | Chem_157A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP345/Chem_157A | 7O4OP345 | Euxyl K 701 | Euxyl K 701 | Chem_157A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP345/Chem_157A | 7O4OP345 | Euxyl K 701 | Euxyl K 701 | Chem_157A | NY-Analytical | 792 | Chemical Inspection | Color | Yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP345/Chem_157A | 7O4OP345 | Euxyl K 701 | Euxyl K 701 | Chem_157A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP345/Chem_157A | 7O4OP345 | Euxyl K 701 | Euxyl K 701 | Chem_157A | NY-Analytical | 792 | Chemical Inspection | Odor | Phenoxyethanol-like (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP567/Chem_NY_BTC 1210-80% | 7O4OP567 | BTC 1210 - 80% | BTC 1210 - 80% | Chem_NY_BTC 1210-80% | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA | COA | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP567/Chem_NY_BTC 1210-80% | 7O4OP567 | BTC 1210 - 80% | BTC 1210 - 80% | Chem_NY_BTC 1210-80% | NY-Analytical | 398 | pH | pH (10% in water) | 7.79 |  | FDA,1 | 6 | 8.5 | TM 200 | TM200 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP567/Chem_NY_BTC 1210-80% | 7O4OP567 | BTC 1210 - 80% | BTC 1210 - 80% | Chem_NY_BTC 1210-80% | NY-Analytical | 397 | Appearance | Appearance | Colorless to slight yellow (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP567/Chem_NY_BTC 1210-80% | 7O4OP567 | BTC 1210 - 80% | BTC 1210 - 80% | Chem_NY_BTC 1210-80% | NY-Analytical | 1015 | Active (BTC 1210-80%) | Active (BTC 1210-80%) | 80.9 | % | FDA,1 | 80 | 82 | TM 3005 | TM3005 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP601/Chem_107A | 7O4OP601 | Fragrance Symrise Moisture Splash 830644 | Fragrance Symrise Moisture Splash 830644 | Chem_107A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | GC Profile |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP601/Chem_107A | 7O4OP601 | Fragrance Symrise Moisture Splash 830644 | Fragrance Symrise Moisture Splash 830644 | Chem_107A | NY-Analytical | 839 | GC Profile | GC Profile | To match standard. Not required when performing IR test. (Done) |  |  |  |  | J&J TM-003030 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP601/Chem_107A | 7O4OP601 | Fragrance Symrise Moisture Splash 830644 | Fragrance Symrise Moisture Splash 830644 | Chem_107A | NY-Analytical | 792 | Chemical Inspection | Appearance | Colorless to pale yellow clear liquid. (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP601/Chem_107A | 7O4OP601 | Fragrance Symrise Moisture Splash 830644 | Fragrance Symrise Moisture Splash 830644 | Chem_107A | NY-Analytical | 792 | Chemical Inspection | Odor | Meets Odor Specification or of Characteristic Odor. [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP601/Chem_107A | 7O4OP601 | Fragrance Symrise Moisture Splash 830644 | Fragrance Symrise Moisture Splash 830644 | Chem_107A | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP601/Chem_107A | 7O4OP601 | Fragrance Symrise Moisture Splash 830644 | Fragrance Symrise Moisture Splash 830644 | Chem_107A | NY-Analytical | 432 | Specific Gravity at 20�C | Specific Gravity at 20°C | 0.9921 |  | FDA,3 | 0.986 | 0.996 | J&J TMWW78012 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP601/Chem_107A | 7O4OP601 | Fragrance Symrise Moisture Splash 830644 | Fragrance Symrise Moisture Splash 830644 | Chem_107A | NY-Analytical | 390 | Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP601/Chem_107A | 7O4OP601 | Fragrance Symrise Moisture Splash 830644 | Fragrance Symrise Moisture Splash 830644 | Chem_107A | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | 1.452 |  | FDA,4 | 1.452 | 1.462 | J&J TMWW78015 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP638/Chem_192A | 7O4OP638 | Enhansys CG-100 | Enhansys Cg-100 | Chem_192A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP638/Chem_192A | 7O4OP638 | Enhansys CG-100 | Enhansys Cg-100 | Chem_192A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear, colorless to light yellow liquid / White solid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP705/Chem_NY_Cetaphil Concentrate | 7O4OP705 | Cetaphil Concentrate | Cetaphil Concentrate | Chem_NY_Cetaphil Concentrate | NY-Analytical | 1051 | Temptale Reading | Temptale Reading |  | ºC | FDA,0 | 5 | 40 | Temptale Reading | 0 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP705/Chem_NY_Cetaphil Concentrate | 7O4OP705 | Cetaphil Concentrate | Cetaphil Concentrate | Chem_NY_Cetaphil Concentrate | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH as is |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP705/Chem_NY_Cetaphil Concentrate | 7O4OP705 | Cetaphil Concentrate | Cetaphil Concentrate | Chem_NY_Cetaphil Concentrate | NY-Analytical | 447 | pH as is | pH | 7 |  | FDA,1 | 5.5 | 7 | TM 200 | TM200 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP705/Chem_NY_Cetaphil Concentrate | 7O4OP705 | Cetaphil Concentrate | Cetaphil Concentrate | Chem_NY_Cetaphil Concentrate | NY-Analytical | 747 | Viscosity | Viscosity (Spindle 3 @ 60RPM in 300 ml beaker) | 915 | cps | FDA,0 | 400 | 1600 | TM 203 | TM203 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP705/Chem_NY_Cetaphil Concentrate | 7O4OP705 | Cetaphil Concentrate | Cetaphil Concentrate | Chem_NY_Cetaphil Concentrate | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP705/Chem_NY_Cetaphil Concentrate | 7O4OP705 | Cetaphil Concentrate | Cetaphil Concentrate | Chem_NY_Cetaphil Concentrate | NY-Analytical | 792 | Chemical Inspection | Description | Viscous, slightly translucent, offwhite liquid: cetyl alcohol odor (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP806/Chem_NY_Carbopol 2020 | 7O4OP806 | Carbopol 2020 | Carbopol 2020 | Chem_NY_Carbopol 2020 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Loss on Drying |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP806/Chem_NY_Carbopol 2020 | 7O4OP806 | Carbopol 2020 | Carbopol 2020 | Chem_NY_Carbopol 2020 | NY-Analytical | 802 | Loss on Drying | Loss on Drying (80C Vacuum 1 hour) | 0.49 | % | FDA,2 |  | 2 | J&J TM7221 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP806/Chem_NY_Carbopol 2020 | 7O4OP806 | Carbopol 2020 | Carbopol 2020 | Chem_NY_Carbopol 2020 | NY-Analytical | 802 | Loss on Drying | Loss on Drying (80C Vacuum 1 hour) Final | 0.49 | % | FDA,2 |  | 2 | J&J TM7221 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP806/Chem_NY_Carbopol 2020 | 7O4OP806 | Carbopol 2020 | Carbopol 2020 | Chem_NY_Carbopol 2020 | NY-Analytical | 747 | Viscosity | Viscosity | 48667 | cps | FDA,0 | 47000 | 77000 | J&J TM7423 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP806/Chem_NY_Carbopol 2020 | 7O4OP806 | Carbopol 2020 | Carbopol 2020 | Chem_NY_Carbopol 2020 | NY-Analytical | 748 | Visual Inspection | Appearance | Fluffy white powder free of lumps (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP806/Chem_NY_Carbopol 2020 | 7O4OP806 | Carbopol 2020 | Carbopol 2020 | Chem_NY_Carbopol 2020 | NY-Analytical | 748 | Visual Inspection | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OP806/Chem_NY_Carbopol 2020 | 7O4OP806 | Carbopol 2020 | Carbopol 2020 | Chem_NY_Carbopol 2020 | NY-Analytical | 390 | Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OQ028/Chem_NY_Hexylene Glycol | 7O4OQ028 | Hexylene Glycol | Hexylene Glycol | Chem_NY_Hexylene Glycol | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis | COA | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4OQ028/Chem_NY_Euxyl K703 | 7O4OQ028 | Hexylene Glycol | Hexylene Glycol | Chem_NY_Euxyl K703 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Density @ 20° C |  | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4OQ028/Chem_NY_Euxyl K703 | 7O4OQ028 | Hexylene Glycol | Hexylene Glycol | Chem_NY_Euxyl K703 | NY-Analytical | 831 | Density @ 20° C | Density | 1.1282 |  | FDA,3 | 1.124 | 1.128 | TM 202 | TM202 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4OQ028/Chem_NY_Hexylene Glycol | 7O4OQ028 | Hexylene Glycol | Hexylene Glycol | Chem_NY_Hexylene Glycol | NY-Analytical | 432 | Specific Gravity at 20�C | Specific Gravity at 20°C |  |  | FDA,4 | 0.92 | 0.923 | TM 202 | TM202 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4OQ028/Chem_NY_Euxyl K703 | 7O4OQ028 | Hexylene Glycol | Hexylene Glycol | Chem_NY_Euxyl K703 | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | 1.539 |  | FDA,3 | 1.539 | 1.543 | TM 204 | TM204 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4OQ028/Chem_NY_Euxyl K703 | 7O4OQ028 | Hexylene Glycol | Hexylene Glycol | Chem_NY_Euxyl K703 | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4OQ028/Chem_NY_Hexylene Glycol | 7O4OQ028 | Hexylene Glycol | Hexylene Glycol | Chem_NY_Hexylene Glycol | NY-Analytical | 390 | Identification | IR Identification | To match standard. (Done) |  |  |  |  | TM 411 | TM411 | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4OQ028/Chem_NY_Hexylene Glycol | 7O4OQ028 | Hexylene Glycol | Hexylene Glycol | Chem_NY_Hexylene Glycol | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4OQ028/Chem_NY_Euxyl K703 | 7O4OQ028 | Hexylene Glycol | Hexylene Glycol | Chem_NY_Euxyl K703 | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4OQ028/Chem_NY_Hexylene Glycol | 7O4OQ028 | Hexylene Glycol | Hexylene Glycol | Chem_NY_Hexylene Glycol | NY-Analytical | 792 | Chemical Inspection | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4OQ028/Chem_NY_Euxyl K703 | 7O4OQ028 | Hexylene Glycol | Hexylene Glycol | Chem_NY_Euxyl K703 | NY-Analytical | 792 | Chemical Inspection | Color | Yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4OQ028/Chem_NY_Hexylene Glycol | 7O4OQ028 | Hexylene Glycol | Hexylene Glycol | Chem_NY_Hexylene Glycol | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4OQ028/Chem_NY_Euxyl K703 | 7O4OQ028 | Hexylene Glycol | Hexylene Glycol | Chem_NY_Euxyl K703 | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4OQ028/Chem_NY_Hexylene Glycol | 7O4OQ028 | Hexylene Glycol | Hexylene Glycol | Chem_NY_Hexylene Glycol | NY-Analytical | 792 | Chemical Inspection | Odor | Match standard (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4OQ028/Chem_NY_Euxyl K703 | 7O4OQ028 | Hexylene Glycol | Hexylene Glycol | Chem_NY_Euxyl K703 | NY-Analytical | 792 | Chemical Inspection | Odor | Phenoxyethanol-like (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | IN | RAW MATERIAL - INDIANA | RM_CHEMICAL | both
7O4OQ243/Chem_165A | 7O4OQ243 | Germall II | Germall Ii | Chem_165A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OQ243/Chem_165A | 7O4OQ243 | Germall II | Germall Ii | Chem_165A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OQ243/Chem_165A | 7O4OQ243 | Germall II | Germall Ii | Chem_165A | NY-Analytical | 792 | Chemical Inspection | Appearance | Fine, free-flowing powder (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OQ243/Chem_165A | 7O4OQ243 | Germall II | Germall Ii | Chem_165A | NY-Analytical | 792 | Chemical Inspection | Color | White (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OQ243/Chem_165A | 7O4OQ243 | Germall II | Germall Ii | Chem_165A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OQ243/Chem_165A | 7O4OQ243 | Germall II | Germall Ii | Chem_165A | NY-Analytical | 792 | Chemical Inspection | Odor | Characteristically mild (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OQ342/Chem_166A | 7O4OQ342 | Sport Wipe Fragrance | Sport Wipe Fragrance | Chem_166A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25�C |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OQ342/Chem_166A | 7O4OQ342 | Sport Wipe Fragrance | Sport Wipe Fragrance | Chem_166A | NY-Analytical | 402 | Specific Gravity at 25�C | Specific Gravity at 25°C | 0.9558 |  | FDA,4 | 0.9455 | 0.9655 | TM 202 | TM202 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OQ342/Chem_166A | 7O4OQ342 | Sport Wipe Fragrance | Sport Wipe Fragrance | Chem_166A | NY-Analytical | 451 | Refractive Index at 25�C | Refractive Index | 1.458 |  | FDA,4 | 1.4505 | 1.4705 | TM 204 | TM204 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OQ342/Chem_166A | 7O4OQ342 | Sport Wipe Fragrance | Sport Wipe Fragrance | Chem_166A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear, mobile liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OQ342/Chem_166A | 7O4OQ342 | Sport Wipe Fragrance | Sport Wipe Fragrance | Chem_166A | NY-Analytical | 792 | Chemical Inspection | Color | Very pale yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OQ342/Chem_166A | 7O4OQ342 | Sport Wipe Fragrance | Sport Wipe Fragrance | Chem_166A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OQ342/Chem_166A | 7O4OQ342 | Sport Wipe Fragrance | Sport Wipe Fragrance | Chem_166A | NY-Analytical | 792 | Chemical Inspection | Odor | Citrus Fresh (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OQ437/Chem_168A | 7O4OQ437 | Aminomethyl Propanol | Aminomethyl Propanol | Chem_168A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Specific Gravity at 25�C |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OQ437/Chem_168A | 7O4OQ437 | Aminomethyl Propanol | Aminomethyl Propanol | Chem_168A | NY-Analytical | 402 | Specific Gravity at 25�C | Specific Gravity at 25°C | 0.944 |  | FDA,4 | 0.938 | 0.9465 | TM 202 | TM202 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OQ437/Chem_168A | 7O4OQ437 | Aminomethyl Propanol | Aminomethyl Propanol | Chem_168A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OQ437/Chem_168A | 7O4OQ437 | Aminomethyl Propanol | Aminomethyl Propanol | Chem_168A | NY-Analytical | 792 | Chemical Inspection | Appearance | Mobile Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OQ437/Chem_168A | 7O4OQ437 | Aminomethyl Propanol | Aminomethyl Propanol | Chem_168A | NY-Analytical | 792 | Chemical Inspection | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OQ437/Chem_168A | 7O4OQ437 | Aminomethyl Propanol | Aminomethyl Propanol | Chem_168A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OQ437/Chem_168A | 7O4OQ437 | Aminomethyl Propanol | Aminomethyl Propanol | Chem_168A | NY-Analytical | 792 | Chemical Inspection | Odor | Amine (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OQ890/Chem_170A | 7O4OQ890 | Butoxyethanol | Butoxyethanol | Chem_170A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Identification |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OQ890/Chem_170A | 7O4OQ890 | Butoxyethanol | Butoxyethanol | Chem_170A | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OQ890/Chem_170A | 7O4OQ890 | Butoxyethanol | Butoxyethanol | Chem_170A | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear Liquid (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OQ890/Chem_170A | 7O4OQ890 | Butoxyethanol | Butoxyethanol | Chem_170A | NY-Analytical | 792 | Chemical Inspection | Color | Colorless (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OQ890/Chem_170A | 7O4OQ890 | Butoxyethanol | Butoxyethanol | Chem_170A | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OQ890/Chem_170A | 7O4OQ890 | Butoxyethanol | Butoxyethanol | Chem_170A | NY-Analytical | 792 | Chemical Inspection | Odor | Ether-like (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 792 | Chemical Inspection | Description | An almost colorless or paleyellowish liquid (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Any individual unspecified impuritv | 0.077 | % | FDA,1 |  | 0.1 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Ch!orhexidine oxazinone analog | 0.042 | % | FDA,1 |  | 0.2 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Chlorhexidine | 93.338 | % | FDA,1 |  |  | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Chlorhexidine urea | 0.085 | % | FDA,1 |  | 0.2 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Chlorhexidine amine | 0.047 | % | FDA,1 |  | 0.3 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Chlorhexidine dimer | 0.09 | % | FDA,1 |  | 0.5 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Chlorhexidine glucityl triazine | 0.083 | % | FDA,1 |  | 0.4 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Chlorhexidine guanidine | 0.083 | % | FDA,1 |  | 0.3 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Chlorhexidine nitrile | 0.114 | % | FDA,1 |  | 0.4 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 390 | Identification | Identification A | Conforms to USP (Done) |  |  |  |  | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 390 | Identification | Identification B | Conforms to USP (Done) |  |  |  |  | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | o-Chlorhexidine and specified unidentified impurit | 0.164 | % | FDA,1 |  | 0.4 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Oxochlorhexidine | 0.288 | % | FDA,1 |  | 0.4 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 812 | Organic Impurites | p-Chloroaniline | 0 | ug/mL | FDA,0 |  | 100 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 812 | Organic Impurites | p-Chloroaniline Final | 0 | ug/mL | FDA,0 |  | 100 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | p-Chlorophenvl urea | 0 | % | FDA,1 |  | 0.2 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Any individual unspecified impuritv | 1.13 |  | FDA,2 |  |  | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Ch!orhexidine oxazinone analog | 0.16 |  | FDA,2 |  |  | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Chlorhexidine | 1.1 |  | FDA,2 |  |  | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Chlorhexidine urea | 0.32 |  | FDA,2 |  |  | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Chlorhexidine amine | 0.21 |  | FDA,2 |  |  | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Chlorhexidine dimer | 0.87 |  | FDA,2 |  |  | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Chlorhexidine glucityl triazine | 0.95 |  | FDA,2 |  |  | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Chlorhexidine guanidine | 0.31 |  | FDA,2 |  |  | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Chlorhexidine nitrile | 0.52 |  | FDA,2 |  |  | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of o-Chlorhexidine and specified unidentified | 0.9 |  | FDA,2 |  |  | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Oxochlorhexidine | 1.43 |  | FDA,2 |  |  | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of p-Chlorophenvl urea | 0 |  | FDA,2 |  |  | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Specified unidentified impuritv 1 | 0.17 |  | FDA,2 |  |  | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Specified unidentified impuritv 1 | 0.011 | % | FDA,1 |  | 0.2 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Total Impurities | 0.983 | % | FDA,1 |  | 3 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 398 | pH | pH (Diluted 1 in 20) | 5.55 |  | FDA,1 | 5.5 | 7 | USP <791> | USP_791 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 796 | Assay | Assay | 20.1 | % | FDA,1 | 19 | 21 | USP-HPLC | USP_HPLC | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 796 | Assay | Assay Final | 20.1 | % | FDA,1 | 19 | 21 | USP-HPLC | USP_HPLC | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OS407/Chem_173A | 7O4OS407 | 20% Chlorhexidine Gluconate Solution Ph. Eur., BP | 20% Chlorhexidine Gluconate Solution pH. Eur., Bp | Chem_173A | NY-Analytical | 402 | Specific Gravity at 25�C | Specific Gravity at 25°C | 1.0648 |  | FDA,2 | 1.06 | 1.07 | USP/NF <841> | USP_NF_841 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU181/Chem_NY_Isopropanol 70% | 7O4OU181 | Isopropanol 70% | Isopropanol 70% | Chem_NY_Isopropanol 70% | IN-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA | COA | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU181/Chem_NY_Isopropanol 70% | 7O4OU181 | Isopropanol 70% | Isopropanol 70% | Chem_NY_Isopropanol 70% | IN-Analytical | 792 | Chemical Inspection | Appearance | Clear (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU181/Chem_NY_Isopropanol 70% | 7O4OU181 | Isopropanol 70% | Isopropanol 70% | Chem_NY_Isopropanol 70% | IN-Analytical | 792 | Chemical Inspection | Color | Colorless (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU181/Chem_NY_Isopropanol 70% | 7O4OU181 | Isopropanol 70% | Isopropanol 70% | Chem_NY_Isopropanol 70% | IN-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU181/Chem_NY_Isopropanol 70% | 7O4OU181 | Isopropanol 70% | Isopropanol 70% | Chem_NY_Isopropanol 70% | IN-Analytical | 792 | Chemical Inspection | Odor | Match standard (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU181/Chem_NY_Isopropanol 70% | 7O4OU181 | Isopropanol 70% | Isopropanol 70% | Chem_NY_Isopropanol 70% | IN-Analytical | 390 | Identification | IR Identification | Matches that of the Reference Standard (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 7O4OU225 | Sodium Hypochlorite 12.5% | Sodium Hypochlorite 12.5% | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | pH |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 7O4OU225 | Sodium Hypochlorite 12.5% | Sodium Hypochlorite 12.5% | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 398 | pH | pH (10% Solution) | 11.8 |  | FDA,1 | 10.5 | 13 | TM 200 | TM200 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 7O4OU225 | Sodium Hypochlorite 12.5% | Sodium Hypochlorite 12.5% | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 807 | % NaOCl | % NaOCl by weight | 9.9316 | % | FDA,2 | 11.5 | 14 | TM 366 | TM366 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 7O4OU225 | Sodium Hypochlorite 12.5% | Sodium Hypochlorite 12.5% | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 807 | % NaOCl | % NaOCl by weight Final | 10.0533 | % | FDA,2 | 11.5 | 14 | TM 366 | TM366 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 7O4OU225 | Sodium Hypochlorite 12.5% | Sodium Hypochlorite 12.5% | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 807 | % NaOCl | Sample Weight | 1.0023 | g | FDA,4 |  |  | TM 366 | TM366 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 7O4OU225 | Sodium Hypochlorite 12.5% | Sodium Hypochlorite 12.5% | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 807 | % NaOCl | Titrant End Volume | 27.5 | mL | FDA,1 |  |  | TM 366 | TM366 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 7O4OU225 | Sodium Hypochlorite 12.5% | Sodium Hypochlorite 12.5% | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 807 | % NaOCl | Titrant Start Volume | 0 | mL | FDA,1 |  |  | TM 366 | TM366 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 7O4OU225 | Sodium Hypochlorite 12.5% | Sodium Hypochlorite 12.5% | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 807 | % NaOCl | Total Volume of Titrant | 27.5 | mL | FDA,1 |  |  | TM 366 | TM366 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 7O4OU225 | Sodium Hypochlorite 12.5% | Sodium Hypochlorite 12.5% | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear liquid with a chlorine odor (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | 7O4OU225 | Sodium Hypochlorite 12.5% | Sodium Hypochlorite 12.5% | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 792 | Chemical Inspection | Color | straw-yellow (Done) |  |  |  |  | TM208/220/230/USP/Kao LM | TM208_TM220_TM230_TMUSP_Kao_LM | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU351/Chem_NY_Glycopyrronium Tosylate Monohydrate | 7O4OU351 | Glycopyrronium Tosylate Monohydrate | Glycopyrronium Tosylate Monohydrate | Chem_NY_Glycopyrronium Tosylate Monohydrate | NY-Analytical | 1051 | Temptale Reading | Shipment Temperatutre | Send Temperature data to Dermira (Done) |  |  |  |  | Shipment Temperatutre |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU351/Chem_NY_Glycopyrronium Tosylate Monohydrate | 7O4OU351 | Glycopyrronium Tosylate Monohydrate | Glycopyrronium Tosylate Monohydrate | Chem_NY_Glycopyrronium Tosylate Monohydrate | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | COA | COA | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU351/Chem_NY_Glycopyrronium Tosylate Monohydrate | 7O4OU351 | Glycopyrronium Tosylate Monohydrate | Glycopyrronium Tosylate Monohydrate | Chem_NY_Glycopyrronium Tosylate Monohydrate | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | COA Meets Specification (Done) |  |  |  |  | COA | COA | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU555/Chem_NY_Lidocaine Reduced testing | 7O4OU555 | Lidocaine Base USP | Lidocaine Base USP | Chem_NY_Lidocaine Reduced testing | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Certificate of Analysis | COA | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Lidocaine Base USP | Lidocaine Base USP | Chem_NY_Lidocaine Base USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Lidocaine Base USP | Lidocaine Base USP | Chem_NY_Lidocaine Base USP | NY-Analytical | 792 | Chemical Inspection | Appearance | Match standard (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU555/Chem_NY_Lidocaine Reduced testing | 7O4OU555 | Lidocaine Base USP | Lidocaine Base USP | Chem_NY_Lidocaine Reduced testing | NY-Analytical | 792 | Chemical Inspection | Appearance | Match standard (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Lidocaine Base USP | Lidocaine Base USP | Chem_NY_Lidocaine Base USP | NY-Analytical | 792 | Chemical Inspection | Color | Match standard (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU555/Chem_NY_Lidocaine Reduced testing | 7O4OU555 | Lidocaine Base USP | Lidocaine Base USP | Chem_NY_Lidocaine Reduced testing | NY-Analytical | 792 | Chemical Inspection | Color | Match standard (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Lidocaine Base USP | Lidocaine Base USP | Chem_NY_Lidocaine Base USP | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU555/Chem_NY_Lidocaine Reduced testing | 7O4OU555 | Lidocaine Base USP | Lidocaine Base USP | Chem_NY_Lidocaine Reduced testing | NY-Analytical | 792 | Chemical Inspection | Contamination | None (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Lidocaine Base USP | Lidocaine Base USP | Chem_NY_Lidocaine Base USP | NY-Analytical | 792 | Chemical Inspection | Odor | Match standard (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU555/Chem_NY_Lidocaine Reduced testing | 7O4OU555 | Lidocaine Base USP | Lidocaine Base USP | Chem_NY_Lidocaine Reduced testing | NY-Analytical | 792 | Chemical Inspection | Odor | Match standard (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU555/Chem_NY_Lidocaine Reduced testing | 7O4OU555 | Lidocaine Base USP | Lidocaine Base USP | Chem_NY_Lidocaine Reduced testing | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Lidocaine Base USP | Lidocaine Base USP | Chem_NY_Lidocaine Base USP | NY-Analytical | 390 | Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) |  |  |  |  | TM 411 | TM411 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU555/Chem_NY_Lidocaine Reduced testing | 7O4OU555 | Lidocaine Base USP | Lidocaine Base USP | Chem_NY_Lidocaine Reduced testing | NY-Analytical | 796 | Assay | Assay | 100.8 | % | FDA,1 | 97.5 | 102.5 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Lidocaine Base USP | Lidocaine Base USP | Chem_NY_Lidocaine Base USP | NY-Analytical | 796 | Assay | Assay | 100.4 | % | FDA,1 | 97.5 | 102.5 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Lidocaine Base USP | Lidocaine Base USP | Chem_NY_Lidocaine Base USP | NY-Analytical | 832 | Chloride | Chloride |  | % | FDA,4 |  | 0.0035 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Lidocaine Base USP | Lidocaine Base USP | Chem_NY_Lidocaine Base USP | NY-Analytical | 833 | Heavy Metals | Heavy Metals |  | ug/g | FDA,3 |  | 0.002 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Lidocaine Base USP | Lidocaine Base USP | Chem_NY_Lidocaine Base USP | NY-Analytical | 809 | Melting Range | Melting Range - End |  | C | FDA,1 | 66 | 69 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Lidocaine Base USP | Lidocaine Base USP | Chem_NY_Lidocaine Base USP | NY-Analytical | 809 | Melting Range | Melting Range - Start |  | C | FDA,1 | 66 | 69 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Lidocaine Base USP | Lidocaine Base USP | Chem_NY_Lidocaine Base USP | NY-Analytical | 834 | Residue on Ignition | Residue on Ignition |  | % | FDA,1 |  | 0.1 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OU555/Chem_NY_Lidocaine Base USP | 7O4OU555 | Lidocaine Base USP | Lidocaine Base USP | Chem_NY_Lidocaine Base USP | NY-Analytical | 835 | Sulfate Content | Sulfate |  | % | FDA,2 |  | 0.1 | USP | USP | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OX231/Chem_NY_FD&C Red 40 | 7O4OX231 | FD&C Red 40 (FD&C Grade) | Fd&C Red 40 (Fd&C Grade) | Chem_NY_FD&C Red 40 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Appearance |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OX231/Chem_NY_FD&C Red 40 | 7O4OX231 | FD&C Red 40 (FD&C Grade) | Fd&C Red 40 (Fd&C Grade) | Chem_NY_FD&C Red 40 | NY-Analytical | 397 | Appearance | Physical Properties | Brown Powder with Purple/red tint Powder (Done) |  |  |  |  | TM 208 | TM208 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7O4OX231/Chem_NY_FD&C Red 40 | 7O4OX231 | FD&C Red 40 (FD&C Grade) | Fd&C Red 40 (Fd&C Grade) | Chem_NY_FD&C Red 40 | NY-Analytical | 1061 | Total Color | Spectrophotometric | 89.4 | % | FDA,0 | 85 |  | TM 3014 | TM3014 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7R4OP485/Chem_NY_Phenoxyethanol_7R4OP485 | 7R4OP485 | Phenoxyethanol | Phenoxyethanol | Chem_NY_Phenoxyethanol_7R4OP485 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Chemical Inspection |  | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7R4OP485/Chem_NY_Phenoxyethanol_7R4OP485 | 7R4OP485 | Phenoxyethanol | Phenoxyethanol | Chem_NY_Phenoxyethanol_7R4OP485 | NY-Analytical | 792 | Chemical Inspection | Appearance | Clear, thin colorless liquid (Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7R4OP485/Chem_NY_Phenoxyethanol_7R4OP485 | 7R4OP485 | Phenoxyethanol | Phenoxyethanol | Chem_NY_Phenoxyethanol_7R4OP485 | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) |  |  |  |  | J&J TMWW78010/J&J TM7252 | TM7252 | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7R4OP485/Chem_NY_Phenoxyethanol_7R4OP485 | 7R4OP485 | Phenoxyethanol | Phenoxyethanol | Chem_NY_Phenoxyethanol_7R4OP485 | NY-Analytical | 390 | Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) |  |  |  |  | J&J TMWW78013 | J&J | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7R4OP485/Chem_NY_Phenoxyethanol_7R4OP485 | 7R4OP485 | Phenoxyethanol | Phenoxyethanol | Chem_NY_Phenoxyethanol_7R4OP485 | NY-Analytical | 892 | Limit of Phenol | Phenol | 18 | ppm | FDA,0 |  | 25 | USP/NF | USP_NF | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7R4OP485/Chem_NY_Phenoxyethanol_7R4OP485 | 7R4OP485 | Phenoxyethanol | Phenoxyethanol | Chem_NY_Phenoxyethanol_7R4OP485 | NY-Analytical | 849 | Purity | Purity | 99.4 | % | FDA,1 | 98 | 102 | USP/NF | USP_NF | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
7R4OP485/Chem_NY_Phenoxyethanol_7R4OP485 | 7R4OP485 | Phenoxyethanol | Phenoxyethanol | Chem_NY_Phenoxyethanol_7R4OP485 | NY-Analytical | 432 | Specific Gravity at 20�C | Specific Gravity at 20°C | 1.1077 |  | FDA,3 | 1.105 | 1.11 | USP/NF | USP_NF | NY | RAW MATERIAL - NY | RM_CHEMICAL | both
8AA654SI/Label_A10 | 8AA654SI | LBL Remove Completely Label | LBL SOLO INR REMV FW W/TRIANGL | Label_A10 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  | IN | RAW MATERIAL - INDIANA | RM_LABEL | both
8AA654SI/Label_A10 | 8AA654SI | LBL Remove Completely Label | LBL SOLO INR REMV FW W/TRIANGL | Label_A10 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 | IN | RAW MATERIAL - INDIANA | RM_LABEL | both
8AA654SI/Label_A10 | 8AA654SI | LBL Remove Completely Label | LBL SOLO INR REMV FW W/TRIANGL | Label_A10 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 | IN | RAW MATERIAL - INDIANA | RM_LABEL | both
8AA654SI/Label_A10 | 8AA654SI | LBL Remove Completely Label | LBL SOLO INR REMV FW W/TRIANGL | Label_A10 | NY-Analytical | 770 | Label Inspection | Length | 1.375 | in | FDA,4 | 1.3125 | 1.4375 | TM 143 / Visual | TM143 | IN | RAW MATERIAL - INDIANA | RM_LABEL | both
8AA654SI/Label_A10 | 8AA654SI | LBL Remove Completely Label | LBL SOLO INR REMV FW W/TRIANGL | Label_A10 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 | IN | RAW MATERIAL - INDIANA | RM_LABEL | both
8AA654SI/Label_A10 | 8AA654SI | LBL Remove Completely Label | LBL SOLO INR REMV FW W/TRIANGL | Label_A10 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 | IN | RAW MATERIAL - INDIANA | RM_LABEL | both
8AA654SI/Label_A10 | 8AA654SI | LBL Remove Completely Label | LBL SOLO INR REMV FW W/TRIANGL | Label_A10 | NY-Analytical | 770 | Label Inspection | Width | 2.5625 | in | FDA,3 | 2.5 | 2.625 | TM 143 / Visual | TM143 | IN | RAW MATERIAL - INDIANA | RM_LABEL | both
8AA70900/Label_NY_NP9990_2.5-4.625 | 8AA70900 | Preparation H Medicated Wipes for Women | Prep H Medicated Wipes For Women | Label_NY_NP9990_2.5-4.625 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  | NY | RAW MATERIAL - NY | RM_LABEL | both
8AA70900/Label_NY_NP9990_2.5-4.625 | 8AA70900 | Preparation H Medicated Wipes for Women | Prep H Medicated Wipes For Women | Label_NY_NP9990_2.5-4.625 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8AA70900/Label_NY_NP9990_2.5-4.625 | 8AA70900 | Preparation H Medicated Wipes for Women | Prep H Medicated Wipes For Women | Label_NY_NP9990_2.5-4.625 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8AA70900/Label_NY_NP9990_2.5-4.625 | 8AA70900 | Preparation H Medicated Wipes for Women | Prep H Medicated Wipes For Women | Label_NY_NP9990_2.5-4.625 | NY-Analytical | 770 | Label Inspection | Length | 2.5 | in | FDA,4 | 2.4375 | 2.5625 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8AA70900/Label_NY_NP9990_2.5-4.625 | 8AA70900 | Preparation H Medicated Wipes for Women | Prep H Medicated Wipes For Women | Label_NY_NP9990_2.5-4.625 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8AA70900/Label_NY_NP9990_2.5-4.625 | 8AA70900 | Preparation H Medicated Wipes for Women | Prep H Medicated Wipes For Women | Label_NY_NP9990_2.5-4.625 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8AA70900/Label_NY_NP9990_2.5-4.625 | 8AA70900 | Preparation H Medicated Wipes for Women | Prep H Medicated Wipes For Women | Label_NY_NP9990_2.5-4.625 | NY-Analytical | 770 | Label Inspection | Width | 4.625 | in | FDA,4 | 4.5625 | 4.6875 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
| 8AH12880 | CVS Hand Sanitizer Wipes packet Bulk |  | Analytical - Label_NY_NP7030_9-2.625 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  |  |  |  | left_only
| 8AH12880 | CVS Hand Sanitizer Wipes packet Bulk |  | Analytical - Label_NY_NP7030_9-2.625 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 |  |  |  | left_only
| 8AH12880 | CVS Hand Sanitizer Wipes packet Bulk |  | Analytical - Label_NY_NP7030_9-2.625 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 |  |  |  | left_only
| 8AH12880 | CVS Hand Sanitizer Wipes packet Bulk |  | Analytical - Label_NY_NP7030_9-2.625 | NY-Analytical | 770 | Label Inspection | Length | 9 | in | FDA,4 | 8.9375 | 9.0625 | TM 143 / Visual | TM143 |  |  |  | left_only
| 8AH12880 | CVS Hand Sanitizer Wipes packet Bulk |  | Analytical - Label_NY_NP7030_9-2.625 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 |  |  |  | left_only
| 8AH12880 | CVS Hand Sanitizer Wipes packet Bulk |  | Analytical - Label_NY_NP7030_9-2.625 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 |  |  |  | left_only
| 8AH12880 | CVS Hand Sanitizer Wipes packet Bulk |  | Analytical - Label_NY_NP7030_9-2.625 | NY-Analytical | 770 | Label Inspection | Width | 2.625 | in | FDA,4 | 2.5625 | 2.6875 | TM 143 / Visual | TM143 |  |  |  | left_only
8AP034RSL/Label_NY_NP9800_3.375-2.3437 | 8AP034RSL | Cetaphil Gentle Skin Cleansing Cloths Reseal | Cetaphil Gentle Skin Cleansing Cloths Reseal | Label_NY_NP9800_3.375-2.3437 | IN-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8AP034RSL/Label_NY_NP9800_3.375-2.3437 | 8AP034RSL | Cetaphil Gentle Skin Cleansing Cloths Reseal | Cetaphil Gentle Skin Cleansing Cloths Reseal | Label_NY_NP9800_3.375-2.3437 | IN-Analytical | 770 | Label Inspection | Certificate of Conformance | CofC meets specifications (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8AP034RSL/Label_NY_NP9800_3.375-2.3437 | 8AP034RSL | Cetaphil Gentle Skin Cleansing Cloths Reseal | Cetaphil Gentle Skin Cleansing Cloths Reseal | Label_NY_NP9800_3.375-2.3437 | IN-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8AP034RSL/Label_NY_NP9800_3.375-2.3437 | 8AP034RSL | Cetaphil Gentle Skin Cleansing Cloths Reseal | Cetaphil Gentle Skin Cleansing Cloths Reseal | Label_NY_NP9800_3.375-2.3437 | IN-Analytical | 770 | Label Inspection | Length | 3.375 | in | FDA,4 | 3.3125 | 3.4375 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8AP034RSL/Label_NY_NP9800_3.375-2.3437 | 8AP034RSL | Cetaphil Gentle Skin Cleansing Cloths Reseal | Cetaphil Gentle Skin Cleansing Cloths Reseal | Label_NY_NP9800_3.375-2.3437 | IN-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8AP034RSL/Label_NY_NP9800_3.375-2.3437 | 8AP034RSL | Cetaphil Gentle Skin Cleansing Cloths Reseal | Cetaphil Gentle Skin Cleansing Cloths Reseal | Label_NY_NP9800_3.375-2.3437 | IN-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8AP034RSL/Label_NY_NP9800_3.375-2.3437 | 8AP034RSL | Cetaphil Gentle Skin Cleansing Cloths Reseal | Cetaphil Gentle Skin Cleansing Cloths Reseal | Label_NY_NP9800_3.375-2.3437 | IN-Analytical | 770 | Label Inspection | Width | 2.3437 | in | FDA,4 | 2.2812 | 2.4062 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
| 8AP13800 | PDI Sani Cloth AF3 Front Label |  | Analytical - Label_NY_NP7200_6.25-4 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  |  |  |  | left_only
| 8AP13800 | PDI Sani Cloth AF3 Front Label |  | Analytical - Label_NY_NP7200_6.25-4 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 |  |  |  | left_only
| 8AP13800 | PDI Sani Cloth AF3 Front Label |  | Analytical - Label_NY_NP7200_6.25-4 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 |  |  |  | left_only
| 8AP13800 | PDI Sani Cloth AF3 Front Label |  | Analytical - Label_NY_NP7200_6.25-4 | NY-Analytical | 770 | Label Inspection | Length | 6.1875 | in | FDA,4 | 6.1875 | 6.3125 | TM 143 / Visual | TM143 |  |  |  | left_only
| 8AP13800 | PDI Sani Cloth AF3 Front Label |  | Analytical - Label_NY_NP7200_6.25-4 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 |  |  |  | left_only
| 8AP13800 | PDI Sani Cloth AF3 Front Label |  | Analytical - Label_NY_NP7200_6.25-4 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 |  |  |  | left_only
| 8AP13800 | PDI Sani Cloth AF3 Front Label |  | Analytical - Label_NY_NP7200_6.25-4 | NY-Analytical | 770 | Label Inspection | Width | 3.9375 | in | FDA,4 | 3.9375 | 4.0625 | TM 143 / Visual | TM143 |  |  |  | left_only
8AP22800/Label_NY_NP7200_5.8125-6.5 | 8AP22800 | Sani pro - Sani Cloth Disinfecting Multi-surface Wipes | Sani Professional - Sani-Cloth DisinfeCt. ing Multi-Surface Wipes | Label_NY_NP7200_5.8125-6.5 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | [5 empty] | Label Inspection |  | NY | RAW MATERIAL - NY | RM_LABEL | both
8AP22800/Label_NY_NP7200_5.8125-6.5 | 8AP22800 | Sani pro - Sani Cloth Disinfecting Multi-surface Wipes | Sani Professional - Sani-Cloth DisinfeCt. ing Multi-Surface Wipes | Label_NY_NP7200_5.8125-6.5 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8AP22800/Label_NY_NP7200_5.8125-6.5 | 8AP22800 | Sani pro - Sani Cloth Disinfecting Multi-surface Wipes | Sani Professional - Sani-Cloth DisinfeCt. ing Multi-Surface Wipes | Label_NY_NP7200_5.8125-6.5 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8AP22800/Label_NY_NP7200_5.8125-6.5 | 8AP22800 | Sani pro - Sani Cloth Disinfecting Multi-surface Wipes | Sani Professional - Sani-Cloth DisinfeCt. ing Multi-Surface Wipes | Label_NY_NP7200_5.8125-6.5 | NY-Analytical | 770 | Label Inspection | Length | 5.75 | in | FDA,4 | 5.75 | 5.875 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8AP22800/Label_NY_NP7200_5.8125-6.5 | 8AP22800 | Sani pro - Sani Cloth Disinfecting Multi-surface Wipes | Sani Professional - Sani-Cloth DisinfeCt. ing Multi-Surface Wipes | Label_NY_NP7200_5.8125-6.5 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8AP22800/Label_NY_NP7200_5.8125-6.5 | 8AP22800 | Sani pro - Sani Cloth Disinfecting Multi-surface Wipes | Sani Professional - Sani-Cloth DisinfeCt. ing Multi-Surface Wipes | Label_NY_NP7200_5.8125-6.5 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8AP22800/Label_NY_NP7200_5.8125-6.5 | 8AP22800 | Sani pro - Sani Cloth Disinfecting Multi-surface Wipes | Sani Professional - Sani-Cloth DisinfeCt. ing Multi-Surface Wipes | Label_NY_NP7200_5.8125-6.5 | NY-Analytical | 770 | Label Inspection | Width | 6.5 | in | FDA,4 | 6.4375 | 6.5625 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8AU120CK/Label_NY_NP7030_6-4 | 8AU120CK | PDI Hygea Personal Cleansing Washcloths | PDI Hygea Personal Cleansing Washcloths | Label_NY_NP7030_6-4 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  | NY | RAW MATERIAL - NY | RM_LABEL | both
8AU120CK/Label_NY_NP7030_6-4 | 8AU120CK | PDI Hygea Personal Cleansing Washcloths | PDI Hygea Personal Cleansing Washcloths | Label_NY_NP7030_6-4 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8AU120CK/Label_NY_NP7030_6-4 | 8AU120CK | PDI Hygea Personal Cleansing Washcloths | PDI Hygea Personal Cleansing Washcloths | Label_NY_NP7030_6-4 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8AU120CK/Label_NY_NP7030_6-4 | 8AU120CK | PDI Hygea Personal Cleansing Washcloths | PDI Hygea Personal Cleansing Washcloths | Label_NY_NP7030_6-4 | NY-Analytical | 770 | Label Inspection | Length | 6 | in | FDA,4 | 5.9375 | 6.0625 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8AU120CK/Label_NY_NP7030_6-4 | 8AU120CK | PDI Hygea Personal Cleansing Washcloths | PDI Hygea Personal Cleansing Washcloths | Label_NY_NP7030_6-4 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8AU120CK/Label_NY_NP7030_6-4 | 8AU120CK | PDI Hygea Personal Cleansing Washcloths | PDI Hygea Personal Cleansing Washcloths | Label_NY_NP7030_6-4 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8AU120CK/Label_NY_NP7030_6-4 | 8AU120CK | PDI Hygea Personal Cleansing Washcloths | PDI Hygea Personal Cleansing Washcloths | Label_NY_NP7030_6-4 | NY-Analytical | 770 | Label Inspection | Width | 4 | in | FDA,4 | 3.9375 | 4.0625 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8BB196CK/Label_NY_NP7030_6-5 | 8BB196CK | PDI Prevantics Device Swab Case label | Prevantics Device Swab Case Label | Label_NY_NP7030_6-5 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  | NY | RAW MATERIAL - NY | RM_LABEL | both
8BB196CK/Label_NY_NP7030_6-5 | 8BB196CK | PDI Prevantics Device Swab Case label | Prevantics Device Swab Case Label | Label_NY_NP7030_6-5 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8BB196CK/Label_NY_NP7030_6-5 | 8BB196CK | PDI Prevantics Device Swab Case label | Prevantics Device Swab Case Label | Label_NY_NP7030_6-5 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8BB196CK/Label_NY_NP7030_6-5 | 8BB196CK | PDI Prevantics Device Swab Case label | Prevantics Device Swab Case Label | Label_NY_NP7030_6-5 | NY-Analytical | 770 | Label Inspection | Length | 5.99 | in | FDA,4 | 5.9375 | 6.0625 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8BB196CK/Label_NY_NP7030_6-5 | 8BB196CK | PDI Prevantics Device Swab Case label | Prevantics Device Swab Case Label | Label_NY_NP7030_6-5 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8BB196CK/Label_NY_NP7030_6-5 | 8BB196CK | PDI Prevantics Device Swab Case label | Prevantics Device Swab Case Label | Label_NY_NP7030_6-5 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8BB196CK/Label_NY_NP7030_6-5 | 8BB196CK | PDI Prevantics Device Swab Case label | Prevantics Device Swab Case Label | Label_NY_NP7030_6-5 | NY-Analytical | 770 | Label Inspection | Width | 5.01 | in | FDA,4 | 4.9375 | 5.0625 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8BB512CL/Label_NY_NP7030_6-4 | 8BB512CL | PDI Povidone Iodine Prep Pad | Povidone-Iodine Prep Pad | Label_NY_NP7030_6-4 | NY-Analytical | 770 | Label Inspection | AQL Sample Size | 50 |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8BB512CL/Label_NY_NP7030_6-4 | 8BB512CL | PDI Povidone Iodine Prep Pad | Povidone-Iodine Prep Pad | Label_NY_NP7030_6-4 | NY-Analytical | 770 | Label Inspection | Artwork / Copy | Sample(s) all meet specifications (Done) | in |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8BP12801/Label_NY_NP7200_4.5-4 | 8BP12801 | PDI Sani Hands Sanitizing Wipes | Sani-Hands Sanitizing Wipes | Label_NY_NP7200_4.5-4 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  | NY | RAW MATERIAL - NY | RM_LABEL | both
8BP12801/Label_NY_NP7200_4.5-4 | 8BP12801 | PDI Sani Hands Sanitizing Wipes | Sani-Hands Sanitizing Wipes | Label_NY_NP7200_4.5-4 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8BP12801/Label_NY_NP7200_4.5-4 | 8BP12801 | PDI Sani Hands Sanitizing Wipes | Sani-Hands Sanitizing Wipes | Label_NY_NP7200_4.5-4 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8BP12801/Label_NY_NP7200_4.5-4 | 8BP12801 | PDI Sani Hands Sanitizing Wipes | Sani-Hands Sanitizing Wipes | Label_NY_NP7200_4.5-4 | NY-Analytical | 770 | Label Inspection | Length | 4.48 | in | FDA,4 | 4.4375 | 4.5625 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8BP12801/Label_NY_NP7200_4.5-4 | 8BP12801 | PDI Sani Hands Sanitizing Wipes | Sani-Hands Sanitizing Wipes | Label_NY_NP7200_4.5-4 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8BP12801/Label_NY_NP7200_4.5-4 | 8BP12801 | PDI Sani Hands Sanitizing Wipes | Sani-Hands Sanitizing Wipes | Label_NY_NP7200_4.5-4 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8BP12801/Label_NY_NP7200_4.5-4 | 8BP12801 | PDI Sani Hands Sanitizing Wipes | Sani-Hands Sanitizing Wipes | Label_NY_NP7200_4.5-4 | NY-Analytical | 770 | Label Inspection | Width | 3.98 | in | FDA,4 | 3.9375 | 4.0625 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8BP24500/Label_NY_NP7200_4-4.5 | 8BP24500 | Sani Professional Sani-Hands Anti-Bac Wipes | Sani Professional Sani-Hands Anti-Bac Wipes | Label_NY_NP7200_4-4.5 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  | NY | RAW MATERIAL - NY | RM_LABEL | both
8BP24500/Label_NY_NP7200_4-4.5 | 8BP24500 | Sani Professional Sani-Hands Anti-Bac Wipes | Sani Professional Sani-Hands Anti-Bac Wipes | Label_NY_NP7200_4-4.5 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8BP24500/Label_NY_NP7200_4-4.5 | 8BP24500 | Sani Professional Sani-Hands Anti-Bac Wipes | Sani Professional Sani-Hands Anti-Bac Wipes | Label_NY_NP7200_4-4.5 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8BP24500/Label_NY_NP7200_4-4.5 | 8BP24500 | Sani Professional Sani-Hands Anti-Bac Wipes | Sani Professional Sani-Hands Anti-Bac Wipes | Label_NY_NP7200_4-4.5 | NY-Analytical | 770 | Label Inspection | Length | 4 | in | FDA,4 | 3.9375 | 4.0625 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8BP24500/Label_NY_NP7200_4-4.5 | 8BP24500 | Sani Professional Sani-Hands Anti-Bac Wipes | Sani Professional Sani-Hands Anti-Bac Wipes | Label_NY_NP7200_4-4.5 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8BP24500/Label_NY_NP7200_4-4.5 | 8BP24500 | Sani Professional Sani-Hands Anti-Bac Wipes | Sani Professional Sani-Hands Anti-Bac Wipes | Label_NY_NP7200_4-4.5 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8BP24500/Label_NY_NP7200_4-4.5 | 8BP24500 | Sani Professional Sani-Hands Anti-Bac Wipes | Sani Professional Sani-Hands Anti-Bac Wipes | Label_NY_NP7200_4-4.5 | NY-Analytical | 770 | Label Inspection | Width | 4.5 | in | FDA,4 | 4.4375 | 4.5625 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8BQ76500/Label_NY_NP7200_6-3.75 | 8BQ76500 | PDI Sani Hands Hand Sanitizing Wipes | Sani-Hands Hand Sanitizing Wipes | Label_NY_NP7200_6-3.75 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  | NY | RAW MATERIAL - NY | RM_LABEL | both
8BQ76500/Label_NY_NP7200_6-3.75 | 8BQ76500 | PDI Sani Hands Hand Sanitizing Wipes | Sani-Hands Hand Sanitizing Wipes | Label_NY_NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8BQ76500/Label_NY_NP7200_6-3.75 | 8BQ76500 | PDI Sani Hands Hand Sanitizing Wipes | Sani-Hands Hand Sanitizing Wipes | Label_NY_NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8BQ76500/Label_NY_NP7200_6-3.75 | 8BQ76500 | PDI Sani Hands Hand Sanitizing Wipes | Sani-Hands Hand Sanitizing Wipes | Label_NY_NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Length | 6 | in | FDA,4 | 5.9375 | 6.0625 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8BQ76500/Label_NY_NP7200_6-3.75 | 8BQ76500 | PDI Sani Hands Hand Sanitizing Wipes | Sani-Hands Hand Sanitizing Wipes | Label_NY_NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8BQ76500/Label_NY_NP7200_6-3.75 | 8BQ76500 | PDI Sani Hands Hand Sanitizing Wipes | Sani-Hands Hand Sanitizing Wipes | Label_NY_NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8BQ76500/Label_NY_NP7200_6-3.75 | 8BQ76500 | PDI Sani Hands Hand Sanitizing Wipes | Sani-Hands Hand Sanitizing Wipes | Label_NY_NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Width | 3.6875 | in | FDA,4 | 3.6875 | 3.8125 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8BQ76501/Label_NY_NP7200_3.75-6 | 8BQ76501 | PDI Sani-Hands Hand Sanitizing Wipes CDA | Sani-Hands Hand Sanitizing Wipes Cda | Label_NY_NP7200_3.75-6 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  | NY | RAW MATERIAL - NY | RM_LABEL | both
8BQ76501/Label_NY_NP7200_3.75-6 | 8BQ76501 | PDI Sani-Hands Hand Sanitizing Wipes CDA | Sani-Hands Hand Sanitizing Wipes Cda | Label_NY_NP7200_3.75-6 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8BQ76501/Label_NY_NP7200_3.75-6 | 8BQ76501 | PDI Sani-Hands Hand Sanitizing Wipes CDA | Sani-Hands Hand Sanitizing Wipes Cda | Label_NY_NP7200_3.75-6 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8BQ76501/Label_NY_NP7200_3.75-6 | 8BQ76501 | PDI Sani-Hands Hand Sanitizing Wipes CDA | Sani-Hands Hand Sanitizing Wipes Cda | Label_NY_NP7200_3.75-6 | NY-Analytical | 770 | Label Inspection | Length | 3.6875 | in | FDA,4 | 3.6875 | 3.8125 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8BQ76501/Label_NY_NP7200_3.75-6 | 8BQ76501 | PDI Sani-Hands Hand Sanitizing Wipes CDA | Sani-Hands Hand Sanitizing Wipes Cda | Label_NY_NP7200_3.75-6 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8BQ76501/Label_NY_NP7200_3.75-6 | 8BQ76501 | PDI Sani-Hands Hand Sanitizing Wipes CDA | Sani-Hands Hand Sanitizing Wipes Cda | Label_NY_NP7200_3.75-6 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8BQ76501/Label_NY_NP7200_3.75-6 | 8BQ76501 | PDI Sani-Hands Hand Sanitizing Wipes CDA | Sani-Hands Hand Sanitizing Wipes Cda | Label_NY_NP7200_3.75-6 | NY-Analytical | 770 | Label Inspection | Width | 6 | in | FDA,4 | 5.9375 | 6.0625 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8CH211CK/Label_A12 | 8CH211CK | CVS pharmacy Total Home Antibac. Moist Wipes | CVS pHarmacy Total Home Antibac. Moist Wipes | Label_A12 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  | NY | RAW MATERIAL - NY | RM_LABEL | both
8CH211CK/Label_A12 | 8CH211CK | CVS pharmacy Total Home Antibac. Moist Wipes | CVS pHarmacy Total Home Antibac. Moist Wipes | Label_A12 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8CH211CK/Label_A12 | 8CH211CK | CVS pharmacy Total Home Antibac. Moist Wipes | CVS pHarmacy Total Home Antibac. Moist Wipes | Label_A12 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8CH211CK/Label_A12 | 8CH211CK | CVS pharmacy Total Home Antibac. Moist Wipes | CVS pHarmacy Total Home Antibac. Moist Wipes | Label_A12 | NY-Analytical | 770 | Label Inspection | Length | 9 | in | FDA,4 | 8.9375 | 9.0625 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8CH211CK/Label_A12 | 8CH211CK | CVS pharmacy Total Home Antibac. Moist Wipes | CVS pHarmacy Total Home Antibac. Moist Wipes | Label_A12 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8CH211CK/Label_A12 | 8CH211CK | CVS pharmacy Total Home Antibac. Moist Wipes | CVS pHarmacy Total Home Antibac. Moist Wipes | Label_A12 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8CH211CK/Label_A12 | 8CH211CK | CVS pharmacy Total Home Antibac. Moist Wipes | CVS pHarmacy Total Home Antibac. Moist Wipes | Label_A12 | NY-Analytical | 770 | Label Inspection | Width | 2.625 | in | FDA,4 | 2.5625 | 2.6875 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8CP13400/Label_NY_NP7100_4.5-4 | 8CP13400 | PDI Sani-Hands 135ct Front Label | Sani-Hands 135Ct. Front Label | Label_NY_NP7100_4.5-4 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  | NY | RAW MATERIAL - NY | RM_LABEL | both
8CP13400/Label_NY_NP7100_4.5-4 | 8CP13400 | PDI Sani-Hands 135ct Front Label | Sani-Hands 135Ct. Front Label | Label_NY_NP7100_4.5-4 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8CP13400/Label_NY_NP7100_4.5-4 | 8CP13400 | PDI Sani-Hands 135ct Front Label | Sani-Hands 135Ct. Front Label | Label_NY_NP7100_4.5-4 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8CP13400/Label_NY_NP7100_4.5-4 | 8CP13400 | PDI Sani-Hands 135ct Front Label | Sani-Hands 135Ct. Front Label | Label_NY_NP7100_4.5-4 | NY-Analytical | 770 | Label Inspection | Length | 4.5 | in | FDA,4 | 4.4375 | 4.5625 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8CP13400/Label_NY_NP7100_4.5-4 | 8CP13400 | PDI Sani-Hands 135ct Front Label | Sani-Hands 135Ct. Front Label | Label_NY_NP7100_4.5-4 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8CP13400/Label_NY_NP7100_4.5-4 | 8CP13400 | PDI Sani-Hands 135ct Front Label | Sani-Hands 135Ct. Front Label | Label_NY_NP7100_4.5-4 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8CP13400/Label_NY_NP7100_4.5-4 | 8CP13400 | PDI Sani-Hands 135ct Front Label | Sani-Hands 135Ct. Front Label | Label_NY_NP7100_4.5-4 | NY-Analytical | 770 | Label Inspection | Width | 4 | in | FDA,4 | 3.9375 | 4.0625 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8CP15900/Label_NY_NP7100/NP7200_6-3.75 | 8CP15900 | PDI Sani-Hands 220ct Front Label | Sani-Hands 220Ct. Front Label | Label_NY_NP7100/NP7200_6-3.75 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  | NY | RAW MATERIAL - NY | RM_LABEL | both
8CP15900/Label_NY_NP7100/NP7200_6-3.75 | 8CP15900 | PDI Sani-Hands 220ct Front Label | Sani-Hands 220Ct. Front Label | Label_NY_NP7100/NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8CP15900/Label_NY_NP7100/NP7200_6-3.75 | 8CP15900 | PDI Sani-Hands 220ct Front Label | Sani-Hands 220Ct. Front Label | Label_NY_NP7100/NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8CP15900/Label_NY_NP7100/NP7200_6-3.75 | 8CP15900 | PDI Sani-Hands 220ct Front Label | Sani-Hands 220Ct. Front Label | Label_NY_NP7100/NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Length | 5.99 | in | FDA,4 | 5.9375 | 6.0625 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8CP15900/Label_NY_NP7100/NP7200_6-3.75 | 8CP15900 | PDI Sani-Hands 220ct Front Label | Sani-Hands 220Ct. Front Label | Label_NY_NP7100/NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8CP15900/Label_NY_NP7100/NP7200_6-3.75 | 8CP15900 | PDI Sani-Hands 220ct Front Label | Sani-Hands 220Ct. Front Label | Label_NY_NP7100/NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8CP15900/Label_NY_NP7100/NP7200_6-3.75 | 8CP15900 | PDI Sani-Hands 220ct Front Label | Sani-Hands 220Ct. Front Label | Label_NY_NP7100/NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Width | 3.705 | in | FDA,4 | 3.6875 | 3.8125 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8CP63801/Label_A20 | 8CP63801 | PDI Sani Cloth AF3 65ct | Sani-Cloth AF3 65Ct. | Label_A20 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  | NY | RAW MATERIAL - NY | RM_LABEL | both
8CP63801/Label_A20 | 8CP63801 | PDI Sani Cloth AF3 65ct | Sani-Cloth AF3 65Ct. | Label_A20 | NY-Analytical | 770 | Label Inspection | Appearance | [5 empty] | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8CP63801/Label_A20 | 8CP63801 | PDI Sani Cloth AF3 65ct | Sani-Cloth AF3 65Ct. | Label_A20 | NY-Analytical | 770 | Label Inspection | Color/Copy | [5 empty] | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8CP63801/Label_A20 | 8CP63801 | PDI Sani Cloth AF3 65ct | Sani-Cloth AF3 65Ct. | Label_A20 | NY-Analytical | 770 | Label Inspection | Length |  | in | FDA,4 | 5.9375 | 6.0625 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8CP63801/Label_A20 | 8CP63801 | PDI Sani Cloth AF3 65ct | Sani-Cloth AF3 65Ct. | Label_A20 | NY-Analytical | 770 | Label Inspection | Odor | [5 empty] | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8CP63801/Label_A20 | 8CP63801 | PDI Sani Cloth AF3 65ct | Sani-Cloth AF3 65Ct. | Label_A20 | NY-Analytical | 770 | Label Inspection | Print | [5 empty] | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8CP63801/Label_A20 | 8CP63801 | PDI Sani Cloth AF3 65ct | Sani-Cloth AF3 65Ct. | Label_A20 | NY-Analytical | 770 | Label Inspection | Width |  | in | FDA,4 | 3.6875 | 3.8125 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8DP92000/Label_NY_NP7200_6.5-5.8125 | 8DP92000 | Sani Professional Sani-Hands 300ct | Sani Professional Sani-Hands 300Ct. | Label_NY_NP7200_6.5-5.8125 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | [5 empty] | Label Inspection |  | NY | RAW MATERIAL - NY | RM_LABEL | both
8DP92000/Label_NY_NP7200_6.5-5.8125 | 8DP92000 | Sani Professional Sani-Hands 300ct | Sani Professional Sani-Hands 300Ct. | Label_NY_NP7200_6.5-5.8125 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8DP92000/Label_NY_NP7200_6.5-5.8125 | 8DP92000 | Sani Professional Sani-Hands 300ct | Sani Professional Sani-Hands 300Ct. | Label_NY_NP7200_6.5-5.8125 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8DP92000/Label_NY_NP7200_6.5-5.8125 | 8DP92000 | Sani Professional Sani-Hands 300ct | Sani Professional Sani-Hands 300Ct. | Label_NY_NP7200_6.5-5.8125 | NY-Analytical | 770 | Label Inspection | Length | 6.5 | in | FDA,4 | 6.4375 | 6.5625 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8DP92000/Label_NY_NP7200_6.5-5.8125 | 8DP92000 | Sani Professional Sani-Hands 300ct | Sani Professional Sani-Hands 300Ct. | Label_NY_NP7200_6.5-5.8125 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8DP92000/Label_NY_NP7200_6.5-5.8125 | 8DP92000 | Sani Professional Sani-Hands 300ct | Sani Professional Sani-Hands 300Ct. | Label_NY_NP7200_6.5-5.8125 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8DP92000/Label_NY_NP7200_6.5-5.8125 | 8DP92000 | Sani Professional Sani-Hands 300ct | Sani Professional Sani-Hands 300Ct. | Label_NY_NP7200_6.5-5.8125 | NY-Analytical | 770 | Label Inspection | Width | 5.8125 | in | FDA,4 | 5.75 | 5.875 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8DQ85400/Label_A41 | 8DQ85400 | LBL PDI SANI CLOTH HB 65CT CAN FRT NP7100 | Sani-Cloth Hb 65Ct. Can Frt Np7100 | Label_A41 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  | NY | RAW MATERIAL - NY | RM_LABEL | both
8DQ85400/Label_A41 | 8DQ85400 | LBL PDI SANI CLOTH HB 65CT CAN FRT NP7100 | Sani-Cloth Hb 65Ct. Can Frt Np7100 | Label_A41 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8DQ85400/Label_A41 | 8DQ85400 | LBL PDI SANI CLOTH HB 65CT CAN FRT NP7100 | Sani-Cloth Hb 65Ct. Can Frt Np7100 | Label_A41 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8DQ85400/Label_A41 | 8DQ85400 | LBL PDI SANI CLOTH HB 65CT CAN FRT NP7100 | Sani-Cloth Hb 65Ct. Can Frt Np7100 | Label_A41 | NY-Analytical | 770 | Label Inspection | Length | 5.5625 | in | FDA,4 | 5.5 | 5.625 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8DQ85400/Label_A41 | 8DQ85400 | LBL PDI SANI CLOTH HB 65CT CAN FRT NP7100 | Sani-Cloth Hb 65Ct. Can Frt Np7100 | Label_A41 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8DQ85400/Label_A41 | 8DQ85400 | LBL PDI SANI CLOTH HB 65CT CAN FRT NP7100 | Sani-Cloth Hb 65Ct. Can Frt Np7100 | Label_A41 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8DQ85400/Label_A41 | 8DQ85400 | LBL PDI SANI CLOTH HB 65CT CAN FRT NP7100 | Sani-Cloth Hb 65Ct. Can Frt Np7100 | Label_A41 | NY-Analytical | 770 | Label Inspection | Width | 6.25 | in | FDA,4 | 6.1875 | 6.3125 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8DQ85900/Label_NY_NP7200_5.5625-6.25 | 8DQ85900 | PDI Sani Cloth Plus | Sani-Cloth Plus | Label_NY_NP7200_5.5625-6.25 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  | NY | RAW MATERIAL - NY | RM_LABEL | both
8DQ85900/Label_NY_NP7200_5.5625-6.25 | 8DQ85900 | PDI Sani Cloth Plus | Sani-Cloth Plus | Label_NY_NP7200_5.5625-6.25 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8DQ85900/Label_NY_NP7200_5.5625-6.25 | 8DQ85900 | PDI Sani Cloth Plus | Sani-Cloth Plus | Label_NY_NP7200_5.5625-6.25 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8DQ85900/Label_NY_NP7200_5.5625-6.25 | 8DQ85900 | PDI Sani Cloth Plus | Sani-Cloth Plus | Label_NY_NP7200_5.5625-6.25 | NY-Analytical | 770 | Label Inspection | Length | 5.5625 | in | FDA,4 | 5.5 | 5.625 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8DQ85900/Label_NY_NP7200_5.5625-6.25 | 8DQ85900 | PDI Sani Cloth Plus | Sani-Cloth Plus | Label_NY_NP7200_5.5625-6.25 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8DQ85900/Label_NY_NP7200_5.5625-6.25 | 8DQ85900 | PDI Sani Cloth Plus | Sani-Cloth Plus | Label_NY_NP7200_5.5625-6.25 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8DQ85900/Label_NY_NP7200_5.5625-6.25 | 8DQ85900 | PDI Sani Cloth Plus | Sani-Cloth Plus | Label_NY_NP7200_5.5625-6.25 | NY-Analytical | 770 | Label Inspection | Width | 6.25 | in | FDA,4 | 6.1875 | 6.3125 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8DQ914SL/Label_NY_NP7030_9-7 | 8DQ914SL | Ecolab Asepti-Wipe II Case Label | Ecolab Asepti-Wipe Ii Case Label | Label_NY_NP7030_9-7 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | [5 empty] | Label Inspection |  | NY | RAW MATERIAL - NY | RM_LABEL | both
8DQ914SL/Label_NY_NP7030_9-7 | 8DQ914SL | Ecolab Asepti-Wipe II Case Label | Ecolab Asepti-Wipe Ii Case Label | Label_NY_NP7030_9-7 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8DQ914SL/Label_NY_NP7030_9-7 | 8DQ914SL | Ecolab Asepti-Wipe II Case Label | Ecolab Asepti-Wipe Ii Case Label | Label_NY_NP7030_9-7 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8DQ914SL/Label_NY_NP7030_9-7 | 8DQ914SL | Ecolab Asepti-Wipe II Case Label | Ecolab Asepti-Wipe Ii Case Label | Label_NY_NP7030_9-7 | NY-Analytical | 770 | Label Inspection | Length | 9 | in | FDA,4 | 8.9375 | 9.0625 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8DQ914SL/Label_NY_NP7030_9-7 | 8DQ914SL | Ecolab Asepti-Wipe II Case Label | Ecolab Asepti-Wipe Ii Case Label | Label_NY_NP7030_9-7 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8DQ914SL/Label_NY_NP7030_9-7 | 8DQ914SL | Ecolab Asepti-Wipe II Case Label | Ecolab Asepti-Wipe Ii Case Label | Label_NY_NP7030_9-7 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8DQ914SL/Label_NY_NP7030_9-7 | 8DQ914SL | Ecolab Asepti-Wipe II Case Label | Ecolab Asepti-Wipe Ii Case Label | Label_NY_NP7030_9-7 | NY-Analytical | 770 | Label Inspection | Width | 7 | in | FDA,4 | 6.9375 | 7.0625 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8GQ772SK/Label_B03 | 8GQ772SK | LBL PDI BILINE NOT A SKIN WIPE | PDI Biline Not A Skin Wipe | Label_B03 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  | NY | RAW MATERIAL - NY | RM_LABEL | both
8GQ772SK/Label_B03 | 8GQ772SK | LBL PDI BILINE NOT A SKIN WIPE | PDI Biline Not A Skin Wipe | Label_B03 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8GQ772SK/Label_B03 | 8GQ772SK | LBL PDI BILINE NOT A SKIN WIPE | PDI Biline Not A Skin Wipe | Label_B03 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8GQ772SK/Label_B03 | 8GQ772SK | LBL PDI BILINE NOT A SKIN WIPE | PDI Biline Not A Skin Wipe | Label_B03 | NY-Analytical | 770 | Label Inspection | Diameter | 1.875 | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8GQ772SK/Label_B03 | 8GQ772SK | LBL PDI BILINE NOT A SKIN WIPE | PDI Biline Not A Skin Wipe | Label_B03 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8GQ772SK/Label_B03 | 8GQ772SK | LBL PDI BILINE NOT A SKIN WIPE | PDI Biline Not A Skin Wipe | Label_B03 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OH217UP/Label_A25 | 8OH217UP | LBL SAFEWAY H.S. WPS. CS LBL | Safeway H.S. Wps. Cs Lbl | Label_A25 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  | NY | RAW MATERIAL - NY | RM_LABEL | both
8OH217UP/Label_A25 | 8OH217UP | LBL SAFEWAY H.S. WPS. CS LBL | Safeway H.S. Wps. Cs Lbl | Label_A25 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OH217UP/Label_A25 | 8OH217UP | LBL SAFEWAY H.S. WPS. CS LBL | Safeway H.S. Wps. Cs Lbl | Label_A25 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OH217UP/Label_A25 | 8OH217UP | LBL SAFEWAY H.S. WPS. CS LBL | Safeway H.S. Wps. Cs Lbl | Label_A25 | NY-Analytical | 770 | Label Inspection | Length | 10 | in | FDA,4 | 9.9375 | 10.0625 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OH217UP/Label_A25 | 8OH217UP | LBL SAFEWAY H.S. WPS. CS LBL | Safeway H.S. Wps. Cs Lbl | Label_A25 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OH217UP/Label_A25 | 8OH217UP | LBL SAFEWAY H.S. WPS. CS LBL | Safeway H.S. Wps. Cs Lbl | Label_A25 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OH217UP/Label_A25 | 8OH217UP | LBL SAFEWAY H.S. WPS. CS LBL | Safeway H.S. Wps. Cs Lbl | Label_A25 | NY-Analytical | 770 | Label Inspection | Width | 4.0625 | in | FDA,4 | 3.9375 | 4.0625 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OJ22750/Label_A11 | 8OJ22750 | LBL PDI HYGEA MULTI PURPOSE NP9800 | PDI Hygea Multi Purpose Np9800 | Label_A11 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  | NY | RAW MATERIAL - NY | RM_LABEL | both
8OJ22750/Label_A11 | 8OJ22750 | LBL PDI HYGEA MULTI PURPOSE NP9800 | PDI Hygea Multi Purpose Np9800 | Label_A11 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OJ22750/Label_A11 | 8OJ22750 | LBL PDI HYGEA MULTI PURPOSE NP9800 | PDI Hygea Multi Purpose Np9800 | Label_A11 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OJ22750/Label_A11 | 8OJ22750 | LBL PDI HYGEA MULTI PURPOSE NP9800 | PDI Hygea Multi Purpose Np9800 | Label_A11 | NY-Analytical | 770 | Label Inspection | Length | 5.35 | in | FDA,4 | 5.3125 | 5.4375 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OJ22750/Label_A11 | 8OJ22750 | LBL PDI HYGEA MULTI PURPOSE NP9800 | PDI Hygea Multi Purpose Np9800 | Label_A11 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OJ22750/Label_A11 | 8OJ22750 | LBL PDI HYGEA MULTI PURPOSE NP9800 | PDI Hygea Multi Purpose Np9800 | Label_A11 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OJ22750/Label_A11 | 8OJ22750 | LBL PDI HYGEA MULTI PURPOSE NP9800 | PDI Hygea Multi Purpose Np9800 | Label_A11 | NY-Analytical | 770 | Label Inspection | Width | 2.61 | in | FDA,4 | 2.5625 | 2.6875 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP03700/Label_NY_NP9840_87-75 | 8OP03700 | Simple Sensitive Skin Experts Cleans. Facial Wipes 25ct US | Simple Sensitive Skin Experts Cleans. Facial Wipes 25Ct. US | Label_NY_NP9840_87-75 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP03700/Label_NY_NP9840_87-75 | 8OP03700 | Simple Sensitive Skin Experts Cleans. Facial Wipes 25ct US | Simple Sensitive Skin Experts Cleans. Facial Wipes 25Ct. US | Label_NY_NP9840_87-75 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP03700/Label_NY_NP9840_87-75 | 8OP03700 | Simple Sensitive Skin Experts Cleans. Facial Wipes 25ct US | Simple Sensitive Skin Experts Cleans. Facial Wipes 25Ct. US | Label_NY_NP9840_87-75 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP03700/Label_NY_NP9840_87-75 | 8OP03700 | Simple Sensitive Skin Experts Cleans. Facial Wipes 25ct US | Simple Sensitive Skin Experts Cleans. Facial Wipes 25Ct. US | Label_NY_NP9840_87-75 | NY-Analytical | 770 | Label Inspection | Length | 86.97 | mm | FDA,4 | 85.4125 | 88.5875 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP03700/Label_NY_NP9840_87-75 | 8OP03700 | Simple Sensitive Skin Experts Cleans. Facial Wipes 25ct US | Simple Sensitive Skin Experts Cleans. Facial Wipes 25Ct. US | Label_NY_NP9840_87-75 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP03700/Label_NY_NP9840_87-75 | 8OP03700 | Simple Sensitive Skin Experts Cleans. Facial Wipes 25ct US | Simple Sensitive Skin Experts Cleans. Facial Wipes 25Ct. US | Label_NY_NP9840_87-75 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP03700/Label_NY_NP9840_87-75 | 8OP03700 | Simple Sensitive Skin Experts Cleans. Facial Wipes 25ct US | Simple Sensitive Skin Experts Cleans. Facial Wipes 25Ct. US | Label_NY_NP9840_87-75 | NY-Analytical | 770 | Label Inspection | Width | 75.06 | mm | FDA,4 | 73.4125 | 76.5875 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP065RSL/Label_NY_NP9995_2.625-3.375 | 8OP065RSL | Ponds Evening Soothe Twlts. | Pond's Evening Soothe Twlts. | Label_NY_NP9995_2.625-3.375 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP065RSL/Label_NY_NP9995_2.625-3.375 | 8OP065RSL | Ponds Evening Soothe Twlts. | Pond's Evening Soothe Twlts. | Label_NY_NP9995_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP065RSL/Label_NY_NP9995_2.625-3.375 | 8OP065RSL | Ponds Evening Soothe Twlts. | Pond's Evening Soothe Twlts. | Label_NY_NP9995_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP065RSL/Label_NY_NP9995_2.625-3.375 | 8OP065RSL | Ponds Evening Soothe Twlts. | Pond's Evening Soothe Twlts. | Label_NY_NP9995_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Length | 2.625 | in | FDA,4 | 2.5625 | 2.6875 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP065RSL/Label_NY_NP9995_2.625-3.375 | 8OP065RSL | Ponds Evening Soothe Twlts. | Pond's Evening Soothe Twlts. | Label_NY_NP9995_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP065RSL/Label_NY_NP9995_2.625-3.375 | 8OP065RSL | Ponds Evening Soothe Twlts. | Pond's Evening Soothe Twlts. | Label_NY_NP9995_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP065RSL/Label_NY_NP9995_2.625-3.375 | 8OP065RSL | Ponds Evening Soothe Twlts. | Pond's Evening Soothe Twlts. | Label_NY_NP9995_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Width | 3.375 | in | FDA,4 | 3.3125 | 3.4375 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP08700/Label_NY_NP9800_3.375-2.625 | 8OP08700 | Cetaphil Gentle Makeup Removing Wipes | Cetaphil Gentle Makeup Removing Wipes | Label_NY_NP9800_3.375-2.625 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP08700/Label_NY_NP9800_3.375-2.625 | 8OP08700 | Cetaphil Gentle Makeup Removing Wipes | Cetaphil Gentle Makeup Removing Wipes | Label_NY_NP9800_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP08700/Label_NY_NP9800_3.375-2.625 | 8OP08700 | Cetaphil Gentle Makeup Removing Wipes | Cetaphil Gentle Makeup Removing Wipes | Label_NY_NP9800_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP08700/Label_NY_NP9800_3.375-2.625 | 8OP08700 | Cetaphil Gentle Makeup Removing Wipes | Cetaphil Gentle Makeup Removing Wipes | Label_NY_NP9800_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Length | 3.375 | in | FDA,4 | 3.3125 | 3.4375 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP08700/Label_NY_NP9800_3.375-2.625 | 8OP08700 | Cetaphil Gentle Makeup Removing Wipes | Cetaphil Gentle Makeup Removing Wipes | Label_NY_NP9800_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP08700/Label_NY_NP9800_3.375-2.625 | 8OP08700 | Cetaphil Gentle Makeup Removing Wipes | Cetaphil Gentle Makeup Removing Wipes | Label_NY_NP9800_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP08700/Label_NY_NP9800_3.375-2.625 | 8OP08700 | Cetaphil Gentle Makeup Removing Wipes | Cetaphil Gentle Makeup Removing Wipes | Label_NY_NP9800_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Width | 2.5625 | in | FDA,4 | 2.5625 | 2.6875 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP09100/Label_NY_NP9991_3.375-2.625 | 8OP09100 | Cetaphil Gentle Makeup Removing Wipes | Cetaphil Gentle Makeup Removing Wipes | Label_NY_NP9991_3.375-2.625 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Label Inspection |  | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP09100/Label_NY_NP9991_3.375-2.625 | 8OP09100 | Cetaphil Gentle Makeup Removing Wipes | Cetaphil Gentle Makeup Removing Wipes | Label_NY_NP9991_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP09100/Label_NY_NP9991_3.375-2.625 | 8OP09100 | Cetaphil Gentle Makeup Removing Wipes | Cetaphil Gentle Makeup Removing Wipes | Label_NY_NP9991_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP09100/Label_NY_NP9991_3.375-2.625 | 8OP09100 | Cetaphil Gentle Makeup Removing Wipes | Cetaphil Gentle Makeup Removing Wipes | Label_NY_NP9991_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Length | 3.375 | in | FDA,4 | 3.3125 | 3.4375 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP09100/Label_NY_NP9991_3.375-2.625 | 8OP09100 | Cetaphil Gentle Makeup Removing Wipes | Cetaphil Gentle Makeup Removing Wipes | Label_NY_NP9991_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP09100/Label_NY_NP9991_3.375-2.625 | 8OP09100 | Cetaphil Gentle Makeup Removing Wipes | Cetaphil Gentle Makeup Removing Wipes | Label_NY_NP9991_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP09100/Label_NY_NP9991_3.375-2.625 | 8OP09100 | Cetaphil Gentle Makeup Removing Wipes | Cetaphil Gentle Makeup Removing Wipes | Label_NY_NP9991_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Width | 2.625 | in | FDA,4 | 2.5625 | 2.6875 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP09300/Label_NY_NP7110_2.55-3 | 8OP09300 | Noxzema Ultimate Clear Anti-Blemish Pads -Front | Noxzema Ultimate Clear Anti-Blemish Pads -Front | Label_NY_NP7110_2.55-3 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP09300/Label_NY_NP7110_2.55-3 | 8OP09300 | Noxzema Ultimate Clear Anti-Blemish Pads -Front | Noxzema Ultimate Clear Anti-Blemish Pads -Front | Label_NY_NP7110_2.55-3 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP09300/Label_NY_NP7110_2.55-3 | 8OP09300 | Noxzema Ultimate Clear Anti-Blemish Pads -Front | Noxzema Ultimate Clear Anti-Blemish Pads -Front | Label_NY_NP7110_2.55-3 | NY-Analytical | 770 | Label Inspection | Length | 2.5 | in | FDA,4 | 2.4875 | 2.6125 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP09300/Label_NY_NP7110_2.55-3 | 8OP09300 | Noxzema Ultimate Clear Anti-Blemish Pads -Front | Noxzema Ultimate Clear Anti-Blemish Pads -Front | Label_NY_NP7110_2.55-3 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP09300/Label_NY_NP7110_2.55-3 | 8OP09300 | Noxzema Ultimate Clear Anti-Blemish Pads -Front | Noxzema Ultimate Clear Anti-Blemish Pads -Front | Label_NY_NP7110_2.55-3 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP09300/Label_NY_NP7110_2.55-3 | 8OP09300 | Noxzema Ultimate Clear Anti-Blemish Pads -Front | Noxzema Ultimate Clear Anti-Blemish Pads -Front | Label_NY_NP7110_2.55-3 | NY-Analytical | 770 | Label Inspection | Width | 3 | in | FDA,4 | 2.9375 | 3.0625 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP09300/Label_NY_NP7110_2.55-3 | 8OP09300 | Noxzema Ultimate Clear Anti-Blemish Pads -Front | Noxzema Ultimate Clear Anti-Blemish Pads -Front | Label_NY_NP7110_2.55-3 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Visual | Visual | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP09301/Label_NY_NP7110_2.55-5 | 8OP09301 | Noxzema Ultimate Clear Anti-Blemish Pads -Back | Noxzema Ultimate Clear Anti-Blemish Pads -Back | Label_NY_NP7110_2.55-5 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP09301/Label_NY_NP7110_2.55-5 | 8OP09301 | Noxzema Ultimate Clear Anti-Blemish Pads -Back | Noxzema Ultimate Clear Anti-Blemish Pads -Back | Label_NY_NP7110_2.55-5 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP09301/Label_NY_NP7110_2.55-5 | 8OP09301 | Noxzema Ultimate Clear Anti-Blemish Pads -Back | Noxzema Ultimate Clear Anti-Blemish Pads -Back | Label_NY_NP7110_2.55-5 | NY-Analytical | 770 | Label Inspection | Length | 2.5313 | in | FDA,4 | 2.4875 | 2.6125 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP09301/Label_NY_NP7110_2.55-5 | 8OP09301 | Noxzema Ultimate Clear Anti-Blemish Pads -Back | Noxzema Ultimate Clear Anti-Blemish Pads -Back | Label_NY_NP7110_2.55-5 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP09301/Label_NY_NP7110_2.55-5 | 8OP09301 | Noxzema Ultimate Clear Anti-Blemish Pads -Back | Noxzema Ultimate Clear Anti-Blemish Pads -Back | Label_NY_NP7110_2.55-5 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP09301/Label_NY_NP7110_2.55-5 | 8OP09301 | Noxzema Ultimate Clear Anti-Blemish Pads -Back | Noxzema Ultimate Clear Anti-Blemish Pads -Back | Label_NY_NP7110_2.55-5 | NY-Analytical | 770 | Label Inspection | Width | 4.9687 | in | FDA,4 | 4.9375 | 5.0625 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP09301/Label_NY_NP7110_2.55-5 | 8OP09301 | Noxzema Ultimate Clear Anti-Blemish Pads -Back | Noxzema Ultimate Clear Anti-Blemish Pads -Back | Label_NY_NP7110_2.55-5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Visual | Visual | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP227RSL/Label_NY_NP9997_2.625-3.375 | 8OP227RSL | LBL Rigid L' Oreal CCL NP9997 | Rigid L'Oreal Ccl Np9997 | Label_NY_NP9997_2.625-3.375 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | [5 empty] | Label Inspection |  | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP227RSL/Label_NY_NP9997_2.625-3.375 | 8OP227RSL | LBL Rigid L' Oreal CCL NP9997 | Rigid L'Oreal Ccl Np9997 | Label_NY_NP9997_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP227RSL/Label_NY_NP9997_2.625-3.375 | 8OP227RSL | LBL Rigid L' Oreal CCL NP9997 | Rigid L'Oreal Ccl Np9997 | Label_NY_NP9997_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP227RSL/Label_NY_NP9997_2.625-3.375 | 8OP227RSL | LBL Rigid L' Oreal CCL NP9997 | Rigid L'Oreal Ccl Np9997 | Label_NY_NP9997_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Length | 2.63 | in | FDA,4 | 2.5625 | 2.6875 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP227RSL/Label_NY_NP9997_2.625-3.375 | 8OP227RSL | LBL Rigid L' Oreal CCL NP9997 | Rigid L'Oreal Ccl Np9997 | Label_NY_NP9997_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP227RSL/Label_NY_NP9997_2.625-3.375 | 8OP227RSL | LBL Rigid L' Oreal CCL NP9997 | Rigid L'Oreal Ccl Np9997 | Label_NY_NP9997_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP227RSL/Label_NY_NP9997_2.625-3.375 | 8OP227RSL | LBL Rigid L' Oreal CCL NP9997 | Rigid L'Oreal Ccl Np9997 | Label_NY_NP9997_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Width | 3.37 | in | FDA,4 | 3.3125 | 3.4375 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP260RSL/Label_NY_NP7700_3.375-2.3437 | 8OP260RSL | Avon ANEW Clean Facial Cleans. Wipes Reseal. | Avon Anew Clean Facial Cleans. Wipes Reseal. | Label_NY_NP7700_3.375-2.3437 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP260RSL/Label_NY_NP7700_3.375-2.3437 | 8OP260RSL | Avon ANEW Clean Facial Cleans. Wipes Reseal. | Avon Anew Clean Facial Cleans. Wipes Reseal. | Label_NY_NP7700_3.375-2.3437 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP260RSL/Label_NY_NP7700_3.375-2.3437 | 8OP260RSL | Avon ANEW Clean Facial Cleans. Wipes Reseal. | Avon Anew Clean Facial Cleans. Wipes Reseal. | Label_NY_NP7700_3.375-2.3437 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP260RSL/Label_NY_NP7700_3.375-2.3437 | 8OP260RSL | Avon ANEW Clean Facial Cleans. Wipes Reseal. | Avon Anew Clean Facial Cleans. Wipes Reseal. | Label_NY_NP7700_3.375-2.3437 | NY-Analytical | 770 | Label Inspection | Length | 3.3125 | in | FDA,4 | 3.3125 | 3.4375 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP260RSL/Label_NY_NP7700_3.375-2.3437 | 8OP260RSL | Avon ANEW Clean Facial Cleans. Wipes Reseal. | Avon Anew Clean Facial Cleans. Wipes Reseal. | Label_NY_NP7700_3.375-2.3437 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP260RSL/Label_NY_NP7700_3.375-2.3437 | 8OP260RSL | Avon ANEW Clean Facial Cleans. Wipes Reseal. | Avon Anew Clean Facial Cleans. Wipes Reseal. | Label_NY_NP7700_3.375-2.3437 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP260RSL/Label_NY_NP7700_3.375-2.3437 | 8OP260RSL | Avon ANEW Clean Facial Cleans. Wipes Reseal. | Avon Anew Clean Facial Cleans. Wipes Reseal. | Label_NY_NP7700_3.375-2.3437 | NY-Analytical | 770 | Label Inspection | Width | 2.2812 | in | FDA,4 | 2.2812 | 2.4062 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP263CK/Label_NY_NP7030_8.5-6.5 | 8OP263CK | Ecolab Eco-Wipe Duo Case Label | Ecolab Eco-Wipe Duo Case Label | Label_NY_NP7030_8.5-6.5 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP263CK/Label_NY_NP7030_8.5-6.5 | 8OP263CK | Ecolab Eco-Wipe Duo Case Label | Ecolab Eco-Wipe Duo Case Label | Label_NY_NP7030_8.5-6.5 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP263CK/Label_NY_NP7030_8.5-6.5 | 8OP263CK | Ecolab Eco-Wipe Duo Case Label | Ecolab Eco-Wipe Duo Case Label | Label_NY_NP7030_8.5-6.5 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP263CK/Label_NY_NP7030_8.5-6.5 | 8OP263CK | Ecolab Eco-Wipe Duo Case Label | Ecolab Eco-Wipe Duo Case Label | Label_NY_NP7030_8.5-6.5 | NY-Analytical | 770 | Label Inspection | Length | 8.5 | in | FDA,4 | 8.4375 | 8.5625 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP263CK/Label_NY_NP7030_8.5-6.5 | 8OP263CK | Ecolab Eco-Wipe Duo Case Label | Ecolab Eco-Wipe Duo Case Label | Label_NY_NP7030_8.5-6.5 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP263CK/Label_NY_NP7030_8.5-6.5 | 8OP263CK | Ecolab Eco-Wipe Duo Case Label | Ecolab Eco-Wipe Duo Case Label | Label_NY_NP7030_8.5-6.5 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP263CK/Label_NY_NP7030_8.5-6.5 | 8OP263CK | Ecolab Eco-Wipe Duo Case Label | Ecolab Eco-Wipe Duo Case Label | Label_NY_NP7030_8.5-6.5 | NY-Analytical | 770 | Label Inspection | Width | 6.5 | in | FDA,4 | 6.4375 | 6.5625 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP308SIRF/Label_NY_NP9300_2-1.375 | 8OP308SIRF | Remove Completely Label | Remove Completely Label | Label_NY_NP9300_2-1.375 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP308SIRF/Label_NY_NP9300_2-1.375 | 8OP308SIRF | Remove Completely Label | Remove Completely Label | Label_NY_NP9300_2-1.375 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP308SIRF/Label_NY_NP9300_2-1.375 | 8OP308SIRF | Remove Completely Label | Remove Completely Label | Label_NY_NP9300_2-1.375 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP308SIRF/Label_NY_NP9300_2-1.375 | 8OP308SIRF | Remove Completely Label | Remove Completely Label | Label_NY_NP9300_2-1.375 | NY-Analytical | 770 | Label Inspection | Length | 2 | in | FDA,4 | 1.9375 | 2.0625 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP308SIRF/Label_NY_NP9300_2-1.375 | 8OP308SIRF | Remove Completely Label | Remove Completely Label | Label_NY_NP9300_2-1.375 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP308SIRF/Label_NY_NP9300_2-1.375 | 8OP308SIRF | Remove Completely Label | Remove Completely Label | Label_NY_NP9300_2-1.375 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP308SIRF/Label_NY_NP9300_2-1.375 | 8OP308SIRF | Remove Completely Label | Remove Completely Label | Label_NY_NP9300_2-1.375 | NY-Analytical | 770 | Label Inspection | Width | 1.375 | in | FDA,4 | 1.3125 | 1.4375 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP585SIRF/Label_NY_NP9300_1.375-2 | 8OP585SIRF | Remove Completely Label | Remove Completely Label | Label_NY_NP9300_1.375-2 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP585SIRF/Label_NY_NP9300_1.375-2 | 8OP585SIRF | Remove Completely Label | Remove Completely Label | Label_NY_NP9300_1.375-2 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP585SIRF/Label_NY_NP9300_1.375-2 | 8OP585SIRF | Remove Completely Label | Remove Completely Label | Label_NY_NP9300_1.375-2 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP585SIRF/Label_NY_NP9300_1.375-2 | 8OP585SIRF | Remove Completely Label | Remove Completely Label | Label_NY_NP9300_1.375-2 | NY-Analytical | 770 | Label Inspection | Length | 1.37 | in | FDA,4 | 1.3125 | 1.4375 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP585SIRF/Label_NY_NP9300_1.375-2 | 8OP585SIRF | Remove Completely Label | Remove Completely Label | Label_NY_NP9300_1.375-2 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP585SIRF/Label_NY_NP9300_1.375-2 | 8OP585SIRF | Remove Completely Label | Remove Completely Label | Label_NY_NP9300_1.375-2 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OP585SIRF/Label_NY_NP9300_1.375-2 | 8OP585SIRF | Remove Completely Label | Remove Completely Label | Label_NY_NP9300_1.375-2 | NY-Analytical | 770 | Label Inspection | Width | 1.99 | in | FDA,4 | 1.9375 | 2.0625 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OX351CK/Label_NY_50#W_5_4 | 8OX351CK | Prevantics -OR Tri-Fold IFU | Prevantics-OR Tri-Fold Ifu | Label_NY_50#W_5_4 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) |  |  |  |  | Label Inspection |  | NY | RAW MATERIAL - NY | RM_LABEL | both
8OX351CK/Label_NY_50#W_5_4 | 8OX351CK | Prevantics -OR Tri-Fold IFU | Prevantics-OR Tri-Fold Ifu | Label_NY_50#W_5_4 | NY-Analytical | 770 | Label Inspection | Appearance | Material must be clean and free of particulate matter (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OX351CK/Label_NY_50#W_5_4 | 8OX351CK | Prevantics -OR Tri-Fold IFU | Prevantics-OR Tri-Fold Ifu | Label_NY_50#W_5_4 | NY-Analytical | 770 | Label Inspection | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OX351CK/Label_NY_50#W_5_4 | 8OX351CK | Prevantics -OR Tri-Fold IFU | Prevantics-OR Tri-Fold Ifu | Label_NY_50#W_5_4 | NY-Analytical | 770 | Label Inspection | Length | 5.01 | in | FDA,4 | 4.9375 | 5.0625 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OX351CK/Label_NY_50#W_5_4 | 8OX351CK | Prevantics -OR Tri-Fold IFU | Prevantics-OR Tri-Fold Ifu | Label_NY_50#W_5_4 | NY-Analytical | 770 | Label Inspection | NP Structure | 50# White Paper Stock (Done) | in |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OX351CK/Label_NY_50#W_5_4 | 8OX351CK | Prevantics -OR Tri-Fold IFU | Prevantics-OR Tri-Fold Ifu | Label_NY_50#W_5_4 | NY-Analytical | 770 | Label Inspection | Odor | No solvent odor is to be present (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OX351CK/Label_NY_50#W_5_4 | 8OX351CK | Prevantics -OR Tri-Fold IFU | Prevantics-OR Tri-Fold Ifu | Label_NY_50#W_5_4 | NY-Analytical | 770 | Label Inspection | Print | No smears, streaks, spots, or voids (Done) |  |  |  |  | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
8OX351CK/Label_NY_50#W_5_4 | 8OX351CK | Prevantics -OR Tri-Fold IFU | Prevantics-OR Tri-Fold Ifu | Label_NY_50#W_5_4 | NY-Analytical | 770 | Label Inspection | Width | 3.99 | in | FDA,4 | 3.9375 | 4.0625 | TM 143 / Visual | TM143 | NY | RAW MATERIAL - NY | RM_LABEL | both
| 9OX465GL | Hot Melt Adhesive, BAM 418 |  | Analytical - NY_Glue_01 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) |  |  |  |  | Visual | Visual |  |  |  | left_only
| 9OX465GL | Hot Melt Adhesive, BAM 418 |  | Analytical - NY_Glue_01 | NY-Analytical | 395 | Color | Color | Very Pale Amber (Done) |  |  |  |  | Visual | Visual |  |  |  | left_only

## MatchedMainPlans
Mat_MatCode | plan.f | SP_TestPlan | Test_Lab | Res_TestCode | TestName | j | Analyte.C | Numeric_Result | Unnamed: 9 | Res_Units | Rounding_Rule | Low_Release_Limit | High_Release_Limit | METHOD | First_merge
2AA284SLI | 2AA284SLI/Lid_01B | Lid_01B | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, | [5 empty] | Canister & Lid Inspection | Main_Plans
2AA284SLI | 2AA284SLI/Lid_01B | Lid_01B | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2AA284SLI | 2AA284SLI/Lid_01B | Lid_01B | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2AA284SLI | 2AA284SLI/Lid_01B | Lid_01B | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2AA284SLI | 2AA284SLI/Lid_01B | Lid_01B | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.6875 |  | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | Main_Plans
2AA284SLI | 2AA284SLI/Lid_01B | Lid_01B | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.875 |  | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | Main_Plans
2AA284SLI | 2AA284SLI/Lid_01B | Lid_01B | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2AA347 | 2AA347/Web_NP8550_10.25 | Web_NP8550_10.25 | NY-Analytical | 885 | Web Inspection | AQL Sample Size | AQL Sample Size | 315 |  |  | FDA,0 | 0 |  | TM101/TM107/TM132/Measure | Main_Plans
2AA347 | 2AA347/Web_NP8550_10.25 | Web_NP8550_10.25 | NY-Analytical | 885 | Web Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Sample(s) all meet specifications (Done) | [5 empty] | TM101/TM107/TM132/Measure | Main_Plans
2AA347 | 2AA347/Web_NP8550_10.25 | Web_NP8550_10.25 | NY-Analytical | 885 | Web Inspection | Basis_Weight | Basis Weight | 41.3 |  | #/Ream | FDA,1 | 37.8 | 48.2 | TM101/TM107/TM132/Measure | Main_Plans
2AA347 | 2AA347/Web_NP8550_10.25 | Web_NP8550_10.25 | NY-Analytical | 885 | Web Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | TM101/TM107/TM132/Measure | Main_Plans
2AA347 | 2AA347/Web_NP8550_10.25 | Web_NP8550_10.25 | NY-Analytical | 885 | Web Inspection | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 5762.95 |  | g/in | FDA,0 | 3500 |  | TM101/TM107/TM132/Measure | Main_Plans
2AA347 | 2AA347/Web_NP8550_10.25 | Web_NP8550_10.25 | NY-Analytical | 885 | Web Inspection | Web Width (Inches) | Web Width | 10.25 |  | in | FDA,4 | 10.1875 | 10.3125 | TM101/TM107/TM132/Measure | Main_Plans
2AA592 | 2AA592/Web_NP8504_11.0 | Web_NP8504_11.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Dyne Level | Main_Plans
2AA592 | 2AA592/Web_NP8504_11.0 | Web_NP8504_11.0 | NY-Analytical | 836 | Dyne Level | Dyne Level (outside film) | Dyne Level (outside film) | 36 |  |  | FDA,0 | 36 | 38 | Dyne Test Kit | Main_Plans
2AA592 | 2AA592/Web_NP8504_11.0 | Web_NP8504_11.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 10.968 |  | in | FDA,4 | 10.9375 | 11.0625 | Measurement | Main_Plans
2AA592 | 2AA592/Web_NP8504_11.0 | Web_NP8504_11.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 57.54 |  | #/Ream | FDA,1 | 53.3 | 65.1 | TM 101 | Main_Plans
2AA592 | 2AA592/Web_NP8504_11.0 | Web_NP8504_11.0 | NY-Analytical | 762 | Seal Strength | Seal Strength(g/in@350F/30psi/0.5s dwell Top Jaw) | Seal Strength (@350°F/30psi/0.5s dwell Top Jaw) | 5339.63 |  | g/in | FDA,0 | 2700 |  | TM 107 | Main_Plans
2AA592 | 2AA592/Web_NP8504_11.0 | Web_NP8504_11.0 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 3.65 |  | mil | FDA,1 | 3.2 | 3.9 | TM 132 | Main_Plans
2AA592SLI | 2AA592SLI/Lid_NY_NP7600_22R1 | Lid_NY_NP7600_22R1 | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, | [5 empty] | Canister & Lid Inspection | Main_Plans
2AA592SLI | 2AA592SLI/Lid_NY_NP7600_22R1 | Lid_NY_NP7600_22R1 | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) | [5 empty] | Solo Lid Inspection | Main_Plans
2AA592SLI | 2AA592SLI/Lid_NY_NP7600_22R1 | Lid_NY_NP7600_22R1 | NY-Analytical | 818 | Solo Lid Inspection | Color | Color | Meets the specification (Done) | [5 empty] | Solo Lid Inspection | Main_Plans
2AA592SLI | 2AA592SLI/Lid_NY_NP7600_22R1 | Lid_NY_NP7600_22R1 | NY-Analytical | 818 | Solo Lid Inspection | Functionality | Functionality | Meet specification (Done) | [5 empty] | Solo Lid Inspection | Main_Plans
2AA592SLI | 2AA592SLI/Lid_NY_NP7600_22R1 | Lid_NY_NP7600_22R1 | NY-Analytical | 818 | Solo Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.028 |  | in | FDA,3 | 0.025 | 0.035 | Solo Lid Inspection | Main_Plans
2AA592SLI | 2AA592SLI/Lid_NY_NP7600_22R1 | Lid_NY_NP7600_22R1 | NY-Analytical | 818 | Solo Lid Inspection | Open Lid Overall Width | Open Lid Overall Width | 4.36 |  | in | FDA,3 | 4.351 | 4.401 | Solo Lid Inspection | Main_Plans
2AA592SLI | 2AA592SLI/Lid_NY_NP7600_22R1 | Lid_NY_NP7600_22R1 | NY-Analytical | 818 | Solo Lid Inspection | Overall closed lid height | Overall closed lid height | 0.37 |  | in | FDA,3 | 0.355 | 0.385 | Solo Lid Inspection | Main_Plans
2AA592SLI | 2AA592SLI/Lid_NY_NP7600_22R1 | Lid_NY_NP7600_22R1 | NY-Analytical | 818 | Solo Lid Inspection | Weight | Weight | 11.241 |  | g | FDA,1 | 10.5 | 12.5 | Measurement | Main_Plans
2AA592SLI | 2AA592SLI/Lid_NY_NP7600_22R1 | Lid_NY_NP7600_22R1 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2AA592SLI | 2AA592SLI/Lid_NY_NP7600_22R1 | Lid_NY_NP7600_22R1 | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2AA592SLI | 2AA592SLI/Lid_NY_NP7600_22R1 | Lid_NY_NP7600_22R1 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.7025 |  | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | Main_Plans
2AA592SLI | 2AA592SLI/Lid_NY_NP7600_22R1 | Lid_NY_NP7600_22R1 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.8705 |  | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | Main_Plans
2AA592SLI | 2AA592SLI/Lid_NY_NP7600_22R1 | Lid_NY_NP7600_22R1 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2AA632 | 2AA632/Web_NP8504_11.75 | Web_NP8504_11.75 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Dyne Level | Main_Plans
2AA632 | 2AA632/Web_NP8504_11.75 | Web_NP8504_11.75 | NY-Analytical | 836 | Dyne Level | Dyne Level (outside film) | Dyne Level (outside film) | 38 |  |  | FDA,0 | 36 | 38 | Dyne Test Kit | Main_Plans
2AA632 | 2AA632/Web_NP8504_11.75 | Web_NP8504_11.75 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 11.75 |  | in | FDA,4 | 11.6875 | 11.8125 | Measurement | Main_Plans
2AA632 | 2AA632/Web_NP8504_11.75 | Web_NP8504_11.75 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 59.05 |  | #/Ream | FDA,1 | 53.3 | 65.1 | TM 101 | Main_Plans
2AA632 | 2AA632/Web_NP8504_11.75 | Web_NP8504_11.75 | NY-Analytical | 762 | Seal Strength | Seal Strength(g/in@350F/30psi/0.5s dwell Top Jaw) | Seal Strength (@350°F/30psi/0.5s dwell Top Jaw) | 5664 |  | g/in | FDA,0 | 2700 |  | TM 107 | Main_Plans
2AA632 | 2AA632/Web_NP8504_11.75 | Web_NP8504_11.75 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 3.7 |  | mil | FDA,1 | 3.2 | 3.9 | TM 132 | Main_Plans
2AA714SLI | 2AA714SLI/Lid_01V | Lid_01V | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, | [5 empty] | Canister & Lid Inspection | Main_Plans
2AA714SLI | 2AA714SLI/Lid_01V | Lid_01V | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2AA714SLI | 2AA714SLI/Lid_01V | Lid_01V | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2AA714SLI | 2AA714SLI/Lid_01V | Lid_01V | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2AA714SLI | 2AA714SLI/Lid_01V | Lid_01V | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.687 |  | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | Main_Plans
2AA714SLI | 2AA714SLI/Lid_01V | Lid_01V | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.8125 |  | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | Main_Plans
2AA714SLI | 2AA714SLI/Lid_01V | Lid_01V | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2AA922SLI | 2AA922SLI/Lid_01V | Lid_01V | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Certificate of Analysis | Main_Plans
2AD173 | 2AD173/Web_NP5115_4.5 | Web_NP5115_4.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2AD173 | 2AD173/Web_NP5115_4.5 | Web_NP5115_4.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 4.5 |  | in | FDA,4 | 4.4375 | 4.5625 | Measurement | Main_Plans
2AD173 | 2AD173/Web_NP5115_4.5 | Web_NP5115_4.5 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Printed/Varnished) | Basis Weight | 59.5 |  | #/Ream | FDA,1 | 55.6 | 68 | TM 101 | Main_Plans
2AD173 | 2AD173/Web_NP5115_4.5 | Web_NP5115_4.5 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@325DegF/40PSI/0.5Sec_Dwell) | Seal Strength (@325°F/40PSI/0.5Sec Dwell) | 2057.76 |  | g/in | FDA,0 | 1500 |  | TM 107 | Main_Plans
2AJ143SLI | 2AJ143SLI/Lid_NY_NP7600_01C | Lid_NY_NP7600_01C | IN-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, | [5 empty] | Canister & Lid Inspection | Main_Plans
2AJ143SLI | 2AJ143SLI/Lid_NY_NP7600_01C | Lid_NY_NP7600_01C | IN-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2AJ143SLI | 2AJ143SLI/Lid_NY_NP7600_01C | Lid_NY_NP7600_01C | IN-Analytical | 765 | Canister & Lid Inspection | AQL Sample Size | AQL Sample Size | 200 |  |  |  | 0 |  | TM 143 / Visual | Main_Plans
2AJ143SLI | 2AJ143SLI/Lid_NY_NP7600_01C | Lid_NY_NP7600_01C | IN-Analytical | 765 | Canister & Lid Inspection | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2AJ143SLI | 2AJ143SLI/Lid_NY_NP7600_01C | Lid_NY_NP7600_01C | IN-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2AJ143SLI | 2AJ143SLI/Lid_NY_NP7600_01C | Lid_NY_NP7600_01C | IN-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.7088 |  | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | Main_Plans
2AJ143SLI | 2AJ143SLI/Lid_NY_NP7600_01C | Lid_NY_NP7600_01C | IN-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.8438 |  | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | Main_Plans
2AJ143SLI | 2AJ143SLI/Lid_NY_NP7600_01C | Lid_NY_NP7600_01C | IN-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2AJ170 | 2AJ170/Web_NP8530_11.0 | Web_NP8530_11.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2AJ170 | 2AJ170/Web_NP8530_11.0 | Web_NP8530_11.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 11 |  | in | FDA,4 | 10.9375 | 11.0625 | Measurement | Main_Plans
2AJ170 | 2AJ170/Web_NP8530_11.0 | Web_NP8530_11.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 34.39 |  | #/Ream | FDA,2 | 29.4 | 37.42 | TM 101 | Main_Plans
2AJ170 | 2AJ170/Web_NP8530_11.0 | Web_NP8530_11.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 5442 |  | g/in | FDA,0 | 3500 |  | TM 107 | Main_Plans
2AJ170 | 2AJ170/Web_NP8530_11.0 | Web_NP8530_11.0 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 1.9 |  | mil | FDA,1 | 1.8 | 2.2 | TM 132 | Main_Plans
2AJ170SLI | 2AJ170SLI/Lid_NY_NP7600_22FA | Lid_NY_NP7600_22FA | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, | [5 empty] | Canister & Lid Inspection | Main_Plans
2AJ170SLI | 2AJ170SLI/Lid_NY_NP7600_22FA | Lid_NY_NP7600_22FA | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) | [5 empty] | Solo Lid Inspection | Main_Plans
2AJ170SLI | 2AJ170SLI/Lid_NY_NP7600_22FA | Lid_NY_NP7600_22FA | NY-Analytical | 818 | Solo Lid Inspection | Color | Color | Meets the specification (Done) | [5 empty] | Solo Lid Inspection | Main_Plans
2AJ170SLI | 2AJ170SLI/Lid_NY_NP7600_22FA | Lid_NY_NP7600_22FA | NY-Analytical | 818 | Solo Lid Inspection | Functionality | Functionality | Meet specification (Done) | [5 empty] | Solo Lid Inspection | Main_Plans
2AJ170SLI | 2AJ170SLI/Lid_NY_NP7600_22FA | Lid_NY_NP7600_22FA | NY-Analytical | 818 | Solo Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.028 |  | in | FDA,3 | 0.025 | 0.035 | Solo Lid Inspection | Main_Plans
2AJ170SLI | 2AJ170SLI/Lid_NY_NP7600_22FA | Lid_NY_NP7600_22FA | NY-Analytical | 818 | Solo Lid Inspection | Open Lid Overall Width | Open Lid Overall Width | 4.36 |  | in | FDA,3 | 4.351 | 4.401 | Solo Lid Inspection | Main_Plans
2AJ170SLI | 2AJ170SLI/Lid_NY_NP7600_22FA | Lid_NY_NP7600_22FA | NY-Analytical | 818 | Solo Lid Inspection | Overall closed lid height | Overall closed lid height | 0.37 |  | in | FDA,3 | 0.355 | 0.385 | Solo Lid Inspection | Main_Plans
2AJ170SLI | 2AJ170SLI/Lid_NY_NP7600_22FA | Lid_NY_NP7600_22FA | NY-Analytical | 818 | Solo Lid Inspection | Weight | Weight | 11.25 |  | g | FDA,1 | 10.5 | 12.5 | Measurement | Main_Plans
2AJ170SLI | 2AJ170SLI/Lid_NY_NP7600_22FA | Lid_NY_NP7600_22FA | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2AJ170SLI | 2AJ170SLI/Lid_NY_NP7600_22FA | Lid_NY_NP7600_22FA | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2AJ170SLI | 2AJ170SLI/Lid_NY_NP7600_22FA | Lid_NY_NP7600_22FA | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.67 |  | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | Main_Plans
2AJ170SLI | 2AJ170SLI/Lid_NY_NP7600_22FA | Lid_NY_NP7600_22FA | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.84 |  | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | Main_Plans
2AJ170SLI | 2AJ170SLI/Lid_NY_NP7600_22FA | Lid_NY_NP7600_22FA | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2AJ190SLI | 2AJ190SLI/Lid_NY_NP7600_22V | Lid_NY_NP7600_22V | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, | [5 empty] | Canister & Lid Inspection | Main_Plans
2AJ190SLI | 2AJ190SLI/Lid_NY_NP7600_22V | Lid_NY_NP7600_22V | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) | [5 empty] | Solo Lid Inspection | Main_Plans
2AJ190SLI | 2AJ190SLI/Lid_NY_NP7600_22V | Lid_NY_NP7600_22V | NY-Analytical | 818 | Solo Lid Inspection | Appearance | Appearance | Meet specification (Done) | [5 empty] | Solo Lid Inspection | Main_Plans
2AJ190SLI | 2AJ190SLI/Lid_NY_NP7600_22V | Lid_NY_NP7600_22V | NY-Analytical | 818 | Solo Lid Inspection | Color | Color | Meets the specification (Done) | [5 empty] | Solo Lid Inspection | Main_Plans
2AJ190SLI | 2AJ190SLI/Lid_NY_NP7600_22V | Lid_NY_NP7600_22V | NY-Analytical | 818 | Solo Lid Inspection | Functionality | Functionality | Meet specification (Done) | [5 empty] | Solo Lid Inspection | Main_Plans
2AJ190SLI | 2AJ190SLI/Lid_NY_NP7600_22V | Lid_NY_NP7600_22V | NY-Analytical | 818 | Solo Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.0295 |  | in | FDA,3 | 0.025 | 0.035 | Solo Lid Inspection | Main_Plans
2AJ190SLI | 2AJ190SLI/Lid_NY_NP7600_22V | Lid_NY_NP7600_22V | NY-Analytical | 818 | Solo Lid Inspection | Odor | Odor | None detectable (Done) | [5 empty] | Solo Lid Inspection | Main_Plans
2AJ190SLI | 2AJ190SLI/Lid_NY_NP7600_22V | Lid_NY_NP7600_22V | NY-Analytical | 818 | Solo Lid Inspection | Open Lid Overall Width | Open Lid Overall Width | 4.374 |  | in | FDA,3 | 4.351 | 4.401 | Solo Lid Inspection | Main_Plans
2AJ190SLI | 2AJ190SLI/Lid_NY_NP7600_22V | Lid_NY_NP7600_22V | NY-Analytical | 818 | Solo Lid Inspection | Overall closed lid height | Overall closed lid height | 0.373 |  | in | FDA,3 | 0.355 | 0.385 | Solo Lid Inspection | Main_Plans
2AJ190SLI | 2AJ190SLI/Lid_NY_NP7600_22V | Lid_NY_NP7600_22V | NY-Analytical | 818 | Solo Lid Inspection | Weight | Weight | 11.19 |  | g | FDA,1 | 10.5 | 12.5 | Measurement | Main_Plans
2AJ190SLI | 2AJ190SLI/Lid_NY_NP7600_22V | Lid_NY_NP7600_22V | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2AJ190SLI | 2AJ190SLI/Lid_NY_NP7600_22V | Lid_NY_NP7600_22V | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.7188 |  | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | Main_Plans
2AJ190SLI | 2AJ190SLI/Lid_NY_NP7600_22V | Lid_NY_NP7600_22V | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.875 |  | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | Main_Plans
2AJ324SLI | 2AJ324SLI/Lid_NY_NP7600_22X | Lid_NY_NP7600_22X | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, | [5 empty] | Canister & Lid Inspection | Main_Plans
2AJ324SLI | 2AJ324SLI/Lid_NY_NP7600_22X | Lid_NY_NP7600_22X | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) | [5 empty] | Solo Lid Inspection | Main_Plans
2AJ324SLI | 2AJ324SLI/Lid_NY_NP7600_22X | Lid_NY_NP7600_22X | NY-Analytical | 818 | Solo Lid Inspection | Color | Color | Meets the specification (Done) | [5 empty] | Solo Lid Inspection | Main_Plans
2AJ324SLI | 2AJ324SLI/Lid_NY_NP7600_22X | Lid_NY_NP7600_22X | NY-Analytical | 818 | Solo Lid Inspection | Functionality | Functionality | Meet specification (Done) | [5 empty] | Solo Lid Inspection | Main_Plans
2AJ324SLI | 2AJ324SLI/Lid_NY_NP7600_22X | Lid_NY_NP7600_22X | NY-Analytical | 818 | Solo Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.0297 |  | in | FDA,3 | 0.025 | 0.035 | Solo Lid Inspection | Main_Plans
2AJ324SLI | 2AJ324SLI/Lid_NY_NP7600_22X | Lid_NY_NP7600_22X | NY-Analytical | 818 | Solo Lid Inspection | Open Lid Overall Width | Open Lid Overall Width | 4.371 |  | in | FDA,3 | 4.351 | 4.401 | Solo Lid Inspection | Main_Plans
2AJ324SLI | 2AJ324SLI/Lid_NY_NP7600_22X | Lid_NY_NP7600_22X | NY-Analytical | 818 | Solo Lid Inspection | Overall closed lid height | Overall closed lid height | 0.3705 |  | in | FDA,3 | 0.355 | 0.385 | Solo Lid Inspection | Main_Plans
2AJ324SLI | 2AJ324SLI/Lid_NY_NP7600_22X | Lid_NY_NP7600_22X | NY-Analytical | 818 | Solo Lid Inspection | Weight | Weight | 11.385 |  | g | FDA,1 | 10.5 | 12.5 | Measurement | Main_Plans
2AJ324SLI | 2AJ324SLI/Lid_NY_NP7600_22X | Lid_NY_NP7600_22X | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2AJ324SLI | 2AJ324SLI/Lid_NY_NP7600_22X | Lid_NY_NP7600_22X | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2AJ324SLI | 2AJ324SLI/Lid_NY_NP7600_22X | Lid_NY_NP7600_22X | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.7088 |  | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | Main_Plans
2AJ324SLI | 2AJ324SLI/Lid_NY_NP7600_22X | Lid_NY_NP7600_22X | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.875 |  | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | Main_Plans
2AJ324SLI | 2AJ324SLI/Lid_NY_NP7600_22X | Lid_NY_NP7600_22X | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2AP027 | 2AP027/Web_NP8512_310 | Web_NP8512_310 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2AP027 | 2AP027/Web_NP8512_310 | Web_NP8512_310 | NY-Analytical | 761 | Web Measurement | Web Width (mm) | Web Width | 310 |  | mm | FDA,0 | 308 | 312 | Measurement | Main_Plans
2AP027 | 2AP027/Web_NP8512_310 | Web_NP8512_310 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/inch@130DegC/14psi/1.0 sec dwell) | Seal Strength (@130°C/14psi/1.0 sec dwell) | 8798 |  | g/in | FDA,0 | 2000 |  | TM 107 | Main_Plans
2AP027 | 2AP027/Web_NP8512_310 | Web_NP8512_310 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 3.5 |  | mil | FDA,2 | 3.06 | 3.74 | TM 132 | Main_Plans
2AP027 | 2AP027/Web_NP8512_310 | Web_NP8512_310 | NY-Analytical | 748 | Visual Inspection | Color/Copy | Color/Copy | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
2AP075 | 2AP075/Web_NP8542_12.205 | Web_NP8542_12.205 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2AP075 | 2AP075/Web_NP8542_12.205 | Web_NP8542_12.205 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 12.25 |  | in | FDA,4 | 12.1425 | 12.2675 | Measurement | Main_Plans
2AP075 | 2AP075/Web_NP8542_12.205 | Web_NP8542_12.205 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 60.85 |  | #/Ream | FDA,2 | 52.53 | 64.21 | TM 101 | Main_Plans
2AP075 | 2AP075/Web_NP8542_12.205 | Web_NP8542_12.205 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/mm@130DegC/1bar/0.5sec dwell) | Seal Strength (@130°C/1bar/0.5sec dwell) | 7718 |  | g/mm | FDA,0 | 25 |  | TM 107 | Main_Plans
2AP087 | 2AP087/Web_NP8550_12.205 | Web_NP8550_12.205 | NY-Analytical | 885 | Web Inspection | AQL Sample Size | AQL Sample Size | 200 |  |  | FDA,0 | 0 |  | TM101/TM107/TM132/Measure | Main_Plans
2AP087 | 2AP087/Web_NP8550_12.205 | Web_NP8550_12.205 | NY-Analytical | 885 | Web Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Sample(s) all meet specifications (Done) | [5 empty] | TM101/TM107/TM132/Measure | Main_Plans
2AP087 | 2AP087/Web_NP8550_12.205 | Web_NP8550_12.205 | NY-Analytical | 885 | Web Inspection | Basis_Weight | Basis Weight | 43.3 |  | #/Ream | FDA,1 | 37.8 | 48.2 | TM101/TM107/TM132/Measure | Main_Plans
2AP087 | 2AP087/Web_NP8550_12.205 | Web_NP8550_12.205 | NY-Analytical | 885 | Web Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | TM101/TM107/TM132/Measure | Main_Plans
2AP087 | 2AP087/Web_NP8550_12.205 | Web_NP8550_12.205 | NY-Analytical | 885 | Web Inspection | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 5039.27 |  | g/in | FDA,0 | 3500 |  | TM101/TM107/TM132/Measure | Main_Plans
2AP087 | 2AP087/Web_NP8550_12.205 | Web_NP8550_12.205 | NY-Analytical | 885 | Web Inspection | UPC Verification | UPC Verification | UPC Scan is acceptable (Scan into comment section below - Done) | [5 empty] | TM101/TM107/TM132/Measure | Main_Plans
2AP087 | 2AP087/Web_NP8550_12.205 | Web_NP8550_12.205 | NY-Analytical | 885 | Web Inspection | Web Width (Inches) | Web Width | 12.1425 |  | in | FDA,4 | 12.1425 | 12.2675 | TM101/TM107/TM132/Measure | Main_Plans
2AP087 | 2AP087/Web_NP8550_12.205 | Web_NP8550_12.205 | NY-Analytical | 885 | Web Inspection | Wind Direction | Wind Direction | Meets specification (Indicate numeric value - Done) | [5 empty] | TM101/TM107/TM132/Measure | Main_Plans
2AP113 | 2AP113/Web_NP8542_11.0 | Web_NP8542_11.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2AP113 | 2AP113/Web_NP8542_11.0 | Web_NP8542_11.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 11 |  | in | FDA,4 | 10.9375 | 11.0625 | Measurement | Main_Plans
2AP113 | 2AP113/Web_NP8542_11.0 | Web_NP8542_11.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 61.11 |  | #/Ream | FDA,2 | 52.53 | 64.21 | TM 101 | Main_Plans
2AP113 | 2AP113/Web_NP8542_11.0 | Web_NP8542_11.0 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/mm@130DegC/1bar/0.5sec dwell) | Seal Strength (@130°C/1bar/0.5sec dwell) | 180.6 |  | g/mm | FDA,0 | 25 |  | TM 107 | Main_Plans
2AP211 | 2AP211/Web_NP8512_12.205 | Web_NP8512_12.205 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2AP211 | 2AP211/Web_NP8512_12.205 | Web_NP8512_12.205 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 12.25 |  | in | FDA,4 | 12.1425 | 12.2675 | Measurement | Main_Plans
2AP211 | 2AP211/Web_NP8512_12.205 | Web_NP8512_12.205 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/inch@130DegC/14psi/1.0 sec dwell) | Seal Strength (@130°C/14psi/1.0 sec dwell) | 7698.56 |  | g/in | FDA,0 | 2000 |  | TM 107 | Main_Plans
2AP211 | 2AP211/Web_NP8512_12.205 | Web_NP8512_12.205 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 3.5 |  | mil | FDA,2 | 3.06 | 3.74 | TM 132 | Main_Plans
2AP211 | 2AP211/Web_NP8512_12.205 | Web_NP8512_12.205 | NY-Analytical | 748 | Visual Inspection | Color/Copy | Color/Copy | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
2AP229 | 2AP229/Web_NP8512_9.5 | Web_NP8512_9.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2AP229 | 2AP229/Web_NP8512_9.5 | Web_NP8512_9.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 9.5 |  | in | FDA,4 | 9.4375 | 9.5625 | Measurement | Main_Plans
2AP229 | 2AP229/Web_NP8512_9.5 | Web_NP8512_9.5 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/inch@130DegC/14psi/1.0 sec dwell) | Seal Strength (@130°C/14psi/1.0 sec dwell) | 7419.71 |  | g/in | FDA,0 | 2000 |  | TM 107 | Main_Plans
2AP229 | 2AP229/Web_NP8512_9.5 | Web_NP8512_9.5 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 3.6 |  | mil | FDA,2 | 3.06 | 3.74 | TM 132 | Main_Plans
2AP229 | 2AP229/Web_NP8512_9.5 | Web_NP8512_9.5 | NY-Analytical | 748 | Visual Inspection | Color/Copy | Color/Copy | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
2AP260 | 2AP260/Web_NP8535_12.75 | Web_NP8535_12.75 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2AP260 | 2AP260/Web_NP8535_12.75 | Web_NP8535_12.75 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 12.75 |  | in | FDA,4 | 12.6875 | 12.8125 | Measurement | Main_Plans
2AP260 | 2AP260/Web_NP8535_12.75 | Web_NP8535_12.75 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 41.14 |  | #/Ream | FDA,2 | 35.82 | 45.6 | TM 101 | Main_Plans
2AP260 | 2AP260/Web_NP8535_12.75 | Web_NP8535_12.75 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 6659.74 |  | g/in | FDA,0 | 3500 |  | TM 107 | Main_Plans
2AP260 | 2AP260/Web_NP8535_12.75 | Web_NP8535_12.75 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 2.4 |  | mil | FDA,1 | 2.2 | 2.8 | TM 132 | Main_Plans
2AP806 | 2AP806/Web_NP8430_310 | Web_NP8430_310 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2AP806 | 2AP806/Web_NP8430_310 | Web_NP8430_310 | NY-Analytical | 761 | Web Measurement | Web Width (mm) | Web Width | 311 |  | mm | FDA,0 | 308 | 312 | Measurement | Main_Plans
2AP806 | 2AP806/Web_NP8430_310 | Web_NP8430_310 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 51.9 |  | #/Ream | FDA,2 | 44.88 | 57.12 | TM 101 | Main_Plans
2AP806 | 2AP806/Web_NP8430_310 | Web_NP8430_310 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 7699.25 |  | g/in | FDA,0 | 3500 |  | TM 107 | Main_Plans
2AP806 | 2AP806/Web_NP8430_310 | Web_NP8430_310 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 3.05 |  | mil | FDA,2 | 2.64 | 3.36 | TM 132 | Main_Plans
2AU191 | 2AU191/Web_NP5115_7.5 | Web_NP5115_7.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2AU191 | 2AU191/Web_NP5115_7.5 | Web_NP5115_7.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 7.4375 |  | in | FDA,4 | 7.4375 | 7.5625 | Measurement | Main_Plans
2AU191 | 2AU191/Web_NP5115_7.5 | Web_NP5115_7.5 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Printed/Varnished) | Basis Weight | 60.5 |  | #/Ream | FDA,1 | 55.6 | 68 | TM 101 | Main_Plans
2AU191 | 2AU191/Web_NP5115_7.5 | Web_NP5115_7.5 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@325DegF/40PSI/0.5Sec_Dwell) | Seal Strength (@325°F/40PSI/0.5Sec Dwell) | 1823.34 |  | g/in | FDA,0 | 1500 |  | TM 107 | Main_Plans
2AU252 | 2AU252/Web_NP3459_8.0 | Web_NP3459_8.0 |  | 1044 | Basis Weight (lbs/ream) | Basis Weight | Basis Weight | 63.4 |  | Lb/Ream | FDA,2 | 57.56 | 70.4 | ASTM D 646 | Main_Plans
2AU252 | 2AU252/Web_NP3459_8.0 | Web_NP3459_8.0 |  | 757 | Caliper | caliper (COA) | Caliper | 3.8 |  | mil | FDA,3 | 3.515 | 3.515 | Caliper | Main_Plans
2AU252 | 2AU252/Web_NP3459_8.0 | Web_NP3459_8.0 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/inch@290DegF/40psi/0.5 sec dwell) | Seal Strength (@290°F/40psi/0.5 sec dwell) |  |  | g/in | FDA,0 | 7500 |  | TM 107 | Main_Plans
2AU252 | 2AU252/Web_NP3459_8.0 | Web_NP3459_8.0 |  | 761 | Web Measurement | Web Width (Inches) | Web Width | 8 |  | in | FDA,4 | 7.9375 | 8.0625 | TM 143 | Main_Plans
2AU252 | 2AU252/Web_NP3459_8.0 | Web_NP3459_8.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [6 empty] | Visual | Main_Plans
2AU252 | 2AU252/Web_NP3459_8.0 | Web_NP3459_8.0 |  | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | Visual Observation | Main_Plans
2AU252 | 2AU252/Web_NP3459_8.0 | Web_NP3459_8.0 |  | 748 | Visual Inspection | Color/Copy | Color/Copy | Meets specifications (Done) | [5 empty] | Visual Observation | Main_Plans
2AU555 | 2AU555/Web_NP3460_7.5 | Web_NP3460_7.5 | NY-Analytical | 1044 | Basis Weight (lbs/ream) | Basis Weight | Basis Weight | 72.89 |  | Lb/Ream | FDA,1 | 66.6 | 81.4 | ASTM D 646 | Main_Plans
2AU555 | 2AU555/Web_NP3460_7.5 | Web_NP3460_7.5 | NY-Analytical | 757 | Caliper | caliper (COA) | Caliper | 4.22 |  | mil | FDA,4 | 3.995 | 3.995 | Caliper | Main_Plans
2AU555 | 2AU555/Web_NP3460_7.5 | Web_NP3460_7.5 | NY-Analytical | 836 | Dyne Level | Dyne Level (outside film) | Dyne Level (outside film) | 36 |  |  | FDA,0 |  |  | Dyne Test Kit | Main_Plans
2AU555 | 2AU555/Web_NP3460_7.5 | Web_NP3460_7.5 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/inch@290DegF/40psi/0.5 sec dwell) | Seal Strength (@290°F/40psi/0.5 sec dwell) | 8720.24 |  | g/in | FDA,0 | 5850 |  | TM 107 | Main_Plans
2AU555 | 2AU555/Web_NP3460_7.5 | Web_NP3460_7.5 | NY-Analytical | 761 | Web Measurement | Core ID (Inches) | Core ID | 3 |  | in | FDA,4 | 2.9375 | 3.0625 | TM 143 | Main_Plans
2AU555 | 2AU555/Web_NP3460_7.5 | Web_NP3460_7.5 | NY-Analytical | 761 | Web Measurement | Roll OD | Roll OD | 15 |  | in | FDA,4 | 14.9375 | 15.0625 | TM 143 | Main_Plans
2AU555 | 2AU555/Web_NP3460_7.5 | Web_NP3460_7.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 7.5 |  | in | FDA,4 | 7.4375 | 7.5625 | TM 143 | Main_Plans
2AU555 | 2AU555/Web_NP3460_7.5 | Web_NP3460_7.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Visual | Main_Plans
2AU555 | 2AU555/Web_NP3460_7.5 | Web_NP3460_7.5 | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | Visual Observation | Main_Plans
2AU555 | 2AU555/Web_NP3460_7.5 | Web_NP3460_7.5 | NY-Analytical | 748 | Visual Inspection | Color/Copy | Color/Copy | Meets specifications (Done) | [5 empty] | Visual Observation | Main_Plans
2BA585SLI | 2BA585SLI/Lid_19AB | Lid_19AB | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, | [5 empty] | Canister & Lid Inspection | Main_Plans
2BA585SLI | 2BA585SLI/Lid_19AB | Lid_19AB | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2BA585SLI | 2BA585SLI/Lid_19AB | Lid_19AB | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2BA585SLI | 2BA585SLI/Lid_19AB | Lid_19AB | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2BA585SLI | 2BA585SLI/Lid_19AB | Lid_19AB | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 73 |  | mm | FDA,0 | 71 | 75 | TM 143 / Visual | Main_Plans
2BA585SLI | 2BA585SLI/Lid_19AB | Lid_19AB | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 54 |  | mm | FDA,0 | 52 | 56 | TM 143 / Visual | Main_Plans
2BA585SLI | 2BA585SLI/Lid_19AB | Lid_19AB | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2BA972 | 2BA972/Web_NP8540_13.5 | Web_NP8540_13.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [6 empty] | Web Measurement | Main_Plans
2BA972 | 2BA972/Web_NP8540_13.5 | Web_NP8540_13.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 13.5 |  | in | FDA,4 | 13.4375 | 13.5625 | Measurement | Main_Plans
2BA972 | 2BA972/Web_NP8540_13.5 | Web_NP8540_13.5 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 53.97 |  | #/Ream | FDA,1 | 45.4 | 57.8 | TM 101 | Main_Plans
2BA972 | 2BA972/Web_NP8540_13.5 | Web_NP8540_13.5 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) |  |  | g/in | FDA,0 | 3500 |  | TM 107 | Main_Plans
2BA972 | 2BA972/Web_NP8540_13.5 | Web_NP8540_13.5 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 3.1 |  | mil | FDA,2 | 2.86 | 3.64 | TM 132 | Main_Plans
2BC801 | 2BC801/Web_NP3870_6.0 | Web_NP3870_6.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2BC801 | 2BC801/Web_NP3870_6.0 | Web_NP3870_6.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 6 |  | in | FDA,4 | 5.9375 | 6.0625 | Measurement | Main_Plans
2BC801 | 2BC801/Web_NP3870_6.0 | Web_NP3870_6.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 92.46 |  | #/Ream | FDA,1 | 81 | 99 | TM 101 | Main_Plans
2BC801 | 2BC801/Web_NP3870_6.0 | Web_NP3870_6.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@350DegF/40psi/1.0 sec dwell) | Seal Strength (@350°F/40psi/1.0 sec dwell) | 10097 |  | g/in | FDA,0 | 5000 |  | TM 107 | Main_Plans
2BH592 | 2BH592/Web_NP3895_6.0 | Web_NP3895_6.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2BH592 | 2BH592/Web_NP3895_6.0 | Web_NP3895_6.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 6 |  | in | FDA,4 | 5.9375 | 6.0625 | Measurement | Main_Plans
2BH592 | 2BH592/Web_NP3895_6.0 | Web_NP3895_6.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/3000_sq_ft) | Basis Weight | 69.71 |  | #/3000sqft | FDA,1 | 65.3 | 79.8 | TM 101 | Main_Plans
2BH592 | 2BH592/Web_NP3895_6.0 | Web_NP3895_6.0 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/inch@250DegF/30psi/0.5 sec dwell) | Seal Strength (@250°F/30psi/0.5 sec dwell) | 5740.1 |  | g/in | FDA,0 | 2270 |  | TM 107 | Main_Plans
2BH592 | 2BH592/Web_NP3895_6.0 | Web_NP3895_6.0 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 4.5 |  | mil | FDA,1 | 3.9 | 4.7 | TM 132 | Main_Plans
2BJ123 | 2BJ123/Web_NP8550_11 | Web_NP8550_11 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2BJ123 | 2BJ123/Web_NP8550_11 | Web_NP8550_11 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 11 |  | in | FDA,4 | 10.9375 | 11.0625 | Measurement | Main_Plans
2BJ123 | 2BJ123/Web_NP8550_11 | Web_NP8550_11 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 43.12 |  | #/Ream | FDA,1 | 37.8 | 48.2 | TM 101 | Main_Plans
2BJ123 | 2BJ123/Web_NP8550_11 | Web_NP8550_11 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 6855 |  | g/in | FDA,0 | 3500 |  | TM 107 | Main_Plans
2BJ143 | 2BJ143/Web_NP8530_16.25 | Web_NP8530_16.25 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2BJ143 | 2BJ143/Web_NP8530_16.25 | Web_NP8530_16.25 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 16.25 |  | in | FDA,4 | 16.1875 | 16.3125 | Measurement | Main_Plans
2BJ143 | 2BJ143/Web_NP8530_16.25 | Web_NP8530_16.25 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 34.26 |  | #/Ream | FDA,2 | 29.4 | 37.42 | TM 101 | Main_Plans
2BJ143 | 2BJ143/Web_NP8530_16.25 | Web_NP8530_16.25 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 7619 |  | g/in | FDA,0 | 3500 |  | TM 107 | Main_Plans
2BJ143 | 2BJ143/Web_NP8530_16.25 | Web_NP8530_16.25 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 2.1 |  | mil | FDA,1 | 1.8 | 2.2 | TM 132 | Main_Plans
2BJ213 | 2BJ213/Web_NP8530_18 | Web_NP8530_18 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2BJ213 | 2BJ213/Web_NP8530_18 | Web_NP8530_18 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 18 |  | in | FDA,4 | 17.9375 | 18.0625 | Measurement | Main_Plans
2BJ213 | 2BJ213/Web_NP8530_18 | Web_NP8530_18 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 33.33 |  | #/Ream | FDA,2 | 29.4 | 37.42 | TM 101 | Main_Plans
2BJ213 | 2BJ213/Web_NP8530_18 | Web_NP8530_18 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 6908 |  | g/in | FDA,0 | 3500 |  | TM 107 | Main_Plans
2BJ213 | 2BJ213/Web_NP8530_18 | Web_NP8530_18 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 1.85 |  | mil | FDA,1 | 1.8 | 2.2 | TM 132 | Main_Plans
2BJ227 | 2BJ227/Web_NP8530_15.75 | Web_NP8530_15.75 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2BJ227 | 2BJ227/Web_NP8530_15.75 | Web_NP8530_15.75 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 15.75 |  | in | FDA,4 | 15.6875 | 15.8125 | Measurement | Main_Plans
2BJ227 | 2BJ227/Web_NP8530_15.75 | Web_NP8530_15.75 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 33.9 |  | #/Ream | FDA,2 | 29.4 | 37.42 | TM 101 | Main_Plans
2BJ227 | 2BJ227/Web_NP8530_15.75 | Web_NP8530_15.75 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 5068.96 |  | g/in | FDA,0 | 3500 |  | TM 107 | Main_Plans
2BJ227 | 2BJ227/Web_NP8530_15.75 | Web_NP8530_15.75 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 2.18 |  | mil | FDA,1 | 1.8 | 2.2 | TM 132 | Main_Plans
2BP308SLI | 2BP308SLI/Lid_Aveeno Ultra-Calming MUR | Lid_Aveeno Ultra-Calming MUR | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, | [5 empty] | Canister & Lid Inspection | Main_Plans
2BP308SLI | 2BP308SLI/Lid_Aveeno Ultra-Calming MUR | Lid_Aveeno Ultra-Calming MUR | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2BP308SLI | 2BP308SLI/Lid_Aveeno Ultra-Calming MUR | Lid_Aveeno Ultra-Calming MUR | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2BP308SLI | 2BP308SLI/Lid_Aveeno Ultra-Calming MUR | Lid_Aveeno Ultra-Calming MUR | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2BP308SLI | 2BP308SLI/Lid_Aveeno Ultra-Calming MUR | Lid_Aveeno Ultra-Calming MUR | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 72 |  | mm | FDA,0 | 71 | 75 | TM 143 / Visual | Main_Plans
2BP308SLI | 2BP308SLI/Lid_Aveeno Ultra-Calming MUR | Lid_Aveeno Ultra-Calming MUR | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 53 |  | mm | FDA,0 | 52 | 56 | TM 143 / Visual | Main_Plans
2BP308SLI | 2BP308SLI/Lid_Aveeno Ultra-Calming MUR | Lid_Aveeno Ultra-Calming MUR | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2BQ084LI | 2BQ084LI/Lid_11A | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid fits with a tight seal on | [5 empty] | Canister & Lid Inspection | Main_Plans
2BQ084LI | 2BQ084LI/Lid_11A | Lid_11A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Flip Lid Force | Main_Plans
2BQ084LI | 2BQ084LI/Lid_11A | Lid_11A | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | Flip Lid Force | 2.32 |  | lb | FDA,2 | 0.75 |  | TM 143 | Main_Plans
2BQ084LI | 2BQ084LI/Lid_11A | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2BQ084LI | 2BQ084LI/Lid_11A | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | Bead Diameter | 4.85 |  | in | FDA,3 | 4.847 | 4.867 | TM 143 / Visual | Main_Plans
2BQ084LI | 2BQ084LI/Lid_11A | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Translucent Teal (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2BQ084LI | 2BQ084LI/Lid_11A | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Insert Fin Gap (per specification) | Insert Fin Gap | 0.071 |  | in | FDA,3 | 0.065 | 0.075 | TM 143 / Visual | Main_Plans
2BQ084LI | 2BQ084LI/Lid_11A | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter (as per specification) | Lid Inside Diameter | 4.92 |  | in | FDA,3 | 4.917 | 4.937 | TM 143 / Visual | Main_Plans
2BQ084LI | 2BQ084LI/Lid_11A | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.044 |  | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual | Main_Plans
2BQ084LI | 2BQ084LI/Lid_11A | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2BQ084LI | 2BQ084LI/Lid_11A | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness (per specification) | Orifice Flap Thickness | 0.023 |  | in | FDA,3 | 0.018 | 0.028 | TM 143 / Visual | Main_Plans
2BQ084LI | 2BQ084LI/Lid_11A | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.48 |  | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual | Main_Plans
2BQ084LI | 2BQ084LI/Lid_11A | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Outer Diameter (per specification) | Pop-Up Hole Outer Diameter | 0.166 |  | in | FDA,3 | 0.16 | 0.17 | TM 143 / Visual | Main_Plans
2BQ084LI | 2BQ084LI/Lid_11A | Lid_11A | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 34.11 |  | g | FDA,1 | 32 | 42 | TM 143 / Visual | Main_Plans
2BQ471LI | 2BQ471LI/Lid_04F_560 | Lid_04F_560 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2BQ471LI | 2BQ471LI/Lid_04F_560 | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2BQ471LI | 2BQ471LI/Lid_04F_560 | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Cap Thickness | Cap Thickness | 0.151 |  | in | FDA,3 | 0.147 | 0.157 | TM 143 / Visual | Main_Plans
2BQ471LI | 2BQ471LI/Lid_04F_560 | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Translucent Blue (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2BQ471LI | 2BQ471LI/Lid_04F_560 | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2BQ471LI | 2BQ471LI/Lid_04F_560 | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter (as per specification) | Lid Inside Diameter | 3.7695 |  | in | FDA,3 | 3.76 | 3.77 | TM 143 / Visual | Main_Plans
2BQ471LI | 2BQ471LI/Lid_04F_560 | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.046 |  | in | FDA,3 | 0.04 | 0.05 | TM 143 / Visual | Main_Plans
2BQ471LI | 2BQ471LI/Lid_04F_560 | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2BQ471LI | 2BQ471LI/Lid_04F_560 | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.486 |  | in | FDA,3 | 1.473 | 1.493 | TM 143 / Visual | Main_Plans
2BQ471LI | 2BQ471LI/Lid_04F_560 | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Outer Diameter (per specification) | Pop-Up Hole Outer Diameter | 0.555 |  | in | FDA,3 | 0.555 | 0.565 | TM 143 / Visual | Main_Plans
2BQ471LI | 2BQ471LI/Lid_04F_560 | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Thicknes | Pop-Up Hole Thicknes | 0.015 |  | in | FDA,3 | 0.015 | 0.025 | TM 143 / Visual | Main_Plans
2BQ471LI | 2BQ471LI/Lid_04F_560 | Lid_04F_560 | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 22.425 |  | g | FDA,1 | 20 | 24 | TM 143 / Visual | Main_Plans
2BQ551LI | 2BQ551LI/Lid_13A | Lid_13A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Flip Lid Force | Main_Plans
2BQ551LI | 2BQ551LI/Lid_13A | Lid_13A | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | Flip Lid Force | 1.61 |  | lb | FDA,2 | 0.75 |  | TM 143 | Main_Plans
2BQ551LI | 2BQ551LI/Lid_13A | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2BQ551LI | 2BQ551LI/Lid_13A | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | Bead Diameter | 4.8545 |  | in | FDA,3 | 4.847 | 4.867 | TM 143 / Visual | Main_Plans
2BQ551LI | 2BQ551LI/Lid_13A | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Center to Center Distance | Center to Center Distance | 0.45 |  | in | FDA,3 | 0.445 | 0.455 | TM 143 / Visual | Main_Plans
2BQ551LI | 2BQ551LI/Lid_13A | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Translucent Purple (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2BQ551LI | 2BQ551LI/Lid_13A | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | End Circle Diameter | End Circle Diameter | 0.123 |  | in | FDA,3 | 0.12 | 0.13 | TM 143 / Visual | Main_Plans
2BQ551LI | 2BQ551LI/Lid_13A | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2BQ551LI | 2BQ551LI/Lid_13A | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Insert Fin Gap (per specification) | Insert Fin Gap | 0.076 |  | in | FDA,3 | 0.07 | 0.08 | TM 143 / Visual | Main_Plans
2BQ551LI | 2BQ551LI/Lid_13A | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter (as per specification) | Lid Inside Diameter | 4.923 |  | in | FDA,3 | 4.917 | 4.937 | TM 143 / Visual | Main_Plans
2BQ551LI | 2BQ551LI/Lid_13A | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.04 |  | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual | Main_Plans
2BQ551LI | 2BQ551LI/Lid_13A | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2BQ551LI | 2BQ551LI/Lid_13A | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness (per specification) | Orifice Flap Thickness | 0.04 |  | in | FDA,3 | 0.03 | 0.04 | TM 143 / Visual | Main_Plans
2BQ551LI | 2BQ551LI/Lid_13A | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.483 |  | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual | Main_Plans
2BQ551LI | 2BQ551LI/Lid_13A | Lid_13A | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 34.73 |  | g | FDA,1 | 32 | 42 | TM 143 / Visual | Main_Plans
2BS420 | 2BS420/Web_NP3870_11.5 | Web_NP3870_11.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2BS420 | 2BS420/Web_NP3870_11.5 | Web_NP3870_11.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 11.5 |  | in | FDA,4 | 11.4375 | 11.5625 | Measurement | Main_Plans
2BS420 | 2BS420/Web_NP3870_11.5 | Web_NP3870_11.5 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 91.8 |  | #/Ream | FDA,1 | 81 | 99 | TM 101 | Main_Plans
2BS420 | 2BS420/Web_NP3870_11.5 | Web_NP3870_11.5 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@350DegF/40psi/1.0 sec dwell) | Seal Strength (@350°F/40psi/1.0 sec dwell) | 10192.62 |  | g/in | FDA,0 | 5000 |  | TM 107 | Main_Plans
2BU275 | 2BU275/Web_NP3895_8.5 | Web_NP3895_8.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2BU275 | 2BU275/Web_NP3895_8.5 | Web_NP3895_8.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 8.5 |  | in | FDA,4 | 8.4375 | 8.5625 | Measurement | Main_Plans
2BU275 | 2BU275/Web_NP3895_8.5 | Web_NP3895_8.5 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/3000_sq_ft) | Basis Weight | 71.69 |  | #/3000sqft | FDA,1 | 65.3 | 79.8 | TM 101 | Main_Plans
2BU275 | 2BU275/Web_NP3895_8.5 | Web_NP3895_8.5 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/inch@250DegF/30psi/0.5 sec dwell) | Seal Strength (@250°F/30psi/0.5 sec dwell) | 5589 |  | g/in | FDA,0 | 2270 |  | TM 107 | Main_Plans
2BU275 | 2BU275/Web_NP3895_8.5 | Web_NP3895_8.5 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 4.05 |  | mil | FDA,1 | 3.9 | 4.7 | TM 132 | Main_Plans
2CA580SLI | 2CA580SLI/Lid_01R | Lid_01R | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, | [5 empty] | Canister & Lid Inspection | Main_Plans
2CA580SLI | 2CA580SLI/Lid_01R | Lid_01R | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2CA580SLI | 2CA580SLI/Lid_01R | Lid_01R | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2CA580SLI | 2CA580SLI/Lid_01R | Lid_01R | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2CA580SLI | 2CA580SLI/Lid_01R | Lid_01R | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.73 |  | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | Main_Plans
2CA580SLI | 2CA580SLI/Lid_01R | Lid_01R | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.875 |  | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | Main_Plans
2CA580SLI | 2CA580SLI/Lid_01R | Lid_01R | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2CA735 | 2CA735/Web_NP8810_12.125 | Web_NP8810_12.125 | NY-Analytical | 1019 | Moisture Vapor Transmission rate | Report result from COA | Report result from COA | 5 |  | GSM | FDA,0 |  | 10 | Report result from COA | Main_Plans
2CA735 | 2CA735/Web_NP8810_12.125 | Web_NP8810_12.125 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Dyne Level | Main_Plans
2CA735 | 2CA735/Web_NP8810_12.125 | Web_NP8810_12.125 | NY-Analytical | 836 | Dyne Level | Dyne Level (outside film) | Dyne Level (outside film) | 36 |  |  | FDA,0 | 36 | 38 | Dyne Test Kit | Main_Plans
2CA735 | 2CA735/Web_NP8810_12.125 | Web_NP8810_12.125 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 12.125 |  | in | FDA,4 | 12.0625 | 12.1875 | Measurement | Main_Plans
2CA735 | 2CA735/Web_NP8810_12.125 | Web_NP8810_12.125 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 42.33 |  | #/Ream | FDA,1 | 39.4 | 48.2 | TM 101 | Main_Plans
2CA735 | 2CA735/Web_NP8810_12.125 | Web_NP8810_12.125 | NY-Analytical | 762 | Seal Strength | Seal Strength(g/in@350F/30psi/0.5s dwell Top Jaw) | Seal Strength (@350°F/30psi/0.5s dwell Top Jaw) | 5303 |  | gms/f | FDA,0 | 500 |  | TM 107 | Main_Plans
2CA735 | 2CA735/Web_NP8810_12.125 | Web_NP8810_12.125 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 2.7 |  | mil | FDA,2 | 2.48 | 3.03 | TM 132 | Main_Plans
2CA735 | 2CA735/Web_NP8810_12.125 | Web_NP8810_12.125 | NY-Analytical | 397 | Appearance | Appearance | Appearance | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
2CB196 | 2CB196/Web_NP5114_41.0 | Web_NP5114_41.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2CB196 | 2CB196/Web_NP5114_41.0 | Web_NP5114_41.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 40.9375 |  | in | FDA,4 | 40.9375 | 41.0625 | Measurement | Main_Plans
2CB196 | 2CB196/Web_NP5114_41.0 | Web_NP5114_41.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 59.79 |  | #/Ream | FDA,1 | 53.1 | 64.9 | TM 101 | Main_Plans
2CB196 | 2CB196/Web_NP5114_41.0 | Web_NP5114_41.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/40psi/0.5 sec dewll) | Seal Strength (@300°F/40psi/0.5 sec dewll) | 2891.81 |  | g/in | FDA,0 | 1300 |  | TM 107 | Main_Plans
2CB196 | 2CB196/Web_NP5114_41.0 | Web_NP5114_41.0 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 3.2 |  | mil | FDA,2 | 3.15 | 3.85 | TM 132 | Main_Plans
2CB512 | 2CB512/Web_NP3870_4.5 | Web_NP3870_4.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2CB512 | 2CB512/Web_NP3870_4.5 | Web_NP3870_4.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 4.5 |  | in | FDA,4 | 4.4375 | 4.5625 | Measurement | Main_Plans
2CB512 | 2CB512/Web_NP3870_4.5 | Web_NP3870_4.5 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 92.6 |  | #/Ream | FDA,1 | 81 | 99 | TM 101 | Main_Plans
2CB512 | 2CB512/Web_NP3870_4.5 | Web_NP3870_4.5 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@350DegF/40psi/1.0 sec dwell) | Seal Strength (@350°F/40psi/1.0 sec dwell) | 9864.87 |  | g/in | FDA,0 | 5000 |  | TM 107 | Main_Plans
2DH211 | 2DH211/Web_NP5116_6.0 | Web_NP5116_6.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2DH211 | 2DH211/Web_NP5116_6.0 | Web_NP5116_6.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 6 |  | in | FDA,4 | 5.9375 | 6.0625 | Measurement | Main_Plans
2DH211 | 2DH211/Web_NP5116_6.0 | Web_NP5116_6.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Printed/Varnished) | Basis Weight | 57.67 |  | #/Ream | FDA,1 | 55.8 | 68.2 | TM 101 | Main_Plans
2DH211 | 2DH211/Web_NP5116_6.0 | Web_NP5116_6.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/40psi/0.5 sec dewll) | Seal Strength (@300°F/40psi/0.5 sec dewll) | 1800.28 |  | g/in | FDA,0 | 1500 |  | TM 107 | Main_Plans
2ED110 | 2ED110/Web_NP5410_4.5 | Web_NP5410_4.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2ED110 | 2ED110/Web_NP5410_4.5 | Web_NP5410_4.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 4.5 |  | in | FDA,4 | 4.4375 | 4.5625 | Measurement | Main_Plans
2ED110 | 2ED110/Web_NP5410_4.5 | Web_NP5410_4.5 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Printed) | Basis Weight | 54.5 |  | #/Ream | FDA,1 | 49.8 | 60.8 | TM 101 | Main_Plans
2ED110 | 2ED110/Web_NP5410_4.5 | Web_NP5410_4.5 | NY-Analytical | 762 | Seal Strength | Destruct Bond | Destruct Bond | Break occurs off the seal (Done) |  | g/in |  |  |  | TM 107 | Main_Plans
2ED110 | 2ED110/Web_NP5410_4.5 | Web_NP5410_4.5 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/40psi/0.5 sec dewll) | Seal Strength (@300°F/40psi/0.5 sec dewll) | 1814.1 |  | g/in | FDA,0 | 1500 |  | TM 107 | Main_Plans
2EU351 | 2EU351/Web_NP5211_7.0 | Web_NP5211_7.0 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Eyespot Repeat | Main_Plans
2EU351 | 2EU351/Web_NP5211_7.0 | Web_NP5211_7.0 | NY-Analytical | 999 | Eyespot Repeat | Eye spot repeat over 7 impressions | Eye spot repeat over 7 impressions | 488 |  | mm | FDA,0 | 488 | 490 | TM 208 | Main_Plans
2EU351 | 2EU351/Web_NP5211_7.0 | Web_NP5211_7.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 7 |  | in | FDA,4 | 6.9375 | 7.0625 | TM 143 | Main_Plans
2EU351 | 2EU351/Web_NP5211_7.0 | Web_NP5211_7.0 | NY-Analytical | 1031 | 2D Matrix (Barcode) | 2D Matrix (Barcode) | 2D Matrix (Barcode) | Meet Specifications (Done) | [5 empty] | Visual | Main_Plans
2EU351 | 2EU351/Web_NP5211_7.0 | Web_NP5211_7.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Visual | Main_Plans
2EU351 | 2EU351/Web_NP5211_7.0 | Web_NP5211_7.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis (ELLI) | Certificate of Analysis | COA Meets Specification (Done) | [5 empty] | Visual | Main_Plans
2EU351 | 2EU351/Web_NP5211_7.0 | Web_NP5211_7.0 | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | Visual Observation | Main_Plans
2EU351 | 2EU351/Web_NP5211_7.0 | Web_NP5211_7.0 | NY-Analytical | 748 | Visual Inspection | Color/Copy | Color/Copy | Meets specifications (Done) | [5 empty] | Visual Observation | Main_Plans
2FB6038 | 2FB6038/Web_NP5410_8.0 | Web_NP5410_8.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2FB6038 | 2FB6038/Web_NP5410_8.0 | Web_NP5410_8.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width |  |  | in | FDA,4 | 7.9375 | 8.0625 | Measurement | Main_Plans
2FB6038 | 2FB6038/Web_NP5410_8.0 | Web_NP5410_8.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Printed) | Basis Weight | 52.91 |  | #/Ream | FDA,1 | 49.8 | 60.8 | TM 101 | Main_Plans
2FB6038 | 2FB6038/Web_NP5410_8.0 | Web_NP5410_8.0 | NY-Analytical | 762 | Seal Strength | Destruct Bond | Destruct Bond |  |  | g/in |  |  |  | TM 107 | Main_Plans
2FB6038 | 2FB6038/Web_NP5410_8.0 | Web_NP5410_8.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/40psi/0.5 sec dewll) | Seal Strength (@300°F/40psi/0.5 sec dewll) |  |  | g/in | FDA,0 | 1500 |  | TM 107 | Main_Plans
2FH070 | 2FH070/Web_NP5115_6.0 | Web_NP5115_6.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2FH070 | 2FH070/Web_NP5115_6.0 | Web_NP5115_6.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 6 |  | in | FDA,4 | 5.9375 | 6.0625 | Measurement | Main_Plans
2FH070 | 2FH070/Web_NP5115_6.0 | Web_NP5115_6.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Printed/Varnished) | Basis Weight | 59.66 |  | #/Ream | FDA,1 | 55.6 | 68 | TM 101 | Main_Plans
2FH070 | 2FH070/Web_NP5115_6.0 | Web_NP5115_6.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@325DegF/40PSI/0.5Sec_Dwell) | Seal Strength (@325°F/40PSI/0.5Sec Dwell) | 2304.16 |  | g/in | FDA,0 | 1500 |  | TM 107 | Main_Plans
2FJ141R | 2FJ141R/Web_NP8535_16 | Web_NP8535_16 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2FJ141R | 2FJ141R/Web_NP8535_16 | Web_NP8535_16 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 16 |  | in | FDA,4 | 15.9375 | 16.0625 | Measurement | Main_Plans
2FJ141R | 2FJ141R/Web_NP8535_16 | Web_NP8535_16 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 42.06 |  | #/Ream | FDA,2 | 35.82 | 45.6 | TM 101 | Main_Plans
2FJ141R | 2FJ141R/Web_NP8535_16 | Web_NP8535_16 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 7567 |  | g/in | FDA,0 | 3500 |  | TM 107 | Main_Plans
2FJ141R | 2FJ141R/Web_NP8535_16 | Web_NP8535_16 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 2.45 |  | mil | FDA,1 | 2.2 | 2.8 | TM 132 | Main_Plans
2HA580 | 2HA580/Web_NP8540_17.0 | Web_NP8540_17.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2HA580 | 2HA580/Web_NP8540_17.0 | Web_NP8540_17.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 17 |  | in | FDA,4 | 16.9375 | 17.0625 | Measurement | Main_Plans
2HA580 | 2HA580/Web_NP8540_17.0 | Web_NP8540_17.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 52.91 |  | #/Ream | FDA,1 | 45.4 | 57.8 | TM 101 | Main_Plans
2HA580 | 2HA580/Web_NP8540_17.0 | Web_NP8540_17.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 7254 |  | g/in | FDA,0 | 3500 |  | TM 107 | Main_Plans
2HA580 | 2HA580/Web_NP8540_17.0 | Web_NP8540_17.0 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 3 |  | mil | FDA,2 | 2.86 | 3.64 | TM 132 | Main_Plans
2HU120 | 2HU120/Web_NP5115_8.5 | Web_NP5115_8.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2HU120 | 2HU120/Web_NP5115_8.5 | Web_NP5115_8.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 8.5 |  | in | FDA,4 | 8.4375 | 8.5625 | Measurement | Main_Plans
2HU120 | 2HU120/Web_NP5115_8.5 | Web_NP5115_8.5 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Printed/Varnished) | Basis Weight | 58.47 |  | #/Ream | FDA,1 | 55.6 | 68 | TM 101 | Main_Plans
2HU120 | 2HU120/Web_NP5115_8.5 | Web_NP5115_8.5 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@325DegF/40PSI/0.5Sec_Dwell) | Seal Strength (@325°F/40PSI/0.5Sec Dwell) | 2199 |  | g/in | FDA,0 | 1500 |  | TM 107 | Main_Plans
2JC699 | 2JC699/Web_NP5410_6.0 | Web_NP5410_6.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2JC699 | 2JC699/Web_NP5410_6.0 | Web_NP5410_6.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 6 |  | in | FDA,4 | 5.9375 | 6.0625 | Measurement | Main_Plans
2JC699 | 2JC699/Web_NP5410_6.0 | Web_NP5410_6.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Printed) | Basis Weight | 55.6 |  | #/Ream | FDA,1 | 49.8 | 60.8 | TM 101 | Main_Plans
2JC699 | 2JC699/Web_NP5410_6.0 | Web_NP5410_6.0 | NY-Analytical | 762 | Seal Strength | Destruct Bond | Destruct Bond | Break occurs off the seal (Done) |  | g/in |  |  |  | TM 107 | Main_Plans
2JC699 | 2JC699/Web_NP5410_6.0 | Web_NP5410_6.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/40psi/0.5 sec dewll) | Seal Strength (@300°F/40psi/0.5 sec dewll) | 1686.57 |  | g/in | FDA,0 | 1500 |  | TM 107 | Main_Plans
2LB406 | 2LB406/Web_NP3870_10.0 | Web_NP3870_10.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2LB406 | 2LB406/Web_NP3870_10.0 | Web_NP3870_10.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 10 |  | in | FDA,4 | 9.9375 | 10.0625 | Measurement | Main_Plans
2LB406 | 2LB406/Web_NP3870_10.0 | Web_NP3870_10.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 93.53 |  | #/Ream | FDA,1 | 83.7 | 96.3 | TM 101 | Main_Plans
2LB406 | 2LB406/Web_NP3870_10.0 | Web_NP3870_10.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@350DegF/40psi/1.0 sec dwell) | Seal Strength (@350°F/40psi/1.0 sec dwell) | 10196.54 |  | g/in | FDA,0 | 5000 |  | TM 107 | Main_Plans
2OA347SLI | 2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, | [5 empty] | Canister & Lid Inspection | Main_Plans
2OA347SLI | 2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) | [5 empty] | Solo Lid Inspection | Main_Plans
2OA347SLI | 2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 818 | Solo Lid Inspection | Appearance | Appearance | Meet specification (Done) | [5 empty] | Solo Lid Inspection | Main_Plans
2OA347SLI | 2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 818 | Solo Lid Inspection | Color | Color | Meets the specification (Done) | [5 empty] | Solo Lid Inspection | Main_Plans
2OA347SLI | 2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 818 | Solo Lid Inspection | Functionality | Functionality | Meet specification (Done) | [5 empty] | Solo Lid Inspection | Main_Plans
2OA347SLI | 2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 818 | Solo Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.0285 |  | in | FDA,3 | 0.025 | 0.035 | Solo Lid Inspection | Main_Plans
2OA347SLI | 2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 818 | Solo Lid Inspection | Odor | Odor | None detectable (Done) | [5 empty] | Solo Lid Inspection | Main_Plans
2OA347SLI | 2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 818 | Solo Lid Inspection | Open Lid Overall Width | Open Lid Overall Width | 4.367 |  | in | FDA,3 | 4.351 | 4.401 | Solo Lid Inspection | Main_Plans
2OA347SLI | 2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 818 | Solo Lid Inspection | Overall closed lid height | Overall closed lid height | 0.3688 |  | in | FDA,3 | 0.355 | 0.385 | Solo Lid Inspection | Main_Plans
2OA347SLI | 2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 818 | Solo Lid Inspection | Weight | Weight | 11.33 |  | g | FDA,1 | 10.5 | 12.5 | Measurement | Main_Plans
2OA347SLI | 2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OA347SLI | 2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.7187 |  | in | FDA,4 | 3.6875 | 3.8125 | TM 143 / Visual | Main_Plans
2OA347SLI | 2OA347SLI/Lid_NY_Summer's Eve Fresh Cycle | Lid_NY_Summer's Eve Fresh Cycle | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.875 |  | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | Main_Plans
2OA592SLI | 2OA592SLI/Lid_NY_NP7600_22S | Lid_NY_NP7600_22S | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, | [5 empty] | Canister & Lid Inspection | Main_Plans
2OA592SLI | 2OA592SLI/Lid_NY_NP7600_22S | Lid_NY_NP7600_22S | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) | [5 empty] | Solo Lid Inspection | Main_Plans
2OA592SLI | 2OA592SLI/Lid_NY_NP7600_22S | Lid_NY_NP7600_22S | NY-Analytical | 818 | Solo Lid Inspection | Appearance | Appearance | Meet specification (Done) | [5 empty] | Solo Lid Inspection | Main_Plans
2OA592SLI | 2OA592SLI/Lid_NY_NP7600_22S | Lid_NY_NP7600_22S | NY-Analytical | 818 | Solo Lid Inspection | Color | Color | Teal Pantone 3145C | [5 empty] | Solo Lid Inspection | Main_Plans
2OA592SLI | 2OA592SLI/Lid_NY_NP7600_22S | Lid_NY_NP7600_22S | NY-Analytical | 818 | Solo Lid Inspection | Functionality | Functionality | Meet specification (Done) | [5 empty] | Solo Lid Inspection | Main_Plans
2OA592SLI | 2OA592SLI/Lid_NY_NP7600_22S | Lid_NY_NP7600_22S | NY-Analytical | 818 | Solo Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.028 |  | in | FDA,3 | 0.025 | 0.035 | Solo Lid Inspection | Main_Plans
2OA592SLI | 2OA592SLI/Lid_NY_NP7600_22S | Lid_NY_NP7600_22S | NY-Analytical | 818 | Solo Lid Inspection | Odor | Odor | None detectable (Done) | [5 empty] | Solo Lid Inspection | Main_Plans
2OA592SLI | 2OA592SLI/Lid_NY_NP7600_22S | Lid_NY_NP7600_22S | NY-Analytical | 818 | Solo Lid Inspection | Open Lid Overall Width | Open Lid Overall Width | 4.361 |  | in | FDA,3 | 4.351 | 4.401 | Solo Lid Inspection | Main_Plans
2OA592SLI | 2OA592SLI/Lid_NY_NP7600_22S | Lid_NY_NP7600_22S | NY-Analytical | 818 | Solo Lid Inspection | Overall closed lid height | Overall closed lid height | 0.3815 |  | in | FDA,3 | 0.355 | 0.385 | Solo Lid Inspection | Main_Plans
2OA592SLI | 2OA592SLI/Lid_NY_NP7600_22S | Lid_NY_NP7600_22S | NY-Analytical | 818 | Solo Lid Inspection | Weight | Weight | 11.26 |  | g | FDA,1 | 10.5 | 12.5 | Measurement | Main_Plans
2OA592SLI | 2OA592SLI/Lid_NY_NP7600_22S | Lid_NY_NP7600_22S | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OA592SLI | 2OA592SLI/Lid_NY_NP7600_22S | Lid_NY_NP7600_22S | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.75 |  | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | Main_Plans
2OA592SLI | 2OA592SLI/Lid_NY_NP7600_22S | Lid_NY_NP7600_22S | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.875 |  | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | Main_Plans
2OA632SLI | 2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, | [5 empty] | Canister & Lid Inspection | Main_Plans
2OA632SLI | 2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) | [5 empty] | Solo Lid Inspection | Main_Plans
2OA632SLI | 2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 818 | Solo Lid Inspection | Color | Color | Meets the specification (Done) | [5 empty] | Solo Lid Inspection | Main_Plans
2OA632SLI | 2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 818 | Solo Lid Inspection | Functionality | Functionality | Meet specification (Done) | [5 empty] | Solo Lid Inspection | Main_Plans
2OA632SLI | 2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 818 | Solo Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.027 |  | in | FDA,3 | 0.025 | 0.035 | Solo Lid Inspection | Main_Plans
2OA632SLI | 2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 818 | Solo Lid Inspection | Open Lid Overall Width | Open Lid Overall Width | 4.358 |  | in | FDA,3 | 4.351 | 4.401 | Solo Lid Inspection | Main_Plans
2OA632SLI | 2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 818 | Solo Lid Inspection | Overall closed lid height | Overall closed lid height | 0.3785 |  | in | FDA,3 | 0.355 | 0.385 | Solo Lid Inspection | Main_Plans
2OA632SLI | 2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 818 | Solo Lid Inspection | Weight | Weight | 11.235 |  | g | FDA,1 | 10.5 | 12.5 | Measurement | Main_Plans
2OA632SLI | 2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OA632SLI | 2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OA632SLI | 2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.69 |  | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | Main_Plans
2OA632SLI | 2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.83 |  | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | Main_Plans
2OA632SLI | 2OA632SLI/Lid_NY_Prep H medicated Hemorrhoidal Wipes | Lid_NY_Prep H medicated Hemorrhoidal Wipes | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OJ124SLI | 2OJ124SLI/Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, | [5 empty] | Canister & Lid Inspection | Main_Plans
2OJ124SLI | 2OJ124SLI/Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2OJ124SLI | 2OJ124SLI/Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OJ124SLI | 2OJ124SLI/Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OJ124SLI | 2OJ124SLI/Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.7187 |  | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | Main_Plans
2OJ124SLI | 2OJ124SLI/Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.875 |  | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | Main_Plans
2OJ124SLI | 2OJ124SLI/Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | Lid_NY_Cetaphil Gentle Skin cleansing wipes solo | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OJ130SLI | 2OJ130SLI/Lid_NY_NP7600_22Y | Lid_NY_NP7600_22Y | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, | [5 empty] | Canister & Lid Inspection | Main_Plans
2OJ130SLI | 2OJ130SLI/Lid_NY_NP7600_22Y | Lid_NY_NP7600_22Y | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) | [5 empty] | Solo Lid Inspection | Main_Plans
2OJ130SLI | 2OJ130SLI/Lid_NY_NP7600_22Y | Lid_NY_NP7600_22Y | NY-Analytical | 818 | Solo Lid Inspection | Color | Color | Meets the specification (Done) | [5 empty] | Solo Lid Inspection | Main_Plans
2OJ130SLI | 2OJ130SLI/Lid_NY_NP7600_22Y | Lid_NY_NP7600_22Y | NY-Analytical | 818 | Solo Lid Inspection | Functionality | Functionality | Meet specification (Done) | [5 empty] | Solo Lid Inspection | Main_Plans
2OJ130SLI | 2OJ130SLI/Lid_NY_NP7600_22Y | Lid_NY_NP7600_22Y | NY-Analytical | 818 | Solo Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.0285 |  | in | FDA,3 | 0.025 | 0.035 | Solo Lid Inspection | Main_Plans
2OJ130SLI | 2OJ130SLI/Lid_NY_NP7600_22Y | Lid_NY_NP7600_22Y | NY-Analytical | 818 | Solo Lid Inspection | Open Lid Overall Width | Open Lid Overall Width | 4.373 |  | in | FDA,3 | 4.351 | 4.401 | Solo Lid Inspection | Main_Plans
2OJ130SLI | 2OJ130SLI/Lid_NY_NP7600_22Y | Lid_NY_NP7600_22Y | NY-Analytical | 818 | Solo Lid Inspection | Overall closed lid height | Overall closed lid height | 0.3715 |  | in | FDA,3 | 0.355 | 0.385 | Solo Lid Inspection | Main_Plans
2OJ130SLI | 2OJ130SLI/Lid_NY_NP7600_22Y | Lid_NY_NP7600_22Y | NY-Analytical | 818 | Solo Lid Inspection | Weight | Weight | 11.4 |  | g | FDA,1 | 10.5 | 12.5 | Measurement | Main_Plans
2OJ130SLI | 2OJ130SLI/Lid_NY_NP7600_22Y | Lid_NY_NP7600_22Y | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OJ130SLI | 2OJ130SLI/Lid_NY_NP7600_22Y | Lid_NY_NP7600_22Y | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OJ130SLI | 2OJ130SLI/Lid_NY_NP7600_22Y | Lid_NY_NP7600_22Y | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.7187 |  | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | Main_Plans
2OJ130SLI | 2OJ130SLI/Lid_NY_NP7600_22Y | Lid_NY_NP7600_22Y | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.875 |  | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | Main_Plans
2OJ130SLI | 2OJ130SLI/Lid_NY_NP7600_22Y | Lid_NY_NP7600_22Y | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OJ213SLI | 2OJ213SLI/Lid_01A | Lid_01A | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, | [5 empty] | Canister & Lid Inspection | Main_Plans
2OJ213SLI | 2OJ213SLI/Lid_01A | Lid_01A | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2OJ213SLI | 2OJ213SLI/Lid_01A | Lid_01A | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OJ213SLI | 2OJ213SLI/Lid_01A | Lid_01A | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OJ213SLI | 2OJ213SLI/Lid_01A | Lid_01A | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 1.875 |  | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | Main_Plans
2OJ213SLI | 2OJ213SLI/Lid_01A | Lid_01A | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 3.6875 |  | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | Main_Plans
2OJ213SLI | 2OJ213SLI/Lid_01A | Lid_01A | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OM304 | 2OM304/Web_NP8540_16.0 | Web_NP8540_16.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2OM304 | 2OM304/Web_NP8540_16.0 | Web_NP8540_16.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 16 |  | in | FDA,4 | 15.9375 | 16.0625 | Measurement | Main_Plans
2OM304 | 2OM304/Web_NP8540_16.0 | Web_NP8540_16.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 53.176 |  | #/Ream | FDA,1 | 45.4 | 57.8 | TM 101 | Main_Plans
2OM304 | 2OM304/Web_NP8540_16.0 | Web_NP8540_16.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 7038.28 |  | g/in | FDA,0 | 3500 |  | TM 107 | Main_Plans
2OM304 | 2OM304/Web_NP8540_16.0 | Web_NP8540_16.0 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 3.3 |  | mil | FDA,2 | 2.86 | 3.64 | TM 132 | Main_Plans
2OM304SLI | 2OM304SLI/Lid_No-Rinse sanitizing Solo | Lid_No-Rinse sanitizing Solo | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, | [5 empty] | Canister & Lid Inspection | Main_Plans
2OM304SLI | 2OM304SLI/Lid_No-Rinse sanitizing Solo | Lid_No-Rinse sanitizing Solo | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2OM304SLI | 2OM304SLI/Lid_No-Rinse sanitizing Solo | Lid_No-Rinse sanitizing Solo | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OM304SLI | 2OM304SLI/Lid_No-Rinse sanitizing Solo | Lid_No-Rinse sanitizing Solo | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OM304SLI | 2OM304SLI/Lid_No-Rinse sanitizing Solo | Lid_No-Rinse sanitizing Solo | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.7187 |  | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | Main_Plans
2OM304SLI | 2OM304SLI/Lid_No-Rinse sanitizing Solo | Lid_No-Rinse sanitizing Solo | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.875 |  | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | Main_Plans
2OM304SLI | 2OM304SLI/Lid_No-Rinse sanitizing Solo | Lid_No-Rinse sanitizing Solo | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OM701SLI | 2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | Lid_No-Rinse sanitizing Solo_01 | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, | [5 empty] | Canister & Lid Inspection | Main_Plans
2OM701SLI | 2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | Lid_No-Rinse sanitizing Solo_01 | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2OM701SLI | 2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | Lid_No-Rinse sanitizing Solo_01 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OM701SLI | 2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | Lid_No-Rinse sanitizing Solo_01 | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OM701SLI | 2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | Lid_No-Rinse sanitizing Solo_01 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.687 |  | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | Main_Plans
2OM701SLI | 2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | Lid_No-Rinse sanitizing Solo_01 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.8125 |  | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | Main_Plans
2OM701SLI | 2OM701SLI/Lid_No-Rinse sanitizing Solo_01 | Lid_No-Rinse sanitizing Solo_01 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OP037 | 2OP037/Web_NP8511_245 | Web_NP8511_245 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2OP037 | 2OP037/Web_NP8511_245 | Web_NP8511_245 | NY-Analytical | 761 | Web Measurement | Web Width (mm) | Web Width | 245 |  | mm | FDA,0 | 243 | 247 | Measurement | Main_Plans
2OP037 | 2OP037/Web_NP8511_245 | Web_NP8511_245 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 47.88 |  | Lbs/Ream | FDA,1 | 43.1 | 52.7 | TM 101 | Main_Plans
2OP037 | 2OP037/Web_NP8511_245 | Web_NP8511_245 | NY-Analytical | 1048 | Seal strength (NP8511) | Seal strength 2500g/in@140DegC/1sec/14psi | Seal Strength (@140°C/1sec/14psi | 6296.05 |  | g/in | FDA,0 | 2500 |  | TM 107 | Main_Plans
2OP037 | 2OP037/Web_NP8511_245 | Web_NP8511_245 | NY-Analytical | 763 | Web Thickness | Thickness | Thickness | 71 |  | micron | FDA,1 | 64.8 | 79.2 | TM 132 | Main_Plans
2OP037 | 2OP037/Web_NP8511_245 | Web_NP8511_245 | NY-Analytical | 885 | Web Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Sample(s) all meet specifications (Done) | [5 empty] | Visual Observation | Main_Plans
2OP063 | 2OP063/Web_NP8511_320 | Web_NP8511_320 | NY-Analytical | 836 | Dyne Level | Dyne Level (outside film) | Dyne Level (outside film) | 36 |  |  | FDA,0 |  |  | Dyne Level | Main_Plans
2OP063 | 2OP063/Web_NP8511_320 | Web_NP8511_320 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2OP063 | 2OP063/Web_NP8511_320 | Web_NP8511_320 | NY-Analytical | 761 | Web Measurement | Web Width (mm) | Web Width | 320 |  | mm | FDA,0 | 318 | 322 | Measurement | Main_Plans
2OP063 | 2OP063/Web_NP8511_320 | Web_NP8511_320 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 47.8 |  | Lbs/Ream | FDA,1 | 43.1 | 52.7 | TM 101 | Main_Plans
2OP063 | 2OP063/Web_NP8511_320 | Web_NP8511_320 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/inch@140C/14psi/1 s dwellTopJaw) | Seal Strength (@140°C/14psi/1 s dwellTopJaw) | 7207.05 |  | g/in | FDA,0 | 2500 |  | TM 107 | Main_Plans
2OP063 | 2OP063/Web_NP8511_320 | Web_NP8511_320 | NY-Analytical | 1048 | Seal strength (NP8511) | Seal strength 2500g/in@140DegC/1sec/14psi | Seal Strength (@140°C/1sec/14psi | 5948 |  | g/in | FDA,0 | 2500 |  | TM 107 | Main_Plans
2OP063 | 2OP063/Web_NP8511_320 | Web_NP8511_320 | NY-Analytical | 763 | Web Thickness | Thickness | Thickness | 74 |  | micron | FDA,1 | 64.8 | 79.2 | TM 132 | Main_Plans
2OP065 | 2OP065/Web_NP8536_11.875 | Web_NP8536_11.875 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2OP065 | 2OP065/Web_NP8536_11.875 | Web_NP8536_11.875 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 11.875 |  | in | FDA,4 | 11.8125 | 11.9375 | Measurement | Main_Plans
2OP065 | 2OP065/Web_NP8536_11.875 | Web_NP8536_11.875 | NY-Analytical | 1044 | Basis Weight (lbs/ream) | Basis Weight | Basis Weight | 41.4 |  | Lb/Ream | FDA,2 | 40.1 | 49.02 | TM 101 | Main_Plans
2OP065 | 2OP065/Web_NP8536_11.875 | Web_NP8536_11.875 | NY-Analytical | 1049 | Seal strength (NP8536) | Seal Strength (g/inch@320DegF/40psi/0.5 sec dwell) | Seal Strength (@320°F/40psi/0.5 sec dwell) | 8624 |  | g/in | FDA,0 | 2000 |  | TM 107 | Main_Plans
2OP065 | 2OP065/Web_NP8536_11.875 | Web_NP8536_11.875 | NY-Analytical | 763 | Web Thickness | Thickness | Thickness | 2.4 |  | mil | FDA,3 | 2.367 | 2.893 | TM 132 | Main_Plans
2OP065 | 2OP065/Web_NP8536_11.875 | Web_NP8536_11.875 | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | Visual | Main_Plans
2OP065 | 2OP065/Web_NP8536_11.875 | Web_NP8536_11.875 | NY-Analytical | 748 | Visual Inspection | Barcode Legibility/Quality | Barcode Legibility/Quality | Barcode should scan and code should be matched under the barcode (Done) | [5 empty] | Visual | Main_Plans
2OP065 | 2OP065/Web_NP8536_11.875 | Web_NP8536_11.875 | NY-Analytical | 748 | Visual Inspection | Color/Copy | Color/Copy | Meets specifications (Done) | [5 empty] | Visual | Main_Plans
2OP072 | 2OP072/Web_NP8511_245 | Web_NP8511_245 | NY-Analytical | 836 | Dyne Level | Dyne Level (outside film) | Dyne Level (outside film) | 34 |  |  | FDA,0 |  |  | Dyne Level | Main_Plans
2OP073 | 2OP073/Web_NP8536_11.875 | Web_NP8536_11.875 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight |  |  | lbs/rm | FDA,2 |  |  | TM 101 | Main_Plans
2OP073 | 2OP073/Web_NP8536_11.875 | Web_NP8536_11.875 | NY-Analytical | 762 | Seal Strength | Seal Strength(g/in@350DegF/40psi/0.5sec dwell) | Seal Strength (@350°F/40psi/0.5sec dwell) |  |  | g/in | FDA,0 | 2500 |  | TM 107 | Main_Plans
2OP075 | 2OP075/Web_NP8541_12.205 | Web_NP8541_12.205 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2OP075 | 2OP075/Web_NP8541_12.205 | Web_NP8541_12.205 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 12.187 |  | in | FDA,4 | 12.1425 | 12.2675 | Measurement | Main_Plans
2OP075 | 2OP075/Web_NP8541_12.205 | Web_NP8541_12.205 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 86.11 |  | GSM | FDA,1 | 74.3 | 90.9 | TM 101 | Main_Plans
2OP075 | 2OP075/Web_NP8541_12.205 | Web_NP8541_12.205 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/inch@130DegC/0.5 sec/1 bar) | Seal Strength (@130°C/0.5 sec/1 bar) | 8916 |  | g/in | FDA,0 | 1000 |  | TM 107 | Main_Plans
2OP075 | 2OP075/Web_NP8541_12.205 | Web_NP8541_12.205 | NY-Analytical | 763 | Web Thickness | Thickness | Thickness | 0.084 |  | mm | FDA,3 | 0.074 | 0.094 | TM 132 | Main_Plans
2OP075 | 2OP075/Web_NP8541_12.205 | Web_NP8541_12.205 | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | Visual | Main_Plans
2OP075 | 2OP075/Web_NP8541_12.205 | Web_NP8541_12.205 | NY-Analytical | 748 | Visual Inspection | Barcode Legibility/Quality | Barcode Legibility/Quality | Barcode should scan and code should be matched under the barcode (Done) | [5 empty] | Visual | Main_Plans
2OP075 | 2OP075/Web_NP8541_12.205 | Web_NP8541_12.205 | NY-Analytical | 748 | Visual Inspection | Color/Copy | Color/Copy | Meets specifications (Done) | [5 empty] | Visual | Main_Plans
2OP092 | 2OP092/Web_NP8536_10 | Web_NP8536_10 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2OP092 | 2OP092/Web_NP8536_10 | Web_NP8536_10 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 10 |  | in | FDA,4 | 9.9375 | 10.0625 | Measurement | Main_Plans
2OP092 | 2OP092/Web_NP8536_10 | Web_NP8536_10 | NY-Analytical | 1044 | Basis Weight (lbs/ream) | Basis Weight | Basis Weight | 42.3 |  | Lb/Ream | FDA,2 | 40.1 | 49.02 | TM 101 | Main_Plans
2OP092 | 2OP092/Web_NP8536_10 | Web_NP8536_10 | NY-Analytical | 1049 | Seal strength (NP8536) | Seal Strength (g/inch@320DegF/40psi/0.5 sec dwell) | Seal Strength (@320°F/40psi/0.5 sec dwell) | 8116.37 |  | g/in | FDA,0 | 2000 |  | TM 107 | Main_Plans
2OP092 | 2OP092/Web_NP8536_10 | Web_NP8536_10 | NY-Analytical | 763 | Web Thickness | Thickness | Thickness | 2.8 |  | mil | FDA,3 | 2.367 | 2.893 | TM 132 | Main_Plans
2OP092 | 2OP092/Web_NP8536_10 | Web_NP8536_10 | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | Visual | Main_Plans
2OP092 | 2OP092/Web_NP8536_10 | Web_NP8536_10 | NY-Analytical | 748 | Visual Inspection | Barcode Legibility/Quality | Barcode Legibility/Quality | Barcode should scan and code should be matched under the barcode (Done) | [5 empty] | Visual | Main_Plans
2OP092 | 2OP092/Web_NP8536_10 | Web_NP8536_10 | NY-Analytical | 748 | Visual Inspection | Color/Copy | Color/Copy | Meets specifications (Done) | [5 empty] | Visual | Main_Plans
2OP093LI | 2OP093LI/Lid_NY_Threaded Blue Cap | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid fits with a tight seal on | [5 empty] | Canister & Lid Inspection | Main_Plans
2OP093LI | 2OP093LI/Lid_NY_Threaded Blue Cap | Lid_NY_Threaded Blue Cap | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2OP093LI | 2OP093LI/Lid_NY_Threaded Blue Cap | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OP093LI | 2OP093LI/Lid_NY_Threaded Blue Cap | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Matches Specification (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OP093LI | 2OP093LI/Lid_NY_Threaded Blue Cap | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | Diameter | Diameter | 2.7515 |  | in | FDA,3 | 2.748 | 2.764 | TM 143 / Visual | Main_Plans
2OP093LI | 2OP093LI/Lid_NY_Threaded Blue Cap | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | E Diameter | E Diameter | 2.5495 |  | in | FDA,3 | 2.548 | 2.562 | TM 143 / Visual | Main_Plans
2OP093LI | 2OP093LI/Lid_NY_Threaded Blue Cap | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | Height | Height | 0.5625 |  | in | FDA,3 | 0.56 | 0.58 | TM 143 / Visual | Main_Plans
2OP093LI | 2OP093LI/Lid_NY_Threaded Blue Cap | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OP093LI | 2OP093LI/Lid_NY_Threaded Blue Cap | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | T Diameter | T Diameter | 2.619 |  | in | FDA,3 | 2.619 | 2.639 | TM 143 / Visual | Main_Plans
2OP093LI | 2OP093LI/Lid_NY_Threaded Blue Cap | Lid_NY_Threaded Blue Cap | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 10.465 |  | g | FDA,2 |  |  | TM 143 / Visual | Main_Plans
2OP134 | 2OP134/Reduced Testing_NY_Can 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Cap Closure | Functionality - Cap Closure | Cap of lid | [5 empty] | Canister & Lid Inspection | Main_Plans
2OP134 | 2OP134/Reduced Testing_NY_Can 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid fits with a tight seal on | [5 empty] | Canister & Lid Inspection | Main_Plans
2OP134 | 2OP134/Reduced Testing_NY_Can 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) | [5 empty] | COA | Main_Plans
2OP134 | 2OP134/Reduced Testing_NY_Can 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2OP134 | 2OP134/Reduced Testing_NY_Can 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2OP134 | 2OP134/Reduced Testing_NY_Can 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | White (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2OP134 | 2OP134/Reduced Testing_NY_Can 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Bead OD | Bead OD | 3.77 |  | in | FDA,3 | 3.768 | 3.792 | Canister & Lid Inspection | Main_Plans
2OP134 | 2OP134/Reduced Testing_NY_Can 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Canister Height OD | Canister Height OD | 6.66 |  | in | FDA,3 | 6.608 | 6.732 | Canister & Lid Inspection | Main_Plans
2OP134 | 2OP134/Reduced Testing_NY_Can 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Lip (width) | Lip Width | 0.138 |  | in | FDA,3 | 0.127 | 0.143 | Canister & Lid Inspection | Main_Plans
2OP134 | 2OP134/Reduced Testing_NY_Can 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Neck OD | Neck OD | 3.68 |  | in | FDA,3 | 3.673 | 3.697 | Canister & Lid Inspection | Main_Plans
2OP134 | 2OP134/Reduced Testing_NY_Can 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | HDPE Resin Identification Marker | Meets the Specification (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2OP134 | 2OP134/Reduced Testing_NY_Can 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2OP134 | 2OP134/Reduced Testing_NY_Can 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight | 53.94 |  | g | FDA,1 | 52 | 56 | Measurement | Main_Plans
2OP134 | 2OP134/Reduced Testing_NY_Can 2OP134 | Reduced Testing_NY_Can 2OP134 | NY-Analytical | 766 | Lid Pull Force | Lid Pull-Off Force | Lid Pull-Off Force | 2997 |  | g | FDA,0 | 2000 | 3500 | TM 148 | Main_Plans
2OP138 | 2OP138/Reduced Testing_NY_Can 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) | [5 empty] | COA | Main_Plans
2OP138 | 2OP138/Reduced Testing_NY_Can 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2OP138 | 2OP138/Reduced Testing_NY_Can 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | [6 empty] | Canister & Lid Inspection | Main_Plans
2OP138 | 2OP138/Reduced Testing_NY_Can 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | [6 empty] | Canister & Lid Inspection | Main_Plans
2OP138 | 2OP138/Reduced Testing_NY_Can 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Bead OD | Bead OD |  |  | in | FDA,3 | 4.922 | 4.958 | Canister & Lid Inspection | Main_Plans
2OP138 | 2OP138/Reduced Testing_NY_Can 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Canister OD | Canister OD |  |  | in | FDA,3 | 4.956 | 5.076 | Canister & Lid Inspection | Main_Plans
2OP138 | 2OP138/Reduced Testing_NY_Can 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Neck OD | Neck OD |  |  | in | FDA,3 | 4.824 | 4.86 | Canister & Lid Inspection | Main_Plans
2OP138 | 2OP138/Reduced Testing_NY_Can 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Total Height OD | Total Height OD |  |  | in | FDA,3 | 6.745 | 6.865 | Canister & Lid Inspection | Main_Plans
2OP138 | 2OP138/Reduced Testing_NY_Can 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | [6 empty] | Canister & Lid Inspection | Main_Plans
2OP138 | 2OP138/Reduced Testing_NY_Can 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | HDPE Resin Identification Marker | [6 empty] | Canister & Lid Inspection | Main_Plans
2OP138 | 2OP138/Reduced Testing_NY_Can 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Lip Width (for reference only) | Lip Width (for reference only) |  |  | in | FDA,3 | 0.14 |  | Canister & Lid Inspection | Main_Plans
2OP138 | 2OP138/Reduced Testing_NY_Can 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | [6 empty] | Canister & Lid Inspection | Main_Plans
2OP138 | 2OP138/Reduced Testing_NY_Can 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Wall Thickness (Bottom base area) | Wall Thickness (Bottom base area) |  |  | in | FDA,3 | 0.01 |  | Canister & Lid Inspection | Main_Plans
2OP138 | 2OP138/Reduced Testing_NY_Can 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Wall Thickness (Top area) | Wall Thickness (Top area) |  |  | in | FDA,3 | 0.02 |  | Canister & Lid Inspection | Main_Plans
2OP138 | 2OP138/Reduced Testing_NY_Can 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight |  |  | g | FDA,1 | 66.5 | 73.5 | Measurement | Main_Plans
2OP138 | 2OP138/Reduced Testing_NY_Can 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 1052 | Top Load Compression Strength | Top load Compression Strength (0.5"/mSR,0.450"TD) | Top load Compression Strength (0.5"/mSR,0.450"TD) | 150.3 |  | lb | FDA,0 | 84 |  | NPTM 1019 | Main_Plans
2OP138 | 2OP138/Reduced Testing_NY_Can 2OP138 | Reduced Testing_NY_Can 2OP138 | NY-Analytical | 1053 | Cavity | cavity Inserts | cavity Inserts | [6 empty] | Visual | Main_Plans
2OP138C | 2OP138C/Reduced Testing_NY_Can 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2OP138C | 2OP138C/Reduced Testing_NY_Can 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | [6 empty] | Canister & Lid Inspection | Main_Plans
2OP138C | 2OP138C/Reduced Testing_NY_Can 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | [6 empty] | Canister & Lid Inspection | Main_Plans
2OP138C | 2OP138C/Reduced Testing_NY_Can 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Bead OD | Bead OD |  |  | in | FDA,3 | 4.922 | 4.958 | Canister & Lid Inspection | Main_Plans
2OP138C | 2OP138C/Reduced Testing_NY_Can 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Canister OD | Canister OD |  |  | in | FDA,3 | 4.956 | 5.076 | Canister & Lid Inspection | Main_Plans
2OP138C | 2OP138C/Reduced Testing_NY_Can 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Neck OD | Neck OD |  |  | in | FDA,3 | 4.824 | 4.86 | Canister & Lid Inspection | Main_Plans
2OP138C | 2OP138C/Reduced Testing_NY_Can 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Total Height OD | Total Height OD | 6.785 |  | in | FDA,3 | 6.745 | 6.865 | Canister & Lid Inspection | Main_Plans
2OP138C | 2OP138C/Reduced Testing_NY_Can 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | [6 empty] | Canister & Lid Inspection | Main_Plans
2OP138C | 2OP138C/Reduced Testing_NY_Can 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | HDPE Resin Identification Marker | [6 empty] | Canister & Lid Inspection | Main_Plans
2OP138C | 2OP138C/Reduced Testing_NY_Can 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | [6 empty] | Canister & Lid Inspection | Main_Plans
2OP138C | 2OP138C/Reduced Testing_NY_Can 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Wall Thickness (Bottom base area) | Wall Thickness (Bottom base area) |  |  | in | FDA,3 | 0.01 |  | Canister & Lid Inspection | Main_Plans
2OP138C | 2OP138C/Reduced Testing_NY_Can 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Wall Thickness (Top area) | Wall Thickness (Top area) |  |  | in | FDA,3 | 0.02 |  | Canister & Lid Inspection | Main_Plans
2OP138C | 2OP138C/Reduced Testing_NY_Can 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight |  |  | g | FDA,1 | 66.5 | 73.5 | Measurement | Main_Plans
2OP138C | 2OP138C/Reduced Testing_NY_Can 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 1052 | Top Load Compression Strength | Top load Compression Strength (0.5"/mSR,0.450"TD) | Top load Compression Strength (0.5"/mSR,0.450"TD) |  |  | lb | FDA,0 | 84 |  | NPTM 1019 | Main_Plans
2OP138C | 2OP138C/Reduced Testing_NY_Can 2OP138C | Reduced Testing_NY_Can 2OP138C | NY-Analytical | 1053 | Cavity | cavity Inserts | cavity Inserts | [6 empty] | Visual | Main_Plans
2OP159LI | 2OP159LI/Lid_03B | Lid_03B | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2OP159LI | 2OP159LI/Lid_03B | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OP159LI | 2OP159LI/Lid_03B | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | Bead Diameter | 4.856 |  | in | FDA,3 | 4.847 | 4.867 | TM 143 / Visual | Main_Plans
2OP159LI | 2OP159LI/Lid_03B | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Cap Thickness | Cap Thickness | 0.049 |  | in | FDA,3 | 0.045 | 0.055 | TM 143 / Visual | Main_Plans
2OP159LI | 2OP159LI/Lid_03B | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Blue (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OP159LI | 2OP159LI/Lid_03B | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OP159LI | 2OP159LI/Lid_03B | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Insert Fin Gap (per specification) | Insert Fin Gap | 0.07 |  | in | FDA,3 | 0.065 | 0.075 | TM 143 / Visual | Main_Plans
2OP159LI | 2OP159LI/Lid_03B | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter (as per specification) | Lid Inside Diameter | 4.929 |  | in | FDA,3 | 4.922 | 4.932 | TM 143 / Visual | Main_Plans
2OP159LI | 2OP159LI/Lid_03B | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.042 |  | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual | Main_Plans
2OP159LI | 2OP159LI/Lid_03B | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OP159LI | 2OP159LI/Lid_03B | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.486 |  | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual | Main_Plans
2OP159LI | 2OP159LI/Lid_03B | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Outer Diameter (per specification) | Pop-Up Hole Outer Diameter | 0.163 |  | in | FDA,3 | 0.16 | 0.17 | TM 143 / Visual | Main_Plans
2OP159LI | 2OP159LI/Lid_03B | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Thicknes | Pop-Up Hole Thicknes | 0.02 |  | in | FDA,3 | 0.012 | 0.022 | TM 143 / Visual | Main_Plans
2OP159LI | 2OP159LI/Lid_03B | Lid_03B | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 34.335 |  | g | FDA,1 | 32 | 42 | TM 143 / Visual | Main_Plans
2OP213 | 2OP213/Web_NP8541_11 | Web_NP8541_11 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2OP213 | 2OP213/Web_NP8541_11 | Web_NP8541_11 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 11 |  | in | FDA,4 | 10.9375 | 11.0625 | Measurement | Main_Plans
2OP213 | 2OP213/Web_NP8541_11 | Web_NP8541_11 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 86.76 |  | GSM | FDA,1 | 74.3 | 90.9 | TM 101 | Main_Plans
2OP213 | 2OP213/Web_NP8541_11 | Web_NP8541_11 | NY-Analytical | 762 | Seal Strength | Seal Strength (g/inch@130DegC/0.5 sec/1 bar) | Seal Strength (@130°C/0.5 sec/1 bar) | 8561 |  | g/in | FDA,0 | 1000 |  | TM 107 | Main_Plans
2OP213 | 2OP213/Web_NP8541_11 | Web_NP8541_11 | NY-Analytical | 763 | Web Thickness | Thickness | Thickness | 0.088 |  | mm | FDA,3 | 0.074 | 0.094 | TM 132 | Main_Plans
2OP213 | 2OP213/Web_NP8541_11 | Web_NP8541_11 | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | Visual | Main_Plans
2OP213 | 2OP213/Web_NP8541_11 | Web_NP8541_11 | NY-Analytical | 748 | Visual Inspection | Barcode Legibility/Quality | Barcode Legibility/Quality | Barcode should scan and code should be matched under the barcode (Done) | [5 empty] | Visual | Main_Plans
2OP213 | 2OP213/Web_NP8541_11 | Web_NP8541_11 | NY-Analytical | 748 | Visual Inspection | Color/Copy | Color/Copy | Meets specifications (Done) | [5 empty] | Visual | Main_Plans
2OP223 | 2OP223/Web_NP8511_320 | Web_NP8511_320 | NY-Analytical | 885 | Web Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Sample(s) all meet specifications (Done) | [5 empty] | Visual Observation | Main_Plans
2OP242LI | 2OP242LI/Lid_17C | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Cap Closure | Functionality - Cap Closure | Cap of lid | [5 empty] | Canister & Lid Inspection | Main_Plans
2OP242LI | 2OP242LI/Lid_17C | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid fits with a tight seal on | [5 empty] | Canister & Lid Inspection | Main_Plans
2OP242LI | 2OP242LI/Lid_17C | Lid_17C | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Flip Lid Force | Main_Plans
2OP242LI | 2OP242LI/Lid_17C | Lid_17C | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | Flip Lid Force | 1.11 |  | lb | FDA,2 | 0.75 |  | TM 143 | Main_Plans
2OP242LI | 2OP242LI/Lid_17C | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OP242LI | 2OP242LI/Lid_17C | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | Bead Diameter | 4.835 |  | in | FDA,3 | 4.831 | 4.851 | TM 143 / Visual | Main_Plans
2OP242LI | 2OP242LI/Lid_17C | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Translucent Plum (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OP242LI | 2OP242LI/Lid_17C | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.045 |  | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual | Main_Plans
2OP242LI | 2OP242LI/Lid_17C | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OP242LI | 2OP242LI/Lid_17C | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Diameter | Orifice Diameter | 0.62 |  | in | FDA,3 | 0.62 | 0.63 | TM 143 / Visual | Main_Plans
2OP242LI | 2OP242LI/Lid_17C | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness (per specification) | Orifice Flap Thickness | 0.021 |  | in | FDA,3 | 0.017 | 0.023 | TM 143 / Visual | Main_Plans
2OP242LI | 2OP242LI/Lid_17C | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.483 |  | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual | Main_Plans
2OP242LI | 2OP242LI/Lid_17C | Lid_17C | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 38.41 |  | g | FDA,1 | 32.5 | 41.5 | TM 143 / Visual | Main_Plans
2OP280LI | 2OP280LI/Lid_17A | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Cap Closure | Functionality - Cap Closure | Cap of lid | [5 empty] | Canister & Lid Inspection | Main_Plans
2OP280LI | 2OP280LI/Lid_17A | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid fits with a tight seal on | [5 empty] | Canister & Lid Inspection | Main_Plans
2OP280LI | 2OP280LI/Lid_17A | Lid_17A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Flip Lid Force | Main_Plans
2OP280LI | 2OP280LI/Lid_17A | Lid_17A | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | Flip Lid Force | 2.88 |  | lb | FDA,2 | 0.75 |  | TM 143 | Main_Plans
2OP280LI | 2OP280LI/Lid_17A | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OP280LI | 2OP280LI/Lid_17A | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | Bead Diameter | 4.851 |  | in | FDA,3 | 4.831 | 4.851 | TM 143 / Visual | Main_Plans
2OP280LI | 2OP280LI/Lid_17A | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Translucent Orange (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OP280LI | 2OP280LI/Lid_17A | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.035 |  | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual | Main_Plans
2OP280LI | 2OP280LI/Lid_17A | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OP280LI | 2OP280LI/Lid_17A | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Diameter | Orifice Diameter | 1.468 |  | in | FDA,3 | 1.468 | 1.478 | TM 143 / Visual | Main_Plans
2OP280LI | 2OP280LI/Lid_17A | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness (per specification) | Orifice Flap Thickness | 0.035 |  | in | FDA,3 | 0.03 | 0.04 | TM 143 / Visual | Main_Plans
2OP280LI | 2OP280LI/Lid_17A | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.484 |  | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual | Main_Plans
2OP280LI | 2OP280LI/Lid_17A | Lid_17A | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 37.555 |  | g | FDA,1 | 36 | 42 | TM 143 / Visual | Main_Plans
2OP300LI | 2OP300LI/Lid_13E | Lid_13E | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2OP300LI | 2OP300LI/Lid_13E | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OP300LI | 2OP300LI/Lid_13E | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | Bead Diameter | 4.852 |  | in | FDA,3 | 4.847 | 4.867 | TM 143 / Visual | Main_Plans
2OP300LI | 2OP300LI/Lid_13E | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Center to Center Distance | Center to Center Distance | 0.455 |  | in | FDA,3 | 0.445 | 0.455 | TM 143 / Visual | Main_Plans
2OP300LI | 2OP300LI/Lid_13E | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Translucent Blue (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OP300LI | 2OP300LI/Lid_13E | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | End Circle Diameter | End Circle Diameter | 0.128 |  | in | FDA,3 | 0.12 | 0.13 | TM 143 / Visual | Main_Plans
2OP300LI | 2OP300LI/Lid_13E | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OP300LI | 2OP300LI/Lid_13E | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Insert Fin Gap (per specification) | Insert Fin Gap | 0.08 |  | in | FDA,3 | 0.07 | 0.08 | TM 143 / Visual | Main_Plans
2OP300LI | 2OP300LI/Lid_13E | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.04 |  | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual | Main_Plans
2OP300LI | 2OP300LI/Lid_13E | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OP300LI | 2OP300LI/Lid_13E | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness (per specification) | Orifice Flap Thickness | 0.031 |  | in | FDA,3 | 0.03 | 0.04 | TM 143 / Visual | Main_Plans
2OP300LI | 2OP300LI/Lid_13E | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.4885 |  | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual | Main_Plans
2OP300LI | 2OP300LI/Lid_13E | Lid_13E | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 35.535 |  | g | FDA,1 | 32 | 42 | TM 143 / Visual | Main_Plans
2OP300LI | 2OP300LI/Lid_13E | Lid_13E | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | Flip Lid Force | 1.05 |  | lb | FDA,2 | 0.75 |  | TM 148 | Main_Plans
2OP435LI | 2OP435LI/Lid_04C_560 | Lid_04C_560 | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2OP435LI | 2OP435LI/Lid_04C_560 | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OP435LI | 2OP435LI/Lid_04C_560 | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Cap Thickness | Cap Thickness | 0.157 |  | in | FDA,3 | 0.147 | 0.157 | TM 143 / Visual | Main_Plans
2OP435LI | 2OP435LI/Lid_04C_560 | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Translucent Purple (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OP435LI | 2OP435LI/Lid_04C_560 | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OP435LI | 2OP435LI/Lid_04C_560 | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter (as per specification) | Lid Inside Diameter | 3.765 |  | in | FDA,3 | 3.76 | 3.77 | TM 143 / Visual | Main_Plans
2OP435LI | 2OP435LI/Lid_04C_560 | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.04 |  | in | FDA,3 | 0.04 | 0.05 | TM 143 / Visual | Main_Plans
2OP435LI | 2OP435LI/Lid_04C_560 | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OP435LI | 2OP435LI/Lid_04C_560 | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.492 |  | in | FDA,3 | 1.473 | 1.493 | TM 143 / Visual | Main_Plans
2OP435LI | 2OP435LI/Lid_04C_560 | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Outer Diameter (per specification) | Pop-Up Hole Outer Diameter | 0.565 |  | in | FDA,3 | 0.555 | 0.565 | TM 143 / Visual | Main_Plans
2OP435LI | 2OP435LI/Lid_04C_560 | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Thicknes | Pop-Up Hole Thicknes | 0.02 |  | in | FDA,3 | 0.015 | 0.025 | TM 143 / Visual | Main_Plans
2OP435LI | 2OP435LI/Lid_04C_560 | Lid_04C_560 | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 21.1 |  | g | FDA,1 | 20 | 24 | TM 143 / Visual | Main_Plans
2OP630 | 2OP630/Web_NP8440_310 | Web_NP8440_310 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2OP630 | 2OP630/Web_NP8440_310 | Web_NP8440_310 | NY-Analytical | 761 | Web Measurement | Web Width (mm) | Web Width | 310 |  | mm | FDA,0 | 308 | 312 | Measurement | Main_Plans
2OP630 | 2OP630/Web_NP8440_310 | Web_NP8440_310 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 50.66 |  | #/Ream | FDA,2 | 44.88 | 57.12 | TM 101 | Main_Plans
2OP630 | 2OP630/Web_NP8440_310 | Web_NP8440_310 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 6841.04 |  | g/in | FDA,0 | 3500 |  | TM 107 | Main_Plans
2OP630 | 2OP630/Web_NP8440_310 | Web_NP8440_310 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 3.2 |  | mil | FDA,0 | 2.64 | 3.36 | TM 132 | Main_Plans
2OP638 | 2OP638/Reduced Testing_NY_Can 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) | [5 empty] | COA | Main_Plans
2OP638 | 2OP638/Reduced Testing_NY_Can 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2OP638 | 2OP638/Reduced Testing_NY_Can 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Bead OD | Bead OD | 4.9545 |  | in | FDA,3 | 4.922 | 4.958 | Canister & Lid Inspection | Main_Plans
2OP638 | 2OP638/Reduced Testing_NY_Can 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Canister OD | Canister OD | 5.031 |  | in | FDA,3 | 4.956 | 5.076 | Canister & Lid Inspection | Main_Plans
2OP638 | 2OP638/Reduced Testing_NY_Can 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Neck OD | Neck OD | 4.8425 |  | in | FDA,3 | 4.824 | 4.86 | Canister & Lid Inspection | Main_Plans
2OP638 | 2OP638/Reduced Testing_NY_Can 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Total Height OD | Total Height OD | 8.75 |  | in | FDA,3 | 8.688 | 8.808 | Canister & Lid Inspection | Main_Plans
2OP638 | 2OP638/Reduced Testing_NY_Can 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | [6 empty] | Canister & Lid Inspection | Main_Plans
2OP638 | 2OP638/Reduced Testing_NY_Can 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | HDPE Resin Identification Marker | [6 empty] | Canister & Lid Inspection | Main_Plans
2OP638 | 2OP638/Reduced Testing_NY_Can 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Lip Width (for reference only) | Lip Width (for reference only) | 0.15 |  | in | FDA,3 | 0.14 |  | Canister & Lid Inspection | Main_Plans
2OP638 | 2OP638/Reduced Testing_NY_Can 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Wall Thickness (Bottom base area) | Wall Thickness (Bottom base area) |  |  | in | FDA,3 | 0.01 |  | Canister & Lid Inspection | Main_Plans
2OP638 | 2OP638/Reduced Testing_NY_Can 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Wall Thickness (Top area) | Wall Thickness (Top area) |  |  | in | FDA,3 | 0.02 |  | Canister & Lid Inspection | Main_Plans
2OP638 | 2OP638/Reduced Testing_NY_Can 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight | 90.2 |  | g | FDA,1 | 85.5 | 94.5 | Measurement | Main_Plans
2OP638 | 2OP638/Reduced Testing_NY_Can 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 1052 | Top Load Compression Strength | Top load Compression Strength (0.5"/mSR,0.450"TD) | Top load Compression Strength (0.5"/mSR,0.450"TD) |  |  | lb | FDA,0 | 97 |  | NPTM 1019 | Main_Plans
2OP638 | 2OP638/Reduced Testing_NY_Can 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OP638 | 2OP638/Reduced Testing_NY_Can 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | White (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OP638 | 2OP638/Reduced Testing_NY_Can 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OP638 | 2OP638/Reduced Testing_NY_Can 2OP638 | Reduced Testing_NY_Can 2OP638 | NY-Analytical | 1053 | Cavity | cavity Inserts | cavity Inserts | [6 empty] | Visual | Main_Plans
2OP638LI | 2OP638LI/Lid_13D | Lid_13D | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2OP638LI | 2OP638LI/Lid_13D | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OP638LI | 2OP638LI/Lid_13D | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | Bead Diameter | 4.8615 |  | in | FDA,3 | 4.847 | 4.867 | TM 143 / Visual | Main_Plans
2OP638LI | 2OP638LI/Lid_13D | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Center to Center Distance | Center to Center Distance | 0.4535 |  | in | FDA,3 | 0.445 | 0.455 | TM 143 / Visual | Main_Plans
2OP638LI | 2OP638LI/Lid_13D | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Translucent Gray (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OP638LI | 2OP638LI/Lid_13D | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | End Circle Diameter | End Circle Diameter | 0.122 |  | in | FDA,3 | 0.12 | 0.13 | TM 143 / Visual | Main_Plans
2OP638LI | 2OP638LI/Lid_13D | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OP638LI | 2OP638LI/Lid_13D | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Insert Fin Gap (per specification) | Insert Fin Gap | 0.0745 |  | in | FDA,3 | 0.07 | 0.08 | TM 143 / Visual | Main_Plans
2OP638LI | 2OP638LI/Lid_13D | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter (as per specification) | Lid Inside Diameter | 4.9218 |  | in | FDA,3 | 4.917 | 4.937 | TM 143 / Visual | Main_Plans
2OP638LI | 2OP638LI/Lid_13D | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.041 |  | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual | Main_Plans
2OP638LI | 2OP638LI/Lid_13D | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OP638LI | 2OP638LI/Lid_13D | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness (per specification) | Orifice Flap Thickness | 0.0345 |  | in | FDA,3 | 0.03 | 0.04 | TM 143 / Visual | Main_Plans
2OP638LI | 2OP638LI/Lid_13D | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.487 |  | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual | Main_Plans
2OP638LI | 2OP638LI/Lid_13D | Lid_13D | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 33.92 |  | g | FDA,1 | 32 | 42 | TM 143 / Visual | Main_Plans
2OP638LI | 2OP638LI/Lid_13D | Lid_13D | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | Flip Lid Force | 1.79 |  | lb | FDA,2 | 0.75 |  | TM 148 | Main_Plans
2OQ753 | 2OQ753/Reduced Testing_NY_Can 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid has a tight friction fit | [5 empty] | Canister & Lid Inspection | Main_Plans
2OQ753 | 2OQ753/Reduced Testing_NY_Can 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) | [5 empty] | COA | Main_Plans
2OQ753 | 2OQ753/Reduced Testing_NY_Can 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Certificate of Analysis | Main_Plans
2OQ753 | 2OQ753/Reduced Testing_NY_Can 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2OQ753 | 2OQ753/Reduced Testing_NY_Can 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | White (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2OQ753 | 2OQ753/Reduced Testing_NY_Can 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Bead OD | Bead OD | 4.9355 |  | in | FDA,3 | 4.908 | 4.944 | Canister & Lid Inspection | Main_Plans
2OQ753 | 2OQ753/Reduced Testing_NY_Can 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Canister OD | Canister OD | 5.0145 |  | in | FDA,3 | 4.998 | 5.034 | Canister & Lid Inspection | Main_Plans
2OQ753 | 2OQ753/Reduced Testing_NY_Can 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Lip (width) | Lip Width | 0.17 |  | in | FDA,3 | 0.142 | 0.178 | Canister & Lid Inspection | Main_Plans
2OQ753 | 2OQ753/Reduced Testing_NY_Can 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Total Height OD | Total Height OD | 6.748 |  | in | FDA,3 | 6.706 | 6.826 | Canister & Lid Inspection | Main_Plans
2OQ753 | 2OQ753/Reduced Testing_NY_Can 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | HDPE Resin Identification Marker | Meets the Specification (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2OQ753 | 2OQ753/Reduced Testing_NY_Can 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2OQ753 | 2OQ753/Reduced Testing_NY_Can 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight | 69.21 |  | g | FDA,1 | 67.5 | 72.5 | Measurement | Main_Plans
2OQ753 | 2OQ753/Reduced Testing_NY_Can 2OQ753 | Reduced Testing_NY_Can 2OQ753 | NY-Analytical | 766 | Lid Pull Force | Lid Pull-Off Force | Lid Pull-Off Force | 2859.39 |  | g | FDA,0 | 2000 | 3500 | TM 148 | Main_Plans
2OQ753X | 2OQ753X/Reduced Testing_NY_Can 2OQ753 X | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid has a tight friction fit | [5 empty] | Canister & Lid Inspection | Main_Plans
2OQ753X | 2OQ753X/Reduced Testing_NY_Can 2OQ753 X | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) | [5 empty] | COA | Main_Plans
2OQ753X | 2OQ753X/Reduced Testing_NY_Can 2OQ753 X | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [6 empty] | Canister & Lid Inspection | Main_Plans
2OQ753X | 2OQ753X/Reduced Testing_NY_Can 2OQ753 X | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2OQ753X | 2OQ753X/Reduced Testing_NY_Can 2OQ753 X | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | White (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2OQ753X | 2OQ753X/Reduced Testing_NY_Can 2OQ753 X | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | HDPE Resin Identification Marker | Meets the Specification (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2OQ753X | 2OQ753X/Reduced Testing_NY_Can 2OQ753 X | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2OQ753X | 2OQ753X/Reduced Testing_NY_Can 2OQ753 X | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight | 70.655 |  | g | FDA,1 | 66.5 | 73.5 | Measurement | Main_Plans
2OQ753X | 2OQ753X/Reduced Testing_NY_Can 2OQ753 X | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 1052 | Top Load Compression Strength | Top load Compression Strength (0.5"/mSR,0.450"TD) | Top load Compression Strength (0.5"/mSR,0.450"TD) |  |  | lb | FDA,0 | 84 |  | NPTM 1019 | Main_Plans
2OQ753X | 2OQ753X/Reduced Testing_NY_Can 2OQ753 X | Reduced Testing_NY_Can 2OQ753 X | NY-Analytical | 1053 | Cavity | cavity Inserts | cavity Inserts | Meets specification (Done) | [5 empty] | Visual | Main_Plans
2OQ850 | 2OQ850/Reduced Testing_NY_Can 2OQ850 | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid has a tight friction fit | [5 empty] | Canister & Lid Inspection | Main_Plans
2OQ850 | 2OQ850/Reduced Testing_NY_Can 2OQ850 | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) | [5 empty] | COA | Main_Plans
2OQ850 | 2OQ850/Reduced Testing_NY_Can 2OQ850 | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Certificate of Analysis | Main_Plans
2OQ850 | 2OQ850/Reduced Testing_NY_Can 2OQ850 | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2OQ850 | 2OQ850/Reduced Testing_NY_Can 2OQ850 | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | White (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2OQ850 | 2OQ850/Reduced Testing_NY_Can 2OQ850 | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | HDPE Resin Identification Marker | Meets the Specification (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2OQ850 | 2OQ850/Reduced Testing_NY_Can 2OQ850 | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2OQ850 | 2OQ850/Reduced Testing_NY_Can 2OQ850 | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight | 88.41 |  | g | FDA,1 | 87 | 93 | Measurement | Main_Plans
2OQ850 | 2OQ850/Reduced Testing_NY_Can 2OQ850 | Reduced Testing_NY_Can 2OQ850 | NY-Analytical | 766 | Lid Pull Force | Lid Pull-Off Force | Lid Pull-Off Force |  |  | g | FDA,0 | 2000 | 3500 | TM 148 | Main_Plans
2OQ850LI | 2OQ850LI/Lid_11B | Lid_11B | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Flip Lid Force | Main_Plans
2OQ850LI | 2OQ850LI/Lid_11B | Lid_11B | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | Flip Lid Force | 2.41 |  | lb | FDA,2 | 0.75 |  | TM 143 | Main_Plans
2OQ850LI | 2OQ850LI/Lid_11B | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OQ850LI | 2OQ850LI/Lid_11B | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | Bead Diameter | 4.861 |  | in | FDA,3 | 4.847 | 4.867 | TM 143 / Visual | Main_Plans
2OQ850LI | 2OQ850LI/Lid_11B | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Translucent Red (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OQ850LI | 2OQ850LI/Lid_11B | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OQ850LI | 2OQ850LI/Lid_11B | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Insert Fin Gap (per specification) | Insert Fin Gap | 0.073 |  | in | FDA,3 | 0.065 | 0.075 | TM 143 / Visual | Main_Plans
2OQ850LI | 2OQ850LI/Lid_11B | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter (as per specification) | Lid Inside Diameter | 4.919 |  | in | FDA,3 | 4.917 | 4.937 | TM 143 / Visual | Main_Plans
2OQ850LI | 2OQ850LI/Lid_11B | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.0425 |  | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual | Main_Plans
2OQ850LI | 2OQ850LI/Lid_11B | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OQ850LI | 2OQ850LI/Lid_11B | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness (per specification) | Orifice Flap Thickness | 0.024 |  | in | FDA,3 | 0.018 | 0.028 | TM 143 / Visual | Main_Plans
2OQ850LI | 2OQ850LI/Lid_11B | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.488 |  | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual | Main_Plans
2OQ850LI | 2OQ850LI/Lid_11B | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Outer Diameter (per specification) | Pop-Up Hole Outer Diameter | 0.165 |  | in | FDA,3 | 0.16 | 0.17 | TM 143 / Visual | Main_Plans
2OQ850LI | 2OQ850LI/Lid_11B | Lid_11B | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 34.04 |  | g | FDA,1 | 33 | 42 | TM 143 / Visual | Main_Plans
2OQ882LI | 2OQ882LI/Lid_12A | Lid_12A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Flip Lid Force | Main_Plans
2OQ882LI | 2OQ882LI/Lid_12A | Lid_12A | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | Flip Lid Force | 1.61 |  | lb | FDA,2 | 0.75 |  | TM 143 | Main_Plans
2OQ882LI | 2OQ882LI/Lid_12A | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OQ882LI | 2OQ882LI/Lid_12A | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | Bead Diameter | 4.854 |  | in | FDA,3 | 4.852 | 4.862 | TM 143 / Visual | Main_Plans
2OQ882LI | 2OQ882LI/Lid_12A | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Cap Thickness | Cap Thickness | 0.048 |  | in | FDA,3 | 0.045 | 0.055 | TM 143 / Visual | Main_Plans
2OQ882LI | 2OQ882LI/Lid_12A | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Dark Green (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OQ882LI | 2OQ882LI/Lid_12A | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OQ882LI | 2OQ882LI/Lid_12A | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Insert Fin Gap (per specification) | Insert Fin Gap | 0.074 |  | in | FDA,3 | 0.065 | 0.075 | TM 143 / Visual | Main_Plans
2OQ882LI | 2OQ882LI/Lid_12A | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.045 |  | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual | Main_Plans
2OQ882LI | 2OQ882LI/Lid_12A | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) | [5 empty] | TM 143 / Visual | Main_Plans
2OQ882LI | 2OQ882LI/Lid_12A | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness (per specification) | Orifice Flap Thickness | 0.021 |  | in | FDA,3 | 0.018 | 0.028 | TM 143 / Visual | Main_Plans
2OQ882LI | 2OQ882LI/Lid_12A | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.488 |  | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual | Main_Plans
2OQ882LI | 2OQ882LI/Lid_12A | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Outer Diameter (per specification) | Pop-Up Hole Outer Diameter | 0.168 |  | in | FDA,3 | 0.16 | 0.17 | TM 143 / Visual | Main_Plans
2OQ882LI | 2OQ882LI/Lid_12A | Lid_12A | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 35.675 |  | g | FDA,1 | 32 | 42 | TM 143 / Visual | Main_Plans
2OU130 | 2OU130/Web_NP5115_9.0 | Web_NP5115_9.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2OU130 | 2OU130/Web_NP5115_9.0 | Web_NP5115_9.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 8.9375 |  | in | FDA,4 | 8.9375 | 9.0625 | Measurement | Main_Plans
2OU130 | 2OU130/Web_NP5115_9.0 | Web_NP5115_9.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Printed/Varnished) | Basis Weight | 60.8 |  | #/Ream | FDA,1 | 55.6 | 68 | TM 101 | Main_Plans
2OU130 | 2OU130/Web_NP5115_9.0 | Web_NP5115_9.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@325DegF/40PSI/0.5Sec_Dwell) | Seal Strength (@325°F/40PSI/0.5Sec Dwell) | 2198.05 |  | g/in | FDA,0 | 1500 |  | TM 107 | Main_Plans
2OX465B | 2OX465B/Bottle, 26ml | Bottle, 26ml | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2OX465B | 2OX465B/Bottle, 26ml | Bottle, 26ml | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2OX465B | 2OX465B/Bottle, 26ml | Bottle, 26ml | NY-Analytical | 765 | Canister & Lid Inspection | Bottle Internal Diameter at Flange | Bottle Internal Diameter at Flange | 0.722 |  | in | FDA,3 | 0.709 | 0.729 | Canister & Lid Inspection | Main_Plans
2OX465B | 2OX465B/Bottle, 26ml | Bottle, 26ml | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Natural (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2OX465B | 2OX465B/Bottle, 26ml | Bottle, 26ml | NY-Analytical | 765 | Canister & Lid Inspection | Overall Length | Overall Length | 4.855 |  | in | FDA,3 | 4.825 | 4.875 | Canister & Lid Inspection | Main_Plans
2OX465B | 2OX465B/Bottle, 26ml | Bottle, 26ml | NY-Analytical | 765 | Canister & Lid Inspection | Seal Flange Diameter | Seal Flange Diameter | 0.8735 |  | in | FDA,3 | 0.864 | 0.874 | Canister & Lid Inspection | Main_Plans
2OX465B | 2OX465B/Bottle, 26ml | Bottle, 26ml | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 2 |  | g | FDA,2 |  |  | TM 610 | Main_Plans
2OX465B | 2OX465B/Bottle, 26ml | Bottle, 26ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer | [5 empty] | TM 610 | Main_Plans
2OX465B | 2OX465B/Bottle, 26ml | Bottle, 26ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA | [5 empty] | TM 610 | Main_Plans
2OX465B | 2OX465B/Bottle, 26ml | Bottle, 26ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
2OX465B | 2OX465B/Bottle, 26ml | Bottle, 26ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
2OX465B | 2OX465B/Bottle, 26ml | Bottle, 26ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
2OX465B | 2OX465B/Bottle, 26ml | Bottle, 26ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 1 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
2OX465B | 2OX465B/Bottle, 26ml | Bottle, 26ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0.0028852 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Main_Plans
2OX465B | 2OX465B/Bottle, 26ml | Bottle, 26ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
2OX465B | 2OX465B/Bottle, 26ml | Bottle, 26ml | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Negative Rods (OOS) | [5 empty] | WI QAS 076 | Main_Plans
2OX465B | 2OX465B/Bottle, 26ml | Bottle, 26ml | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Mold (OOS) | [5 empty] | WI QAS 076 | Main_Plans
2OX510B | 2OX510B/Bottle, 5.1ml | Bottle, 5.1ml | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Certificate of Conformance | Main_Plans
2OX510B | 2OX510B/Bottle, 5.1ml | Bottle, 5.1ml | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2OX510B | 2OX510B/Bottle, 5.1ml | Bottle, 5.1ml | NY-Analytical | 765 | Canister & Lid Inspection | Bottle Internal Diameter at Flange | Bottle Internal Diameter at Flange | 0.443 |  | in | FDA,3 | 0.442 | 0.452 | Canister & Lid Inspection | Main_Plans
2OX510B | 2OX510B/Bottle, 5.1ml | Bottle, 5.1ml | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Natural (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
2OX510B | 2OX510B/Bottle, 5.1ml | Bottle, 5.1ml | NY-Analytical | 765 | Canister & Lid Inspection | Overall Length | Overall Length | 2.8737 |  | in | FDA,3 | 2.848 | 2.888 | Canister & Lid Inspection | Main_Plans
2OX510B | 2OX510B/Bottle, 5.1ml | Bottle, 5.1ml | NY-Analytical | 765 | Canister & Lid Inspection | Seal Flange Diameter | Seal Flange Diameter | 0.55 |  | in | FDA,3 | 0.544 | 0.554 | Canister & Lid Inspection | Main_Plans
2OX510B | 2OX510B/Bottle, 5.1ml | Bottle, 5.1ml | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 4.5 |  | g | FDA,2 |  |  | TM 610 | Main_Plans
2OX510B | 2OX510B/Bottle, 5.1ml | Bottle, 5.1ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA | [5 empty] | TM 610 | Main_Plans
2OX510B | 2OX510B/Bottle, 5.1ml | Bottle, 5.1ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA | [5 empty] | TM 610 | Main_Plans
2OX510B | 2OX510B/Bottle, 5.1ml | Bottle, 5.1ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 1.1 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
2OX510B | 2OX510B/Bottle, 5.1ml | Bottle, 5.1ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0.24444 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
2OX510B | 2OX510B/Bottle, 5.1ml | Bottle, 5.1ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
2OX510B | 2OX510B/Bottle, 5.1ml | Bottle, 5.1ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 1 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
2OX510B | 2OX510B/Bottle, 5.1ml | Bottle, 5.1ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0.022222 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Main_Plans
2OX510B | 2OX510B/Bottle, 5.1ml | Bottle, 5.1ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 1 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
2OX510B | 2OX510B/Bottle, 5.1ml | Bottle, 5.1ml | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Positive Rods (Done) | [5 empty] | WI QAS 076 | Main_Plans
2OX510B | 2OX510B/Bottle, 5.1ml | Bottle, 5.1ml | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Mold (OOS) | [5 empty] | WI QAS 076 | Main_Plans
2OX510FS | 2OX510FS/Foil seal 5.1 ml | Foil seal 5.1 ml | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | COA | Main_Plans
2OX510FS | 2OX510FS/Foil seal 5.1 ml | Foil seal 5.1 ml | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 1.78 |  | g | FDA,2 |  |  | TM 610 | Main_Plans
2OX510FS | 2OX510FS/Foil seal 5.1 ml | Foil seal 5.1 ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA | [5 empty] | TM 610 | Main_Plans
2OX510FS | 2OX510FS/Foil seal 5.1 ml | Foil seal 5.1 ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer | [5 empty] | TM 610 | Main_Plans
2OX510FS | 2OX510FS/Foil seal 5.1 ml | Foil seal 5.1 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 11.6 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
2OX510FS | 2OX510FS/Foil seal 5.1 ml | Foil seal 5.1 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 6.5169 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
2OX510FS | 2OX510FS/Foil seal 5.1 ml | Foil seal 5.1 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 11 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
2OX510FS | 2OX510FS/Foil seal 5.1 ml | Foil seal 5.1 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
2OX510FS | 2OX510FS/Foil seal 5.1 ml | Foil seal 5.1 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Main_Plans
2OX510FS | 2OX510FS/Foil seal 5.1 ml | Foil seal 5.1 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
2OX510FS | 2OX510FS/Foil seal 5.1 ml | Foil seal 5.1 ml | NY-Analytical | 1056 | Functionality | Foil diameter to fit Bottle(Item#2OX510B)opening | Foil diameter to fit Bottle(Item#2OX510B)opening | Pass | [5 empty] | Visual | Main_Plans
2OX510FS | 2OX510FS/Foil seal 5.1 ml | Foil seal 5.1 ml | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Negative Rods (OOS) | [5 empty] | WI QAS 076 | Main_Plans
2PD351 | 2PD351/Web_NP5116_4.5 | Web_NP5116_4.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2PD351 | 2PD351/Web_NP5116_4.5 | Web_NP5116_4.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 4.5 |  | in | FDA,4 | 4.4375 | 4.5625 | Measurement | Main_Plans
2PD351 | 2PD351/Web_NP5116_4.5 | Web_NP5116_4.5 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Printed/Varnished) | Basis Weight | 59.82 |  | #/Ream | FDA,1 | 55.8 | 68.2 | TM 101 | Main_Plans
2PD351 | 2PD351/Web_NP5116_4.5 | Web_NP5116_4.5 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/40psi/0.5 sec dewll) | Seal Strength (@300°F/40psi/0.5 sec dewll) | 2163 |  | g/in | FDA,0 | 1500 |  | TM 107 | Main_Plans
2QB114X | 2QB114X/Web_NP3860_5.0 | Web_NP3860_5.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2QB114X | 2QB114X/Web_NP3860_5.0 | Web_NP3860_5.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 5 |  | in | FDA,4 | 4.9375 | 5.0625 | Measurement | Main_Plans
2QB114X | 2QB114X/Web_NP3860_5.0 | Web_NP3860_5.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 89.29 |  | #/Ream | FDA,1 | 78.9 | 96.5 | TM 101 | Main_Plans
2QB114X | 2QB114X/Web_NP3860_5.0 | Web_NP3860_5.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@350DegF/40psi/1.0 sec dwell) | Seal Strength (@350°F/40psi/1.0 sec dwell) | 9928 |  | g/in | FDA,0 | 7250 |  | TM 107 | Main_Plans
2QB114X | 2QB114X/Web_NP3860_5.0 | Web_NP3860_5.0 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 5 |  | mil | FDA,1 | 4.7 | 5.7 | TM 132 | Main_Plans
2RQ890 | 2RQ890/Web_NP9508_18.125 | Web_NP9508_18.125 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2RQ890 | 2RQ890/Web_NP9508_18.125 | Web_NP9508_18.125 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 18.125 |  | in | FDA,4 | 18.0625 | 18.1875 | Measurement | Main_Plans
2RQ890 | 2RQ890/Web_NP9508_18.125 | Web_NP9508_18.125 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Unprinted) | Basis Weight |  |  | #/Ream | FDA,1 | 37.8 | 46.2 | TM 101 | Main_Plans
2RQ890 | 2RQ890/Web_NP9508_18.125 | Web_NP9508_18.125 | NY-Analytical | 762 | Seal Strength | Seal Strength (@300DegF/40psi/0.5 sec dewll) | Seal Strength (@300°F/40psi/0.5 sec dewll) | Sample peeled apart and showed stress whitening when sealed (Done) | [5 empty] | TM 107 | Main_Plans
2RQ890 | 2RQ890/Web_NP9508_18.125 | Web_NP9508_18.125 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 2.6 |  | mil | FDA,1 | 2.3 | 2.8 | TM 132 | Main_Plans
2RS324 | 2RS324/Web_NP3860_12.5 | Web_NP3860_12.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2RS324 | 2RS324/Web_NP3860_12.5 | Web_NP3860_12.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 12.5 |  | in | FDA,4 | 12.4375 | 12.5625 | Measurement | Main_Plans
2RS324 | 2RS324/Web_NP3860_12.5 | Web_NP3860_12.5 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 89 |  | #/Ream | FDA,1 | 78.9 | 96.5 | TM 101 | Main_Plans
2RS324 | 2RS324/Web_NP3860_12.5 | Web_NP3860_12.5 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@350DegF/40psi/1.0 sec dwell) | Seal Strength (@350°F/40psi/1.0 sec dwell) | 9622 |  | g/in | FDA,0 | 7250 |  | TM 107 | Main_Plans
2RS324 | 2RS324/Web_NP3860_12.5 | Web_NP3860_12.5 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 4.91 |  | mil | FDA,1 | 4.7 | 5.7 | TM 132 | Main_Plans
2RS407 | 2RS407/Web_NP3885_13.0 | Web_NP3885_13.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2RS407 | 2RS407/Web_NP3885_13.0 | Web_NP3885_13.0 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 12.9375 |  | in | FDA,4 | 12.9375 | 13.0625 | Measurement | Main_Plans
2RS407 | 2RS407/Web_NP3885_13.0 | Web_NP3885_13.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 80.69 |  | #/Ream | FDA,0 | 74 | 90 | TM 101 | Main_Plans
2RS407 | 2RS407/Web_NP3885_13.0 | Web_NP3885_13.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@350DegF/50psi/1.0 sec dwell) | Seal Strength (@350°F/50psi/1.0 sec dwell) | 1941 |  | g/in | FDA,0 | 700 |  | TM 107 | Main_Plans
2RS419 | 2RS419/Web_NP3885_16 | Web_NP3885_16 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2RS419 | 2RS419/Web_NP3885_16 | Web_NP3885_16 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 16 |  | in | FDA,4 | 15.9375 | 16.0625 | Measurement | Main_Plans
2RS419 | 2RS419/Web_NP3885_16 | Web_NP3885_16 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 81.85 |  | #/Ream | FDA,0 | 74 | 90 | TM 101 | Main_Plans
2RS419 | 2RS419/Web_NP3885_16 | Web_NP3885_16 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@350DegF/50psi/1.0 sec dwell) | Seal Strength (@350°F/50psi/1.0 sec dwell) | 2030.5 |  | g/in | FDA,0 | 700 |  | TM 107 | Main_Plans
2SS273 | 2SS273/Web_NP3860_15.5 | Web_NP3860_15.5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Main_Plans
2SS273 | 2SS273/Web_NP3860_15.5 | Web_NP3860_15.5 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 15.5 |  | in | FDA,4 | 15.4375 | 15.5625 | Measurement | Main_Plans
2SS273 | 2SS273/Web_NP3860_15.5 | Web_NP3860_15.5 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 87.7 |  | #/Ream | FDA,1 | 78.9 | 96.5 | TM 101 | Main_Plans
2SS273 | 2SS273/Web_NP3860_15.5 | Web_NP3860_15.5 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@350DegF/40psi/1.0 sec dwell) | Seal Strength (@350°F/40psi/1.0 sec dwell) | 10111.34 |  | g/in | FDA,0 | 7250 |  | TM 107 | Main_Plans
2SS273 | 2SS273/Web_NP3860_15.5 | Web_NP3860_15.5 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 5.2 |  | mil | FDA,1 | 4.7 | 5.7 | TM 132 | Main_Plans
2SS428 | 2SS428/Web_NP3860_8.0 | Web_NP3860_8.0 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Basis Weight | Main_Plans
2SS428 | 2SS428/Web_NP3860_8.0 | Web_NP3860_8.0 | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 87.36 |  | #/Ream | FDA,1 | 78.9 | 96.5 | TM 101 | Main_Plans
2SS428 | 2SS428/Web_NP3860_8.0 | Web_NP3860_8.0 | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@350DegF/40psi/1.0 sec dwell) | Seal Strength (@350°F/40psi/1.0 sec dwell) | 9981.29 |  | g/in | FDA,0 | 7250 |  | TM 107 | Main_Plans
2SS428 | 2SS428/Web_NP3860_8.0 | Web_NP3860_8.0 | NY-Analytical | 763 | Web Thickness | Web_Thickness_(mils) | Web Thickness | 5.2 |  | mil | FDA,1 | 4.7 | 5.7 | TM 132 | Main_Plans
3AP211 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Applicator Measurements | Main_Plans
3AP211 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 44 |  | in | FDA,0 | 44 | 46 | Applicator Measurements | Main_Plans
3AP211 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 37.5 |  | in | FDA,1 | 37.3 | 37.7 | Applicator Measurements | Main_Plans
3AP211 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 55.54 |  | GSM | FDA,1 | 49.5 | 60.5 | TM 101 | Main_Plans
3AP211 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 1147.63 |  | g/in | FDA,0 | 692 | 1438 | TM 103 | Main_Plans
3AP211 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 4131.08 |  | g/in | FDA,0 | 2785 | 5785 | TM 103 | Main_Plans
3AP211 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3AP211 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 748 | Visual Inspection | Color | Color | White to Off-White (Done) | [5 empty] | TM 208 | Main_Plans
3AP211 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3AP211 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 509 | J&J_Pathogens | Biggy or SDA | Biggy or SDA | No Growth_SDA | [5 empty] | TM 869 | Main_Plans
3AP211 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 509 | J&J_Pathogens | Cetrimide | Cetrimide | No Growth | [5 empty] | TM 869 | Main_Plans
3AP211 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 524 | J&J Fungal | JJ Fungal_RM | JJ Fungal RM | 0 |  | CFU/g |  |  | 200 | TM 869 | Main_Plans
3AP211 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 509 | J&J_Pathogens | MacConkey | MacConkey | No Growth | [5 empty] | TM 869 | Main_Plans
3AP211 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 509 | J&J_Pathogens | MSA or VJ | MSA or VJ | No Growth_MSA | [5 empty] | TM 869 | Main_Plans
3AP211 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 700 | J&J SDA Controls | Negative Control | Negative Control | No Growth SDA | [5 empty] | TM 869 | Main_Plans
3AP211 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | NegativeControl | No Growth TSB | [5 empty] | TM 869 | Main_Plans
3AP211 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 509 | J&J_Pathogens | TSA | TSA | No Growth | [5 empty] | TM 869 | Main_Plans
3AP211 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ RM | TSA JJ RM | 0 |  | CFU/g |  |  | 200 | TM 869 | Main_Plans
3AP211 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 517 | J&J_Enrichment | TSB | TSB | Not Turbid | [5 empty] | TM 869 | Main_Plans
3AP211 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 573 | GramStain_Candida | Gram Stain_J&JCandida | Gram Stain J&JCandida | Non-Subculturable Growth (Done) | [5 empty] | WI QAS 076 | Main_Plans
3AP211 | 3AP211/Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | Applicator_NY_Spunlaced Nonwoven 55gsm 37 1/2" | NY-Microbiology | 545 | GramStain_J&J_TSA | GramStain_J&J_TSA | GramStain J&J TSA | Gram Positive Rods (Done) | [5 empty] | WI QAS 076 | Main_Plans
3AQ470-A | 3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading | [5 empty] | Color (TM 208) | Main_Plans
3AQ470-A | 3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Free of hard wrinkles. Formation uniform (Done) | [5 empty] | See Analyte for Method # | Main_Plans
3AQ470-A | 3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | See Analyte for Method # | Main_Plans
3AQ470-A | 3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 49.51 |  | GSM | FDA,1 | 45 | 55 | See Analyte for Method # | Main_Plans
3AQ470-A | 3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (mm) (CofA) | Caliper | 0.59 |  | mm | FDA,2 | 0.51 | 0.71 | See Analyte for Method # | Main_Plans
3AQ470-A | 3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | See Analyte for Method # | Main_Plans
3AQ470-A | 3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | May have slight odor when wet (Done) | [5 empty] | See Analyte for Method # | Main_Plans
3AQ470-A | 3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter | 50 |  | in | FDA,0 | 45 | 60 | See Analyte for Method # | Main_Plans
3AQ470-A | 3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Width (Inches) (Measurement) | Roll Width | 37.75 |  | in | FDA,2 | 37.5 | 38 | See Analyte for Method # | Main_Plans
3AQ470-A | 3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Wet CD (g/in) (TM 103) | Wet CD | 1634 |  | g/in | FDA,0 | 1100 | 2300 | See Analyte for Method # | Main_Plans
3AQ470-A | 3AQ470-A/Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | Applicator_NY_Rolls Spunlaced Nonwoven - 37.75"-50gsm | NY-Analytical | 884 | Applicator_Inspection | Wet MD (g/in) (TM 103) | Wet MD | 6147 |  | g/in | FDA,0 | 3000 | 7000 | See Analyte for Method # | Main_Plans
3AU228 | 3AU228/Applicator_06I | Applicator_06I | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Applicator Measurements | Main_Plans
3AU228 | 3AU228/Applicator_06I | Applicator_06I | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 30 |  | in | FDA,0 | 28 | 32 | Applicator Measurements | Main_Plans
3AU228 | 3AU228/Applicator_06I | Applicator_06I | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 7.5 |  | in | FDA,3 | 7.375 | 7.625 | Applicator Measurements | Main_Plans
3AU228 | 3AU228/Applicator_06I | Applicator_06I | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 50.29 |  | GSM | FDA,1 | 49.5 | 60.5 | TM 101 | Main_Plans
3AU228 | 3AU228/Applicator_06I | Applicator_06I | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper | 0.52 |  | mm | FDA,2 | 0.5 | 0.64 | TM 102 | Main_Plans
3AU228 | 3AU228/Applicator_06I | Applicator_06I | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 1925.71 |  | g/in | FDA,0 | 1443 | 2463 | TM 103 | Main_Plans
3AU228 | 3AU228/Applicator_06I | Applicator_06I | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 6181.64 |  | g/in | FDA,0 | 4883 | 7967 | TM 103 | Main_Plans
3AU228 | 3AU228/Applicator_06I | Applicator_06I | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3AU228 | 3AU228/Applicator_06I | Applicator_06I | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3AU228 | 3AU228/Applicator_06I | Applicator_06I | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3AU228 | 3AU228/Applicator_06I | Applicator_06I | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 1.03 |  | g | FDA,2 |  |  | TM 610 | Main_Plans
3AU228 | 3AU228/Applicator_06I | Applicator_06I | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer | [5 empty] | TM 610 | Main_Plans
3AU228 | 3AU228/Applicator_06I | Applicator_06I | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer | [5 empty] | TM 610 | Main_Plans
3AU228 | 3AU228/Applicator_06I | Applicator_06I | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3AU228 | 3AU228/Applicator_06I | Applicator_06I | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
3AU228 | 3AU228/Applicator_06I | Applicator_06I | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3AU228 | 3AU228/Applicator_06I | Applicator_06I | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3AU228 | 3AU228/Applicator_06I | Applicator_06I | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Main_Plans
3AU228 | 3AU228/Applicator_06I | Applicator_06I | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3BJ141 | 3BJ141/Applicator_07K | Applicator_07K | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Applicator Measurements | Main_Plans
3BJ141 | 3BJ141/Applicator_07K | Applicator_07K | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 11.625 |  | in | FDA,2 | 11.5 | 11.75 | Applicator Measurements | Main_Plans
3BJ141 | 3BJ141/Applicator_07K | Applicator_07K | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 41.33 |  | GSM | FDA,1 | 38.7 | 47.3 | TM 101 | Main_Plans
3BJ141 | 3BJ141/Applicator_07K | Applicator_07K | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(inch) | Caliper | 0.0225 |  | in | FDA,3 | 0.022 | 0.03 | TM 102 | Main_Plans
3BJ141 | 3BJ141/Applicator_07K | Applicator_07K | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 1064.23 |  | g/in | FDA,0 | 921 | 1535 | TM 103 | Main_Plans
3BJ141 | 3BJ141/Applicator_07K | Applicator_07K | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 5060.98 |  | g/in | FDA,0 | 3623 | 5650 | TM 103 | Main_Plans
3BJ141 | 3BJ141/Applicator_07K | Applicator_07K | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3BJ141 | 3BJ141/Applicator_07K | Applicator_07K | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3BJ141 | 3BJ141/Applicator_07K | Applicator_07K | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3DQ595 | 3DQ595/Applicator_10D | Applicator_10D | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Applicator Measurements | Main_Plans
3DQ595 | 3DQ595/Applicator_10D | Applicator_10D | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 46 |  | in | FDA,0 | 44 | 47 | Applicator Measurements | Main_Plans
3DQ595 | 3DQ595/Applicator_10D | Applicator_10D | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 30 |  | in | FDA,3 | 29.875 | 30.125 | Applicator Measurements | Main_Plans
3DQ595 | 3DQ595/Applicator_10D | Applicator_10D | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 39.18 |  | GSM | FDA,1 | 38 | 42 | TM 101 | Main_Plans
3DQ595 | 3DQ595/Applicator_10D | Applicator_10D | NY-Analytical | 757 | Caliper | Caliper-1_Ply_(mm) | Caliper | 0.36 |  | mm | FDA,2 | 0.35 | 0.45 | TM 102 | Main_Plans
3DQ595 | 3DQ595/Applicator_10D | Applicator_10D | NY-Analytical | 758 | Tensile Strength | Dry_CD_(g/in) | Dry CD | 667.15 |  | g/in | FDA,0 | 513 | 1130 | TM 103 | Main_Plans
3DQ595 | 3DQ595/Applicator_10D | Applicator_10D | NY-Analytical | 758 | Tensile Strength | Dry_MD_(g/in) | Dry CD | 2695.76 |  | g/in | FDA,0 | 2054 | 4108 | TM 103 | Main_Plans
3DQ595 | 3DQ595/Applicator_10D | Applicator_10D | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3DQ595 | 3DQ595/Applicator_10D | Applicator_10D | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3DQ595 | 3DQ595/Applicator_10D | Applicator_10D | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3DQ595N | 3DQ595N/Applicator_10I | Applicator_10I | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Applicator Measurements | Main_Plans
3DQ595N | 3DQ595N/Applicator_10I | Applicator_10I | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 46 |  | in | FDA,0 | 44 | 47 | Applicator Measurements | Main_Plans
3DQ595N | 3DQ595N/Applicator_10I | Applicator_10I | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 30 |  | in | FDA,3 | 29.875 | 30.125 | Applicator Measurements | Main_Plans
3DQ595N | 3DQ595N/Applicator_10I | Applicator_10I | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 39.83 |  | GSM | FDA,0 | 38 | 42 | TM 101 | Main_Plans
3DQ595N | 3DQ595N/Applicator_10I | Applicator_10I | NY-Analytical | 757 | Caliper | Caliper-1_Ply_(mm) | Caliper | 0.41 |  | mm | FDA,2 | 0.35 | 0.45 | TM 102 | Main_Plans
3DQ595N | 3DQ595N/Applicator_10I | Applicator_10I | NY-Analytical | 758 | Tensile Strength | Dry_CD_(g/in) | Dry CD | 573.11 |  | g/in | FDA,0 | 513 | 1130 | TM 103 | Main_Plans
3DQ595N | 3DQ595N/Applicator_10I | Applicator_10I | NY-Analytical | 758 | Tensile Strength | Dry_MD_(g/in) | Dry CD | 3431.99 |  | g/in | FDA,0 | 2054 | 4108 | TM 103 | Main_Plans
3DQ595N | 3DQ595N/Applicator_10I | Applicator_10I | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3DQ595N | 3DQ595N/Applicator_10I | Applicator_10I | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3DQ595N | 3DQ595N/Applicator_10I | Applicator_10I | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OA426 | 3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading | [5 empty] | Color (TM 208) | Main_Plans
3OA426 | 3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Free of hard wrinkles, formation uniform, large dot embossing pattern. (Done) | [5 empty] | See Analyte for Method # | Main_Plans
3OA426 | 3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 200 |  |  | FDA,0 |  |  | See Analyte for Method # | Main_Plans
3OA426 | 3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 52.01 |  | GSM | FDA,0 | 47 | 57 | See Analyte for Method # | Main_Plans
3OA426 | 3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (mm) (CofA) | Caliper | 0.79 |  | mm | FDA,2 | 0.74 | 0.83 | See Analyte for Method # | Main_Plans
3OA426 | 3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | See Analyte for Method # | Main_Plans
3OA426 | 3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | May have slight odor when wet (Done) | [5 empty] | See Analyte for Method # | Main_Plans
3OA426 | 3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter | 45 |  | in | FDA,0 | 43 | 48 | See Analyte for Method # | Main_Plans
3OA426 | 3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Width (Inches) (Measurement) | Roll Width | 8 |  | in | FDA,3 | 7.875 | 8.125 | See Analyte for Method # | Main_Plans
3OA426 | 3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Wet CD (g/in) (TM 103) | Wet CD | 1337.35 |  | g/in | FDA,0 | 1200 | 2100 | See Analyte for Method # | Main_Plans
3OA426 | 3OA426/Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | Applicator_NY_Rolls Spunlaced Nonwoven – 8” – 52 gsm | NY-Analytical | 884 | Applicator_Inspection | Wet MD (g/in) (TM 103) | Wet MD | 5490.37 |  | g/in | FDA,0 | 3800 | 6270 | See Analyte for Method # | Main_Plans
3OA580 | 3OA580/Applicator_38A | Applicator_38A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Applicator Measurements | Main_Plans
3OA580 | 3OA580/Applicator_38A | Applicator_38A | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 45 |  | in | FDA,0 | 43 | 47 | Applicator Measurements | Main_Plans
3OA580 | 3OA580/Applicator_38A | Applicator_38A | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 11.5 |  | in | FDA,3 | 11.375 | 11.625 | Applicator Measurements | Main_Plans
3OA580 | 3OA580/Applicator_38A | Applicator_38A | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 44.35 |  | GSM | FDA,0 | 40 | 50 | TM 101 | Main_Plans
3OA580 | 3OA580/Applicator_38A | Applicator_38A | NY-Analytical | 757 | Caliper | Caliper_at_Pt_of_Mfr_(micron) | Caliper | 500 |  | um | FDA,0 | 416 | 534 | TM 102 | Main_Plans
3OA580 | 3OA580/Applicator_38A | Applicator_38A | NY-Analytical | 758 | Tensile Strength | Dry_CD_CoA_(g/in) | Dry CD (CoA) | 1326 |  | g/in | FDA,0 | 610 |  | TM 103 | Main_Plans
3OA580 | 3OA580/Applicator_38A | Applicator_38A | NY-Analytical | 758 | Tensile Strength | Dry_MD_CoA_(g/in) | Dry MD (CoA) | 3086 |  | g/in | FDA,0 | 1825 |  | TM 103 | Main_Plans
3OA580 | 3OA580/Applicator_38A | Applicator_38A | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OA580 | 3OA580/Applicator_38A | Applicator_38A | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OA580 | 3OA580/Applicator_38A | Applicator_38A | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OA592 | 3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Applicator Measurements | Main_Plans
3OA592 | 3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 45 |  | in | FDA,0 | 45 | 48 | Applicator Measurements | Main_Plans
3OA592 | 3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 6 |  | in | FDA,3 | 6 | 6.125 | Applicator Measurements | Main_Plans
3OA592 | 3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 69.3 |  | GSM | FDA,0 | 65 | 75 | TM 101 | Main_Plans
3OA592 | 3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper | 0.6 |  | mm | FDA,2 | 0.54 | 0.66 | TM 102 | Main_Plans
3OA592 | 3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 387.69 |  | g/in | FDA,0 | 300 | 500 | TM 103 | Main_Plans
3OA592 | 3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 793.49 |  | g/in | FDA,0 | 610 | 1040 | TM 103 | Main_Plans
3OA592 | 3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OA592 | 3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OA592 | 3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OA592 | 3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 2.76 |  | g | FDA,2 |  |  | TM 610 | Main_Plans
3OA592 | 3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer | [5 empty] | TM 610 | Main_Plans
3OA592 | 3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer | [5 empty] | TM 610 | Main_Plans
3OA592 | 3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OA592 | 3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
3OA592 | 3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OA592 | 3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OA592 | 3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Main_Plans
3OA592 | 3OA592/Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | Applicator_NY_Nonwoven Rolls Dispersible Unembossed - 6" - 70gsm | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OA911 | 3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Applicator Measurements | Main_Plans
3OA911 | 3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 7.9 |  | in | FDA,3 | 7.9 | 8.025 | Applicator Measurements | Main_Plans
3OA911 | 3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 54.9 |  | GSM | FDA,0 | 50 | 60 | TM 101 | Main_Plans
3OA911 | 3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_CoA_(mm) | Caliper | 0.72 |  | mm | FDA,2 | 0.63 | 0.87 | TM 102 | Main_Plans
3OA911 | 3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 758 | Tensile Strength | Dry_CD_(g/in) | Dry CD | 1382.58 |  | g/in | FDA,0 | 1036 |  | TM 103 | Main_Plans
3OA911 | 3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 758 | Tensile Strength | Dry_MD_(g/in) | Dry CD | 6608.28 |  | g/in | FDA,0 | 4404 |  | TM 103 | Main_Plans
3OA911 | 3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OA911 | 3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OA911 | 3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OA911 | 3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 1.05 |  | g | FDA,2 |  |  | TM 610 | Main_Plans
3OA911 | 3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | [6 empty] | TM 610 | Main_Plans
3OA911 | 3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | [6 empty] | TM 610 | Main_Plans
3OA911 | 3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OA911 | 3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
3OA911 | 3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OA911 | 3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OA911 | 3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Main_Plans
3OA911 | 3OA911/Applicator_NY_55gsm Big Dot Embossed Nonwoven | Applicator_NY_55gsm Big Dot Embossed Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OA972 | 3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Applicator Measurements | Main_Plans
3OA972 | 3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 44 |  | in | FDA,0 | 43 | 47 | Applicator Measurements | Main_Plans
3OA972 | 3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 9 |  | in | FDA,3 | 8.875 | 9.125 | Applicator Measurements | Main_Plans
3OA972 | 3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 40.9 |  | GSM | FDA,1 | 35.2 | 44.8 | TM 101 | Main_Plans
3OA972 | 3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 757 | Caliper | Caliper_CoA_(uM) | Caliper | 437.8 |  | micron | FDA,0 | 340 | 540 | TM 102 | Main_Plans
3OA972 | 3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 758 | Tensile Strength | Wet_CD_CoA_(g/in) | Wet CD CoA | 1145 |  | g/in | FDA,0 | 400 | 1800 | TM 103 | Main_Plans
3OA972 | 3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 758 | Tensile Strength | Wet_MD_CoA_(g/in) | Wet MD CoA | 2050 |  | g/in | FDA,0 | 1000 | 3400 | TM 103 | Main_Plans
3OA972 | 3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OA972 | 3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OA972 | 3OA972/Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | Applicator_NY_40gsm Hydroentangled Nonwoven 9" with Diamond Emboss | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OC585 | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading | [5 empty] | Color (TM 208) | Main_Plans
3OC585 | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | May have residual sulfur odor when wet | [5 empty] | Odor (TM 208) | Main_Plans
3OC585 | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Free of hard wrinkles, and formation uniform, with micro crepe texture (Done) | [5 empty] | See Analyte for Method # | Main_Plans
3OC585 | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | See Analyte for Method # | Main_Plans
3OC585 | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 54.04 |  | GSM | FDA,0 | 49.5 | 60.5 | See Analyte for Method # | Main_Plans
3OC585 | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (mm) (TM 102) | Caliper | 0.55 |  | mm | FDA,2 | 0.52 | 0.74 | See Analyte for Method # | Main_Plans
3OC585 | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | See Analyte for Method # | Main_Plans
3OC585 | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter | 32 |  | in | FDA,0 | 28 | 36 | See Analyte for Method # | Main_Plans
3OC585 | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Roll Width (Inches) (Measurement) | Roll Width | 6 |  | in | FDA,0 | 5.875 | 6.125 | See Analyte for Method # | Main_Plans
3OC585 | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Wet CD (g/in) (TM 103) | Wet CD | 1004.84 |  | g/in | FDA,0 | 992 | 2208 | See Analyte for Method # | Main_Plans
3OC585 | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Analytical | 884 | Applicator_Inspection | Wet MD (g/in) (TM 103) | Wet MD | 4435.12 |  | g/in | FDA,0 | 4392 | 7808 | See Analyte for Method # | Main_Plans
3OC585 | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 1.5 |  | g | FDA,2 |  |  | TM 610 | Main_Plans
3OC585 | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA | [5 empty] | TM 610 | Main_Plans
3OC585 | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA | [5 empty] | TM 610 | Main_Plans
3OC585 | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OC585 | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
3OC585 | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OC585 | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OC585 | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Main_Plans
3OC585 | 3OC585/Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | Applicator_NY_55 gsm Unembossed Spunlaced Nonwoven – 6” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OD186 | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading | [5 empty] | Color (TM 208) | Main_Plans
3OD186 | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | May have residual sulfur odor when wet | [5 empty] | Odor (TM 208) | Main_Plans
3OD186 | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Meets specifications (Done) | [5 empty] | See Analyte for Method # | Main_Plans
3OD186 | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 200 |  |  | FDA,0 |  |  | See Analyte for Method # | Main_Plans
3OD186 | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (lbs/ream) (TM 101) | Basis Weight | 23.81 |  | #/Ream | FDA,1 | 21.5 | 25 | See Analyte for Method # | Main_Plans
3OD186 | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Caliper at point of manufacture (mils) (TM 102) | Caliper | 4.9 |  | mil | FDA,1 | 4.2 | 5.6 | See Analyte for Method # | Main_Plans
3OD186 | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | See Analyte for Method # | Main_Plans
3OD186 | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter | 32 |  | in | FDA,0 | 31 | 33 | See Analyte for Method # | Main_Plans
3OD186 | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Roll Width (Inches) (Measurement) | Roll Width | 7.75 |  | in | FDA,4 | 7.6875 | 7.8125 | See Analyte for Method # | Main_Plans
3OD186 | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Wet CD (g/in) (TM 103) | Wet CD | 1261.89 |  | g/in | FDA,0 | 1000 | 2500 | See Analyte for Method # | Main_Plans
3OD186 | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Analytical | 884 | Applicator_Inspection | Wet MD (g/in) (TM 103) | Wet MD | 2429.95 |  | g/in | FDA,0 | 1500 | 3850 | See Analyte for Method # | Main_Plans
3OD186 | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 1.1 |  | g | FDA,2 |  |  | TM 610 | Main_Plans
3OD186 | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA | [5 empty] | TM 610 | Main_Plans
3OD186 | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer | [5 empty] | TM 610 | Main_Plans
3OD186 | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OD186 | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
3OD186 | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OD186 | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OD186 | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Main_Plans
3OD186 | 3OD186/Applicator_NY_Paper Crepe 23 lb., 7 3/4" | Applicator_NY_Paper Crepe 23 lb., 7 3/4" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OD438 | 3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading | [5 empty] | Color (TM 208) | Main_Plans
3OD438 | 3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | May have residual sulfur odor when wet | [5 empty] | Odor (TM 208) | Main_Plans
3OD438 | 3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Smooth free of hard wrinkles, formation uniform matches embossing pattern (Done) | [5 empty] | See Analyte for Method # | Main_Plans
3OD438 | 3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 125 |  |  | FDA,0 |  |  | See Analyte for Method # | Main_Plans
3OD438 | 3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 34.66 |  | GSM | FDA,1 | 31.5 | 38.5 | See Analyte for Method # | Main_Plans
3OD438 | 3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (inches) (TM 102) | Caliper | 0.0077 |  | in | FDA,4 | 0.006 | 0.0087 | See Analyte for Method # | Main_Plans
3OD438 | 3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | See Analyte for Method # | Main_Plans
3OD438 | 3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Dry CD (g/in) (TM103) | Dry CD | 1868.87 |  | g/in | FDA,0 | 900 | 2200 | See Analyte for Method # | Main_Plans
3OD438 | 3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Dry MD (g/in) (TM 103) | Dry CD | 4655.69 |  | g/in | FDA,0 | 2000 | 5000 | See Analyte for Method # | Main_Plans
3OD438 | 3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter | 32 |  | in | FDA,0 | 30 | 34 | See Analyte for Method # | Main_Plans
3OD438 | 3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Roll Width (Inches) (Measurement) | Roll Width | 7.875 |  | in | FDA,2 | 7.75 | 8 | See Analyte for Method # | Main_Plans
3OD438 | 3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Wet CD (g/in) (TM 103) | Wet CD | 789.96 |  | g/in | FDA,0 | 250 | 1250 | See Analyte for Method # | Main_Plans
3OD438 | 3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Analytical | 884 | Applicator_Inspection | Wet MD (g/in) (TM 103) | Wet MD | 2141.89 |  | g/in | FDA,0 | 800 | 2600 | See Analyte for Method # | Main_Plans
3OD438 | 3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 0.97 |  | g | FDA,2 |  |  | TM 610 | Main_Plans
3OD438 | 3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA | [5 empty] | TM 610 | Main_Plans
3OD438 | 3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA | [5 empty] | TM 610 | Main_Plans
3OD438 | 3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OD438 | 3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
3OD438 | 3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OD438 | 3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OD438 | 3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Main_Plans
3OD438 | 3OD438/Applicator_NY_35 gsm Unembossed Nonwoven | Applicator_NY_35 gsm Unembossed Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OH143 | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading | [5 empty] | Color (TM 208) | Main_Plans
3OH143 | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | May have residual sulfur odor when wet | [5 empty] | Odor (TM 208) | Main_Plans
3OH143 | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Free of hard wrinkles and formation uniform (Done) | [5 empty] | See Analyte for Method # | Main_Plans
3OH143 | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 315 |  |  | FDA,0 |  |  | See Analyte for Method # | Main_Plans
3OH143 | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 46.72 |  | GSM | FDA,1 | 40.5 | 49.5 | See Analyte for Method # | Main_Plans
3OH143 | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (mm) (TM 102) | Caliper | 0.64 |  | mm | FDA,2 | 0.55 | 0.67 | See Analyte for Method # | Main_Plans
3OH143 | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | See Analyte for Method # | Main_Plans
3OH143 | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter | 32 |  | in | FDA,0 | 31 | 33 | See Analyte for Method # | Main_Plans
3OH143 | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Roll Width (Inches) (Measurement) | Roll Width | 7 |  | in | FDA,0 | 6.875 | 7.125 | See Analyte for Method # | Main_Plans
3OH143 | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Wet CD (g/in) (TM 103) | Wet CD | 2598 |  | g/in | FDA,0 | 1131 | 1593 | See Analyte for Method # | Main_Plans
3OH143 | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Analytical | 884 | Applicator_Inspection | Wet MD (g/in) (TM 103) | Wet MD | 3711 |  | g/in | FDA,0 | 3906 | 4934 | See Analyte for Method # | Main_Plans
3OH143 | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 1.75 |  | g | FDA,2 |  |  | TM 610 | Main_Plans
3OH143 | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA | [5 empty] | TM 610 | Main_Plans
3OH143 | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA | [5 empty] | TM 610 | Main_Plans
3OH143 | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OH143 | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
3OH143 | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OH143 | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OH143 | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Main_Plans
3OH143 | 3OH143/Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | Applicator_NY_Embossed Spunlaced Nonwoven 45 gsm – 7” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OH581 | 3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading | [5 empty] | Color (TM 208) | Main_Plans
3OH581 | 3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Smooth free of hard wrinkles, formation uniform matches embossing pattern (Done) | [5 empty] | See Analyte for Method # | Main_Plans
3OH581 | 3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 200 |  |  | FDA,0 |  |  | See Analyte for Method # | Main_Plans
3OH581 | 3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 52.74 |  | GSM | FDA,1 | 45 | 55 | See Analyte for Method # | Main_Plans
3OH581 | 3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Caliper 1-ply (mm) (TM 102) | Caliper | 0.55 |  | mm | FDA,2 | 0.51 | 0.63 | See Analyte for Method # | Main_Plans
3OH581 | 3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | See Analyte for Method # | Main_Plans
3OH581 | 3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Dry CD (g/in) (TM103) | Dry CD | 2389.39 |  | g/in | FDA,0 | 1801 | 2821 | See Analyte for Method # | Main_Plans
3OH581 | 3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Dry MD (g/in) (TM 103) | Dry CD | 8511.74 |  | g/in | FDA,0 | 6675 | 9447 | See Analyte for Method # | Main_Plans
3OH581 | 3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | No objectionable odor detectable wet or dry (Done) | [5 empty] | See Analyte for Method # | Main_Plans
3OH581 | 3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter | 30 |  | in | FDA,0 | 30 | 32 | See Analyte for Method # | Main_Plans
3OH581 | 3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Analytical | 884 | Applicator_Inspection | Roll Width (Inches) (Measurement) | Roll Width | 7 |  | in | FDA,3 | 6.875 | 7.125 | See Analyte for Method # | Main_Plans
3OH581 | 3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 1.07 |  | g | FDA,2 |  |  | TM 610 | Main_Plans
3OH581 | 3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer | [5 empty] | TM 610 | Main_Plans
3OH581 | 3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA | [5 empty] | TM 610 | Main_Plans
3OH581 | 3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OH581 | 3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
3OH581 | 3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OH581 | 3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OH581 | 3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Main_Plans
3OH581 | 3OH581/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester – 7” | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OH592 | 3OH592/Applicator_36A | Applicator_36A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Applicator Measurements | Main_Plans
3OH592 | 3OH592/Applicator_36A | Applicator_36A | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 32 |  | in | FDA,0 | 30 | 34 | Applicator Measurements | Main_Plans
3OH592 | 3OH592/Applicator_36A | Applicator_36A | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 7.875 |  | in | FDA,3 | 7.75 | 8 | Applicator Measurements | Main_Plans
3OH592 | 3OH592/Applicator_36A | Applicator_36A | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 39.181 |  | GSM | FDA,0 | 35 | 45 | TM 101 | Main_Plans
3OH592 | 3OH592/Applicator_36A | Applicator_36A | NY-Analytical | 757 | Caliper | Caliper_CoA_(uM) | Caliper | 424 |  | um | FDA,0 | 340 | 540 | TM 102 | Main_Plans
3OH592 | 3OH592/Applicator_36A | Applicator_36A | NY-Analytical | 758 | Tensile Strength | Wet_CD_CoA_(g/in) | Wet CD CoA | 1235 |  | g/in | FDA,0 | 400 |  | TM 103 | Main_Plans
3OH592 | 3OH592/Applicator_36A | Applicator_36A | NY-Analytical | 758 | Tensile Strength | Wet_MD_CoA_(g/in) | Wet MD CoA | 2903 |  | g/in | FDA,0 | 1000 |  | TM 103 | Main_Plans
3OH592 | 3OH592/Applicator_36A | Applicator_36A | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OH592 | 3OH592/Applicator_36A | Applicator_36A | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OH592 | 3OH592/Applicator_36A | Applicator_36A | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OJ213 | 3OJ213/Applicator_08A | Applicator_08A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [6 empty] | Applicator Measurements | Main_Plans
3OJ213 | 3OJ213/Applicator_08A | Applicator_08A | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width |  |  | in | FDA,3 | 11.875 | 12.125 | Applicator Measurements | Main_Plans
3OJ213 | 3OJ213/Applicator_08A | Applicator_08A | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 41.76 |  | GSM | FDA,0 | 36 | 44 | TM 101 | Main_Plans
3OJ213 | 3OJ213/Applicator_08A | Applicator_08A | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_CoA_(mm) | Caliper | 0.54 |  | mm | FDA,2 | 0.48 | 0.61 | TM 102 | Main_Plans
3OJ213 | 3OJ213/Applicator_08A | Applicator_08A | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 1096.49 |  | g/in | FDA,0 | 1030 | 1546 | TM 103 | Main_Plans
3OJ213 | 3OJ213/Applicator_08A | Applicator_08A | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 4699.94 |  | g/in | FDA,0 | 3343 | 4350 | TM 103 | Main_Plans
3OJ213 | 3OJ213/Applicator_08A | Applicator_08A | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OJ213 | 3OJ213/Applicator_08A | Applicator_08A | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OJ213 | 3OJ213/Applicator_08A | Applicator_08A | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OJ213 | 3OJ213/Applicator_08A | Applicator_08A | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight |  |  | g | FDA,2 |  |  | TM 610 | Main_Plans
3OJ213 | 3OJ213/Applicator_08A | Applicator_08A | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | [6 empty] | TM 610 | Main_Plans
3OJ213 | 3OJ213/Applicator_08A | Applicator_08A | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | [6 empty] | TM 610 | Main_Plans
3OJ213 | 3OJ213/Applicator_08A | Applicator_08A | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) |  |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OJ213 | 3OJ213/Applicator_08A | Applicator_08A | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) |  |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
3OJ213 | 3OJ213/Applicator_08A | Applicator_08A | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) |  |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OJ213 | 3OJ213/Applicator_08A | Applicator_08A | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) |  |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OJ213 | 3OJ213/Applicator_08A | Applicator_08A | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) |  |  | CFU/g | FDA,1 |  | 15 | TM 610 | Main_Plans
3OJ213 | 3OJ213/Applicator_08A | Applicator_08A | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OM304 | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Meets specifications (Done) | [5 empty] | See Analyte for Method # | Main_Plans
3OM304 | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 315 |  |  | FDA,0 |  |  | See Analyte for Method # | Main_Plans
3OM304 | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 47.36 |  | GSM | FDA,0 | 45 | 55 | See Analyte for Method # | Main_Plans
3OM304 | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Caliper 1-ply (mm) (TM 102) | Caliper | 0.46 |  | mm | FDA,2 | 0.51 | 0.63 | See Analyte for Method # | Main_Plans
3OM304 | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | See Analyte for Method # | Main_Plans
3OM304 | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | Meets specifications (Done) | [5 empty] | See Analyte for Method # | Main_Plans
3OM304 | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Dry CD (g/in) (TM103) | Dry CD | 2142.67 |  | g/in | FDA,0 | 1801 | 2821 | See Analyte for Method # | Main_Plans
3OM304 | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Dry MD (g/in) (TM 103) | Dry CD | 7892 |  | g/in | FDA,0 | 6675 | 9447 | See Analyte for Method # | Main_Plans
3OM304 | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | Meets specifications (Done) | [5 empty] | See Analyte for Method # | Main_Plans
3OM304 | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (COA) | Roll Diameter (COA) | 43 |  |  | FDA,0 | 41 | 45 | See Analyte for Method # | Main_Plans
3OM304 | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Analytical | 884 | Applicator_Inspection | Roll width (COA) | Roll width (COA) | 9 |  | in | FDA,2 | 8.75 | 9.25 | See Analyte for Method # | Main_Plans
3OM304 | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 0.89 |  | g | FDA,2 |  |  | TM 610 | Main_Plans
3OM304 | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA | [5 empty] | TM 610 | Main_Plans
3OM304 | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer | [5 empty] | TM 610 | Main_Plans
3OM304 | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OM304 | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
3OM304 | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OM304 | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OM304 | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Main_Plans
3OM304 | 3OM304/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-9" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OP065 | 3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading | [5 empty] | Color (TM 208) | Main_Plans
3OP065 | 3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Free of hard wrinkles. Formation uniform (Done) | [5 empty] | See Analyte for Method # | Main_Plans
3OP065 | 3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 125 |  |  | FDA,0 |  |  | See Analyte for Method # | Main_Plans
3OP065 | 3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 57.26 |  | GSM | FDA,1 | 54 | 66 | See Analyte for Method # | Main_Plans
3OP065 | 3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (mm) (CofA) | Caliper | 0.8 |  | mm | FDA,2 | 0.67 | 0.81 | See Analyte for Method # | Main_Plans
3OP065 | 3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | See Analyte for Method # | Main_Plans
3OP065 | 3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | Meets specifications (Done) | [5 empty] | See Analyte for Method # | Main_Plans
3OP065 | 3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter | 46 |  | in | FDA,0 | 45 | 47 | See Analyte for Method # | Main_Plans
3OP065 | 3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Width (Inches) (Measurement) | Roll Width | 30 |  | in | FDA,2 | 29.75 | 30.25 | See Analyte for Method # | Main_Plans
3OP065 | 3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Wet CD (g/in) (TM 103) | Wet CD | 1584 |  | g/in | FDA,0 | 1300 | 2000 | See Analyte for Method # | Main_Plans
3OP065 | 3OP065/Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | Applicator_NY_Rolls spunlaced Nonwoven Diamond - 60gsm | NY-Analytical | 884 | Applicator_Inspection | Wet MD (g/in) (TM 103) | Wet MD | 5195.31 |  | g/in | FDA,0 | 4000 | 5500 | See Analyte for Method # | Main_Plans
3OP081 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Applicator Measurements | Main_Plans
3OP081 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter |  |  | in | FDA,0 | 45 | 47 | Applicator Measurements | Main_Plans
3OP081 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width |  |  | in | FDA,2 | 29.75 | 30.25 | Applicator Measurements | Main_Plans
3OP081 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 57.7 |  | GSM | FDA,1 | 49.5 | 60.5 | TM 101 | Main_Plans
3OP081 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper | 0.51 |  | mm | FDA,2 | 0.4 | 0.6 | TM 102 | Main_Plans
3OP081 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 758 | Tensile Strength | Elongation % Wet CD | Elongation % Wet CD | 103.07 |  | % | FDA,0 | 80 | 130 | TM 103 | Main_Plans
3OP081 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 758 | Tensile Strength | Elongation % Wet MD | Elongation % Wet MD | 24.68 |  | % | FDA,0 | 20 | 35 | TM 103 | Main_Plans
3OP081 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 893.28 |  | g/in | FDA,0 | 650 | 1300 | TM 103 | Main_Plans
3OP081 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 2898.19 |  | g/in | FDA,0 | 2400 | 3400 | TM 103 | Main_Plans
3OP081 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | [6 empty] | TM 208 | Main_Plans
3OP081 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 748 | Visual Inspection | Color | Color | [6 empty] | TM 208 | Main_Plans
3OP081 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Analytical | 748 | Visual Inspection | Odor | Odor | [6 empty] | TM 208 | Main_Plans
3OP081 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 0.84 |  | g | FDA,2 |  |  | TM 610 | Main_Plans
3OP081 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer | [5 empty] | TM 610 | Main_Plans
3OP081 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer | [5 empty] | TM 610 | Main_Plans
3OP081 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OP081 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
3OP081 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OP081 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OP081 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Main_Plans
3OP081 | 3OP081/Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | Applicator_NY_Rolls Flat Spunlaced Nonwoven-55gsm-30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OP093PD | 3OP093PD/Applicator_NY_Die Cut Pads | Applicator_NY_Die Cut Pads | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Applicator Measurements | Main_Plans
3OP093PD | 3OP093PD/Applicator_NY_Die Cut Pads | Applicator_NY_Die Cut Pads | NY-Analytical | 759 | Applicator Measurements | Pad Diameter | Pad Diameter | 2.2188 |  | in | FDA,2 | 2.19 | 2.31 | Applicator Measurements | Main_Plans
3OP093PD | 3OP093PD/Applicator_NY_Die Cut Pads | Applicator_NY_Die Cut Pads | NY-Analytical | 759 | Applicator Measurements | Stack Count | Stack Count | 90 |  | Units | FDA,0 | 90 | 90 | Applicator Measurements | Main_Plans
3OP093PD | 3OP093PD/Applicator_NY_Die Cut Pads | Applicator_NY_Die Cut Pads | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OP093PD | 3OP093PD/Applicator_NY_Die Cut Pads | Applicator_NY_Die Cut Pads | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OP093PD | 3OP093PD/Applicator_NY_Die Cut Pads | Applicator_NY_Die Cut Pads | NY-Analytical | 748 | Visual Inspection | Odor | Odor | No objectionable odor Wet or Dry [Done] | [5 empty] | TM 208 | Main_Plans
3OP093PD | 3OP093PD/Applicator_NY_Die Cut Pads | Applicator_NY_Die Cut Pads | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 2.07 |  | g | FDA,2 |  |  | TM 610 | Main_Plans
3OP093PD | 3OP093PD/Applicator_NY_Die Cut Pads | Applicator_NY_Die Cut Pads | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA | [5 empty] | TM 610 | Main_Plans
3OP093PD | 3OP093PD/Applicator_NY_Die Cut Pads | Applicator_NY_Die Cut Pads | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer | [5 empty] | TM 610 | Main_Plans
3OP093PD | 3OP093PD/Applicator_NY_Die Cut Pads | Applicator_NY_Die Cut Pads | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 1 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OP093PD | 3OP093PD/Applicator_NY_Die Cut Pads | Applicator_NY_Die Cut Pads | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0.042735 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
3OP093PD | 3OP093PD/Applicator_NY_Die Cut Pads | Applicator_NY_Die Cut Pads | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OP093PD | 3OP093PD/Applicator_NY_Die Cut Pads | Applicator_NY_Die Cut Pads | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OP093PD | 3OP093PD/Applicator_NY_Die Cut Pads | Applicator_NY_Die Cut Pads | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Main_Plans
3OP093PD | 3OP093PD/Applicator_NY_Die Cut Pads | Applicator_NY_Die Cut Pads | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OP093PD | 3OP093PD/Applicator_NY_Die Cut Pads | Applicator_NY_Die Cut Pads | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Positive Cocci (OOS) | [5 empty] | WI QAS 076 | Main_Plans
3OP093PD | 3OP093PD/Applicator_NY_Die Cut Pads | Applicator_NY_Die Cut Pads | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Non-Subculturable Growth (Done) | [5 empty] | WI QAS 076 | Main_Plans
3OP134 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Applicator Measurements | Main_Plans
3OP134 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 40 |  | in | FDA,0 | 39 | 40 | Applicator Measurements | Main_Plans
3OP134 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 30 |  | in | FDA,3 | 29.875 | 30.125 | Applicator Measurements | Main_Plans
3OP134 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 34.01 |  | GSM | FDA,1 | 31.5 | 38.5 | TM 101 | Main_Plans
3OP134 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(inch) | Caliper | 0.008 |  | in | FDA,4 | 0.006 | 0.009 | TM 102 | Main_Plans
3OP134 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_CoA_(mm) | Caliper | 0.0079 |  | in |  |  |  | TM 102 | Main_Plans
3OP134 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 812.99 |  | g/in | FDA,0 | 250 | 1250 | TM 103 | Main_Plans
3OP134 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 1821.41 |  | g/in | FDA,0 | 800 | 2600 | TM 103 | Main_Plans
3OP134 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OP134 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 748 | Visual Inspection | Color | Color | White with no off colors or shading [Done] | [5 empty] | TM 208 | Main_Plans
3OP134 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | Applicator_NY_35gsm Stora Flat - 30" | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OP134 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 0.97 |  | g | FDA,2 |  |  | TM 610 | Main_Plans
3OP134 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA | [5 empty] | TM 610 | Main_Plans
3OP134 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer | [5 empty] | TM 610 | Main_Plans
3OP134 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OP134 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
3OP134 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OP134 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OP134 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Main_Plans
3OP134 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OP134 | 3OP134/Applicator_NY_35gsm Stora Flat - 30" | Applicator_NY_35gsm Stora Flat - 30" | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | [6 empty] | WI QAS 076 | Main_Plans
3OP138 | 3OP138/Applicator_36D | Applicator_36D | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [6 empty] | Applicator Measurements | Main_Plans
3OP138 | 3OP138/Applicator_36D | Applicator_36D | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter |  |  | in | FDA,0 | 43 | 47 | Applicator Measurements | Main_Plans
3OP138 | 3OP138/Applicator_36D | Applicator_36D | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 30 |  | in | FDA,3 | 29.875 | 30.125 | Applicator Measurements | Main_Plans
3OP138 | 3OP138/Applicator_36D | Applicator_36D | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 40.26 |  | GSM | FDA,0 | 35 | 45 | TM 101 | Main_Plans
3OP138 | 3OP138/Applicator_36D | Applicator_36D | NY-Analytical | 757 | Caliper | Caliper_CoA_(uM) | Caliper |  |  | um | FDA,0 | 340 | 540 | TM 102 | Main_Plans
3OP138 | 3OP138/Applicator_36D | Applicator_36D | NY-Analytical | 758 | Tensile Strength | Wet_CD_CoA_(g/in) | Wet CD CoA | 1228 |  | g/in | FDA,0 | 400 |  | TM 103 | Main_Plans
3OP138 | 3OP138/Applicator_36D | Applicator_36D | NY-Analytical | 758 | Tensile Strength | Wet_MD_CoA_(g/in) | Wet MD CoA | 2545 |  | g/in | FDA,0 | 1000 |  | TM 103 | Main_Plans
3OP138 | 3OP138/Applicator_36D | Applicator_36D | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OP138 | 3OP138/Applicator_36D | Applicator_36D | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OP138 | 3OP138/Applicator_36D | Applicator_36D | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OP242 | 3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 37.5 |  | in | FDA,3 | 37.375 | 37.625 | Applicator Measurements | Main_Plans
3OP242 | 3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Basis Weight | Main_Plans
3OP242 | 3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 35.95 |  | GSM | FDA,1 | 31.5 | 38.5 | TM 101 | Main_Plans
3OP242 | 3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper | 0.34 |  | mm | FDA,2 | 0.27 | 0.41 | TM 102 | Main_Plans
3OP242 | 3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 758 | Tensile Strength | Wet_CD_CoA_(g/in) | Wet CD CoA | 520.88 |  | g/in | FDA,0 | 300 |  | TM 103 | Main_Plans
3OP242 | 3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 758 | Tensile Strength | Wet_MD_CoA_(g/in) | Wet MD CoA | 3556.57 |  | g/in | FDA,0 | 2000 | 4400 | TM 103 | Main_Plans
3OP242 | 3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OP242 | 3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 748 | Visual Inspection | Color | Color | White with no off colors or shading [Done] | [5 empty] | TM 208 | Main_Plans
3OP242 | 3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OP242 | 3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 0.95 |  | g | FDA,2 |  |  | TM 610 | Main_Plans
3OP242 | 3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA | [5 empty] | TM 610 | Main_Plans
3OP242 | 3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer | [5 empty] | TM 610 | Main_Plans
3OP242 | 3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 1.5 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OP242 | 3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 1.5789 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
3OP242 | 3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 1 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OP242 | 3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OP242 | 3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Main_Plans
3OP242 | 3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OP242 | 3OP242/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 37.5" | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Non-Subculturable Growth (Done) | [5 empty] | WI QAS 076 | Main_Plans
3OP257 | 3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (COA) | Roll Diameter (COA) | 43 |  | in | FDA,0 | 41 | 45 | Applicator Measurements | Main_Plans
3OP257 | 3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 759 | Applicator Measurements | Roll width (COA) | Roll width (COA) | 37.5 |  | in | FDA,2 | 37.25 | 37.75 | Applicator Measurements | Main_Plans
3OP257 | 3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Basis Weight | Main_Plans
3OP257 | 3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 51.02 |  | GSM | FDA,0 | 45 | 55 | TM 101 | Main_Plans
3OP257 | 3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper | 0.52 |  | mm | FDA,2 | 0.51 | 0.63 | TM 102 | Main_Plans
3OP257 | 3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 758 | Tensile Strength | Dry_CD_(g/in) | Dry CD | 2193.6 |  | g/in | FDA,0 | 1801 | 2821 | TM 103 | Main_Plans
3OP257 | 3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 758 | Tensile Strength | Dry_MD_(g/in) | Dry CD | 7519.45 |  | g/in | FDA,0 | 6675 | 9447 | TM 103 | Main_Plans
3OP257 | 3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OP257 | 3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OP257 | 3OP257/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-37.5" | NY-Analytical | 748 | Visual Inspection | Odor | Odor | No objectionable odor Wet or Dry [Done] | [5 empty] | TM 208 | Main_Plans
3OP920 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading | [5 empty] | Color (TM 208) | Main_Plans
3OP920 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 759 | Applicator Measurements | Roll Width (mm) | Roll Width | 953 |  | mm | FDA,0 | 950 | 956 | Applicator Measurements | Main_Plans
3OP920 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Free of hard wrinkles, formation uniform/quilted big dot emboss pattern (Done) | [5 empty] | See Analyte for Method # | Main_Plans
3OP920 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 315 |  |  | FDA,0 |  |  | See Analyte for Method # | Main_Plans
3OP920 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 32.08 |  | GSM | FDA,1 | 31.5 | 38.5 | See Analyte for Method # | Main_Plans
3OP920 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (mm) (TM 102) | Caliper | 0.34 |  | mm | FDA,2 | 0.26 | 0.4 | See Analyte for Method # | Main_Plans
3OP920 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | See Analyte for Method # | Main_Plans
3OP920 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Dry CD (g/in) (TM103) | Dry CD | 1254.21 |  | g/in | FDA,0 | 1100 | 1700 | See Analyte for Method # | Main_Plans
3OP920 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Dry MD (g/in) (TM 103) | Dry CD | 2070.52 |  | g/in | FDA,0 | 2000 | 3000 | See Analyte for Method # | Main_Plans
3OP920 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | May have slight odor when wet (Done) | [5 empty] | See Analyte for Method # | Main_Plans
3OP920 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter | 40 |  | in | FDA,0 | 40 |  | See Analyte for Method # | Main_Plans
3OP920 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 884 | Applicator_Inspection | Roll Width (Inches) (Measurement) | Roll Width | 37.5 |  | in | FDA,2 | 37.25 | 37.75 | See Analyte for Method # | Main_Plans
3OP920 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 34.44 |  | GSM | FDA,1 | 31.5 | 38.5 | TM 101 | Main_Plans
3OP920 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper | 0.29 |  | mm | FDA,2 | 0.26 | 0.4 | TM 102 | Main_Plans
3OP920 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 758 | Tensile Strength | Dry_CD_(g/in) | Dry CD | 1143.94 |  | g/in | FDA,0 | 1100 | 1700 | TM 103 | Main_Plans
3OP920 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 758 | Tensile Strength | Dry_MD_(g/in) | Dry CD | 2220.05 |  | g/in | FDA,0 | 2000 | 3000 | TM 103 | Main_Plans
3OP920 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OP920 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OP920 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OP920 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 0.95 |  | g | FDA,2 |  |  | TM 610 | Main_Plans
3OP920 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer | [5 empty] | TM 610 | Main_Plans
3OP920 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA | [5 empty] | TM 610 | Main_Plans
3OP920 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OP920 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
3OP920 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OP920 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 1 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OP920 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0.21053 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Main_Plans
3OP920 | 3OP920/Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | Applicator_NY_SMS Nonwovens Honeycomb Embossed- 35 gsm - 37.5" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OQ222 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [6 empty] | Applicator Measurements | Main_Plans
3OQ222 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter |  |  | in | FDA,0 | 42 | 46 | Applicator Measurements | Main_Plans
3OQ222 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 30 |  | in | FDA,3 | 29.875 | 30.125 | Applicator Measurements | Main_Plans
3OQ222 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 41.12 |  | GSM | FDA,1 | 35.5 | 44.5 | TM 101 | Main_Plans
3OQ222 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_CoA_(mm) | Caliper | 0.45 |  | mm | FDA,2 | 0.4 | 0.48 | TM 102 | Main_Plans
3OQ222 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 740.59 |  | g/in | FDA,0 | 600 | 1200 | TM 103 | Main_Plans
3OQ222 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD |  |  | g/in | FDA,0 | 2200 | 3400 | TM 103 | Main_Plans
3OQ222 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | [6 empty] | TM 208 | Main_Plans
3OQ222 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 748 | Visual Inspection | Color | Color | [6 empty] | TM 208 | Main_Plans
3OQ222 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Analytical | 748 | Visual Inspection | Odor | Odor | [6 empty] | TM 208 | Main_Plans
3OQ222 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight |  |  | g | FDA,2 |  |  | TM 610 | Main_Plans
3OQ222 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | [6 empty] | TM 610 | Main_Plans
3OQ222 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | [6 empty] | TM 610 | Main_Plans
3OQ222 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) |  |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OQ222 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) |  |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
3OQ222 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) |  |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OQ222 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) |  |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OQ222 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) |  |  | CFU/g | FDA,1 |  | 15 | TM 610 | Main_Plans
3OQ222 | 3OQ222/Applicator_NY_ 30" Embossed Spunlaced Nonwoven | Applicator_NY_ 30" Embossed Spunlaced Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OQ333 | 3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 1.32 |  | g | FDA,2 |  |  | TM 610 | Main_Plans
3OQ333 | 3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA | [5 empty] | TM 610 | Main_Plans
3OQ333 | 3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA | [5 empty] | TM 610 | Main_Plans
3OQ333 | 3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OQ333 | 3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
3OQ333 | 3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OQ333 | 3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OQ333 | 3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Main_Plans
3OQ333 | 3OQ333/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OQ444 | 3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight |  |  | g | FDA,2 |  |  | TM 610 | Main_Plans
3OQ444 | 3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | [6 empty] | TM 610 | Main_Plans
3OQ444 | 3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | [6 empty] | TM 610 | Main_Plans
3OQ444 | 3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) |  |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OQ444 | 3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) |  |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
3OQ444 | 3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) |  |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OQ444 | 3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) |  |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OQ444 | 3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) |  |  | CFU/g | FDA,1 |  | 15 | TM 610 | Main_Plans
3OQ444 | 3OQ444/Applicator_NY_36" Embossed Spunlaced Nonwoven | Applicator_NY_36" Embossed Spunlaced Nonwoven | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OQ551 | 3OQ551/Applicator_15E | Applicator_15E | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (COA) | Roll Diameter (COA) | 45 |  | in | FDA,0 | 43 | 47 | Applicator Measurements | Main_Plans
3OQ551 | 3OQ551/Applicator_15E | Applicator_15E | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 42 |  | in | FDA,3 | 41.875 | 42.125 | Applicator Measurements | Main_Plans
3OQ551 | 3OQ551/Applicator_15E | Applicator_15E | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Basis Weight | Main_Plans
3OQ551 | 3OQ551/Applicator_15E | Applicator_15E | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight |  |  | GSM | FDA,0 | 35 | 45 | TM 101 | Main_Plans
3OQ551 | 3OQ551/Applicator_15E | Applicator_15E | NY-Analytical | 757 | Caliper | Caliper_at_Pt_of_Mfr_(micron) | Caliper | 431.8 |  | um | FDA,0 | 340 | 540 | TM 102 | Main_Plans
3OQ551 | 3OQ551/Applicator_15E | Applicator_15E | NY-Analytical | 758 | Tensile Strength | Wet_CD_CoA_(g/25mm) | Wet CD CoA | 1073 |  | g/25mm | FDA,0 | 400 |  | TM 103 | Main_Plans
3OQ551 | 3OQ551/Applicator_15E | Applicator_15E | NY-Analytical | 758 | Tensile Strength | Wet_MD_CoA_(g/25mm) | Wet MD CoA | 2506 |  | g/25mm | FDA,0 | 1000 |  | TM 103 | Main_Plans
3OQ551 | 3OQ551/Applicator_15E | Applicator_15E | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OQ551 | 3OQ551/Applicator_15E | Applicator_15E | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OQ551 | 3OQ551/Applicator_15E | Applicator_15E | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OQ850N | 3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading | [5 empty] | Color (TM 208) | Main_Plans
3OQ850N | 3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Smooth, free of hard wrinkles and formation uniform (Done) | [5 empty] | See Analyte for Method # | Main_Plans
3OQ850N | 3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 315 |  |  | FDA,0 |  |  | See Analyte for Method # | Main_Plans
3OQ850N | 3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 40.69 |  | GSM | FDA,1 | 38 | 42 | See Analyte for Method # | Main_Plans
3OQ850N | 3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (mm) (TM 102) | Caliper | 0.41 |  | mm |  |  |  | See Analyte for Method # | Main_Plans
3OQ850N | 3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | See Analyte for Method # | Main_Plans
3OQ850N | 3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Dry CD (g/in) (TM103) | Dry CD | 1098.41 |  | g/in | FDA,0 | 513 | 1130 | See Analyte for Method # | Main_Plans
3OQ850N | 3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Dry MD (g/in) (TM 103) | Dry CD | 2976.77 |  | g/in | FDA,0 | 2054 | 4108 | See Analyte for Method # | Main_Plans
3OQ850N | 3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | No objectionable odor detectable wet or dry (Done) | [5 empty] | See Analyte for Method # | Main_Plans
3OQ850N | 3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter | 46 |  | in | FDA,0 | 44 | 47 | See Analyte for Method # | Main_Plans
3OQ850N | 3OQ850N/Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | Applicator_NY_Thermal Bonded Nonwoven-37 1/2" | NY-Analytical | 884 | Applicator_Inspection | Roll Width (Inches) (Measurement) | Roll Width | 37.5 |  | in | FDA,3 | 37.375 | 37.625 | See Analyte for Method # | Main_Plans
3OS407 | 3OS407/Applicator_25B | Applicator_25B | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Swab Measurements | Main_Plans
3OS407 | 3OS407/Applicator_25B | Applicator_25B | NY-Analytical | 760 | Swab Measurements | Mitt_Length_(Inches) | Mitt Length | 0.99 |  | in | FDA,3 | 0.97 | 1.03 | TM 109 | Main_Plans
3OS407 | 3OS407/Applicator_25B | Applicator_25B | NY-Analytical | 760 | Swab Measurements | Mitt_Thickness_(Inches) | Mitt Thickness | 0.314 |  | in | FDA,3 | 0.265 | 0.385 | TM 109 | Main_Plans
3OS407 | 3OS407/Applicator_25B | Applicator_25B | NY-Analytical | 760 | Swab Measurements | Mitt_Width_(Inches) | Mitt Width | 0.49 |  | in | FDA,2 | 0.47 | 0.53 | TM 109 | Main_Plans
3OS407 | 3OS407/Applicator_25B | Applicator_25B | NY-Analytical | 760 | Swab Measurements | Overall_Swab_Length_(Inches) | Overall Swab Length | 4.77 |  | in | FDA,3 | 4.69 | 4.81 | TM 109 | Main_Plans
3OS407 | 3OS407/Applicator_25B | Applicator_25B | NY-Analytical | 760 | Swab Measurements | Scrub_Test | Scrub Test | Meets specifications (Done) | [5 empty] | TM 109 | Main_Plans
3OS407 | 3OS407/Applicator_25B | Applicator_25B | NY-Analytical | 760 | Swab Measurements | Tuft_Adhesion_(Lb.) | Tuft Adhesion | 2.59 |  | lb | FDA,1 | 1.4 |  | TM 109 | Main_Plans
3OS407 | 3OS407/Applicator_25B | Applicator_25B | NY-Analytical | 760 | Swab Measurements | Use_Test-Solvents | Use Test-Solvents | Meets specifications (Done) | [5 empty] | TM 109 | Main_Plans
3OS407 | 3OS407/Applicator_25B | Applicator_25B | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 1.2 |  | g | FDA,2 |  |  | TM 610 | Main_Plans
3OS407 | 3OS407/Applicator_25B | Applicator_25B | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA | [5 empty] | TM 610 | Main_Plans
3OS407 | 3OS407/Applicator_25B | Applicator_25B | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA | [5 empty] | TM 610 | Main_Plans
3OS407 | 3OS407/Applicator_25B | Applicator_25B | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OS407 | 3OS407/Applicator_25B | Applicator_25B | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
3OS407 | 3OS407/Applicator_25B | Applicator_25B | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) |  |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OS407 | 3OS407/Applicator_25B | Applicator_25B | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OS407 | 3OS407/Applicator_25B | Applicator_25B | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) |  |  | CFU/g | FDA,1 |  | 15 | TM 610 | Main_Plans
3OS407 | 3OS407/Applicator_25B | Applicator_25B | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OS407 | 3OS407/Applicator_25B | Applicator_25B | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Positive Rods (Done) | [5 empty] | WI QAS 076 | Main_Plans
3OS407 | 3OS407/Applicator_25B | Applicator_25B | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Non-Subculturable Growth (Done) | [5 empty] | WI QAS 076 | Main_Plans
3OS419 | 3OS419/Applicator_25C | Applicator_25C | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Swab Measurements | Main_Plans
3OS419 | 3OS419/Applicator_25C | Applicator_25C | NY-Analytical | 760 | Swab Measurements | Mitt_Length_(Inches) | Mitt Length | 1.22 |  | in | FDA,2 | 1.16 | 1.36 | TM 109 | Main_Plans
3OS419 | 3OS419/Applicator_25C | Applicator_25C | NY-Analytical | 760 | Swab Measurements | Mitt_Thickness_(Inches) | Mitt Thickness | 0.47 |  | in | FDA,3 | 0.41 | 0.53 | TM 109 | Main_Plans
3OS419 | 3OS419/Applicator_25C | Applicator_25C | NY-Analytical | 760 | Swab Measurements | Mitt_Width_(Inches) | Mitt Width | 1.19 |  | in | FDA,2 | 1.1 | 1.3 | TM 109 | Main_Plans
3OS419 | 3OS419/Applicator_25C | Applicator_25C | NY-Analytical | 760 | Swab Measurements | Overall_Swab_Length_(Inches) | Overall Swab Length | 6.03 |  | in | FDA,2 | 5.85 | 6.15 | TM 109 | Main_Plans
3OS419 | 3OS419/Applicator_25C | Applicator_25C | NY-Analytical | 760 | Swab Measurements | Scrub_Test | Scrub Test | Meets specifications (Done) | [5 empty] | TM 109 | Main_Plans
3OS419 | 3OS419/Applicator_25C | Applicator_25C | NY-Analytical | 760 | Swab Measurements | Tuft_Adhesion_(Lb.) | Tuft Adhesion | 5.12 |  | lb | FDA,1 | 1.4 |  | TM 109 | Main_Plans
3OS419 | 3OS419/Applicator_25C | Applicator_25C | NY-Analytical | 760 | Swab Measurements | Use_Test-Solvents | Use Test-Solvents | Meets specifications (Done) | [5 empty] | TM 109 | Main_Plans
3OS419 | 3OS419/Applicator_25C | Applicator_25C | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 1.26 |  | g | FDA,2 |  |  | TM 610 | Main_Plans
3OS419 | 3OS419/Applicator_25C | Applicator_25C | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | [6 empty] | TM 610 | Main_Plans
3OS419 | 3OS419/Applicator_25C | Applicator_25C | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | [6 empty] | TM 610 | Main_Plans
3OS419 | 3OS419/Applicator_25C | Applicator_25C | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) |  |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OS419 | 3OS419/Applicator_25C | Applicator_25C | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) |  |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
3OS419 | 3OS419/Applicator_25C | Applicator_25C | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) |  |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OS419 | 3OS419/Applicator_25C | Applicator_25C | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OS419 | 3OS419/Applicator_25C | Applicator_25C | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Main_Plans
3OS419 | 3OS419/Applicator_25C | Applicator_25C | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OS419 | 3OS419/Applicator_25C | Applicator_25C | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Positive Rods (Done) | [5 empty] | WI QAS 076 | Main_Plans
3OS484 | 3OS484/Applicator_26A | Applicator_26A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Swab Measurements | Main_Plans
3OS484 | 3OS484/Applicator_26A | Applicator_26A | NY-Analytical | 760 | Swab Measurements | Overall_Swab_Length_(Inches) | Overall Swab Length | 7.875 |  | in | FDA,2 | 7.82 | 7.94 | TM 109 | Main_Plans
3OS484 | 3OS484/Applicator_26A | Applicator_26A | NY-Analytical | 760 | Swab Measurements | Scrub_Test | Scrub Test | Meets specifications (Done) | [5 empty] | TM 109 | Main_Plans
3OS484 | 3OS484/Applicator_26A | Applicator_26A | NY-Analytical | 760 | Swab Measurements | Tuft_Diameter_(Inches) | Tuft Diameter | 0.48 |  | in | FDA,2 | 0.42 | 0.53 | TM 109 | Main_Plans
3OS484 | 3OS484/Applicator_26A | Applicator_26A | NY-Analytical | 760 | Swab Measurements | Use_Test-Solvents | Use Test-Solvents | Meets specifications (Done) | [5 empty] | TM 109 | Main_Plans
3OS484 | 3OS484/Applicator_26A | Applicator_26A | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 1.01 |  | g | FDA,2 |  |  | TM 610 | Main_Plans
3OS484 | 3OS484/Applicator_26A | Applicator_26A | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA | [5 empty] | TM 610 | Main_Plans
3OS484 | 3OS484/Applicator_26A | Applicator_26A | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA | [5 empty] | TM 610 | Main_Plans
3OS484 | 3OS484/Applicator_26A | Applicator_26A | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 1 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OS484 | 3OS484/Applicator_26A | Applicator_26A | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0.20833 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
3OS484 | 3OS484/Applicator_26A | Applicator_26A | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OS484 | 3OS484/Applicator_26A | Applicator_26A | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OS484 | 3OS484/Applicator_26A | Applicator_26A | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Main_Plans
3OS484 | 3OS484/Applicator_26A | Applicator_26A | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OS484 | 3OS484/Applicator_26A | Applicator_26A | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Negative Rods (OOS) | [5 empty] | WI QAS 076 | Main_Plans
3OU252 | 3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | [6 empty] | See Analyte for Method # | Main_Plans
3OU252 | 3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size |  |  |  | FDA,0 |  |  | See Analyte for Method # | Main_Plans
3OU252 | 3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight |  |  | GSM | FDA,1 | 64.4 | 71.2 | See Analyte for Method # | Main_Plans
3OU252 | 3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (mm) (TM 102) | Caliper |  |  | mm | FDA,2 | 0.28 | 0.42 | See Analyte for Method # | Main_Plans
3OU252 | 3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | [6 empty] | See Analyte for Method # | Main_Plans
3OU252 | 3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | [6 empty] | See Analyte for Method # | Main_Plans
3OU252 | 3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | [6 empty] | See Analyte for Method # | Main_Plans
3OU252 | 3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (COA) | Roll Diameter (COA) |  |  |  | FDA,0 | 41 | 45 | See Analyte for Method # | Main_Plans
3OU252 | 3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Roll width (COA) | Roll width (COA) |  |  | in | FDA,2 | 8.94 | 9.13 | See Analyte for Method # | Main_Plans
3OU252 | 3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Wet CD (g/in) (TM 103) | Wet CD |  |  | g/in | FDA,0 | 1500 | 2500 | See Analyte for Method # | Main_Plans
3OU252 | 3OU252/Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | Applicator_NY_Rolls Spunlaced Nonwoven-10"-67.8gsm | NY-Analytical | 884 | Applicator_Inspection | Wet MD (g/in) (TM 103) | Wet MD |  |  | g/in | FDA,0 | 4500 | 6500 | See Analyte for Method # | Main_Plans
3OU265 | 3OU265/Applicator_10H | Applicator_10H | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Applicator Measurements | Main_Plans
3OU265 | 3OU265/Applicator_10H | Applicator_10H | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 44 |  | in | FDA,0 | 41 | 45 | Applicator Measurements | Main_Plans
3OU265 | 3OU265/Applicator_10H | Applicator_10H | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 11.5 |  | in | FDA,2 | 11.25 | 11.75 | Applicator Measurements | Main_Plans
3OU265 | 3OU265/Applicator_10H | Applicator_10H | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 50.6 |  | GSM | FDA,1 | 45 | 55 | TM 101 | Main_Plans
3OU265 | 3OU265/Applicator_10H | Applicator_10H | NY-Analytical | 757 | Caliper | Caliper-1_Ply_(mm) | Caliper | 0.52 |  | mm | FDA,2 | 0.51 | 0.63 | TM 102 | Main_Plans
3OU265 | 3OU265/Applicator_10H | Applicator_10H | NY-Analytical | 758 | Tensile Strength | Dry_CD_(g/in) | Dry CD | 2302.98 |  | g/in | FDA,0 | 1801 | 2821 | TM 103 | Main_Plans
3OU265 | 3OU265/Applicator_10H | Applicator_10H | NY-Analytical | 758 | Tensile Strength | Dry_MD_(g/in) | Dry CD | 8886.78 |  | g/in | FDA,0 | 6675 | 9447 | TM 103 | Main_Plans
3OU265 | 3OU265/Applicator_10H | Applicator_10H | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OU265 | 3OU265/Applicator_10H | Applicator_10H | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OU265 | 3OU265/Applicator_10H | Applicator_10H | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OU265 | 3OU265/Applicator_10H | Applicator_10H | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 1.15 |  | g | FDA,2 |  |  | TM 610 | Main_Plans
3OU265 | 3OU265/Applicator_10H | Applicator_10H | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer | [5 empty] | TM 610 | Main_Plans
3OU265 | 3OU265/Applicator_10H | Applicator_10H | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA | [5 empty] | TM 610 | Main_Plans
3OU265 | 3OU265/Applicator_10H | Applicator_10H | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OU265 | 3OU265/Applicator_10H | Applicator_10H | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
3OU265 | 3OU265/Applicator_10H | Applicator_10H | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OU265 | 3OU265/Applicator_10H | Applicator_10H | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OU265 | 3OU265/Applicator_10H | Applicator_10H | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Main_Plans
3OU265 | 3OU265/Applicator_10H | Applicator_10H | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OU275 | 3OU275/Applicator_36C | Applicator_36C | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Applicator Measurements | Main_Plans
3OU275 | 3OU275/Applicator_36C | Applicator_36C | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 32 |  | in | FDA,0 | 30 | 34 | Applicator Measurements | Main_Plans
3OU275 | 3OU275/Applicator_36C | Applicator_36C | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 11.75 |  | in | FDA,3 | 11.625 | 11.875 | Applicator Measurements | Main_Plans
3OU275 | 3OU275/Applicator_36C | Applicator_36C | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 37.67 |  | GSM | FDA,0 | 35 | 45 | TM 101 | Main_Plans
3OU275 | 3OU275/Applicator_36C | Applicator_36C | NY-Analytical | 757 | Caliper | Caliper_CoA_(uM) | Caliper | 458 |  | um | FDA,0 | 340 | 540 | TM 102 | Main_Plans
3OU275 | 3OU275/Applicator_36C | Applicator_36C | NY-Analytical | 758 | Tensile Strength | Wet_CD_CoA_(g/in) | Wet CD CoA | 1256.91 |  | g/in | FDA,0 | 400 |  | TM 103 | Main_Plans
3OU275 | 3OU275/Applicator_36C | Applicator_36C | NY-Analytical | 758 | Tensile Strength | Wet_MD_CoA_(g/in) | Wet MD CoA | 3006.85 |  | g/in | FDA,0 | 1000 |  | TM 103 | Main_Plans
3OU275 | 3OU275/Applicator_36C | Applicator_36C | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OU275 | 3OU275/Applicator_36C | Applicator_36C | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OU275 | 3OU275/Applicator_36C | Applicator_36C | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3OX105F | 3OX105F/Applicatpr_NY_Foam 10.5ml | Applicatpr_NY_Foam 10.5ml | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Certificate of Conformance | Main_Plans
3OX105F | 3OX105F/Applicatpr_NY_Foam 10.5ml | Applicatpr_NY_Foam 10.5ml | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Bioburden_TAMC | Main_Plans
3OX105F | 3OX105F/Applicatpr_NY_Foam 10.5ml | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 0.91 |  | g | FDA,2 |  |  | TM 610 | Main_Plans
3OX105F | 3OX105F/Applicatpr_NY_Foam 10.5ml | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA | [5 empty] | TM 610 | Main_Plans
3OX105F | 3OX105F/Applicatpr_NY_Foam 10.5ml | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA | [5 empty] | TM 610 | Main_Plans
3OX105F | 3OX105F/Applicatpr_NY_Foam 10.5ml | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OX105F | 3OX105F/Applicatpr_NY_Foam 10.5ml | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
3OX105F | 3OX105F/Applicatpr_NY_Foam 10.5ml | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OX105F | 3OX105F/Applicatpr_NY_Foam 10.5ml | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 1 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OX105F | 3OX105F/Applicatpr_NY_Foam 10.5ml | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0.32967 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Main_Plans
3OX105F | 3OX105F/Applicatpr_NY_Foam 10.5ml | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OX105F | 3OX105F/Applicatpr_NY_Foam 10.5ml | Applicatpr_NY_Foam 10.5ml | NY-Analytical | 397 | Appearance | Appearance | Appearance | Meets specifications (Done) | [5 empty] | Visual | Main_Plans
3OX105F | 3OX105F/Applicatpr_NY_Foam 10.5ml | Applicatpr_NY_Foam 10.5ml | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Mold (OOS) | [5 empty] | WI QAS 076 | Main_Plans
3OX465H | 3OX465H/Handle, 26 ml | Handle, 26 ml | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | COA | Main_Plans
3OX465H | 3OX465H/Handle, 26 ml | Handle, 26 ml | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
3OX465H | 3OX465H/Handle, 26 ml | Handle, 26 ml | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Natural (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
3OX465H | 3OX465H/Handle, 26 ml | Handle, 26 ml | NY-Analytical | 765 | Canister & Lid Inspection | Handle Internal Diameter | Handle Internal Diameter | 0.9575 |  | in | FDA,3 | 0.951 | 0.957 | Canister & Lid Inspection | Main_Plans
3OX465H | 3OX465H/Handle, 26 ml | Handle, 26 ml | NY-Analytical | 765 | Canister & Lid Inspection | Overall Length | Overall Length | 5.23 |  | in | FDA,3 | 5.2 | 5.25 | Canister & Lid Inspection | Main_Plans
3OX465H | 3OX465H/Handle, 26 ml | Handle, 26 ml | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 87.86 |  | g | FDA,2 |  |  | TM 610 | Main_Plans
3OX465H | 3OX465H/Handle, 26 ml | Handle, 26 ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA | [5 empty] | TM 610 | Main_Plans
3OX465H | 3OX465H/Handle, 26 ml | Handle, 26 ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer | [5 empty] | TM 610 | Main_Plans
3OX465H | 3OX465H/Handle, 26 ml | Handle, 26 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 3 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OX465H | 3OX465H/Handle, 26 ml | Handle, 26 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0.034145 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
3OX465H | 3OX465H/Handle, 26 ml | Handle, 26 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OX465H | 3OX465H/Handle, 26 ml | Handle, 26 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OX465H | 3OX465H/Handle, 26 ml | Handle, 26 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Main_Plans
3OX465H | 3OX465H/Handle, 26 ml | Handle, 26 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 13 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OX465H | 3OX465H/Handle, 26 ml | Handle, 26 ml | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Non-Subculturable Growth (Done) | [5 empty] | WI QAS 076 | Main_Plans
3OX465H | 3OX465H/Handle, 26 ml | Handle, 26 ml | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Gram Positive Rods (Done) | [5 empty] | WI QAS 076 | Main_Plans
3OX465Q | 3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | Swab_NY_Swabstick, Cotton-Tipped | NY-Analytical | 760 | Swab Measurements | Mitt_Length_(Inches) | Mitt Length |  |  | in | FDA,3 | 0.593 | 0.781 | Swab Measurements | Main_Plans
3OX465Q | 3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | Swab_NY_Swabstick, Cotton-Tipped | NY-Analytical | 760 | Swab Measurements | Mitt_Width_(Inches) | Mitt Width |  |  | in | FDA,3 | 0.188 | 0.28 | Swab Measurements | Main_Plans
3OX465Q | 3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | Swab_NY_Swabstick, Cotton-Tipped | NY-Analytical | 760 | Swab Measurements | Overall_Swab_Length_(Inches) | Overall Swab Length |  |  | in | FDA,3 | 5.875 | 5.999 | Swab Measurements | Main_Plans
3OX465Q | 3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | Swab_NY_Swabstick, Cotton-Tipped | NY-Analytical | 760 | Swab Measurements | Stick Diameter | Stick Diameter |  |  | in | FDA,3 | 0.093 | 0.107 | Swab Measurements | Main_Plans
3OX465Q | 3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | Swab_NY_Swabstick, Cotton-Tipped | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [6 empty] | Certificate of Conformance | Main_Plans
3OX465Q | 3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | Swab_NY_Swabstick, Cotton-Tipped | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | [6 empty] | Bioburden_TAMC | Main_Plans
3OX465Q | 3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 0.58 |  | g | FDA,2 |  |  | TM 610 | Main_Plans
3OX465Q | 3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA | [5 empty] | TM 610 | Main_Plans
3OX465Q | 3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer | [5 empty] | TM 610 | Main_Plans
3OX465Q | 3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OX465Q | 3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram) | TAMC (Average per Gram) | 0 |  | CFU/g |  |  |  | TM 610 | Main_Plans
3OX465Q | 3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu |  |  |  | TM 610 | Main_Plans
3OX465Q | 3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OX465Q | 3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Main_Plans
3OX465Q | 3OX465Q/Swab_NY_Swabstick, Cotton-Tipped | Swab_NY_Swabstick, Cotton-Tipped | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OX510H | 3OX510H/Handle, 5.1 ml | Handle, 5.1 ml | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | COA | Main_Plans
3OX510H | 3OX510H/Handle, 5.1 ml | Handle, 5.1 ml | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
3OX510H | 3OX510H/Handle, 5.1 ml | Handle, 5.1 ml | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Natural (Done) | [5 empty] | Canister & Lid Inspection | Main_Plans
3OX510H | 3OX510H/Handle, 5.1 ml | Handle, 5.1 ml | NY-Analytical | 765 | Canister & Lid Inspection | Handle Internal Diameter | Handle Internal Diameter | 0.6405 |  | in | FDA,3 | 0.639 | 0.645 | Canister & Lid Inspection | Main_Plans
3OX510H | 3OX510H/Handle, 5.1 ml | Handle, 5.1 ml | NY-Analytical | 765 | Canister & Lid Inspection | Overall Length | Overall Length | 3.156 |  | in | FDA,3 | 3.149 | 3.189 | Canister & Lid Inspection | Main_Plans
3OX510H | 3OX510H/Handle, 5.1 ml | Handle, 5.1 ml | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 1.5 |  | g | FDA,2 |  |  | TM 610 | Main_Plans
3OX510H | 3OX510H/Handle, 5.1 ml | Handle, 5.1 ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer | [5 empty] | TM 610 | Main_Plans
3OX510H | 3OX510H/Handle, 5.1 ml | Handle, 5.1 ml | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA | [5 empty] | TM 610 | Main_Plans
3OX510H | 3OX510H/Handle, 5.1 ml | Handle, 5.1 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OX510H | 3OX510H/Handle, 5.1 ml | Handle, 5.1 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
3OX510H | 3OX510H/Handle, 5.1 ml | Handle, 5.1 ml | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3OX510H | 3OX510H/Handle, 5.1 ml | Handle, 5.1 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3OX510H | 3OX510H/Handle, 5.1 ml | Handle, 5.1 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Main_Plans
3OX510H | 3OX510H/Handle, 5.1 ml | Handle, 5.1 ml | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3XP134 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Applicator Measurements | Main_Plans
3XP134 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 39 |  | in | FDA,0 | 39 | 40 | Applicator Measurements | Main_Plans
3XP134 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 36 |  | in | FDA,3 | 35.875 | 36.125 | Applicator Measurements | Main_Plans
3XP134 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (inches) (TM 102) | Caliper | 0.0082 |  | in | FDA,3 | 0.006 | 0.009 | See Analyte for Method # | Main_Plans
3XP134 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 34.44 |  | GSM | FDA,1 | 31.5 | 38.5 | TM 101 | Main_Plans
3XP134 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 842.87 |  | g/in | FDA,0 | 250 | 1250 | TM 103 | Main_Plans
3XP134 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 1755.59 |  | g/in | FDA,0 | 800 | 2600 | TM 103 | Main_Plans
3XP134 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3XP134 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 748 | Visual Inspection | Color | Color | White with no off colors or shading [Done] | [5 empty] | TM 208 | Main_Plans
3XP134 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | Applicator_NY_35gsm Stora Flat - 36" | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3XP134 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 1.55 |  | g | FDA,2 |  |  | TM 610 | Main_Plans
3XP134 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA | [5 empty] | TM 610 | Main_Plans
3XP134 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer | [5 empty] | TM 610 | Main_Plans
3XP134 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 1 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3XP134 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0.12903 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
3XP134 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 2 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3XP134 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3XP134 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Main_Plans
3XP134 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3XP134 | 3XP134/Applicator_NY_35gsm Stora Flat - 36" | Applicator_NY_35gsm Stora Flat - 36" | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Negative Rods (OOS) | [5 empty] | WI QAS 076 | Main_Plans
3XP138 | 3XP138/Applicator_36B | Applicator_36B | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Applicator Measurements | Main_Plans
3XP138 | 3XP138/Applicator_36B | Applicator_36B | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 45 |  | in | FDA,0 | 43 | 47 | Applicator Measurements | Main_Plans
3XP138 | 3XP138/Applicator_36B | Applicator_36B | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 37.5 |  | in | FDA,3 | 37.375 | 37.625 | Applicator Measurements | Main_Plans
3XP138 | 3XP138/Applicator_36B | Applicator_36B | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 40.04 |  | GSM | FDA,0 | 35 | 45 | TM 101 | Main_Plans
3XP138 | 3XP138/Applicator_36B | Applicator_36B | NY-Analytical | 757 | Caliper | Caliper_CoA_(uM) | Caliper | 419.8 |  | um | FDA,0 | 340 | 540 | TM 102 | Main_Plans
3XP138 | 3XP138/Applicator_36B | Applicator_36B | NY-Analytical | 758 | Tensile Strength | Wet_CD_CoA_(g/in) | Wet CD CoA | 1069 |  | g/in | FDA,0 | 400 |  | TM 103 | Main_Plans
3XP138 | 3XP138/Applicator_36B | Applicator_36B | NY-Analytical | 758 | Tensile Strength | Wet_MD_CoA_(g/in) | Wet MD CoA | 2099 |  | g/in | FDA,0 | 1000 |  | TM 103 | Main_Plans
3XP138 | 3XP138/Applicator_36B | Applicator_36B | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3XP138 | 3XP138/Applicator_36B | Applicator_36B | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3XP138 | 3XP138/Applicator_36B | Applicator_36B | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3XP253 | 3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Applicator Measurements | Main_Plans
3XP253 | 3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 30 |  | in | FDA,3 | 29.875 | 30.125 | Applicator Measurements | Main_Plans
3XP253 | 3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 37.03 |  | GSM | FDA,1 | 31.5 | 38.5 | TM 101 | Main_Plans
3XP253 | 3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_CoA_(mm) | Caliper | 0.33 |  | mm | FDA,2 | 0.27 | 0.41 | TM 102 | Main_Plans
3XP253 | 3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 599.1 |  | g/in | FDA,0 | 300 |  | TM 103 | Main_Plans
3XP253 | 3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 3682.03 |  | g/in | FDA,0 | 2000 | 4400 | TM 103 | Main_Plans
3XP253 | 3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3XP253 | 3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 748 | Visual Inspection | Color | Color | White with no off colors or shading [Done] | [5 empty] | TM 208 | Main_Plans
3XP253 | 3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3XP253 | 3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 0.98 |  | g | FDA,2 |  |  | TM 610 | Main_Plans
3XP253 | 3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA | [5 empty] | TM 610 | Main_Plans
3XP253 | 3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA | [5 empty] | TM 610 | Main_Plans
3XP253 | 3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 1 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3XP253 | 3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0.91837 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
3XP253 | 3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 1 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3XP253 | 3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3XP253 | 3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Main_Plans
3XP253 | 3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3XP253 | 3XP253/Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | Applicator_NY_35gsm Flat Thermoembonded Nonwoven 30" | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Positive Rods (Done) | [5 empty] | WI QAS 076 | Main_Plans
3XP540 | 3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (COA) | Roll Diameter (COA) | 43 |  | in | FDA,0 | 41 | 45 | Applicator Measurements | Main_Plans
3XP540 | 3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 759 | Applicator Measurements | Roll width (COA) | Roll width (COA) | 42 |  | in | FDA,2 | 41.75 | 42.25 | Applicator Measurements | Main_Plans
3XP540 | 3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [6 empty] | Basis Weight | Main_Plans
3XP540 | 3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 52.31 |  | GSM | FDA,0 | 45 | 55 | TM 101 | Main_Plans
3XP540 | 3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper | 0.42 |  | mm | FDA,2 | 0.51 | 0.63 | TM 102 | Main_Plans
3XP540 | 3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 758 | Tensile Strength | Dry_CD_(g/in) | Dry CD | 2372 |  | g/in | FDA,0 | 1801 | 2821 | TM 103 | Main_Plans
3XP540 | 3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 758 | Tensile Strength | Dry_MD_(g/in) | Dry CD | 9752 |  | g/in | FDA,0 | 6675 | 9447 | TM 103 | Main_Plans
3XP540 | 3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3XP540 | 3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3XP540 | 3XP540/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-42" | NY-Analytical | 748 | Visual Inspection | Odor | Odor | No objectionable odor Wet or Dry [Done] | [5 empty] | TM 208 | Main_Plans
3XP667 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter | 43 |  | in | FDA,0 | 41 | 45 | Applicator Measurements | Main_Plans
3XP667 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 38.75 |  | in | FDA,2 | 38.5 | 39 | Applicator Measurements | Main_Plans
3XP667 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Basis Weight | Main_Plans
3XP667 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 49.94 |  | GSM | FDA,1 | 45 | 55 | TM 101 | Main_Plans
3XP667 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_CoA_(mm) | Caliper | 0.56 |  | mm | FDA,2 | 0.51 | 0.63 | TM 102 | Main_Plans
3XP667 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 758 | Tensile Strength | Dry_CD_(g/in) | Dry CD | 2635 |  | g/in | FDA,0 | 1801 | 2821 | TM 103 | Main_Plans
3XP667 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 758 | Tensile Strength | Dry_MD_(g/in) | Dry CD | 8881 |  | g/in | FDA,0 | 6675 | 9447 | TM 103 | Main_Plans
3XP667 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3XP667 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) | [5 empty] | TM 208 | Main_Plans
3XP667 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Analytical | 748 | Visual Inspection | Odor | Odor | No objectionable odor Wet or Dry [Done] | [5 empty] | TM 208 | Main_Plans
3XP667 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 0.94 |  | g | FDA,2 |  |  | TM 610 | Main_Plans
3XP667 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer | [5 empty] | TM 610 | Main_Plans
3XP667 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer | [5 empty] | TM 610 | Main_Plans
3XP667 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3XP667 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
3XP667 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
3XP667 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Main_Plans
3XP667 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Main_Plans
3XP667 | 3XP667/Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | Applicator_NY_Spunlaced Nonwoven 50 gsm 100% Polyester-38.75" | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
5AA347BT | 5AA347BT/Box_NY_0.020SBS_300_3.375-1.313-2.625 | Box_NY_0.020SBS_300_3.375-1.313-2.625 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 315 |  |  | FDA,0 |  |  | Box Inspection | Main_Plans
5AA347BT | 5AA347BT/Box_NY_0.020SBS_300_3.375-1.313-2.625 | Box_NY_0.020SBS_300_3.375-1.313-2.625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5AA347BT | 5AA347BT/Box_NY_0.020SBS_300_3.375-1.313-2.625 | Box_NY_0.020SBS_300_3.375-1.313-2.625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.019 |  | in | FDA,3 | 0.019 | 0.021 | Box Inspection | Main_Plans
5AA347BT | 5AA347BT/Box_NY_0.020SBS_300_3.375-1.313-2.625 | Box_NY_0.020SBS_300_3.375-1.313-2.625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) | [5 empty] | Box Inspection | Main_Plans
5AA347BT | 5AA347BT/Box_NY_0.020SBS_300_3.375-1.313-2.625 | Box_NY_0.020SBS_300_3.375-1.313-2.625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Open Top/Auto Bottom Tray (Done) | [5 empty] | Box Inspection | Main_Plans
5AA347BT | 5AA347BT/Box_NY_0.020SBS_300_3.375-1.313-2.625 | Box_NY_0.020SBS_300_3.375-1.313-2.625 | NY-Analytical | 772 | Box Inspection | Height | Height | 2.625 |  | in | FDA,5 | 2.5938 | 2.6562 | Box Inspection | Main_Plans
5AA347BT | 5AA347BT/Box_NY_0.020SBS_300_3.375-1.313-2.625 | Box_NY_0.020SBS_300_3.375-1.313-2.625 | NY-Analytical | 772 | Box Inspection | Length | Length | 3.375 |  | in | FDA,5 | 3.3438 | 3.4062 | Box Inspection | Main_Plans
5AA347BT | 5AA347BT/Box_NY_0.020SBS_300_3.375-1.313-2.625 | Box_NY_0.020SBS_300_3.375-1.313-2.625 | NY-Analytical | 772 | Box Inspection | Width | Width | 1.3125 |  | in | FDA,5 | 1.2817 | 1.3442 | Box Inspection | Main_Plans
5AA404S48 | 5AA404S48/Box_NY_0.018SBS_111_6-5.875-3.625 | Box_NY_0.018SBS_111_6-5.875-3.625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy |  |  | in |  |  |  | Box Inspection | Main_Plans
5AA404S48 | 5AA404S48/Box_NY_0.018SBS_111_6-5.875-3.625 | Box_NY_0.018SBS_111_6-5.875-3.625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper |  |  | in | FDA,3 | 0.017 | 0.019 | Box Inspection | Main_Plans
5AA404S48 | 5AA404S48/Box_NY_0.018SBS_111_6-5.875-3.625 | Box_NY_0.018SBS_111_6-5.875-3.625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | [6 empty] | Box Inspection | Main_Plans
5AA404S48 | 5AA404S48/Box_NY_0.018SBS_111_6-5.875-3.625 | Box_NY_0.018SBS_111_6-5.875-3.625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | [6 empty] | Box Inspection | Main_Plans
5AA404S48 | 5AA404S48/Box_NY_0.018SBS_111_6-5.875-3.625 | Box_NY_0.018SBS_111_6-5.875-3.625 | NY-Analytical | 772 | Box Inspection | Height | Height |  |  | in | FDA,5 | 3.5938 | 3.6562 | Box Inspection | Main_Plans
5AA404S48 | 5AA404S48/Box_NY_0.018SBS_111_6-5.875-3.625 | Box_NY_0.018SBS_111_6-5.875-3.625 | NY-Analytical | 772 | Box Inspection | Length | Length |  |  | in | FDA,5 | 5.9688 | 6.0312 | Box Inspection | Main_Plans
5AA404S48 | 5AA404S48/Box_NY_0.018SBS_111_6-5.875-3.625 | Box_NY_0.018SBS_111_6-5.875-3.625 | NY-Analytical | 772 | Box Inspection | Width | Width |  |  | in | FDA,5 | 5.8438 | 5.9062 | Box Inspection | Main_Plans
5AA617S96 | 5AA617S96/Box_NY_0.018SBS_47B_3.625-3.875-6 | Box_NY_0.018SBS_47B_3.625-3.875-6 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5AA617S96 | 5AA617S96/Box_NY_0.018SBS_47B_3.625-3.875-6 | Box_NY_0.018SBS_47B_3.625-3.875-6 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.019 |  | in | FDA,3 | 0.017 | 0.019 | Box Inspection | Main_Plans
5AA617S96 | 5AA617S96/Box_NY_0.018SBS_47B_3.625-3.875-6 | Box_NY_0.018SBS_47B_3.625-3.875-6 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) | [5 empty] | Box Inspection | Main_Plans
5AA617S96 | 5AA617S96/Box_NY_0.018SBS_47B_3.625-3.875-6 | Box_NY_0.018SBS_47B_3.625-3.875-6 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Straight Tuck End - French Style (Done) | [5 empty] | Box Inspection | Main_Plans
5AA617S96 | 5AA617S96/Box_NY_0.018SBS_47B_3.625-3.875-6 | Box_NY_0.018SBS_47B_3.625-3.875-6 | NY-Analytical | 772 | Box Inspection | Height | Height | 6 |  | in | FDA,5 | 5.9688 | 6.0312 | Box Inspection | Main_Plans
5AA617S96 | 5AA617S96/Box_NY_0.018SBS_47B_3.625-3.875-6 | Box_NY_0.018SBS_47B_3.625-3.875-6 | NY-Analytical | 772 | Box Inspection | Length | Length | 3.625 |  | in | FDA,5 | 3.5938 | 3.6562 | Box Inspection | Main_Plans
5AA617S96 | 5AA617S96/Box_NY_0.018SBS_47B_3.625-3.875-6 | Box_NY_0.018SBS_47B_3.625-3.875-6 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.875 |  | in | FDA,5 | 3.8438 | 3.9062 | Box Inspection | Main_Plans
5AA632S60 | 5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | Box_NY_0.020SBS_101_6-3.625-7 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 3 | [5 empty] | Box Inspection | Main_Plans
5AA632S60 | 5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | Box_NY_0.020SBS_101_6-3.625-7 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5AA632S60 | 5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | Box_NY_0.020SBS_101_6-3.625-7 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.021 |  | in | FDA,3 | 0.019 | 0.021 | Box Inspection | Main_Plans
5AA632S60 | 5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | Box_NY_0.020SBS_101_6-3.625-7 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) | [5 empty] | Box Inspection | Main_Plans
5AA632S60 | 5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | Box_NY_0.020SBS_101_6-3.625-7 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom - French Style (Done) | [5 empty] | Box Inspection | Main_Plans
5AA632S60 | 5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | Box_NY_0.020SBS_101_6-3.625-7 | NY-Analytical | 772 | Box Inspection | Height | Height | 7 |  | in | FDA,5 | 6.9688 | 7.0312 | Box Inspection | Main_Plans
5AA632S60 | 5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | Box_NY_0.020SBS_101_6-3.625-7 | NY-Analytical | 772 | Box Inspection | Length | Length | 6 |  | in | FDA,5 | 5.9688 | 6.0312 | Box Inspection | Main_Plans
5AA632S60 | 5AA632S60/Box_NY_0.020SBS_101_6-3.625-7 | Box_NY_0.020SBS_101_6-3.625-7 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.625 |  | in | FDA,5 | 3.5938 | 3.6562 | Box Inspection | Main_Plans
5AA911TY | 5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom w/perf | [5 empty] | Box Inspection | Main_Plans
5AA911TY | 5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | Box Inspection | Main_Plans
5AA911TY | 5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5AA911TY | 5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0195 |  | in | FDA,3 | 0.019 | 0.021 | Box Inspection | Main_Plans
5AA911TY | 5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) | [5 empty] | Box Inspection | Main_Plans
5AA911TY | 5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | NY-Analytical | 772 | Box Inspection | Height | Height | 7.94 |  | in | FDA,5 | 7.9062 | 7.9688 | Box Inspection | Main_Plans
5AA911TY | 5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | NY-Analytical | 772 | Box Inspection | Length | Length | 4.51 |  | in | FDA,5 | 4.4688 | 4.5312 | Box Inspection | Main_Plans
5AA911TY | 5AA911TY/Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | Box_NY_0.020SBS_105_4.5-7.9375-7.9375 | NY-Analytical | 772 | Box Inspection | Width | Width | 7.94 |  | in | FDA,5 | 7.9062 | 7.9688 | Certificate of Conformance | Main_Plans
5AB123CP | 5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | Label_NY_5AB123CP_1.9375-1.875 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | [6 empty] | Label Inspection | Main_Plans
5AB123CP | 5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | Label_NY_5AB123CP_1.9375-1.875 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
5AB123CP | 5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | Label_NY_5AB123CP_1.9375-1.875 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
5AB123CP | 5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | Label_NY_5AB123CP_1.9375-1.875 | NY-Analytical | 770 | Label Inspection | Length | Length | 1.875 |  | in | FDA,4 | 1.875 | 2 | TM 143 / Visual | Main_Plans
5AB123CP | 5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | Label_NY_5AB123CP_1.9375-1.875 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
5AB123CP | 5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | Label_NY_5AB123CP_1.9375-1.875 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
5AB123CP | 5AB123CP/Label_NY_5AB123CP_1.9375-1.875 | Label_NY_5AB123CP_1.9375-1.875 | NY-Analytical | 770 | Label Inspection | Width | Width | 1.9375 |  | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | Main_Plans
5AB123ST | 5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | Box_NY_0.018SBS_107_10.25-2.375-2.75 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5AB123ST | 5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | Box_NY_0.018SBS_107_10.25-2.375-2.75 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.018 |  | in | FDA,3 | 0.017 | 0.019 | Box Inspection | Main_Plans
5AB123ST | 5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | Box_NY_0.018SBS_107_10.25-2.375-2.75 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) | [5 empty] | Box Inspection | Main_Plans
5AB123ST | 5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | Box_NY_0.018SBS_107_10.25-2.375-2.75 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | TT/AB w/friction lock (Done) | [5 empty] | Box Inspection | Main_Plans
5AB123ST | 5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | Box_NY_0.018SBS_107_10.25-2.375-2.75 | NY-Analytical | 772 | Box Inspection | Height | Height | 2.7515 |  | in | FDA,5 | 2.7188 | 2.7812 | Box Inspection | Main_Plans
5AB123ST | 5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | Box_NY_0.018SBS_107_10.25-2.375-2.75 | NY-Analytical | 772 | Box Inspection | Length | Length | 10.22 |  | in | FDA,5 | 10.2188 | 10.2812 | Box Inspection | Main_Plans
5AB123ST | 5AB123ST/Box_NY_0.018SBS_107_10.25-2.375-2.75 | Box_NY_0.018SBS_107_10.25-2.375-2.75 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.351 |  | in | FDA,5 | 2.3438 | 2.4062 | Box Inspection | Main_Plans
5AD02424 | 5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | Box_NY_0.018SBS_107_4-2.03125-4.5 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top Auto Bottom w/ Tab Lock | [5 empty] | Box Inspection | Main_Plans
5AD02424 | 5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | Box_NY_0.018SBS_107_4-2.03125-4.5 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5AD02424 | 5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | Box_NY_0.018SBS_107_4-2.03125-4.5 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0175 |  | in | FDA,3 | 0.017 | 0.019 | Box Inspection | Main_Plans
5AD02424 | 5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | Box_NY_0.018SBS_107_4-2.03125-4.5 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) | [5 empty] | Box Inspection | Main_Plans
5AD02424 | 5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | Box_NY_0.018SBS_107_4-2.03125-4.5 | NY-Analytical | 772 | Box Inspection | Height | Height | 4.5 |  | in | FDA,5 | 4.4688 | 4.5312 | Box Inspection | Main_Plans
5AD02424 | 5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | Box_NY_0.018SBS_107_4-2.03125-4.5 | NY-Analytical | 772 | Box Inspection | Length | Length | 4 |  | in | FDA,5 | 3.9688 | 4.0312 | Box Inspection | Main_Plans
5AD02424 | 5AD02424/Box_NY_0.018SBS_107_4-2.03125-4.5 | Box_NY_0.018SBS_107_4-2.03125-4.5 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.0312 |  | in | FDA,5 | 2 | 2.0625 | Box Inspection | Main_Plans
5AH45016 | 5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | Box_NY_0.018SBS_101_3.875-2-3.5 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 200 | [5 empty] | Box Inspection | Main_Plans
5AH45016 | 5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | Box_NY_0.018SBS_101_3.875-2-3.5 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5AH45016 | 5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | Box_NY_0.018SBS_101_3.875-2-3.5 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.018 |  | in | FDA,3 | 0.017 | 0.019 | Box Inspection | Main_Plans
5AH45016 | 5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | Box_NY_0.018SBS_101_3.875-2-3.5 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) | [5 empty] | Box Inspection | Main_Plans
5AH45016 | 5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | Box_NY_0.018SBS_101_3.875-2-3.5 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | French Tuck Top / Auto Bottom (Done) | [5 empty] | Box Inspection | Main_Plans
5AH45016 | 5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | Box_NY_0.018SBS_101_3.875-2-3.5 | NY-Analytical | 772 | Box Inspection | Height | Height | 3.5 |  | in | FDA,5 | 3.4688 | 3.5312 | Box Inspection | Main_Plans
5AH45016 | 5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | Box_NY_0.018SBS_101_3.875-2-3.5 | NY-Analytical | 772 | Box Inspection | Length | Length | 3.875 |  | in | FDA,5 | 3.8438 | 3.9062 | Box Inspection | Main_Plans
5AH45016 | 5AH45016/Box_NY_0.018SBS_101_3.875-2-3.5 | Box_NY_0.018SBS_101_3.875-2-3.5 | NY-Analytical | 772 | Box Inspection | Width | Width | 2 |  | in | FDA,5 | 1.9688 | 2.0312 | Box Inspection | Main_Plans
5AH58195 | 5AH58195/Box_NY_0.020SBS_127_6.25-2.5625-4.625 | Box_NY_0.020SBS_127_6.25-2.5625-4.625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Straight Tuck w/Partition Dispenser Box | [5 empty] | Box Inspection | Main_Plans
5AH58195 | 5AH58195/Box_NY_0.020SBS_127_6.25-2.5625-4.625 | Box_NY_0.020SBS_127_6.25-2.5625-4.625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5AH58195 | 5AH58195/Box_NY_0.020SBS_127_6.25-2.5625-4.625 | Box_NY_0.020SBS_127_6.25-2.5625-4.625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0195 |  | in | FDA,3 | 0.019 | 0.021 | Box Inspection | Main_Plans
5AH58195 | 5AH58195/Box_NY_0.020SBS_127_6.25-2.5625-4.625 | Box_NY_0.020SBS_127_6.25-2.5625-4.625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) | [5 empty] | Box Inspection | Main_Plans
5AH58195 | 5AH58195/Box_NY_0.020SBS_127_6.25-2.5625-4.625 | Box_NY_0.020SBS_127_6.25-2.5625-4.625 | NY-Analytical | 772 | Box Inspection | Height | Height | 4.625 |  | in | FDA,5 | 4.5938 | 4.6562 | Box Inspection | Main_Plans
5AH58195 | 5AH58195/Box_NY_0.020SBS_127_6.25-2.5625-4.625 | Box_NY_0.020SBS_127_6.25-2.5625-4.625 | NY-Analytical | 772 | Box Inspection | Length | Length | 6.25 |  | in | FDA,5 | 6.2188 | 6.2812 | Box Inspection | Main_Plans
5AH58195 | 5AH58195/Box_NY_0.020SBS_127_6.25-2.5625-4.625 | Box_NY_0.020SBS_127_6.25-2.5625-4.625 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.5625 |  | in | FDA,5 | 2.5312 | 2.5938 | Box Inspection | Main_Plans
5AJ130BT | 5AJ130BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size |  |  |  | FDA,0 |  |  | Box Inspection | Main_Plans
5AJ130BT | 5AJ130BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | [6 empty] | Box Inspection | Main_Plans
5AJ130BT | 5AJ130BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper |  |  | in | FDA,3 | 0.019 | 0.021 | Box Inspection | Main_Plans
5AJ130BT | 5AJ130BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | [6 empty] | Box Inspection | Main_Plans
5AJ130BT | 5AJ130BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | [6 empty] | Box Inspection | Main_Plans
5AJ130BT | 5AJ130BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | NY-Analytical | 772 | Box Inspection | Height | Height |  |  | in | FDA,5 | 2.5938 | 2.6562 | Box Inspection | Main_Plans
5AJ130BT | 5AJ130BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | NY-Analytical | 772 | Box Inspection | Length | Length |  |  | in | FDA,5 | 3.2812 | 3.3438 | Box Inspection | Main_Plans
5AJ130BT | 5AJ130BT/Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | Box_NY_0.020SBS_306_3.3125-1.4375-2.625 | NY-Analytical | 772 | Box Inspection | Width | Width |  |  | in | FDA,5 | 1.4062 | 1.4688 | Box Inspection | Main_Plans
5AP027TY | 5AP027TY/Box_57A | Box_57A | NY-Analytical | 772 | Box Inspection | Appearance | Appearance | Meets specification (Done) | [5 empty] | Box Inspection | Main_Plans
5AP027TY | 5AP027TY/Box_57A | Box_57A | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) | [5 empty] | Box Inspection | Main_Plans
5AP027TY | 5AP027TY/Box_57A | Box_57A | NY-Analytical | 772 | Box Inspection | Color | Color | Clear (Done) | [5 empty] | Box Inspection | Main_Plans
5AP027TY | 5AP027TY/Box_57A | Box_57A | NY-Analytical | 772 | Box Inspection | Functionality | Functionality | Meets specification (Done) | [5 empty] | Box Inspection | Main_Plans
5AP027TY | 5AP027TY/Box_57A | Box_57A | NY-Analytical | 772 | Box Inspection | Length | Length | 10.75 |  | in | FDA,3 | 10.66 | 10.84 | Box Inspection | Main_Plans
5AP027TY | 5AP027TY/Box_57A | Box_57A | NY-Analytical | 772 | Box Inspection | Odor | Odor | None Detectable (Done) | [5 empty] | Box Inspection | Main_Plans
5AP027TY | 5AP027TY/Box_57A | Box_57A | NY-Analytical | 772 | Box Inspection | Width | Width | 5.625 |  | in | FDA,3 | 5.535 | 5.715 | Box Inspection | Main_Plans
5AP07530 | 5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | Box_NY_0.020BC_110_6.375-3.4375-4.75 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 125 |  |  | FDA,0 |  |  | Box Inspection | Main_Plans
5AP07530 | 5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | Box_NY_0.020BC_110_6.375-3.4375-4.75 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5AP07530 | 5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | Box_NY_0.020BC_110_6.375-3.4375-4.75 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.02 |  | in | FDA,3 | 0.019 | 0.021 | Box Inspection | Main_Plans
5AP07530 | 5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | Box_NY_0.020BC_110_6.375-3.4375-4.75 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | BC (Done) | [5 empty] | Box Inspection | Main_Plans
5AP07530 | 5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | Box_NY_0.020BC_110_6.375-3.4375-4.75 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | TT/AB w/finger holes (Done) | [5 empty] | Box Inspection | Main_Plans
5AP07530 | 5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | Box_NY_0.020BC_110_6.375-3.4375-4.75 | NY-Analytical | 772 | Box Inspection | Height | Height | 4.771 |  | in | FDA,5 | 4.7188 | 4.7812 | Box Inspection | Main_Plans
5AP07530 | 5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | Box_NY_0.020BC_110_6.375-3.4375-4.75 | NY-Analytical | 772 | Box Inspection | Length | Length | 6.395 |  | in | FDA,5 | 6.3438 | 6.4062 | Box Inspection | Main_Plans
5AP07530 | 5AP07530/Box_NY_0.020BC_110_6.375-3.4375-4.75 | Box_NY_0.020BC_110_6.375-3.4375-4.75 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.4225 |  | in | FDA,5 | 3.4062 | 3.4688 | Box Inspection | Main_Plans
5AP209SL | 5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | Box_NY_0.020BC_220_5.09375-3.6875-4.625 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 30 |  |  | FDA,0 |  |  | Box Inspection | Main_Plans
5AP209SL | 5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | Box_NY_0.020BC_220_5.09375-3.6875-4.625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5AP209SL | 5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | Box_NY_0.020BC_220_5.09375-3.6875-4.625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0195 |  | in | FDA,3 | 0.019 | 0.021 | Box Inspection | Main_Plans
5AP209SL | 5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | Box_NY_0.020BC_220_5.09375-3.6875-4.625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | BC (Done) | [5 empty] | Box Inspection | Main_Plans
5AP209SL | 5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | Box_NY_0.020BC_220_5.09375-3.6875-4.625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Diecut Glued Sleeve w/ 2-1" fingerholes (Done) | [5 empty] | Box Inspection | Main_Plans
5AP209SL | 5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | Box_NY_0.020BC_220_5.09375-3.6875-4.625 | NY-Analytical | 772 | Box Inspection | Height | Height | 4.61 |  | in | FDA,5 | 4.5938 | 4.6562 | Box Inspection | Main_Plans
5AP209SL | 5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | Box_NY_0.020BC_220_5.09375-3.6875-4.625 | NY-Analytical | 772 | Box Inspection | Length | Length | 5.07 |  | in | FDA,5 | 5.0625 | 5.125 | Box Inspection | Main_Plans
5AP209SL | 5AP209SL/Box_NY_0.020BC_220_5.09375-3.6875-4.625 | Box_NY_0.020BC_220_5.09375-3.6875-4.625 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.7 |  | in | FDA,5 | 3.6562 | 3.7188 | Box Inspection | Main_Plans
5AP21195 | 5AP21195/Box_01M | Box_01M | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | Bending Chip | [5 empty] | Box Inspection | Main_Plans
5AP21195 | 5AP21195/Box_01M | Box_01M | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5AP21195 | 5AP21195/Box_01M | Box_01M | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0235 |  | in | FDA,3 | 0.023 | 0.025 | Box Inspection | Main_Plans
5AP21195 | 5AP21195/Box_01M | Box_01M | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) | [5 empty] | Box Inspection | Main_Plans
5AP21195 | 5AP21195/Box_01M | Box_01M | NY-Analytical | 772 | Box Inspection | Height | Height | 5.875 |  | in | FDA,5 | 5.8438 | 5.9062 | Box Inspection | Main_Plans
5AP21195 | 5AP21195/Box_01M | Box_01M | NY-Analytical | 772 | Box Inspection | Length | Length | 4.625 |  | in | FDA,5 | 4.5938 | 4.6562 | Box Inspection | Main_Plans
5AP21195 | 5AP21195/Box_01M | Box_01M | NY-Analytical | 772 | Box Inspection | Width | Width | 4.5 |  | in | FDA,5 | 4.4688 | 4.5312 | Box Inspection | Main_Plans
5AP227TY | 5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | Box_NY_0.020SBS_300_4.625-5-3.625 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 125 |  |  | FDA,0 |  |  | Box Inspection | Main_Plans
5AP227TY | 5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | Box_NY_0.020SBS_300_4.625-5-3.625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5AP227TY | 5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | Box_NY_0.020SBS_300_4.625-5-3.625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.02 |  | in | FDA,3 | 0.019 | 0.021 | Box Inspection | Main_Plans
5AP227TY | 5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | Box_NY_0.020SBS_300_4.625-5-3.625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) | [5 empty] | Box Inspection | Main_Plans
5AP227TY | 5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | Box_NY_0.020SBS_300_4.625-5-3.625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Open Top/Auto Bottom Tray (Done) | [5 empty] | Box Inspection | Main_Plans
5AP227TY | 5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | Box_NY_0.020SBS_300_4.625-5-3.625 | NY-Analytical | 772 | Box Inspection | Height | Height | 3.61 |  | in | FDA,5 | 3.5938 | 3.6562 | Box Inspection | Main_Plans
5AP227TY | 5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | Box_NY_0.020SBS_300_4.625-5-3.625 | NY-Analytical | 772 | Box Inspection | Length | Length | 4.6 |  | in | FDA,5 | 4.5938 | 4.6562 | Box Inspection | Main_Plans
5AP227TY | 5AP227TY/Box_NY_0.020SBS_300_4.625-5-3.625 | Box_NY_0.020SBS_300_4.625-5-3.625 | NY-Analytical | 772 | Box Inspection | Width | Width | 5.01 |  | in | FDA,5 | 4.9688 | 5.0312 | Box Inspection | Main_Plans
5AP229TY | 5AP229TY/Box_NY_0.024SBS_303_8.228-3.661-6.024 | Box_NY_0.024SBS_303_8.228-3.661-6.024 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5AP229TY | 5AP229TY/Box_NY_0.024SBS_303_8.228-3.661-6.024 | Box_NY_0.024SBS_303_8.228-3.661-6.024 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.023 |  | in | FDA,3 | 0.023 | 0.025 | Box Inspection | Main_Plans
5AP229TY | 5AP229TY/Box_NY_0.024SBS_303_8.228-3.661-6.024 | Box_NY_0.024SBS_303_8.228-3.661-6.024 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) | [5 empty] | Box Inspection | Main_Plans
5AP229TY | 5AP229TY/Box_NY_0.024SBS_303_8.228-3.661-6.024 | Box_NY_0.024SBS_303_8.228-3.661-6.024 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | A/B tray w/tear-off lid (Done) | [5 empty] | Box Inspection | Main_Plans
5AP229TY | 5AP229TY/Box_NY_0.024SBS_303_8.228-3.661-6.024 | Box_NY_0.024SBS_303_8.228-3.661-6.024 | NY-Analytical | 772 | Box Inspection | Height | Height | 6.03 |  | in | FDA,5 | 5.9928 | 6.0553 | Box Inspection | Main_Plans
5AP229TY | 5AP229TY/Box_NY_0.024SBS_303_8.228-3.661-6.024 | Box_NY_0.024SBS_303_8.228-3.661-6.024 | NY-Analytical | 772 | Box Inspection | Length | Length | 8.22 |  | in | FDA,5 | 8.1967 | 8.2592 | Box Inspection | Main_Plans
5AP229TY | 5AP229TY/Box_NY_0.024SBS_303_8.228-3.661-6.024 | Box_NY_0.024SBS_303_8.228-3.661-6.024 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.64 |  | in | FDA,5 | 3.6298 | 3.6923 | Box Inspection | Main_Plans
5AP308TY | 5AP308TY/Box_63A | Box_63A | NY-Analytical | 772 | Box Inspection | Appearance | Appearance | Meets specification (Done) | [5 empty] | Box Inspection | Main_Plans
5AP308TY | 5AP308TY/Box_63A | Box_63A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Box Inspection | Main_Plans
5AP308TY | 5AP308TY/Box_63A | Box_63A | NY-Analytical | 772 | Box Inspection | Color | Color | Clear (Done) | [5 empty] | Box Inspection | Main_Plans
5AP308TY | 5AP308TY/Box_63A | Box_63A | NY-Analytical | 772 | Box Inspection | Flange - Sides | Flange - Sides | 0.075 |  | in | FDA,3 | 0.045 | 0.135 | Box Inspection | Main_Plans
5AP308TY | 5AP308TY/Box_63A | Box_63A | NY-Analytical | 772 | Box Inspection | Flange - Top & Bottom | Flange - Top & Bottom | 0.099 |  | in | FDA,3 | 0.045 | 0.135 | Box Inspection | Main_Plans
5AP308TY | 5AP308TY/Box_63A | Box_63A | NY-Analytical | 772 | Box Inspection | Length | Length | 7.2 |  | in | FDA,3 | 7.156 | 7.22 | Box Inspection | Main_Plans
5AP308TY | 5AP308TY/Box_63A | Box_63A | NY-Analytical | 772 | Box Inspection | Odor | Odor | None Detectable (Done) | [5 empty] | Box Inspection | Main_Plans
5AP308TY | 5AP308TY/Box_63A | Box_63A | NY-Analytical | 772 | Box Inspection | Total Height | Total Height | 2.95 |  | in | FDA,3 | 2.899 | 2.963 | Box Inspection | Main_Plans
5AP308TY | 5AP308TY/Box_63A | Box_63A | NY-Analytical | 772 | Box Inspection | Width | Width | 6.1 |  | in | FDA,3 | 6.093 | 6.157 | Box Inspection | Main_Plans
5AU25206 | 5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | Box_NY_0.018SBS_100_101.6-44.45-101.6 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 200 |  |  | FDA,0 |  |  | Box Inspection | Main_Plans
5AU25206 | 5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | Box_NY_0.018SBS_100_101.6-44.45-101.6 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5AU25206 | 5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | Box_NY_0.018SBS_100_101.6-44.45-101.6 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.018 |  | in | FDA,3 | 0.017 | 0.019 | Box Inspection | Main_Plans
5AU25206 | 5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | Box_NY_0.018SBS_100_101.6-44.45-101.6 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) | [5 empty] | Box Inspection | Main_Plans
5AU25206 | 5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | Box_NY_0.018SBS_100_101.6-44.45-101.6 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) | [5 empty] | Box Inspection | Main_Plans
5AU25206 | 5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | Box_NY_0.018SBS_100_101.6-44.45-101.6 | NY-Analytical | 772 | Box Inspection | Height | Height | 102 |  | mm | FDA,1 | 100.8 | 102.4 | Box Inspection | Main_Plans
5AU25206 | 5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | Box_NY_0.018SBS_100_101.6-44.45-101.6 | NY-Analytical | 772 | Box Inspection | Length | Length | 102 |  | mm | FDA,1 | 100.8 | 102.4 | Box Inspection | Main_Plans
5AU25206 | 5AU25206/Box_NY_0.018SBS_100_101.6-44.45-101.6 | Box_NY_0.018SBS_100_101.6-44.45-101.6 | NY-Analytical | 772 | Box Inspection | Width | Width | 45 |  | mm | FDA,1 | 43.7 | 45.2 | Box Inspection | Main_Plans
5AU252IN | 5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 125 |  |  | FDA,0 |  |  | Box Inspection | Main_Plans
5AU252IN | 5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Does not meet specifications (OOS) |  | in |  |  |  | Box Inspection | Main_Plans
5AU252IN | 5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.019 |  | in | FDA,3 | 0.019 | 0.021 | Box Inspection | Main_Plans
5AU252IN | 5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | TBC (Done) | [5 empty] | Box Inspection | Main_Plans
5AU252IN | 5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | RT (Done) | [5 empty] | Box Inspection | Main_Plans
5AU252IN | 5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [6 empty] | Box Inspection | Main_Plans
5AU252IN | 5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 772 | Box Inspection | Height | Height | 4.2812 |  | in | FDA,5 | 4.2812 | 4.3438 | Box Inspection | Main_Plans
5AU252IN | 5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 772 | Box Inspection | Length | Length | 4.125 |  | in | FDA,5 | 4.0625 | 4.125 | Box Inspection | Main_Plans
5AU252IN | 5AU252IN/Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | Box_NY_0.020TBC_123_4.09375-3.625-4.3125 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.625 |  | in | FDA,5 | 3.5938 | 3.6562 | Box Inspection | Main_Plans
5AU31450 | 5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | Box_NY_0.020SBS_169_6-3.125-5.9375 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5AU31450 | 5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | Box_NY_0.020SBS_169_6-3.125-5.9375 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.019 |  | in | FDA,3 | 0.019 | 0.021 | Box Inspection | Main_Plans
5AU31450 | 5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | Box_NY_0.020SBS_169_6-3.125-5.9375 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) | [5 empty] | Box Inspection | Main_Plans
5AU31450 | 5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | Box_NY_0.020SBS_169_6-3.125-5.9375 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Dispenser (Done) | [5 empty] | Box Inspection | Main_Plans
5AU31450 | 5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | Box_NY_0.020SBS_169_6-3.125-5.9375 | NY-Analytical | 772 | Box Inspection | Height | Height | 5.9375 |  | in | FDA,5 | 5.9062 | 5.9688 | Box Inspection | Main_Plans
5AU31450 | 5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | Box_NY_0.020SBS_169_6-3.125-5.9375 | NY-Analytical | 772 | Box Inspection | Length | Length | 6 |  | in | FDA,5 | 5.9688 | 6.0312 | Box Inspection | Main_Plans
5AU31450 | 5AU31450/Box_NY_0.020SBS_169_6-3.125-5.9375 | Box_NY_0.020SBS_169_6-3.125-5.9375 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.125 |  | in | FDA,5 | 3.0938 | 3.1562 | Box Inspection | Main_Plans
5AU44410 | 5AU44410/Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5AU44410 | 5AU44410/Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0155 |  | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Main_Plans
5AU44410 | 5AU44410/Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) | [5 empty] | Box Inspection | Main_Plans
5AU44410 | 5AU44410/Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Straight Tuck End w/1" glued double header (Done) | [5 empty] | Box Inspection | Main_Plans
5AU44410 | 5AU44410/Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Height | Height | 3.8125 |  | in | FDA,5 | 3.7812 | 3.8438 | Box Inspection | Main_Plans
5AU44410 | 5AU44410/Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Length | Length | 2.9375 |  | in | FDA,5 | 2.9062 | 2.9688 | Box Inspection | Main_Plans
5AU44410 | 5AU44410/Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | Box_NY_0.016SBS_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Width | Width | 1.875 |  | in | FDA,5 | 1.8438 | 1.9062 | Box Inspection | Main_Plans
5AU57314 | 5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | Box_NY_0.016CCSBS_100_3-2.5-4 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 200 |  |  | FDA,0 |  |  | Box Inspection | Main_Plans
5AU57314 | 5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | Box_NY_0.016CCSBS_100_3-2.5-4 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5AU57314 | 5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | Box_NY_0.016CCSBS_100_3-2.5-4 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0155 |  | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Main_Plans
5AU57314 | 5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | Box_NY_0.016CCSBS_100_3-2.5-4 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | Clay Coated SBS (Done) | [5 empty] | Box Inspection | Main_Plans
5AU57314 | 5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | Box_NY_0.016CCSBS_100_3-2.5-4 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) | [5 empty] | Box Inspection | Main_Plans
5AU57314 | 5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | Box_NY_0.016CCSBS_100_3-2.5-4 | NY-Analytical | 772 | Box Inspection | Height | Height | 4 |  | in | FDA,5 | 3.9688 | 4.0312 | Box Inspection | Main_Plans
5AU57314 | 5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | Box_NY_0.016CCSBS_100_3-2.5-4 | NY-Analytical | 772 | Box Inspection | Length | Length | 3 |  | in | FDA,5 | 2.9688 | 3.0312 | Box Inspection | Main_Plans
5AU57314 | 5AU57314/Box_NY_0.016CCSBS_100_3-2.5-4 | Box_NY_0.016CCSBS_100_3-2.5-4 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.5 |  | in | FDA,5 | 2.4688 | 2.5312 | Box Inspection | Main_Plans
5BU19101 | 5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | Box_NY_0.020SBS_129_5.75-3-3.875 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5BU19101 | 5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | Box_NY_0.020SBS_129_5.75-3-3.875 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0207 |  | in | FDA,3 | 0.019 | 0.021 | Box Inspection | Main_Plans
5BU19101 | 5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | Box_NY_0.020SBS_129_5.75-3-3.875 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) | [5 empty] | Box Inspection | Main_Plans
5BU19101 | 5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | Box_NY_0.020SBS_129_5.75-3-3.875 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Straight Tuck End French Style Dipsenser w/False Bottom | [5 empty] | Box Inspection | Main_Plans
5BU19101 | 5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | Box_NY_0.020SBS_129_5.75-3-3.875 | NY-Analytical | 772 | Box Inspection | Height | Height | 3.875 |  | in | FDA,5 | 3.8438 | 3.9062 | Box Inspection | Main_Plans
5BU19101 | 5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | Box_NY_0.020SBS_129_5.75-3-3.875 | NY-Analytical | 772 | Box Inspection | Length | Length | 5.75 |  | in | FDA,5 | 5.7188 | 5.7812 | Box Inspection | Main_Plans
5BU19101 | 5BU19101/Box_NY_0.020SBS_129_5.75-3-3.875 | Box_NY_0.020SBS_129_5.75-3-3.875 | NY-Analytical | 772 | Box Inspection | Width | Width | 3 |  | in | FDA,5 | 2.9688 | 3.0312 | Box Inspection | Main_Plans
5BU26595 | 5BU26595/Box_NY_0.024SBS_102_4.390625-4.390625-12 | Box_NY_0.024SBS_102_4.390625-4.390625-12 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom Dispenser Box | [5 empty] | Box Inspection | Main_Plans
5BU26595 | 5BU26595/Box_NY_0.024SBS_102_4.390625-4.390625-12 | Box_NY_0.024SBS_102_4.390625-4.390625-12 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5BU26595 | 5BU26595/Box_NY_0.024SBS_102_4.390625-4.390625-12 | Box_NY_0.024SBS_102_4.390625-4.390625-12 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.024 |  | in | FDA,3 | 0.023 | 0.025 | Box Inspection | Main_Plans
5BU26595 | 5BU26595/Box_NY_0.024SBS_102_4.390625-4.390625-12 | Box_NY_0.024SBS_102_4.390625-4.390625-12 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) | [5 empty] | Box Inspection | Main_Plans
5BU26595 | 5BU26595/Box_NY_0.024SBS_102_4.390625-4.390625-12 | Box_NY_0.024SBS_102_4.390625-4.390625-12 | NY-Analytical | 772 | Box Inspection | Height | Height | 12 |  | in | FDA,5 | 11.9688 | 12.0312 | Box Inspection | Main_Plans
5BU26595 | 5BU26595/Box_NY_0.024SBS_102_4.390625-4.390625-12 | Box_NY_0.024SBS_102_4.390625-4.390625-12 | NY-Analytical | 772 | Box Inspection | Length | Length | 4.41 |  | in | FDA,5 | 4.3594 | 4.4219 | Box Inspection | Main_Plans
5BU26595 | 5BU26595/Box_NY_0.024SBS_102_4.390625-4.390625-12 | Box_NY_0.024SBS_102_4.390625-4.390625-12 | NY-Analytical | 772 | Box Inspection | Width | Width | 4.39 |  | in | FDA,5 | 4.3594 | 4.4219 | Box Inspection | Main_Plans
5BU27500 | 5BU27500/Box_NY_0.024SBS_102_4.3906-4.3906-12 | Box_NY_0.024SBS_102_4.3906-4.3906-12 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom Dispenser Box | [5 empty] | Box Inspection | Main_Plans
5BU27500 | 5BU27500/Box_NY_0.024SBS_102_4.3906-4.3906-12 | Box_NY_0.024SBS_102_4.3906-4.3906-12 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5BU27500 | 5BU27500/Box_NY_0.024SBS_102_4.3906-4.3906-12 | Box_NY_0.024SBS_102_4.3906-4.3906-12 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.024 |  | in | FDA,3 | 0.023 | 0.025 | Box Inspection | Main_Plans
5BU27500 | 5BU27500/Box_NY_0.024SBS_102_4.3906-4.3906-12 | Box_NY_0.024SBS_102_4.3906-4.3906-12 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) | [5 empty] | Box Inspection | Main_Plans
5BU27500 | 5BU27500/Box_NY_0.024SBS_102_4.3906-4.3906-12 | Box_NY_0.024SBS_102_4.3906-4.3906-12 | NY-Analytical | 772 | Box Inspection | Height | Height | 12 |  | in | FDA,5 | 11.9688 | 12.0312 | Box Inspection | Main_Plans
5BU27500 | 5BU27500/Box_NY_0.024SBS_102_4.3906-4.3906-12 | Box_NY_0.024SBS_102_4.3906-4.3906-12 | NY-Analytical | 772 | Box Inspection | Length | Length | 4.4 |  | in | FDA,5 | 4.3594 | 4.4219 | Box Inspection | Main_Plans
5BU27500 | 5BU27500/Box_NY_0.024SBS_102_4.3906-4.3906-12 | Box_NY_0.024SBS_102_4.3906-4.3906-12 | NY-Analytical | 772 | Box Inspection | Width | Width | 4.41 |  | in | FDA,5 | 4.3594 | 4.4219 | Box Inspection | Main_Plans
5BU37230 | 5BU37230/Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 125 | [5 empty] | Box Inspection | Main_Plans
5BU37230 | 5BU37230/Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5BU37230 | 5BU37230/Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.017 |  | in | FDA,3 | 0.017 | 0.019 | Box Inspection | Main_Plans
5BU37230 | 5BU37230/Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) | [5 empty] | Box Inspection | Main_Plans
5BU37230 | 5BU37230/Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom - French Style (Done) | [5 empty] | Box Inspection | Main_Plans
5BU37230 | 5BU37230/Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | NY-Analytical | 772 | Box Inspection | Height | Height | 2.93 |  | in | FDA,5 | 2.9062 | 2.9688 | Box Inspection | Main_Plans
5BU37230 | 5BU37230/Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | NY-Analytical | 772 | Box Inspection | Length | Length | 3.84 |  | in | FDA,5 | 3.7812 | 3.8438 | Box Inspection | Main_Plans
5BU37230 | 5BU37230/Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | Box_NY_0.018SBS_101_3.8125-3.375-2.9375 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.38 |  | in | FDA,5 | 3.3438 | 3.4062 | Box Inspection | Main_Plans
5BU55510 | 5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5BU55510 | 5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0155 |  | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Main_Plans
5BU55510 | 5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | Bright Metalized Poly Board (Done) | [5 empty] | Box Inspection | Main_Plans
5BU55510 | 5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Straight Tuck End w/glued header (Done) | [5 empty] | Box Inspection | Main_Plans
5BU55510 | 5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Height | Height | 3.8125 |  | in | FDA,5 | 3.7812 | 3.8438 | Box Inspection | Main_Plans
5BU55510 | 5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Length | Length | 2.9375 |  | in | FDA,5 | 2.9062 | 2.9688 | Box Inspection | Main_Plans
5BU55510 | 5BU55510/Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | Box_NY_0.016BMPB_124_2.9375-1.875-3.8125 | NY-Analytical | 772 | Box Inspection | Width | Width | 1.875 |  | in | FDA,5 | 1.8438 | 1.9062 | Box Inspection | Main_Plans
5BU71230 | 5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | Box_NY_0.018SBS_122_4.6875-3-3.8125 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5BU71230 | 5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | Box_NY_0.018SBS_122_4.6875-3-3.8125 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.018 |  | in | FDA,3 | 0.017 | 0.019 | Box Inspection | Main_Plans
5BU71230 | 5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | Box_NY_0.018SBS_122_4.6875-3-3.8125 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) | [5 empty] | Box Inspection | Main_Plans
5BU71230 | 5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | Box_NY_0.018SBS_122_4.6875-3-3.8125 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Straight Tuck End (Done) | [5 empty] | Box Inspection | Main_Plans
5BU71230 | 5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | Box_NY_0.018SBS_122_4.6875-3-3.8125 | NY-Analytical | 772 | Box Inspection | Height | Height | 3.82 |  | in | FDA,5 | 3.7812 | 3.8438 | Box Inspection | Main_Plans
5BU71230 | 5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | Box_NY_0.018SBS_122_4.6875-3-3.8125 | NY-Analytical | 772 | Box Inspection | Length | Length | 4.7 |  | in | FDA,5 | 4.6562 | 4.7188 | Box Inspection | Main_Plans
5BU71230 | 5BU71230/Box_NY_0.018SBS_122_4.6875-3-3.8125 | Box_NY_0.018SBS_122_4.6875-3-3.8125 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.01 |  | in | FDA,5 | 2.9688 | 3.0312 | Box Inspection | Main_Plans
5CA284TR | 5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | Box_NY_0.018SBS_111_6-5.3125-3.6875 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5CA284TR | 5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | Box_NY_0.018SBS_111_6-5.3125-3.6875 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0175 |  | in | FDA,3 | 0.017 | 0.019 | Box Inspection | Main_Plans
5CA284TR | 5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | Box_NY_0.018SBS_111_6-5.3125-3.6875 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) | [5 empty] | Box Inspection | Main_Plans
5CA284TR | 5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | Box_NY_0.018SBS_111_6-5.3125-3.6875 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom French Style w/Perfs (Done) | [5 empty] | Box Inspection | Main_Plans
5CA284TR | 5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | Box_NY_0.018SBS_111_6-5.3125-3.6875 | NY-Analytical | 772 | Box Inspection | Height | Height | 3.6875 |  | in | FDA,5 | 3.6562 | 3.7188 | Box Inspection | Main_Plans
5CA284TR | 5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | Box_NY_0.018SBS_111_6-5.3125-3.6875 | NY-Analytical | 772 | Box Inspection | Length | Length | 6 |  | in | FDA,5 | 5.9688 | 6.0312 | Box Inspection | Main_Plans
5CA284TR | 5CA284TR/Box_NY_0.018SBS_111_6-5.3125-3.6875 | Box_NY_0.018SBS_111_6-5.3125-3.6875 | NY-Analytical | 772 | Box Inspection | Width | Width | 5.3125 |  | in | FDA,5 | 5.2812 | 5.3438 | Box Inspection | Main_Plans
5CB19600 | 5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | Box_NY_0.018SBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 125 |  |  | FDA,0 |  |  | Box Inspection | Main_Plans
5CB19600 | 5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | Box_NY_0.018SBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5CB19600 | 5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | Box_NY_0.018SBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.018 |  | in | FDA,3 | 0.017 | 0.019 | Box Inspection | Main_Plans
5CB19600 | 5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | Box_NY_0.018SBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) | [5 empty] | Box Inspection | Main_Plans
5CB19600 | 5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | Box_NY_0.018SBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) | [5 empty] | Box Inspection | Main_Plans
5CB19600 | 5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | Box_NY_0.018SBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Height | Height | 2.0938 |  | in | FDA,5 | 2.0625 | 2.125 | Box Inspection | Main_Plans
5CB19600 | 5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | Box_NY_0.018SBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Length | Length | 5.125 |  | in | FDA,5 | 5.0938 | 5.1562 | Box Inspection | Main_Plans
5CB19600 | 5CB19600/Box_NY_0.018SBS_100_5.125-2.25-2.09375 | Box_NY_0.018SBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.25 |  | in | FDA,5 | 2.2188 | 2.2812 | Box Inspection | Main_Plans
5CB51200 | 5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | UCSBS | [5 empty] | Box Inspection | Main_Plans
5CB51200 | 5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5CB51200 | 5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0155 |  | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Main_Plans
5CB51200 | 5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) | [5 empty] | Box Inspection | Main_Plans
5CB51200 | 5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | NY-Analytical | 772 | Box Inspection | Height | Height | 2.375 |  | in | FDA,5 | 2.3438 | 2.4062 | Box Inspection | Main_Plans
5CB51200 | 5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | NY-Analytical | 772 | Box Inspection | Length | Length | 6.125 |  | in | FDA,5 | 6.0938 | 6.1562 | Box Inspection | Main_Plans
5CB51200 | 5CB51200/Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | Box_NY_0.016UCSBS_100_6.125-2.75-2.375 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.75 |  | in | FDA,5 | 2.7188 | 2.7812 | Box Inspection | Main_Plans
5CC80100 | 5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | UCSBS | [5 empty] | Box Inspection | Main_Plans
5CC80100 | 5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | Box Inspection | Main_Plans
5CC80100 | 5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5CC80100 | 5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.015 |  | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Main_Plans
5CC80100 | 5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) | [5 empty] | Box Inspection | Main_Plans
5CC80100 | 5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Height | Height | 3.07 |  | in | FDA,5 | 3.0312 | 3.0938 | Box Inspection | Main_Plans
5CC80100 | 5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Length | Length | 5.07 |  | in | FDA,5 | 5.0312 | 5.0938 | Box Inspection | Main_Plans
5CC80100 | 5CC80100/Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | Box_NY_0.016UCSBS_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Width | Width | 1.73 |  | in | FDA,5 | 1.7188 | 1.7812 | Box Inspection | Main_Plans
5CD35185 | 5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | UCSBS | [5 empty] | Box Inspection | Main_Plans
5CD35185 | 5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 125 |  |  | FDA,0 |  |  | Box Inspection | Main_Plans
5CD35185 | 5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5CD35185 | 5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.016 |  | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Main_Plans
5CD35185 | 5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) | [5 empty] | Box Inspection | Main_Plans
5CD35185 | 5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 772 | Box Inspection | Height | Height | 6.078 |  | in | FDA,5 | 6.0312 | 6.0938 | Box Inspection | Main_Plans
5CD35185 | 5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 772 | Box Inspection | Length | Length | 6.5 |  | in | FDA,5 | 6.4688 | 6.5312 | Box Inspection | Main_Plans
5CD35185 | 5CD35185/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.312 |  | in | FDA,5 | 2.2812 | 2.3438 | Box Inspection | Main_Plans
5CS42050 | 5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | UCSBS | [5 empty] | Box Inspection | Main_Plans
5CS42050 | 5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | Box Inspection | Main_Plans
5CS42050 | 5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5CS42050 | 5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.016 |  | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Main_Plans
5CS42050 | 5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) | [5 empty] | Box Inspection | Main_Plans
5CS42050 | 5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | NY-Analytical | 772 | Box Inspection | Height | Height | 5.81 |  | in | FDA,5 | 5.7812 | 5.8438 | Box Inspection | Main_Plans
5CS42050 | 5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | NY-Analytical | 772 | Box Inspection | Length | Length | 6.51 |  | in | FDA,5 | 6.4688 | 6.5312 | Box Inspection | Main_Plans
5CS42050 | 5CS42050/Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | Box_NY_0.016UCSBS_100_6.5-3.3125-5.8125 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.32 |  | in | FDA,5 | 3.2812 | 3.3438 | Box Inspection | Main_Plans
5DA941TY | 5DA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5DA941TY | 5DA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | NY-Analytical | 772 | Box Inspection | Back Height | Back Height | 8.625 |  | in | FDA,5 | 8.5938 | 8.6562 | Box Inspection | Main_Plans
5DA941TY | 5DA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.024 |  | in | FDA,3 | 0.023 | 0.025 | Box Inspection | Main_Plans
5DA941TY | 5DA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | WCK (Done) | [5 empty] | Box Inspection | Main_Plans
5DA941TY | 5DA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom French Style w/Perfs (Done) | [5 empty] | Box Inspection | Main_Plans
5DA941TY | 5DA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | NY-Analytical | 772 | Box Inspection | Front Height | Front Height | 3.5 |  | in | FDA,5 | 3.4688 | 3.5312 | Box Inspection | Main_Plans
5DA941TY | 5DA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | NY-Analytical | 772 | Box Inspection | Length | Length | 5.24 |  | in | FDA,5 | 5.2188 | 5.2812 | Box Inspection | Main_Plans
5DA941TY | 5DA941TY/Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | Box_NY_0.024WCK_111_5.25-6.25-8.625/3.5 | NY-Analytical | 772 | Box Inspection | Width | Width | 6.26 |  | in | FDA,5 | 6.2188 | 6.2812 | Box Inspection | Main_Plans
5DH12824 | 5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | Box_NY_0.018SBS_100_4-2.125-6.125 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 30 |  |  | FDA,0 |  |  | Box Inspection | Main_Plans
5DH12824 | 5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | Box_NY_0.018SBS_100_4-2.125-6.125 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5DH12824 | 5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | Box_NY_0.018SBS_100_4-2.125-6.125 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.017 |  | in | FDA,3 | 0.017 | 0.019 | Box Inspection | Main_Plans
5DH12824 | 5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | Box_NY_0.018SBS_100_4-2.125-6.125 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) | [5 empty] | Box Inspection | Main_Plans
5DH12824 | 5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | Box_NY_0.018SBS_100_4-2.125-6.125 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) | [5 empty] | Box Inspection | Main_Plans
5DH12824 | 5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | Box_NY_0.018SBS_100_4-2.125-6.125 | NY-Analytical | 772 | Box Inspection | Height | Height | 6.125 |  | in | FDA,5 | 6.0938 | 6.1562 | Box Inspection | Main_Plans
5DH12824 | 5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | Box_NY_0.018SBS_100_4-2.125-6.125 | NY-Analytical | 772 | Box Inspection | Length | Length | 4 |  | in | FDA,5 | 3.9688 | 4.0312 | Box Inspection | Main_Plans
5DH12824 | 5DH12824/Box_NY_0.018SBS_100_4-2.125-6.125 | Box_NY_0.018SBS_100_4-2.125-6.125 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.125 |  | in | FDA,5 | 2.0938 | 2.1562 | Box Inspection | Main_Plans
5EH21130 | 5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | Box_NY_0.016SBS_100_4.75-2.375-3.25 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | Box Inspection | Main_Plans
5EH21130 | 5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | Box_NY_0.016SBS_100_4.75-2.375-3.25 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5EH21130 | 5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | Box_NY_0.016SBS_100_4.75-2.375-3.25 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.015 |  | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Main_Plans
5EH21130 | 5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | Box_NY_0.016SBS_100_4.75-2.375-3.25 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) | [5 empty] | Box Inspection | Main_Plans
5EH21130 | 5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | Box_NY_0.016SBS_100_4.75-2.375-3.25 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) | [5 empty] | Box Inspection | Main_Plans
5EH21130 | 5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | Box_NY_0.016SBS_100_4.75-2.375-3.25 | NY-Analytical | 772 | Box Inspection | Height | Height | 3.26 |  | in | FDA,5 | 3.2188 | 3.2812 | Box Inspection | Main_Plans
5EH21130 | 5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | Box_NY_0.016SBS_100_4.75-2.375-3.25 | NY-Analytical | 772 | Box Inspection | Length | Length | 4.75 |  | in | FDA,5 | 4.7188 | 4.7812 | Box Inspection | Main_Plans
5EH21130 | 5EH21130/Box_NY_0.016SBS_100_4.75-2.375-3.25 | Box_NY_0.016SBS_100_4.75-2.375-3.25 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.36 |  | in | FDA,5 | 2.3438 | 2.4062 | Certificate of Analysis | Main_Plans
5FD41900 | 5FD41900/Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | Box_NY_0.016UCSBS_100_6.5-2.3125-6.0625 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Box Inspection | Main_Plans
5FH07050 | 5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5FH07050 | 5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.02 |  | in | FDA,3 | 0.019 | 0.021 | Box Inspection | Main_Plans
5FH07050 | 5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) | [5 empty] | Box Inspection | Main_Plans
5FH07050 | 5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Straight Tuck End w/Glued Internal Dividers&15/16" Dummy Platform Bottom (Done) | [5 empty] | Box Inspection | Main_Plans
5FH07050 | 5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | NY-Analytical | 772 | Box Inspection | Height | Height | 6 |  | in | FDA,5 | 5.9844 | 6.0469 | Box Inspection | Main_Plans
5FH07050 | 5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | NY-Analytical | 772 | Box Inspection | Length | Length | 6.14 |  | in | FDA,5 | 6.0938 | 6.1562 | Box Inspection | Main_Plans
5FH07050 | 5FH07050/Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | Box_NY_0.020SBS_125_6.125-2.4375-6.015625 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.44 |  | in | FDA,5 | 2.4062 | 2.4688 | Box Inspection | Main_Plans
5GB10300 | 5GB10300/Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | UCSBS | [5 empty] | Box Inspection | Main_Plans
5GB10300 | 5GB10300/Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5GB10300 | 5GB10300/Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.015 |  | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Main_Plans
5GB10300 | 5GB10300/Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) | [5 empty] | Box Inspection | Main_Plans
5GB10300 | 5GB10300/Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | NY-Analytical | 772 | Box Inspection | Height | Height | 2.375 |  | in | FDA,5 | 2.3438 | 2.4062 | Box Inspection | Main_Plans
5GB10300 | 5GB10300/Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | NY-Analytical | 772 | Box Inspection | Length | Length | 6.125 |  | in | FDA,5 | 6.0938 | 6.1562 | Box Inspection | Main_Plans
5GB10300 | 5GB10300/Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | Box_NY_0.016UCSBS_100_6.125-2.750-2.375 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.75 |  | in | FDA,5 | 2.7188 | 2.7812 | Box Inspection | Main_Plans
5HC12400 | 5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | UCSBS | [5 empty] | Box Inspection | Main_Plans
5HC12400 | 5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 20 |  |  | FDA,0 |  |  | Box Inspection | Main_Plans
5HC12400 | 5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5HC12400 | 5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0169 |  | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Main_Plans
5HC12400 | 5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) | [5 empty] | Box Inspection | Main_Plans
5HC12400 | 5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | NY-Analytical | 772 | Box Inspection | Height | Height | 3.0625 |  | in | FDA,5 | 3.0312 | 3.0938 | Box Inspection | Main_Plans
5HC12400 | 5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | NY-Analytical | 772 | Box Inspection | Length | Length | 5.0625 |  | in | FDA,5 | 5.0312 | 5.0938 | Box Inspection | Main_Plans
5HC12400 | 5HC12400/Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | Box_NY_0.016UCSBS_100_5.0625-4.25-3.0625 | NY-Analytical | 772 | Box Inspection | Width | Width | 4.25 |  | in | FDA,5 | 4.2188 | 4.2812 | Box Inspection | Main_Plans
5HC22324 | 5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5HC22324 | 5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0198 |  | in | FDA,3 | 0.019 | 0.021 | Box Inspection | Main_Plans
5HC22324 | 5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | WCCN (Done) | [5 empty] | Box Inspection | Main_Plans
5HC22324 | 5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) | [5 empty] | Box Inspection | Main_Plans
5HC22324 | 5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Height | Height | 3.0625 |  | in | FDA,5 | 3.0312 | 3.0938 | Box Inspection | Main_Plans
5HC22324 | 5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Length | Length | 5.0625 |  | in | FDA,5 | 5.0312 | 5.0938 | Box Inspection | Main_Plans
5HC22324 | 5HC22324/Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | Box_NY_0.020WCCN_100_5.0625-1.75-3.0625 | NY-Analytical | 772 | Box Inspection | Width | Width | 1.75 |  | in | FDA,5 | 1.7188 | 1.7812 | Box Inspection | Main_Plans
5JB16400 | 5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | Box_NY_0.016UCSBS_100_4-3.125-2.375 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | UCSBS | [5 empty] | Box Inspection | Main_Plans
5JB16400 | 5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | Box_NY_0.016UCSBS_100_4-3.125-2.375 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5JB16400 | 5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | Box_NY_0.016UCSBS_100_4-3.125-2.375 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.015 |  | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Main_Plans
5JB16400 | 5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | Box_NY_0.016UCSBS_100_4-3.125-2.375 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) | [5 empty] | Box Inspection | Main_Plans
5JB16400 | 5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | Box_NY_0.016UCSBS_100_4-3.125-2.375 | NY-Analytical | 772 | Box Inspection | Height | Height | 2.375 |  | in | FDA,5 | 2.3438 | 2.4062 | Box Inspection | Main_Plans
5JB16400 | 5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | Box_NY_0.016UCSBS_100_4-3.125-2.375 | NY-Analytical | 772 | Box Inspection | Length | Length | 4 |  | in | FDA,5 | 3.9688 | 4.0312 | Box Inspection | Main_Plans
5JB16400 | 5JB16400/Box_NY_0.016UCSBS_100_4-3.125-2.375 | Box_NY_0.016UCSBS_100_4-3.125-2.375 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.125 |  | in | FDA,5 | 3.0938 | 3.1562 | Box Inspection | Main_Plans
5LB60302 | 5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | Box_NY_0.016UCSBS_100_6.125-4-2.375 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | UCSBS | [5 empty] | Box Inspection | Main_Plans
5LB60302 | 5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | Box_NY_0.016UCSBS_100_6.125-4-2.375 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | Box Inspection | Main_Plans
5LB60302 | 5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | Box_NY_0.016UCSBS_100_6.125-4-2.375 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5LB60302 | 5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | Box_NY_0.016UCSBS_100_6.125-4-2.375 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.016 |  | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Main_Plans
5LB60302 | 5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | Box_NY_0.016UCSBS_100_6.125-4-2.375 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) | [5 empty] | Box Inspection | Main_Plans
5LB60302 | 5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | Box_NY_0.016UCSBS_100_6.125-4-2.375 | NY-Analytical | 772 | Box Inspection | Height | Height | 2.375 |  | in | FDA,5 | 2.3438 | 2.4062 | Box Inspection | Main_Plans
5LB60302 | 5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | Box_NY_0.016UCSBS_100_6.125-4-2.375 | NY-Analytical | 772 | Box Inspection | Length | Length | 6.0938 |  | in | FDA,5 | 6.0938 | 6.1562 | Box Inspection | Main_Plans
5LB60302 | 5LB60302/Box_NY_0.016UCSBS_100_6.125-4-2.375 | Box_NY_0.016UCSBS_100_6.125-4-2.375 | NY-Analytical | 772 | Box Inspection | Width | Width | 4 |  | in | FDA,5 | 3.9688 | 4.0312 | Box Inspection | Main_Plans
5NB40600 | 5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | UCSBS | [5 empty] | Box Inspection | Main_Plans
5NB40600 | 5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 125 |  |  | FDA,0 |  |  | Box Inspection | Main_Plans
5NB40600 | 5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5NB40600 | 5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0155 |  | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Main_Plans
5NB40600 | 5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) | [5 empty] | Box Inspection | Main_Plans
5NB40600 | 5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Height | Height | 2.07 |  | in | FDA,5 | 2.0625 | 2.125 | Box Inspection | Main_Plans
5NB40600 | 5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Length | Length | 5.13 |  | in | FDA,5 | 5.0938 | 5.1562 | Box Inspection | Main_Plans
5NB40600 | 5NB40600/Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | Box_NY_0.016UCSBS_100_5.125-2.25-2.09375 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.24 |  | in | FDA,5 | 2.2188 | 2.2812 | Box Inspection | Main_Plans
5NU87250 | 5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | Box_NY_0.024SBS_127_4.375-4.375-12 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom Dispenser Box | [5 empty] | Box Inspection | Main_Plans
5NU87250 | 5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | Box_NY_0.024SBS_127_4.375-4.375-12 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5NU87250 | 5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | Box_NY_0.024SBS_127_4.375-4.375-12 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.024 |  | in | FDA,3 | 0.023 | 0.025 | Box Inspection | Main_Plans
5NU87250 | 5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | Box_NY_0.024SBS_127_4.375-4.375-12 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) | [5 empty] | Box Inspection | Main_Plans
5NU87250 | 5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | Box_NY_0.024SBS_127_4.375-4.375-12 | NY-Analytical | 772 | Box Inspection | Height | Height | 12 |  | in | FDA,5 | 11.9688 | 12.0312 | Box Inspection | Main_Plans
5NU87250 | 5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | Box_NY_0.024SBS_127_4.375-4.375-12 | NY-Analytical | 772 | Box Inspection | Length | Length | 4.39 |  | in | FDA,5 | 4.3438 | 4.4062 | Box Inspection | Main_Plans
5NU87250 | 5NU87250/Box_NY_0.024SBS_127_4.375-4.375-12 | Box_NY_0.024SBS_127_4.375-4.375-12 | NY-Analytical | 772 | Box Inspection | Width | Width | 4.39 |  | in | FDA,5 | 4.3438 | 4.4062 | Box Inspection | Main_Plans
5OD125PB | 5OD125PB/Box_80A | Box_80A | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5OD125PB | 5OD125PB/Box_80A | Box_80A | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0005 |  | in | FDA,4 | 0.0005 | 0.0025 | Box Inspection | Main_Plans
5OD125PB | 5OD125PB/Box_80A | Box_80A | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | LDPE (Done) | [5 empty] | Box Inspection | Main_Plans
5OD125PB | 5OD125PB/Box_80A | Box_80A | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Poly Bag w/fold over top on 4” wicketts (Done) | [5 empty] | Box Inspection | Main_Plans
5OD125PB | 5OD125PB/Box_80A | Box_80A | NY-Analytical | 772 | Box Inspection | Length | Length | 6.5 |  | in | FDA,5 | 6.0625 | 6.5 | Box Inspection | Main_Plans
5OD125PB | 5OD125PB/Box_80A | Box_80A | NY-Analytical | 772 | Box Inspection | Outside Flap | Outside Flap | 2.25 |  | in | FDA,5 | 2.2188 | 2.2812 | Box Inspection | Main_Plans
5OD125PB | 5OD125PB/Box_80A | Box_80A | NY-Analytical | 772 | Box Inspection | Width | Width | 8.25 |  | in | FDA,5 | 8.25 | 8.5 | Box Inspection | Main_Plans
5OD239PD | 5OD239PD/Box_51A | Box_51A | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5OD239PD | 5OD239PD/Box_51A | Box_51A | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0175 |  | in | FDA,3 |  |  | Box Inspection | Main_Plans
5OD239PD | 5OD239PD/Box_51A | Box_51A | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | Chip Board (Done) | [5 empty] | Box Inspection | Main_Plans
5OD239PD | 5OD239PD/Box_51A | Box_51A | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Layer Pad (Done) | [5 empty] | Box Inspection | Main_Plans
5OD239PD | 5OD239PD/Box_51A | Box_51A | NY-Analytical | 772 | Box Inspection | Length | Length | 15.375 |  | in | FDA,5 | 15.3438 | 15.4062 | Box Inspection | Main_Plans
5OD239PD | 5OD239PD/Box_51A | Box_51A | NY-Analytical | 772 | Box Inspection | Width | Width | 5.5 |  | in | FDA,5 | 5.4688 | 5.5312 | Box Inspection | Main_Plans
5OH382PD | 5OH382PD/Box_50D | Box_50D | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5OH382PD | 5OH382PD/Box_50D | Box_50D | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0018 |  | in | FDA,3 |  |  | Box Inspection | Main_Plans
5OH382PD | 5OH382PD/Box_50D | Box_50D | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | Any Board (Done) | [5 empty] | Box Inspection | Main_Plans
5OH382PD | 5OH382PD/Box_50D | Box_50D | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Cut Pad (Done) | [5 empty] | Box Inspection | Main_Plans
5OH382PD | 5OH382PD/Box_50D | Box_50D | NY-Analytical | 772 | Box Inspection | Length | Length | 11.5938 |  | in | FDA,5 | 11.5938 | 11.6562 | Box Inspection | Main_Plans
5OH382PD | 5OH382PD/Box_50D | Box_50D | NY-Analytical | 772 | Box Inspection | Width | Width | 9.4688 |  | in | FDA,5 | 9.4688 | 9.5312 | Box Inspection | Main_Plans
5OP067BT | 5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 50 |  |  | FDA,0 |  |  | Box Inspection | Main_Plans
5OP067BT | 5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5OP067BT | 5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.016 |  | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Main_Plans
5OP067BT | 5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS C1S (Done) | [5 empty] | Box Inspection | Main_Plans
5OP067BT | 5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Open Top/Tuck Bottom (Done) | [5 empty] | Box Inspection | Main_Plans
5OP067BT | 5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Height | Height | 2.44 |  | in | FDA,5 | 2.4062 | 2.4688 | Box Inspection | Main_Plans
5OP067BT | 5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Length | Length | 3.2 |  | in | FDA,5 | 3.1488 | 3.2113 | Box Inspection | Main_Plans
5OP067BT | 5OP067BT/Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | Box_NY_0.016SBSC1S_300_3.18-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.19 |  | in | FDA,5 | 2.1562 | 2.2188 | Box Inspection | Main_Plans
5OP073BT | 5OP073BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | Box Inspection | Main_Plans
5OP073BT | 5OP073BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5OP073BT | 5OP073BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.016 |  | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Main_Plans
5OP073BT | 5OP073BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS C1S (Done) | [5 empty] | Box Inspection | Main_Plans
5OP073BT | 5OP073BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Open Top/Tuck Bottom (Done) | [5 empty] | Box Inspection | Main_Plans
5OP073BT | 5OP073BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Height | Height | 2.42 |  | in | FDA,5 | 2.4062 | 2.4688 | Box Inspection | Main_Plans
5OP073BT | 5OP073BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Length | Length | 3.2 |  | in | FDA,5 | 3.1488 | 3.2113 | Box Inspection | Main_Plans
5OP073BT | 5OP073BT/Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | Box_NY_0.016SBSC1S_300_3.125-2.1875-2.4375 | NY-Analytical | 772 | Box Inspection | Width | Width | 2.17 |  | in | FDA,5 | 2.1562 | 2.2188 | Box Inspection | Main_Plans
5OP092BT | 5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 315 |  |  | FDA,0 |  |  | Box Inspection | Main_Plans
5OP092BT | 5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5OP092BT | 5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.016 |  | in | FDA,3 | 0.015 | 0.017 | Box Inspection | Main_Plans
5OP092BT | 5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS C1S (Done) | [5 empty] | Box Inspection | Main_Plans
5OP092BT | 5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Open Top/Tuck Bottom (Done) | [5 empty] | Box Inspection | Main_Plans
5OP092BT | 5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | NY-Analytical | 772 | Box Inspection | Height | Height | 2.6875 |  | in | FDA,5 | 2.6562 | 2.7188 | Box Inspection | Main_Plans
5OP092BT | 5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | NY-Analytical | 772 | Box Inspection | Length | Length | 3.1875 |  | in | FDA,5 | 3.1488 | 3.2113 | Box Inspection | Main_Plans
5OP092BT | 5OP092BT/Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | Box_NY_0.016SBSC1S_300_3.18-1.4375-2.6875 | NY-Analytical | 772 | Box Inspection | Width | Width | 1.4375 |  | in | FDA,5 | 1.4062 | 1.4688 | Box Inspection | Main_Plans
5OP092PB | 5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | Box_NY_LDPE_407_13.5-6.5-6-4 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | Box Inspection | Main_Plans
5OP092PB | 5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | Box_NY_LDPE_407_13.5-6.5-6-4 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5OP092PB | 5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | Box_NY_LDPE_407_13.5-6.5-6-4 | NY-Analytical | 772 | Box Inspection | Back Flip | Back Flip | 5.75 |  | in | FDA,5 | 5.6875 | 6.3125 | Box Inspection | Main_Plans
5OP092PB | 5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | Box_NY_LDPE_407_13.5-6.5-6-4 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | LDPE (Done) | [5 empty] | Box Inspection | Main_Plans
5OP092PB | 5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | Box_NY_LDPE_407_13.5-6.5-6-4 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Poly Bag w/fold over top on 2” wicketts (Done) | [5 empty] | Box Inspection | Main_Plans
5OP092PB | 5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | Box_NY_LDPE_407_13.5-6.5-6-4 | NY-Analytical | 772 | Box Inspection | Gussett | Gussett | 4 |  | in | FDA,5 | 3.6875 | 4.3125 | Box Inspection | Main_Plans
5OP092PB | 5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | Box_NY_LDPE_407_13.5-6.5-6-4 | NY-Analytical | 772 | Box Inspection | Height | Height | 6.5 |  | in | FDA,5 | 6.1875 | 6.8125 | Box Inspection | Main_Plans
5OP092PB | 5OP092PB/Box_NY_LDPE_407_13.5-6.5-6-4 | Box_NY_LDPE_407_13.5-6.5-6-4 | NY-Analytical | 772 | Box Inspection | Width | Width | 13.1875 |  | in | FDA,5 | 13.1875 | 13.8125 | Box Inspection | Main_Plans
5OP308PD | 5OP308PD/Box_50C | Box_50C | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5OP308PD | 5OP308PD/Box_50C | Box_50C | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.02 |  | in | FDA,3 | 0.018 | 0.02 | Box Inspection | Main_Plans
5OP308PD | 5OP308PD/Box_50C | Box_50C | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | Chip Board (Done) | [5 empty] | Box Inspection | Main_Plans
5OP308PD | 5OP308PD/Box_50C | Box_50C | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Cut Pad (Done) | [5 empty] | Box Inspection | Main_Plans
5OP308PD | 5OP308PD/Box_50C | Box_50C | NY-Analytical | 772 | Box Inspection | Length | Length | 7 |  | in | FDA,5 | 6.9688 | 7.0312 | Box Inspection | Main_Plans
5OP308PD | 5OP308PD/Box_50C | Box_50C | NY-Analytical | 772 | Box Inspection | Width | Width | 12.28 |  | in | FDA,5 | 12.2188 | 12.2812 | Box Inspection | Main_Plans
5PB114IS | 5PB114IS/Box_NY_NA50#_402_2.5-3.25 | Box_NY_NA50#_402_2.5-3.25 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 500 |  |  | FDA,0 |  |  | Box Inspection | Main_Plans
5PB114IS | 5PB114IS/Box_NY_NA50#_402_2.5-3.25 | Box_NY_NA50#_402_2.5-3.25 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5PB114IS | 5PB114IS/Box_NY_NA50#_402_2.5-3.25 | Box_NY_NA50#_402_2.5-3.25 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | 50# White Paper Stock (Done) | [5 empty] | Box Inspection | Main_Plans
5PB114IS | 5PB114IS/Box_NY_NA50#_402_2.5-3.25 | Box_NY_NA50#_402_2.5-3.25 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Insert (Done) | [5 empty] | Box Inspection | Main_Plans
5PB114IS | 5PB114IS/Box_NY_NA50#_402_2.5-3.25 | Box_NY_NA50#_402_2.5-3.25 | NY-Analytical | 772 | Box Inspection | Length | Length | 2.5 |  | in | FDA,5 | 2.4688 | 2.5312 | Box Inspection | Main_Plans
5PB114IS | 5PB114IS/Box_NY_NA50#_402_2.5-3.25 | Box_NY_NA50#_402_2.5-3.25 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.25 |  | in | FDA,5 | 3.2188 | 3.2812 | Box Inspection | Main_Plans
5PS40750 | 5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | Box_NY_0.018SBS_100_7.125-3-6.625 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5PS40750 | 5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | Box_NY_0.018SBS_100_7.125-3-6.625 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.0175 |  | in | FDA,3 | 0.017 | 0.019 | Box Inspection | Main_Plans
5PS40750 | 5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | Box_NY_0.018SBS_100_7.125-3-6.625 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) | [5 empty] | Box Inspection | Main_Plans
5PS40750 | 5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | Box_NY_0.018SBS_100_7.125-3-6.625 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) | [5 empty] | Box Inspection | Main_Plans
5PS40750 | 5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | Box_NY_0.018SBS_100_7.125-3-6.625 | NY-Analytical | 772 | Box Inspection | Height | Height | 6.625 |  | in | FDA,5 | 6.5938 | 6.6562 | Box Inspection | Main_Plans
5PS40750 | 5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | Box_NY_0.018SBS_100_7.125-3-6.625 | NY-Analytical | 772 | Box Inspection | Length | Length | 7.125 |  | in | FDA,5 | 7.0938 | 7.1562 | Box Inspection | Main_Plans
5PS40750 | 5PS40750/Box_NY_0.018SBS_100_7.125-3-6.625 | Box_NY_0.018SBS_100_7.125-3-6.625 | NY-Analytical | 772 | Box Inspection | Width | Width | 3 |  | in | FDA,5 | 2.9688 | 3.0312 | Box Inspection | Main_Plans
5QS273IS | 5QS273IS/Box_NY_NA50#_402_5-3.25 | Box_NY_NA50#_402_5-3.25 | NY-Analytical | 772 | Box Inspection | AQL Sample Size | AQL Sample Size | 500 |  |  | FDA,0 |  |  | Box Inspection | Main_Plans
5QS273IS | 5QS273IS/Box_NY_NA50#_402_5-3.25 | Box_NY_NA50#_402_5-3.25 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5QS273IS | 5QS273IS/Box_NY_NA50#_402_5-3.25 | Box_NY_NA50#_402_5-3.25 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | 50# White Paper Stock (Done) | [5 empty] | Box Inspection | Main_Plans
5QS273IS | 5QS273IS/Box_NY_NA50#_402_5-3.25 | Box_NY_NA50#_402_5-3.25 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Insert (Done) | [5 empty] | Box Inspection | Main_Plans
5QS273IS | 5QS273IS/Box_NY_NA50#_402_5-3.25 | Box_NY_NA50#_402_5-3.25 | NY-Analytical | 772 | Box Inspection | Length | Length | 5 |  | in | FDA,5 | 4.9688 | 5.0312 | Box Inspection | Main_Plans
5QS273IS | 5QS273IS/Box_NY_NA50#_402_5-3.25 | Box_NY_NA50#_402_5-3.25 | NY-Analytical | 772 | Box Inspection | Width | Width | 3.2188 |  | in | FDA,5 | 3.2188 | 3.2812 | Box Inspection | Main_Plans
5QS41950 | 5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | Box_NY_0.018SBS_100_6.625-4.125-8.125 | NY-Analytical | 772 | Box Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Meets specifications (Done) |  | in |  |  |  | Box Inspection | Main_Plans
5QS41950 | 5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | Box_NY_0.018SBS_100_6.625-4.125-8.125 | NY-Analytical | 772 | Box Inspection | Caliper | Caliper | 0.017 |  | in | FDA,3 | 0.017 | 0.019 | Box Inspection | Main_Plans
5QS41950 | 5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | Box_NY_0.018SBS_100_6.625-4.125-8.125 | NY-Analytical | 772 | Box Inspection | Carton Material | Carton Material | SBS (Done) | [5 empty] | Box Inspection | Main_Plans
5QS41950 | 5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | Box_NY_0.018SBS_100_6.625-4.125-8.125 | NY-Analytical | 772 | Box Inspection | Carton Style | Carton Style | Tuck Top / Auto Bottom (Done) | [5 empty] | Box Inspection | Main_Plans
5QS41950 | 5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | Box_NY_0.018SBS_100_6.625-4.125-8.125 | NY-Analytical | 772 | Box Inspection | Height | Height | 8.125 |  | in | FDA,5 | 8.0938 | 8.1562 | Box Inspection | Main_Plans
5QS41950 | 5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | Box_NY_0.018SBS_100_6.625-4.125-8.125 | NY-Analytical | 772 | Box Inspection | Length | Length | 6.625 |  | in | FDA,5 | 6.5938 | 6.6562 | Box Inspection | Main_Plans
5QS41950 | 5QS41950/Box_NY_0.018SBS_100_6.625-4.125-8.125 | Box_NY_0.018SBS_100_6.625-4.125-8.125 | NY-Analytical | 772 | Box Inspection | Width | Width | 4.125 |  | in | FDA,5 | 4.0938 | 4.1562 | Box Inspection | Main_Plans
6AA34724 | 6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 20 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6AA34724 | 6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AA34724 | 6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6AA34724 | 6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AA34724 | 6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6.1875 |  | in | FDA,4 | 6.0625 | 6.4375 | TM 141 / Visual | Main_Plans
6AA34724 | 6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 8.4375 |  | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual | Main_Plans
6AA34724 | 6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AA34724 | 6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AA34724 | 6AA34724/SHIP_NY_32K_100_8.25-7-6.25 | SHIP_NY_32K_100_8.25-7-6.25 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.1875 |  | in | FDA,4 | 6.8125 | 7.1875 | TM 141 / Visual | Main_Plans
6AA617S96 | 6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | SHIP_NY_44K_100_12.75-11.625-7.875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AA617S96 | 6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | SHIP_NY_44K_100_12.75-11.625-7.875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | Main_Plans
6AA617S96 | 6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | SHIP_NY_44K_100_12.75-11.625-7.875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AA617S96 | 6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | SHIP_NY_44K_100_12.75-11.625-7.875 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7.75 |  | in | FDA,4 | 7.6875 | 8.0625 | TM 141 / Visual | Main_Plans
6AA617S96 | 6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | SHIP_NY_44K_100_12.75-11.625-7.875 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 12.75 |  | in | FDA,4 | 12.5625 | 12.9375 | TM 141 / Visual | Main_Plans
6AA617S96 | 6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | SHIP_NY_44K_100_12.75-11.625-7.875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AA617S96 | 6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | SHIP_NY_44K_100_12.75-11.625-7.875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AA617S96 | 6AA617S96/SHIP_NY_44K_100_12.75-11.625-7.875 | SHIP_NY_44K_100_12.75-11.625-7.875 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 11.5 |  | in | FDA,4 | 11.4375 | 11.8125 | TM 141 / Visual | Main_Plans
6AA632TY | 6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | SHIP_NY_44K_231_18.5-7.375-14 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AA632TY | 6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | SHIP_NY_44K_231_18.5-7.375-14 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | Main_Plans
6AA632TY | 6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | SHIP_NY_44K_231_18.5-7.375-14 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AA632TY | 6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | SHIP_NY_44K_231_18.5-7.375-14 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 14 |  | in | FDA,4 | 13.875 | 14.125 | TM 141 / Visual | Main_Plans
6AA632TY | 6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | SHIP_NY_44K_231_18.5-7.375-14 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 18.5 |  | in | FDA,4 | 18.375 | 18.625 | TM 141 / Visual | Main_Plans
6AA632TY | 6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | SHIP_NY_44K_231_18.5-7.375-14 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AA632TY | 6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | SHIP_NY_44K_231_18.5-7.375-14 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | D/C (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AA632TY | 6AA632TY/SHIP_NY_44K_231_18.5-7.375-14 | SHIP_NY_44K_231_18.5-7.375-14 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.5 |  | in | FDA,4 | 7.25 | 7.5 | TM 141 / Visual | Main_Plans
6AA709S48 | 6AA709S48/SHIP_NY_40K_100_15-11-6.125 | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 200 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6AA709S48 | 6AA709S48/SHIP_NY_40K_100_15-11-6.125 | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AA709S48 | 6AA709S48/SHIP_NY_40K_100_15-11-6.125 | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 40 |  | Lbs./In. | FDA,0 | 40 |  | TM 141 / Visual | Main_Plans
6AA709S48 | 6AA709S48/SHIP_NY_40K_100_15-11-6.125 | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AA709S48 | 6AA709S48/SHIP_NY_40K_100_15-11-6.125 | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6.125 |  | in | FDA,4 | 5.9375 | 6.3125 | TM 141 / Visual | Main_Plans
6AA709S48 | 6AA709S48/SHIP_NY_40K_100_15-11-6.125 | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15 |  | in | FDA,4 | 14.8125 | 15.1875 | TM 141 / Visual | Main_Plans
6AA709S48 | 6AA709S48/SHIP_NY_40K_100_15-11-6.125 | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AA709S48 | 6AA709S48/SHIP_NY_40K_100_15-11-6.125 | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AA709S48 | 6AA709S48/SHIP_NY_40K_100_15-11-6.125 | SHIP_NY_40K_100_15-11-6.125 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 11 |  | in | FDA,4 | 10.8125 | 11.1875 | TM 141 / Visual | Main_Plans
6AA73530 | 6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6AA73530 | 6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AA73530 | 6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6AA73530 | 6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AA73530 | 6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 8.25 |  | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual | Main_Plans
6AA73530 | 6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 10 |  | in | FDA,4 | 9.8125 | 10.1875 | TM 141 / Visual | Main_Plans
6AA73530 | 6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AA73530 | 6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AA73530 | 6AA73530/SHIP_NY_32K_100_10-8.375-8.25 | SHIP_NY_32K_100_10-8.375-8.25 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 8.375 |  | in | FDA,4 | 8.1875 | 8.5625 | TM 141 / Visual | Main_Plans
6AA76930 | 6AA76930/SHIP_NY_44K_100_9-9-4.75 | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6AA76930 | 6AA76930/SHIP_NY_44K_100_9-9-4.75 | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AA76930 | 6AA76930/SHIP_NY_44K_100_9-9-4.75 | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | Main_Plans
6AA76930 | 6AA76930/SHIP_NY_44K_100_9-9-4.75 | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AA76930 | 6AA76930/SHIP_NY_44K_100_9-9-4.75 | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 4.75 |  | in | FDA,4 | 4.625 | 4.875 | TM 141 / Visual | Main_Plans
6AA76930 | 6AA76930/SHIP_NY_44K_100_9-9-4.75 | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 9 |  | in | FDA,4 | 8.875 | 9.125 | TM 141 / Visual | Main_Plans
6AA76930 | 6AA76930/SHIP_NY_44K_100_9-9-4.75 | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AA76930 | 6AA76930/SHIP_NY_44K_100_9-9-4.75 | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AA76930 | 6AA76930/SHIP_NY_44K_100_9-9-4.75 | SHIP_NY_44K_100_9-9-4.75 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 9 |  | in | FDA,4 | 8.875 | 9.125 | TM 141 / Visual | Main_Plans
6AB16400 | 6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6AB16400 | 6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AB16400 | 6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6AB16400 | 6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AB16400 | 6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 5.25 |  | in | FDA,4 | 4.9375 | 5.3125 | TM 141 / Visual | Main_Plans
6AB16400 | 6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 16.125 |  | in | FDA,4 | 15.8125 | 16.1875 | TM 141 / Visual | Main_Plans
6AB16400 | 6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AB16400 | 6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AB16400 | 6AB16400/SHIP_NY_32K_100_16-4.25-5.125 | SHIP_NY_32K_100_16-4.25-5.125 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 4.25 |  | in | FDA,4 | 4.0625 | 4.4375 | TM 141 / Visual | Main_Plans
6AD11055 | 6AD11055/SHIP_NY_32K_100_10.625-6.375-11.25 | SHIP_NY_32K_100_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AD11055 | 6AD11055/SHIP_NY_32K_100_10.625-6.375-11.25 | SHIP_NY_32K_100_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6AD11055 | 6AD11055/SHIP_NY_32K_100_10.625-6.375-11.25 | SHIP_NY_32K_100_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AD11055 | 6AD11055/SHIP_NY_32K_100_10.625-6.375-11.25 | SHIP_NY_32K_100_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 11.375 |  | in | FDA,4 | 11.0625 | 11.4375 | TM 141 / Visual | Main_Plans
6AD11055 | 6AD11055/SHIP_NY_32K_100_10.625-6.375-11.25 | SHIP_NY_32K_100_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 10.625 |  | in | FDA,4 | 10.4375 | 10.8125 | TM 141 / Visual | Main_Plans
6AD11055 | 6AD11055/SHIP_NY_32K_100_10.625-6.375-11.25 | SHIP_NY_32K_100_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AD11055 | 6AD11055/SHIP_NY_32K_100_10.625-6.375-11.25 | SHIP_NY_32K_100_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AD11055 | 6AD11055/SHIP_NY_32K_100_10.625-6.375-11.25 | SHIP_NY_32K_100_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.375 |  | in | FDA,4 | 6.1875 | 6.5625 | TM 141 / Visual | Main_Plans
6AD41900 | 6AD41900/SHIP_NY_32K_100_13.25-12-6.125 | SHIP_NY_32K_100_13.25-12-6.125 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AD41900 | 6AD41900/SHIP_NY_32K_100_13.25-12-6.125 | SHIP_NY_32K_100_13.25-12-6.125 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6AD41900 | 6AD41900/SHIP_NY_32K_100_13.25-12-6.125 | SHIP_NY_32K_100_13.25-12-6.125 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AD41900 | 6AD41900/SHIP_NY_32K_100_13.25-12-6.125 | SHIP_NY_32K_100_13.25-12-6.125 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6.25 |  | in | FDA,4 | 5.9375 | 6.3125 | TM 141 / Visual | Main_Plans
6AD41900 | 6AD41900/SHIP_NY_32K_100_13.25-12-6.125 | SHIP_NY_32K_100_13.25-12-6.125 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 13.25 |  | in | FDA,4 | 13.0625 | 13.4375 | TM 141 / Visual | Main_Plans
6AD41900 | 6AD41900/SHIP_NY_32K_100_13.25-12-6.125 | SHIP_NY_32K_100_13.25-12-6.125 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AD41900 | 6AD41900/SHIP_NY_32K_100_13.25-12-6.125 | SHIP_NY_32K_100_13.25-12-6.125 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AD41900 | 6AD41900/SHIP_NY_32K_100_13.25-12-6.125 | SHIP_NY_32K_100_13.25-12-6.125 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 11.875 |  | in | FDA,4 | 11.8125 | 12.1875 | TM 141 / Visual | Main_Plans
6AH58195 | 6AH58195/SHIP_NY_32K_100_13.4375-6.25-9.6875 | SHIP_NY_32K_100_13.4375-6.25-9.6875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AH58195 | 6AH58195/SHIP_NY_32K_100_13.4375-6.25-9.6875 | SHIP_NY_32K_100_13.4375-6.25-9.6875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6AH58195 | 6AH58195/SHIP_NY_32K_100_13.4375-6.25-9.6875 | SHIP_NY_32K_100_13.4375-6.25-9.6875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AH58195 | 6AH58195/SHIP_NY_32K_100_13.4375-6.25-9.6875 | SHIP_NY_32K_100_13.4375-6.25-9.6875 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 9.75 |  | in | FDA,4 | 9.5 | 9.875 | TM 141 / Visual | Main_Plans
6AH58195 | 6AH58195/SHIP_NY_32K_100_13.4375-6.25-9.6875 | SHIP_NY_32K_100_13.4375-6.25-9.6875 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 13.4375 |  | in | FDA,4 | 13.25 | 13.625 | TM 141 / Visual | Main_Plans
6AH58195 | 6AH58195/SHIP_NY_32K_100_13.4375-6.25-9.6875 | SHIP_NY_32K_100_13.4375-6.25-9.6875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AH58195 | 6AH58195/SHIP_NY_32K_100_13.4375-6.25-9.6875 | SHIP_NY_32K_100_13.4375-6.25-9.6875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AH58195 | 6AH58195/SHIP_NY_32K_100_13.4375-6.25-9.6875 | SHIP_NY_32K_100_13.4375-6.25-9.6875 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.25 |  | in | FDA,4 | 6.0625 | 6.4375 | TM 141 / Visual | Main_Plans
6AJ19025 | 6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6AJ19025 | 6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AJ19025 | 6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6AJ19025 | 6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AJ19025 | 6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6.75 |  | in | FDA,4 | 6.5625 | 6.9375 | TM 141 / Visual | Main_Plans
6AJ19025 | 6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 12 |  | in | FDA,4 | 11.8125 | 12.1875 | TM 141 / Visual | Main_Plans
6AJ19025 | 6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AJ19025 | 6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AJ19025 | 6AJ19025/SHIP_NY_32K_100_12-7.25-6.75 | SHIP_NY_32K_100_12-7.25-6.75 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.25 |  | in | FDA,4 | 7.0625 | 7.4375 | TM 141 / Visual | Main_Plans
6AM30472 | 6AM30472/SHIP_NY_44K_100_15.5-9-12 | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6AM30472 | 6AM30472/SHIP_NY_44K_100_15.5-9-12 | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AM30472 | 6AM30472/SHIP_NY_44K_100_15.5-9-12 | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | Main_Plans
6AM30472 | 6AM30472/SHIP_NY_44K_100_15.5-9-12 | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AM30472 | 6AM30472/SHIP_NY_44K_100_15.5-9-12 | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 12.125 |  | in | FDA,4 | 11.8125 | 12.1875 | TM 141 / Visual | Main_Plans
6AM30472 | 6AM30472/SHIP_NY_44K_100_15.5-9-12 | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.375 |  | in | FDA,4 | 15.3125 | 15.6875 | TM 141 / Visual | Main_Plans
6AM30472 | 6AM30472/SHIP_NY_44K_100_15.5-9-12 | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AM30472 | 6AM30472/SHIP_NY_44K_100_15.5-9-12 | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AM30472 | 6AM30472/SHIP_NY_44K_100_15.5-9-12 | SHIP_NY_44K_100_15.5-9-12 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 8.875 |  | in | FDA,4 | 8.8125 | 9.1875 | TM 141 / Visual | Main_Plans
6AP03707 | 6AP03707/SHIP_NY_32K_100_13-7.3125-4 | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6AP03707 | 6AP03707/SHIP_NY_32K_100_13-7.3125-4 | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP03707 | 6AP03707/SHIP_NY_32K_100_13-7.3125-4 | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6AP03707 | 6AP03707/SHIP_NY_32K_100_13-7.3125-4 | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP03707 | 6AP03707/SHIP_NY_32K_100_13-7.3125-4 | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 4.1875 |  | in | FDA,4 | 3.8125 | 4.1875 | TM 141 / Visual | Main_Plans
6AP03707 | 6AP03707/SHIP_NY_32K_100_13-7.3125-4 | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 13.062 |  | in | FDA,4 | 12.8125 | 13.1875 | TM 141 / Visual | Main_Plans
6AP03707 | 6AP03707/SHIP_NY_32K_100_13-7.3125-4 | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP03707 | 6AP03707/SHIP_NY_32K_100_13-7.3125-4 | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP03707 | 6AP03707/SHIP_NY_32K_100_13-7.3125-4 | SHIP_NY_32K_100_13-7.3125-4 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.5 |  | in | FDA,4 | 7.125 | 7.5 | TM 141 / Visual | Main_Plans
6AP06528 | 6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | SHIP_NY_32K_100_13.25-7.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP06528 | 6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | SHIP_NY_32K_100_13.25-7.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6AP06528 | 6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | SHIP_NY_32K_100_13.25-7.75-5.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP06528 | 6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | SHIP_NY_32K_100_13.25-7.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 5.9375 |  | in | FDA,4 | 5.5625 | 5.9375 | TM 141 / Visual | Main_Plans
6AP06528 | 6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | SHIP_NY_32K_100_13.25-7.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 13.25 |  | in | FDA,4 | 13.0625 | 13.4375 | TM 141 / Visual | Main_Plans
6AP06528 | 6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | SHIP_NY_32K_100_13.25-7.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP06528 | 6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | SHIP_NY_32K_100_13.25-7.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP06528 | 6AP06528/SHIP_NY_32K_100_13.25-7.75-5.75 | SHIP_NY_32K_100_13.25-7.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.75 |  | in | FDA,4 | 7.5625 | 7.9375 | TM 141 / Visual | Main_Plans
6AP06550 | 6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 20 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6AP06550 | 6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP06550 | 6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6AP06550 | 6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP06550 | 6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 10.375 |  | in | FDA,4 | 10.375 | 10.625 | TM 141 / Visual | Main_Plans
6AP06550 | 6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 17.5 |  | in | FDA,4 | 17.25 | 17.5 | TM 141 / Visual | Main_Plans
6AP06550 | 6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP06550 | 6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP06550 | 6AP06550/SHIP_NY_32K_100_17.375-12.0625-10.5 | SHIP_NY_32K_100_17.375-12.0625-10.5 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 11.9375 |  | in | FDA,4 | 11.9375 | 12.1875 | TM 141 / Visual | Main_Plans
6AP06728 | 6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6AP06728 | 6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP06728 | 6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6AP06728 | 6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP06728 | 6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6 |  | in | FDA,4 | 5.875 | 6.25 | TM 141 / Visual | Main_Plans
6AP06728 | 6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 7 |  | in | FDA,4 | 6.9375 | 7.3125 | TM 141 / Visual | Main_Plans
6AP06728 | 6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP06728 | 6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP06728 | 6AP06728/SHIP_NY_32K_100_7.125-7-6.0625 | SHIP_NY_32K_100_7.125-7-6.0625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.125 |  | in | FDA,4 | 6.8125 | 7.1875 | TM 141 / Visual | Main_Plans
6AP09215 | 6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6AP09215 | 6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP09215 | 6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6AP09215 | 6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP09215 | 6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 5.75 |  | in | FDA,4 | 5.5625 | 5.9375 | TM 141 / Visual | Main_Plans
6AP09215 | 6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 13.5 |  | in | FDA,4 | 13.25 | 13.625 | TM 141 / Visual | Main_Plans
6AP09215 | 6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP09215 | 6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP09215 | 6AP09215/SHIP_NY_32K_100_13.4375-9.25-5.75 | SHIP_NY_32K_100_13.4375-9.25-5.75 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 9.25 |  | in | FDA,4 | 9.0625 | 9.4375 | TM 141 / Visual | Main_Plans
6AP09384 | 6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6AP09384 | 6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP09384 | 6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6AP09384 | 6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP09384 | 6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 3.8125 |  | in | FDA,4 | 3.5625 | 3.8125 | TM 141 / Visual | Main_Plans
6AP09384 | 6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 8.5 |  | in | FDA,4 | 8.375 | 8.625 | TM 141 / Visual | Main_Plans
6AP09384 | 6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP09384 | 6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP09384 | 6AP09384/SHIP_NY_32K_100_8.5-5.75-3.6875 | SHIP_NY_32K_100_8.5-5.75-3.6875 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 5.75 |  | in | FDA,4 | 5.625 | 5.875 | TM 141 / Visual | Main_Plans
6AP11325 | 6AP11325/SHIP_NY_44K_100_10.625-9.75-7.75 | SHIP_NY_44K_100_10.625-9.75-7.75 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP11325 | 6AP11325/SHIP_NY_44K_100_10.625-9.75-7.75 | SHIP_NY_44K_100_10.625-9.75-7.75 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | Main_Plans
6AP11325 | 6AP11325/SHIP_NY_44K_100_10.625-9.75-7.75 | SHIP_NY_44K_100_10.625-9.75-7.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP11325 | 6AP11325/SHIP_NY_44K_100_10.625-9.75-7.75 | SHIP_NY_44K_100_10.625-9.75-7.75 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7.75 |  | in | FDA,4 | 7.5625 | 7.9375 | TM 141 / Visual | Main_Plans
6AP11325 | 6AP11325/SHIP_NY_44K_100_10.625-9.75-7.75 | SHIP_NY_44K_100_10.625-9.75-7.75 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 10.625 |  | in | FDA,4 | 10.4375 | 10.8125 | TM 141 / Visual | Main_Plans
6AP11325 | 6AP11325/SHIP_NY_44K_100_10.625-9.75-7.75 | SHIP_NY_44K_100_10.625-9.75-7.75 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP11325 | 6AP11325/SHIP_NY_44K_100_10.625-9.75-7.75 | SHIP_NY_44K_100_10.625-9.75-7.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP11325 | 6AP11325/SHIP_NY_44K_100_10.625-9.75-7.75 | SHIP_NY_44K_100_10.625-9.75-7.75 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 9.75 |  | in | FDA,4 | 9.5625 | 9.9375 | TM 141 / Visual | Main_Plans
6AP13472 | 6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | SHIP_NY_32K_100_16.3125-12.25-7.5625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP13472 | 6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | SHIP_NY_32K_100_16.3125-12.25-7.5625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6AP13472 | 6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | SHIP_NY_32K_100_16.3125-12.25-7.5625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP13472 | 6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | SHIP_NY_32K_100_16.3125-12.25-7.5625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7.75 |  | in | FDA,4 | 7.375 | 7.75 | TM 141 / Visual | Main_Plans
6AP13472 | 6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | SHIP_NY_32K_100_16.3125-12.25-7.5625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 16.25 |  | in | FDA,4 | 16.125 | 16.5 | TM 141 / Visual | Main_Plans
6AP13472 | 6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | SHIP_NY_32K_100_16.3125-12.25-7.5625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP13472 | 6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | SHIP_NY_32K_100_16.3125-12.25-7.5625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP13472 | 6AP13472/SHIP_NY_32K_100_16.3125-12.25-7.5625 | SHIP_NY_32K_100_16.3125-12.25-7.5625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 12.25 |  | in | FDA,4 | 12.0625 | 12.4375 | TM 141 / Visual | Main_Plans
6AP15984 | 6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | SHIP_NY_44K_100_15.25-10.375-7.6875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP15984 | 6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | SHIP_NY_44K_100_15.25-10.375-7.6875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | Main_Plans
6AP15984 | 6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | SHIP_NY_44K_100_15.25-10.375-7.6875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP15984 | 6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | SHIP_NY_44K_100_15.25-10.375-7.6875 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7.75 |  | in | FDA,4 | 7.5 | 7.875 | TM 141 / Visual | Main_Plans
6AP15984 | 6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | SHIP_NY_44K_100_15.25-10.375-7.6875 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.125 |  | in | FDA,4 | 15.0625 | 15.4375 | TM 141 / Visual | Main_Plans
6AP15984 | 6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | SHIP_NY_44K_100_15.25-10.375-7.6875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP15984 | 6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | SHIP_NY_44K_100_15.25-10.375-7.6875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP15984 | 6AP15984/SHIP_NY_44K_100_15.25-10.375-7.6875 | SHIP_NY_44K_100_15.25-10.375-7.6875 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 10.5 |  | in | FDA,4 | 10.1875 | 10.5625 | TM 141 / Visual | Main_Plans
6AP22325 | 6AP22325/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Shipper Inspection | Main_Plans
6AP22907 | 6AP22907/SHIP_NY_32K_100_8.5625-7.8125-6.4375 | SHIP_NY_32K_100_8.5625-7.8125-6.4375 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6AP22907 | 6AP22907/SHIP_NY_32K_100_8.5625-7.8125-6.4375 | SHIP_NY_32K_100_8.5625-7.8125-6.4375 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP22907 | 6AP22907/SHIP_NY_32K_100_8.5625-7.8125-6.4375 | SHIP_NY_32K_100_8.5625-7.8125-6.4375 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6AP22907 | 6AP22907/SHIP_NY_32K_100_8.5625-7.8125-6.4375 | SHIP_NY_32K_100_8.5625-7.8125-6.4375 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP22907 | 6AP22907/SHIP_NY_32K_100_8.5625-7.8125-6.4375 | SHIP_NY_32K_100_8.5625-7.8125-6.4375 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6.5 |  | in | FDA,4 | 6.25 | 6.625 | TM 141 / Visual | Main_Plans
6AP22907 | 6AP22907/SHIP_NY_32K_100_8.5625-7.8125-6.4375 | SHIP_NY_32K_100_8.5625-7.8125-6.4375 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 8.5 |  | in | FDA,4 | 8.375 | 8.75 | TM 141 / Visual | Main_Plans
6AP22907 | 6AP22907/SHIP_NY_32K_100_8.5625-7.8125-6.4375 | SHIP_NY_32K_100_8.5625-7.8125-6.4375 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP22907 | 6AP22907/SHIP_NY_32K_100_8.5625-7.8125-6.4375 | SHIP_NY_32K_100_8.5625-7.8125-6.4375 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP22907 | 6AP22907/SHIP_NY_32K_100_8.5625-7.8125-6.4375 | SHIP_NY_32K_100_8.5625-7.8125-6.4375 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.75 |  | in | FDA,4 | 7.625 | 8 | TM 141 / Visual | Main_Plans
6AP23025 | 6AP23025/SHIP_NY_32K_100_11.25-5.8125-4.3125 | SHIP_NY_32K_100_11.25-5.8125-4.3125 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6AP23025 | 6AP23025/SHIP_NY_32K_100_11.25-5.8125-4.3125 | SHIP_NY_32K_100_11.25-5.8125-4.3125 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP23025 | 6AP23025/SHIP_NY_32K_100_11.25-5.8125-4.3125 | SHIP_NY_32K_100_11.25-5.8125-4.3125 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6AP23025 | 6AP23025/SHIP_NY_32K_100_11.25-5.8125-4.3125 | SHIP_NY_32K_100_11.25-5.8125-4.3125 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP23025 | 6AP23025/SHIP_NY_32K_100_11.25-5.8125-4.3125 | SHIP_NY_32K_100_11.25-5.8125-4.3125 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 4.375 |  | in | FDA,4 | 4.125 | 4.5 | TM 141 / Visual | Main_Plans
6AP23025 | 6AP23025/SHIP_NY_32K_100_11.25-5.8125-4.3125 | SHIP_NY_32K_100_11.25-5.8125-4.3125 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.25 |  | in | FDA,4 | 11.0625 | 11.4375 | TM 141 / Visual | Main_Plans
6AP23025 | 6AP23025/SHIP_NY_32K_100_11.25-5.8125-4.3125 | SHIP_NY_32K_100_11.25-5.8125-4.3125 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP23025 | 6AP23025/SHIP_NY_32K_100_11.25-5.8125-4.3125 | SHIP_NY_32K_100_11.25-5.8125-4.3125 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP23025 | 6AP23025/SHIP_NY_32K_100_11.25-5.8125-4.3125 | SHIP_NY_32K_100_11.25-5.8125-4.3125 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 5.75 |  | in | FDA,4 | 5.625 | 6 | TM 141 / Visual | Main_Plans
6AP24284 | 6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP24284 | 6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 40 |  | Lbs./In. | FDA,0 | 40 |  | TM 141 / Visual | Main_Plans
6AP24284 | 6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP24284 | 6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 9.625 |  | in | FDA,4 | 9.4375 | 9.8125 | TM 141 / Visual | Main_Plans
6AP24284 | 6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.375 |  | in | FDA,4 | 15.1875 | 15.5625 | TM 141 / Visual | Main_Plans
6AP24284 | 6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP24284 | 6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP24284 | 6AP24284/SHIP_NY_40K_100_15.375-10.5625-9.625 | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 10.5625 |  | in | FDA,4 | 10.375 | 10.75 | TM 141 / Visual | Main_Plans
6AP24372 | 6AP24372/SHIP_NY_44K_100_15.25-10.375-15.4375 | SHIP_NY_44K_100_15.25-10.375-15.4375 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP24372 | 6AP24372/SHIP_NY_44K_100_15.25-10.375-15.4375 | SHIP_NY_44K_100_15.25-10.375-15.4375 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | Main_Plans
6AP24372 | 6AP24372/SHIP_NY_44K_100_15.25-10.375-15.4375 | SHIP_NY_44K_100_15.25-10.375-15.4375 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP24372 | 6AP24372/SHIP_NY_44K_100_15.25-10.375-15.4375 | SHIP_NY_44K_100_15.25-10.375-15.4375 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 15.4375 |  | in | FDA,4 | 15.25 | 15.625 | TM 141 / Visual | Main_Plans
6AP24372 | 6AP24372/SHIP_NY_44K_100_15.25-10.375-15.4375 | SHIP_NY_44K_100_15.25-10.375-15.4375 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.25 |  | in | FDA,4 | 15.0625 | 15.4375 | TM 141 / Visual | Main_Plans
6AP24372 | 6AP24372/SHIP_NY_44K_100_15.25-10.375-15.4375 | SHIP_NY_44K_100_15.25-10.375-15.4375 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP24372 | 6AP24372/SHIP_NY_44K_100_15.25-10.375-15.4375 | SHIP_NY_44K_100_15.25-10.375-15.4375 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP24372 | 6AP24372/SHIP_NY_44K_100_15.25-10.375-15.4375 | SHIP_NY_44K_100_15.25-10.375-15.4375 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 10.375 |  | in | FDA,4 | 10.1875 | 10.5625 | TM 141 / Visual | Main_Plans
6AP24684 | 6AP24684/SHIP_NY_40K_100_15.375-10.5625-9.625 | SHIP_NY_40K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6AP27284 | 6AP27284/SHIP_NY_44K_100_15.375-10.5625-9.625 | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP27284 | 6AP27284/SHIP_NY_44K_100_15.375-10.5625-9.625 | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | Main_Plans
6AP27284 | 6AP27284/SHIP_NY_44K_100_15.375-10.5625-9.625 | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP27284 | 6AP27284/SHIP_NY_44K_100_15.375-10.5625-9.625 | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 9.625 |  | in | FDA,4 | 9.4375 | 9.8125 | TM 141 / Visual | Main_Plans
6AP27284 | 6AP27284/SHIP_NY_44K_100_15.375-10.5625-9.625 | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.25 |  | in | FDA,4 | 15.1875 | 15.5625 | TM 141 / Visual | Main_Plans
6AP27284 | 6AP27284/SHIP_NY_44K_100_15.375-10.5625-9.625 | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP27284 | 6AP27284/SHIP_NY_44K_100_15.375-10.5625-9.625 | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP27284 | 6AP27284/SHIP_NY_44K_100_15.375-10.5625-9.625 | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 10.437 |  | in | FDA,4 | 10.375 | 10.75 | TM 141 / Visual | Main_Plans
6AP30895 | 6AP30895/SHIP_32_1BI | SHIP_32_1BI | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP30895 | 6AP30895/SHIP_32_1BI | SHIP_32_1BI | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6AP30895 | 6AP30895/SHIP_32_1BI | SHIP_32_1BI | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP30895 | 6AP30895/SHIP_32_1BI | SHIP_32_1BI | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7.875 |  | in | FDA,4 | 7.6875 | 8.0625 | TM 141 / Visual | Main_Plans
6AP30895 | 6AP30895/SHIP_32_1BI | SHIP_32_1BI | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 12.5 |  | in | FDA,4 | 12.3125 | 12.6875 | TM 141 / Visual | Main_Plans
6AP30895 | 6AP30895/SHIP_32_1BI | SHIP_32_1BI | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP30895 | 6AP30895/SHIP_32_1BI | SHIP_32_1BI | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP30895 | 6AP30895/SHIP_32_1BI | SHIP_32_1BI | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.25 |  | in | FDA,4 | 7.0625 | 7.4375 | TM 141 / Visual | Main_Plans
6AP32825 | 6AP32825/SHIP_32_1P | SHIP_32_1P | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP32825 | 6AP32825/SHIP_32_1P | SHIP_32_1P | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6AP32825 | 6AP32825/SHIP_32_1P | SHIP_32_1P | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP32825 | 6AP32825/SHIP_32_1P | SHIP_32_1P | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 4.25 |  | in | FDA,4 | 4.125 | 4.5 | TM 141 / Visual | Main_Plans
6AP32825 | 6AP32825/SHIP_32_1P | SHIP_32_1P | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.1875 |  | in | FDA,4 | 11.0625 | 11.4375 | TM 141 / Visual | Main_Plans
6AP32825 | 6AP32825/SHIP_32_1P | SHIP_32_1P | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP32825 | 6AP32825/SHIP_32_1P | SHIP_32_1P | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP32825 | 6AP32825/SHIP_32_1P | SHIP_32_1P | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6 |  | in | FDA,4 | 5.625 | 6 | TM 141 / Visual | Main_Plans
6AP80625 | 6AP80625/SHIP_NY_32K_100_11.25-5.8125-8.625 | SHIP_NY_32K_100_11.25-5.8125-8.625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6AP80625 | 6AP80625/SHIP_NY_32K_100_11.25-5.8125-8.625 | SHIP_NY_32K_100_11.25-5.8125-8.625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP80625 | 6AP80625/SHIP_NY_32K_100_11.25-5.8125-8.625 | SHIP_NY_32K_100_11.25-5.8125-8.625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6AP80625 | 6AP80625/SHIP_NY_32K_100_11.25-5.8125-8.625 | SHIP_NY_32K_100_11.25-5.8125-8.625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP80625 | 6AP80625/SHIP_NY_32K_100_11.25-5.8125-8.625 | SHIP_NY_32K_100_11.25-5.8125-8.625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 8.625 |  | in | FDA,4 | 8.4375 | 8.8125 | TM 141 / Visual | Main_Plans
6AP80625 | 6AP80625/SHIP_NY_32K_100_11.25-5.8125-8.625 | SHIP_NY_32K_100_11.25-5.8125-8.625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.25 |  | in | FDA,4 | 11.0625 | 11.4375 | TM 141 / Visual | Main_Plans
6AP80625 | 6AP80625/SHIP_NY_32K_100_11.25-5.8125-8.625 | SHIP_NY_32K_100_11.25-5.8125-8.625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP80625 | 6AP80625/SHIP_NY_32K_100_11.25-5.8125-8.625 | SHIP_NY_32K_100_11.25-5.8125-8.625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AP80625 | 6AP80625/SHIP_NY_32K_100_11.25-5.8125-8.625 | SHIP_NY_32K_100_11.25-5.8125-8.625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 5.8125 |  | in | FDA,4 | 5.625 | 6 | TM 141 / Visual | Main_Plans
6AS27350 | 6AS27350/SHIP_44_2O | SHIP_44_2O | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AS27350 | 6AS27350/SHIP_44_2O | SHIP_44_2O | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | Main_Plans
6AS27350 | 6AS27350/SHIP_44_2O | SHIP_44_2O | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AS27350 | 6AS27350/SHIP_44_2O | SHIP_44_2O | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 9 |  | in | FDA,4 | 8.8125 | 9.1875 | TM 141 / Visual | Main_Plans
6AS27350 | 6AS27350/SHIP_44_2O | SHIP_44_2O | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.375 |  | in | FDA,4 | 15.1875 | 15.5625 | TM 141 / Visual | Main_Plans
6AS27350 | 6AS27350/SHIP_44_2O | SHIP_44_2O | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AS27350 | 6AS27350/SHIP_44_2O | SHIP_44_2O | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AS27350 | 6AS27350/SHIP_44_2O | SHIP_44_2O | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 11.375 |  | in | FDA,4 | 11.1875 | 11.5625 | TM 141 / Visual | Main_Plans
6AS40750 | 6AS40750/SHIP_32_1AP | SHIP_32_1AP | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AS40750 | 6AS40750/SHIP_32_1AP | SHIP_32_1AP | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6AS40750 | 6AS40750/SHIP_32_1AP | SHIP_32_1AP | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AS40750 | 6AS40750/SHIP_32_1AP | SHIP_32_1AP | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 13.4375 |  | in | FDA,4 | 13.25 | 13.625 | TM 141 / Visual | Main_Plans
6AS40750 | 6AS40750/SHIP_32_1AP | SHIP_32_1AP | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.625 |  | in | FDA,4 | 15.4375 | 15.8125 | TM 141 / Visual | Main_Plans
6AS40750 | 6AS40750/SHIP_32_1AP | SHIP_32_1AP | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AS40750 | 6AS40750/SHIP_32_1AP | SHIP_32_1AP | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AS40750 | 6AS40750/SHIP_32_1AP | SHIP_32_1AP | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.25 |  | in | FDA,4 | 7.0625 | 7.4375 | TM 141 / Visual | Main_Plans
6AS41950 | 6AS41950/SHIP_40_1O | SHIP_40_1O | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AS41950 | 6AS41950/SHIP_40_1O | SHIP_40_1O | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 40 |  | Lbs./In. | FDA,0 | 40 |  | TM 141 / Visual | Main_Plans
6AS41950 | 6AS41950/SHIP_40_1O | SHIP_40_1O | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AS41950 | 6AS41950/SHIP_40_1O | SHIP_40_1O | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 8.3125 |  | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual | Main_Plans
6AS41950 | 6AS41950/SHIP_40_1O | SHIP_40_1O | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 21.125 |  | in | FDA,4 | 20.9375 | 21.3125 | TM 141 / Visual | Main_Plans
6AS41950 | 6AS41950/SHIP_40_1O | SHIP_40_1O | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AS41950 | 6AS41950/SHIP_40_1O | SHIP_40_1O | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AS41950 | 6AS41950/SHIP_40_1O | SHIP_40_1O | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 13.5625 |  | in | FDA,4 | 13.375 | 13.75 | TM 141 / Visual | Main_Plans
6AU12310 | 6AU12310/SHIP_32_1U | SHIP_32_1U | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AU12310 | 6AU12310/SHIP_32_1U | SHIP_32_1U | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6AU12310 | 6AU12310/SHIP_32_1U | SHIP_32_1U | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AU12310 | 6AU12310/SHIP_32_1U | SHIP_32_1U | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 8 |  | in | FDA,4 | 7.8125 | 8.1875 | TM 141 / Visual | Main_Plans
6AU12310 | 6AU12310/SHIP_32_1U | SHIP_32_1U | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 12.125 |  | in | FDA,4 | 11.8125 | 12.1875 | TM 141 / Visual | Main_Plans
6AU12310 | 6AU12310/SHIP_32_1U | SHIP_32_1U | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AU12310 | 6AU12310/SHIP_32_1U | SHIP_32_1U | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AU12310 | 6AU12310/SHIP_32_1U | SHIP_32_1U | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6 |  | in | FDA,4 | 5.9375 | 6.3125 | TM 141 / Visual | Main_Plans
6AU130TY | 6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6AU130TY | 6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AU130TY | 6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6AU130TY | 6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AU130TY | 6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 4.5 |  | in | FDA,4 | 4.375 | 4.625 | TM 141 / Visual | Main_Plans
6AU130TY | 6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 6.625 |  | in | FDA,4 | 6.375 | 6.625 | TM 141 / Visual | Main_Plans
6AU130TY | 6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | White (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AU130TY | 6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | Fold Over Tray w/header (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AU130TY | 6AU130TY/SHIP_NY_32White_203_6.5-9.625-4.5 | SHIP_NY_32White_203_6.5-9.625-4.5 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 9.625 |  | in | FDA,4 | 9.5 | 9.75 | TM 141 / Visual | Main_Plans
6AU25206 | 6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6AU25206 | 6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | No Go - Does not meet Specifications (OOS) | [5 empty] | TM 141 / Visual | Main_Plans
6AU25206 | 6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6AU25206 | 6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AU25206 | 6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 4.75 |  | in | FDA,4 | 4.375 | 4.75 | TM 141 / Visual | Main_Plans
6AU25206 | 6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.375 |  | in | FDA,4 | 11.0625 | 11.4375 | TM 141 / Visual | Main_Plans
6AU25206 | 6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AU25206 | 6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AU25206 | 6AU25206/SHIP_NY_32K_100_11.25-8.4375-4.5625 | SHIP_NY_32K_100_11.25-8.4375-4.5625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 8.5 |  | in | FDA,4 | 8.25 | 8.625 | TM 141 / Visual | Main_Plans
6AU53601 | 6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 50 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6AU53601 | 6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AU53601 | 6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6AU53601 | 6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AU53601 | 6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 8.5 |  | in | FDA,4 | 8.3125 | 8.6875 | TM 141 / Visual | Main_Plans
6AU53601 | 6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.875 |  | in | FDA,4 | 11.625 | 12 | TM 141 / Visual | Main_Plans
6AU53601 | 6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AU53601 | 6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AU53601 | 6AU53601/SHIP_NY_32K_100_11.8125-9.375-8.5 | SHIP_NY_32K_100_11.8125-9.375-8.5 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 9.5 |  | in | FDA,4 | 9.1875 | 9.5625 | TM 141 / Visual | Main_Plans
6AU56400 | 6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 50 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6AU56400 | 6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AU56400 | 6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6AU56400 | 6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AU56400 | 6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7.8125 |  | in | FDA,4 | 7.5005 | 7.8755 | TM 141 / Visual | Main_Plans
6AU56400 | 6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.75 |  | in | FDA,4 | 11.5625 | 11.9375 | TM 141 / Visual | Main_Plans
6AU56400 | 6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AU56400 | 6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AU56400 | 6AU56400/SHIP_NY_32K_100_11.75-9.625-7.688 | SHIP_NY_32K_100_11.75-9.625-7.688 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 9.625 |  | in | FDA,4 | 9.4375 | 9.8125 | TM 141 / Visual | Main_Plans
6AU57314 | 6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6AU57314 | 6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AU57314 | 6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6AU57314 | 6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AU57314 | 6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 8.0625 |  | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual | Main_Plans
6AU57314 | 6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 7.5625 |  | in | FDA,4 | 7.5625 | 7.9375 | TM 141 / Visual | Main_Plans
6AU57314 | 6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AU57314 | 6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AU57314 | 6AU57314/SHIP_NY_32K_100_7.75-6.25-8.25 | SHIP_NY_32K_100_7.75-6.25-8.25 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.375 |  | in | FDA,4 | 6.0625 | 6.4375 | TM 141 / Visual | Main_Plans
6AU71230 | 6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 50 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6AU71230 | 6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AU71230 | 6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6AU71230 | 6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AU71230 | 6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7.75 |  | in | FDA,4 | 7.5625 | 7.9375 | TM 141 / Visual | Main_Plans
6AU71230 | 6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 9.5625 |  | in | FDA,4 | 9.3755 | 9.7505 | TM 141 / Visual | Main_Plans
6AU71230 | 6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AU71230 | 6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6AU71230 | 6AU71230/SHIP_NY_32K_100_9.563-9.188-7.75 | SHIP_NY_32K_100_9.563-9.188-7.75 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 9.188 |  | in | FDA,4 | 9.0005 | 9.3755 | TM 141 / Visual | Main_Plans
6BA709SWN | 6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6BA709SWN | 6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BA709SWN | 6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 40 |  | Lbs./In. | FDA,0 | 40 |  | TM 141 / Visual | Main_Plans
6BA709SWN | 6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BA709SWN | 6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6.125 |  | in | FDA,4 | 5.9375 | 6.3125 | TM 141 / Visual | Main_Plans
6BA709SWN | 6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 14.75 |  | in | FDA,4 | 14.5625 | 14.9375 | TM 141 / Visual | Main_Plans
6BA709SWN | 6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BA709SWN | 6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BA709SWN | 6BA709SWN/SHIP_NY_40K_100_14.75-9.375-6.125 | SHIP_NY_40K_100_14.75-9.375-6.125 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 9.375 |  | in | FDA,4 | 9.1875 | 9.5625 | TM 141 / Visual | Main_Plans
6BB10800 | 6BB10800/SHIP_32_1S | SHIP_32_1S | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BB10800 | 6BB10800/SHIP_32_1S | SHIP_32_1S | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6BB10800 | 6BB10800/SHIP_32_1S | SHIP_32_1S | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BB10800 | 6BB10800/SHIP_32_1S | SHIP_32_1S | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 5.3125 |  | in | FDA,4 | 5.125 | 5.5 | TM 141 / Visual | Main_Plans
6BB10800 | 6BB10800/SHIP_32_1S | SHIP_32_1S | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.875 |  | in | FDA,4 | 11.6875 | 12.0625 | TM 141 / Visual | Main_Plans
6BB10800 | 6BB10800/SHIP_32_1S | SHIP_32_1S | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BB10800 | 6BB10800/SHIP_32_1S | SHIP_32_1S | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BB10800 | 6BB10800/SHIP_32_1S | SHIP_32_1S | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.25 |  | in | FDA,4 | 6.0625 | 6.4375 | TM 141 / Visual | Main_Plans
6BD17385 | 6BD17385/SHIP_44_2AB | SHIP_44_2AB | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BD17385 | 6BD17385/SHIP_44_2AB | SHIP_44_2AB | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | Main_Plans
6BD17385 | 6BD17385/SHIP_44_2AB | SHIP_44_2AB | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BD17385 | 6BD17385/SHIP_44_2AB | SHIP_44_2AB | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6.5 |  | in | FDA,4 | 6.3125 | 6.6875 | TM 141 / Visual | Main_Plans
6BD17385 | 6BD17385/SHIP_44_2AB | SHIP_44_2AB | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 20 |  | in | FDA,4 | 19.8125 | 20.1875 | TM 141 / Visual | Main_Plans
6BD17385 | 6BD17385/SHIP_44_2AB | SHIP_44_2AB | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BD17385 | 6BD17385/SHIP_44_2AB | SHIP_44_2AB | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BD17385 | 6BD17385/SHIP_44_2AB | SHIP_44_2AB | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 12 |  | in | FDA,4 | 11.8125 | 12.1875 | TM 141 / Visual | Main_Plans
6BP21195 | 6BP21195/SHIP_NY_44K_100_14.375-12.1875-9.25 | SHIP_NY_44K_100_14.375-12.1875-9.25 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BP21195 | 6BP21195/SHIP_NY_44K_100_14.375-12.1875-9.25 | SHIP_NY_44K_100_14.375-12.1875-9.25 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | Main_Plans
6BP21195 | 6BP21195/SHIP_NY_44K_100_14.375-12.1875-9.25 | SHIP_NY_44K_100_14.375-12.1875-9.25 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BP21195 | 6BP21195/SHIP_NY_44K_100_14.375-12.1875-9.25 | SHIP_NY_44K_100_14.375-12.1875-9.25 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 9.125 |  | in | FDA,4 | 9.0625 | 9.4375 | TM 141 / Visual | Main_Plans
6BP21195 | 6BP21195/SHIP_NY_44K_100_14.375-12.1875-9.25 | SHIP_NY_44K_100_14.375-12.1875-9.25 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 14.375 |  | in | FDA,4 | 14.1875 | 14.5625 | TM 141 / Visual | Main_Plans
6BP21195 | 6BP21195/SHIP_NY_44K_100_14.375-12.1875-9.25 | SHIP_NY_44K_100_14.375-12.1875-9.25 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BP21195 | 6BP21195/SHIP_NY_44K_100_14.375-12.1875-9.25 | SHIP_NY_44K_100_14.375-12.1875-9.25 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BP21195 | 6BP21195/SHIP_NY_44K_100_14.375-12.1875-9.25 | SHIP_NY_44K_100_14.375-12.1875-9.25 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 12.25 |  | in | FDA,4 | 12 | 12.375 | TM 141 / Visual | Main_Plans
6BP25372 | 6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BP25372 | 6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BP25372 | 6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | Main_Plans
6BP25372 | 6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BP25372 | 6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 15.25 |  | in | FDA,4 | 15 | 15.375 | TM 141 / Visual | Main_Plans
6BP25372 | 6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.125 |  | in | FDA,4 | 15 | 15.375 | TM 141 / Visual | Main_Plans
6BP25372 | 6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BP25372 | 6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BP25372 | 6BP25372/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 10.25 |  | in | FDA,4 | 10 | 10.375 | TM 141 / Visual | Main_Plans
6BP30895 | 6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | SHIP_NY_32K_100_12.5-7.25-7.875 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6BP30895 | 6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | SHIP_NY_32K_100_12.5-7.25-7.875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BP30895 | 6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | SHIP_NY_32K_100_12.5-7.25-7.875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6BP30895 | 6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | SHIP_NY_32K_100_12.5-7.25-7.875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BP30895 | 6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | SHIP_NY_32K_100_12.5-7.25-7.875 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7.6875 |  | in | FDA,4 | 7.6875 | 8.0625 | TM 141 / Visual | Main_Plans
6BP30895 | 6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | SHIP_NY_32K_100_12.5-7.25-7.875 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 12.375 |  | in | FDA,4 | 12.3125 | 12.6875 | TM 141 / Visual | Main_Plans
6BP30895 | 6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | SHIP_NY_32K_100_12.5-7.25-7.875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BP30895 | 6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | SHIP_NY_32K_100_12.5-7.25-7.875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BP30895 | 6BP30895/SHIP_NY_32K_100_12.5-7.25-7.875 | SHIP_NY_32K_100_12.5-7.25-7.875 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.125 |  | in | FDA,4 | 7.0625 | 7.4375 | TM 141 / Visual | Main_Plans
6BP54072 | 6BP54072/SHIP_44_2L | SHIP_44_2L | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BP54072 | 6BP54072/SHIP_44_2L | SHIP_44_2L | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | Main_Plans
6BP54072 | 6BP54072/SHIP_44_2L | SHIP_44_2L | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BP54072 | 6BP54072/SHIP_44_2L | SHIP_44_2L | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 15.4375 |  | in | FDA,4 | 15.25 | 15.625 | TM 141 / Visual | Main_Plans
6BP54072 | 6BP54072/SHIP_44_2L | SHIP_44_2L | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.25 |  | in | FDA,4 | 15.0625 | 15.4375 | TM 141 / Visual | Main_Plans
6BP54072 | 6BP54072/SHIP_44_2L | SHIP_44_2L | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BP54072 | 6BP54072/SHIP_44_2L | SHIP_44_2L | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BP54072 | 6BP54072/SHIP_44_2L | SHIP_44_2L | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 10.375 |  | in | FDA,4 | 10.1875 | 10.5625 | TM 141 / Visual | Main_Plans
6BU13195 | 6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6BU13195 | 6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BU13195 | 6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6BU13195 | 6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BU13195 | 6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 12.25 |  | in | FDA,4 | 12.0625 | 12.4375 | TM 141 / Visual | Main_Plans
6BU13195 | 6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 13.5 |  | in | FDA,4 | 13.4375 | 13.8125 | TM 141 / Visual | Main_Plans
6BU13195 | 6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BU13195 | 6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BU13195 | 6BU13195/SHIP_NY_32K_100_13.625-4.5-12.25 | SHIP_NY_32K_100_13.625-4.5-12.25 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 4.375 |  | in | FDA,4 | 4.3125 | 4.6875 | TM 141 / Visual | Main_Plans
6BU25206 | 6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 10 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6BU25206 | 6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BU25206 | 6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6BU25206 | 6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BU25206 | 6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 4.5 |  | in | FDA,4 | 4.4375 | 4.8125 | TM 141 / Visual | Main_Plans
6BU25206 | 6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.3125 |  | in | FDA,4 | 11.125 | 11.5 | TM 141 / Visual | Main_Plans
6BU25206 | 6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BU25206 | 6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6BU25206 | 6BU25206/SHIP_NY_32K_100_11.3125-8.4375-4.625 | SHIP_NY_32K_100_11.3125-8.4375-4.625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 8.25 |  | in | FDA,4 | 8.25 | 8.625 | TM 141 / Visual | Main_Plans
6CA972CK | 6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 32 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6CA972CK | 6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6CA972CK | 6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | Main_Plans
6CA972CK | 6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6CA972CK | 6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 10.25 |  | in | FDA,4 | 10.0625 | 10.4375 | TM 141 / Visual | Main_Plans
6CA972CK | 6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 18 |  | in | FDA,4 | 17.8125 | 18.1875 | TM 141 / Visual | Main_Plans
6CA972CK | 6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6CA972CK | 6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6CA972CK | 6CA972CK/SHIP_NY_44K_100_18-8.25-10.25 | SHIP_NY_44K_100_18-8.25-10.25 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 8.25 |  | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual | Main_Plans
6CJ21396 | 6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | SHIP_NY_44K_100_16.625-9.375-7 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6CJ21396 | 6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | SHIP_NY_44K_100_16.625-9.375-7 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | Main_Plans
6CJ21396 | 6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | SHIP_NY_44K_100_16.625-9.375-7 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6CJ21396 | 6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | SHIP_NY_44K_100_16.625-9.375-7 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7.1875 |  | in | FDA,4 | 6.8125 | 7.1875 | TM 141 / Visual | Main_Plans
6CJ21396 | 6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | SHIP_NY_44K_100_16.625-9.375-7 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 16.75 |  | in | FDA,4 | 16.4375 | 16.8125 | TM 141 / Visual | Main_Plans
6CJ21396 | 6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | SHIP_NY_44K_100_16.625-9.375-7 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6CJ21396 | 6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | SHIP_NY_44K_100_16.625-9.375-7 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6CJ21396 | 6CJ21396/SHIP_NY_44K_100_16.625-9.375-7 | SHIP_NY_44K_100_16.625-9.375-7 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 9.5 |  | in | FDA,4 | 9.1875 | 9.5625 | TM 141 / Visual | Main_Plans
6CP02724 | 6CP02724/SHIP_NY_32K_100_11-5.8125-4.0625 | SHIP_NY_32K_100_11-5.8125-4.0625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6CP02724 | 6CP02724/SHIP_NY_32K_100_11-5.8125-4.0625 | SHIP_NY_32K_100_11-5.8125-4.0625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6CP02724 | 6CP02724/SHIP_NY_32K_100_11-5.8125-4.0625 | SHIP_NY_32K_100_11-5.8125-4.0625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6CP02724 | 6CP02724/SHIP_NY_32K_100_11-5.8125-4.0625 | SHIP_NY_32K_100_11-5.8125-4.0625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 4.25 |  | in | FDA,4 | 3.875 | 4.25 | TM 141 / Visual | Main_Plans
6CP02724 | 6CP02724/SHIP_NY_32K_100_11-5.8125-4.0625 | SHIP_NY_32K_100_11-5.8125-4.0625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11 |  | in | FDA,4 | 10.8125 | 11.1875 | TM 141 / Visual | Main_Plans
6CP02724 | 6CP02724/SHIP_NY_32K_100_11-5.8125-4.0625 | SHIP_NY_32K_100_11-5.8125-4.0625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6CP02724 | 6CP02724/SHIP_NY_32K_100_11-5.8125-4.0625 | SHIP_NY_32K_100_11-5.8125-4.0625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6CP02724 | 6CP02724/SHIP_NY_32K_100_11-5.8125-4.0625 | SHIP_NY_32K_100_11-5.8125-4.0625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6 |  | in | FDA,4 | 5.625 | 6 | TM 141 / Visual | Main_Plans
6CP22884 | 6CP22884/SHIP_NY_44K_100_15.375-10.5625-9.625 | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [6 empty] | Shipper Inspection | Main_Plans
6DC22370 | 6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6DC22370 | 6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6DC22370 | 6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6DC22370 | 6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6DC22370 | 6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 10.375 |  | in | FDA,4 | 10.1875 | 10.5625 | TM 141 / Visual | Main_Plans
6DC22370 | 6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11 |  | in | FDA,4 | 10.9375 | 11.3125 | TM 141 / Visual | Main_Plans
6DC22370 | 6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6DC22370 | 6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6DC22370 | 6DC22370/SHIP_NY_32K_100_11.125-6.375-10.375 | SHIP_NY_32K_100_11.125-6.375-10.375 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.25 |  | in | FDA,4 | 6.1875 | 6.5625 | TM 141 / Visual | Main_Plans
6DD33333 | 6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | SHIP_NY_44K_100_21-12.5-10.5 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6DD33333 | 6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | SHIP_NY_44K_100_21-12.5-10.5 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | Main_Plans
6DD33333 | 6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | SHIP_NY_44K_100_21-12.5-10.5 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6DD33333 | 6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | SHIP_NY_44K_100_21-12.5-10.5 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 10.5 |  | in | FDA,4 | 10.3125 | 10.6875 | TM 141 / Visual | Main_Plans
6DD33333 | 6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | SHIP_NY_44K_100_21-12.5-10.5 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 21 |  | in | FDA,4 | 20.8125 | 21.1875 | TM 141 / Visual | Main_Plans
6DD33333 | 6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | SHIP_NY_44K_100_21-12.5-10.5 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6DD33333 | 6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | SHIP_NY_44K_100_21-12.5-10.5 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6DD33333 | 6DD33333/SHIP_NY_44K_100_21-12.5-10.5 | SHIP_NY_44K_100_21-12.5-10.5 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 12.625 |  | in | FDA,4 | 12.3125 | 12.6875 | TM 141 / Visual | Main_Plans
6DP63884 | 6DP63884/SHIP_NY_44K_100_15.375-10.5625-9.625 | SHIP_NY_44K_100_15.375-10.5625-9.625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6EP44584 | 6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 50 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6EP44584 | 6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6EP44584 | 6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6EP44584 | 6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6EP44584 | 6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 9.75 |  | in | FDA,4 | 9.4375 | 9.8125 | TM 141 / Visual | Main_Plans
6EP44584 | 6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.25 |  | in | FDA,4 | 15 | 15.375 | TM 141 / Visual | Main_Plans
6EP44584 | 6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6EP44584 | 6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6EP44584 | 6EP44584/SHIP_NY_32K_100_15.1875-10.1875-9.625 | SHIP_NY_32K_100_15.1875-10.1875-9.625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 10.25 |  | in | FDA,4 | 10 | 10.375 | TM 141 / Visual | Main_Plans
6FB71200 | 6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 50 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6FB71200 | 6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6FB71200 | 6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6FB71200 | 6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6FB71200 | 6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 5.125 |  | in | FDA,4 | 4.8125 | 5.1875 | TM 141 / Visual | Main_Plans
6FB71200 | 6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 14.75 |  | in | FDA,4 | 14.4375 | 14.8125 | TM 141 / Visual | Main_Plans
6FB71200 | 6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6FB71200 | 6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC w/ shyed flaps (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6FB71200 | 6FB71200/SHIP_NY_32K_102_14.625-6.375-5 | SHIP_NY_32K_102_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.344 |  | in | FDA,4 | 6.1875 | 6.5625 | TM 141 / Visual | Main_Plans
6FJ14108 | 6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6FJ14108 | 6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6FJ14108 | 6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | Main_Plans
6FJ14108 | 6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6FJ14108 | 6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7 |  | in | FDA,4 | 6.8125 | 7.1875 | TM 141 / Visual | Main_Plans
6FJ14108 | 6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 16.625 |  | in | FDA,4 | 16.4375 | 16.8125 | TM 141 / Visual | Main_Plans
6FJ14108 | 6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6FJ14108 | 6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6FJ14108 | 6FJ14108/SHIP_NY_44K_100_16.625-8.625-7 | SHIP_NY_44K_100_16.625-8.625-7 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 8.5 |  | in | FDA,4 | 8.4375 | 8.8125 | TM 141 / Visual | Main_Plans
6FJ14143 | 6FJ14143/SHIP_NY_44K_100_16.5-10.875-5.5 | SHIP_NY_44K_100_16.5-10.875-5.5 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6FJ14143 | 6FJ14143/SHIP_NY_44K_100_16.5-10.875-5.5 | SHIP_NY_44K_100_16.5-10.875-5.5 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | Main_Plans
6FJ14143 | 6FJ14143/SHIP_NY_44K_100_16.5-10.875-5.5 | SHIP_NY_44K_100_16.5-10.875-5.5 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6FJ14143 | 6FJ14143/SHIP_NY_44K_100_16.5-10.875-5.5 | SHIP_NY_44K_100_16.5-10.875-5.5 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 5.5 |  | in | FDA,4 | 5.3125 | 5.6875 | TM 141 / Visual | Main_Plans
6FJ14143 | 6FJ14143/SHIP_NY_44K_100_16.5-10.875-5.5 | SHIP_NY_44K_100_16.5-10.875-5.5 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 16.5 |  | in | FDA,4 | 16.3125 | 16.6875 | TM 141 / Visual | Main_Plans
6FJ14143 | 6FJ14143/SHIP_NY_44K_100_16.5-10.875-5.5 | SHIP_NY_44K_100_16.5-10.875-5.5 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6FJ14143 | 6FJ14143/SHIP_NY_44K_100_16.5-10.875-5.5 | SHIP_NY_44K_100_16.5-10.875-5.5 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6FJ14143 | 6FJ14143/SHIP_NY_44K_100_16.5-10.875-5.5 | SHIP_NY_44K_100_16.5-10.875-5.5 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 10.875 |  | in | FDA,4 | 10.6875 | 11.0625 | TM 141 / Visual | Main_Plans
6GB10300 | 6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6GB10300 | 6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6GB10300 | 6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6GB10300 | 6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6GB10300 | 6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 5 |  | in | FDA,4 | 4.8125 | 5.1875 | TM 141 / Visual | Main_Plans
6GB10300 | 6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 14.75 |  | in | FDA,4 | 14.4375 | 14.8125 | TM 141 / Visual | Main_Plans
6GB10300 | 6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6GB10300 | 6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC w/ shyed flaps (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6GB10300 | 6GB10300/SHIP_NY_32K_100_14.625-6.375-5 | SHIP_NY_32K_100_14.625-6.375-5 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.25 |  | in | FDA,4 | 6.1875 | 6.5625 | TM 141 / Visual | Main_Plans
6GD35185 | 6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 50 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6GD35185 | 6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6GD35185 | 6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6GD35185 | 6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6GD35185 | 6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 12.4375 |  | in | FDA,4 | 12.25 | 12.625 | TM 141 / Visual | Main_Plans
6GD35185 | 6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 13.25 |  | in | FDA,4 | 13.0625 | 13.4375 | TM 141 / Visual | Main_Plans
6GD35185 | 6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6GD35185 | 6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6GD35185 | 6GD35185/SHIP_NY_32K_100_13.25-12-12.4375 | SHIP_NY_32K_100_13.25-12-12.4375 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 12 |  | in | FDA,4 | 11.8125 | 12.1875 | TM 141 / Visual | Main_Plans
6GS48450 | 6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | SHIP_NY_29K_100_17.125-6.6875-11.75 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6GS48450 | 6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | SHIP_NY_29K_100_17.125-6.6875-11.75 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 29 |  | Lbs./In. | FDA,0 | 29 |  | TM 141 / Visual | Main_Plans
6GS48450 | 6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | SHIP_NY_29K_100_17.125-6.6875-11.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6GS48450 | 6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | SHIP_NY_29K_100_17.125-6.6875-11.75 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 11.875 |  | in | FDA,4 | 11.5625 | 11.9375 | TM 141 / Visual | Main_Plans
6GS48450 | 6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | SHIP_NY_29K_100_17.125-6.6875-11.75 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 17.125 |  | in | FDA,4 | 16.9375 | 17.3125 | TM 141 / Visual | Main_Plans
6GS48450 | 6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | SHIP_NY_29K_100_17.125-6.6875-11.75 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6GS48450 | 6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | SHIP_NY_29K_100_17.125-6.6875-11.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6GS48450 | 6GS48450/SHIP_NY_29K_100_17.125-6.6875-11.75 | SHIP_NY_29K_100_17.125-6.6875-11.75 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.75 |  | in | FDA,4 | 6.5 | 6.875 | TM 141 / Visual | Main_Plans
6HB60307 | 6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | SHIP_NY_40K_100_16.5-12.3125-6.1875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6HB60307 | 6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | SHIP_NY_40K_100_16.5-12.3125-6.1875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 40 |  | Lbs./In. | FDA,0 | 40 |  | TM 141 / Visual | Main_Plans
6HB60307 | 6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | SHIP_NY_40K_100_16.5-12.3125-6.1875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6HB60307 | 6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | SHIP_NY_40K_100_16.5-12.3125-6.1875 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6.312 |  | in | FDA,4 | 6 | 6.375 | TM 141 / Visual | Main_Plans
6HB60307 | 6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | SHIP_NY_40K_100_16.5-12.3125-6.1875 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 16.5 |  | in | FDA,4 | 16.3125 | 16.6875 | TM 141 / Visual | Main_Plans
6HB60307 | 6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | SHIP_NY_40K_100_16.5-12.3125-6.1875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6HB60307 | 6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | SHIP_NY_40K_100_16.5-12.3125-6.1875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6HB60307 | 6HB60307/SHIP_NY_40K_100_16.5-12.3125-6.1875 | SHIP_NY_40K_100_16.5-12.3125-6.1875 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 12.25 |  | in | FDA,4 | 12.125 | 12.5 | TM 141 / Visual | Main_Plans
6HC69900 | 6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 50 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6HC69900 | 6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6HC69900 | 6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6HC69900 | 6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6HC69900 | 6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 5.25 |  | in | FDA,4 | 5 | 5.375 | TM 141 / Visual | Main_Plans
6HC69900 | 6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.75 |  | in | FDA,4 | 15.5625 | 15.9375 | TM 141 / Visual | Main_Plans
6HC69900 | 6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6HC69900 | 6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6HC69900 | 6HC69900/SHIP_NY_32K_100_15.75-8.625-5.1875 | SHIP_NY_32K_100_15.75-8.625-5.1875 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 8.625 |  | in | FDA,4 | 8.4375 | 8.8125 | TM 141 / Visual | Main_Plans
6JA580FW | 6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | SHIP_NY_44K_100_16.750-8.125-10.750 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6JA580FW | 6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | SHIP_NY_44K_100_16.750-8.125-10.750 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | Main_Plans
6JA580FW | 6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | SHIP_NY_44K_100_16.750-8.125-10.750 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6JA580FW | 6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | SHIP_NY_44K_100_16.750-8.125-10.750 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 10.875 |  | in | FDA,4 | 10.5625 | 10.9375 | TM 141 / Visual | Main_Plans
6JA580FW | 6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | SHIP_NY_44K_100_16.750-8.125-10.750 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 16.9375 |  | in | FDA,4 | 16.5625 | 16.9375 | TM 141 / Visual | Main_Plans
6JA580FW | 6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | SHIP_NY_44K_100_16.750-8.125-10.750 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6JA580FW | 6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | SHIP_NY_44K_100_16.750-8.125-10.750 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6JA580FW | 6JA580FW/SHIP_NY_44K_100_16.750-8.125-10.750 | SHIP_NY_44K_100_16.750-8.125-10.750 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 8.1875 |  | in | FDA,4 | 7.9375 | 8.3125 | TM 141 / Visual | Main_Plans
6LB40600 | 6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | SHIP_NY_29K_100_12.375-5.5-4.625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6LB40600 | 6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | SHIP_NY_29K_100_12.375-5.5-4.625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 29 |  | Lbs./In. | FDA,0 | 29 |  | TM 141 / Visual | Main_Plans
6LB40600 | 6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | SHIP_NY_29K_100_12.375-5.5-4.625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6LB40600 | 6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | SHIP_NY_29K_100_12.375-5.5-4.625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 4.75 |  | in | FDA,4 | 4.4375 | 4.8125 | TM 141 / Visual | Main_Plans
6LB40600 | 6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | SHIP_NY_29K_100_12.375-5.5-4.625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 12.437 |  | in | FDA,4 | 12.1875 | 12.5625 | TM 141 / Visual | Main_Plans
6LB40600 | 6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | SHIP_NY_29K_100_12.375-5.5-4.625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6LB40600 | 6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | SHIP_NY_29K_100_12.375-5.5-4.625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6LB40600 | 6LB40600/SHIP_NY_29K_100_12.375-5.5-4.625 | SHIP_NY_29K_100_12.375-5.5-4.625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 5.5 |  | in | FDA,4 | 5.3125 | 5.6875 | TM 141 / Visual | Main_Plans
6OA404S48 | 6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | SHIP_NY_44K_100_15.625-12.5-6.1875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OA404S48 | 6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | SHIP_NY_44K_100_15.625-12.5-6.1875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | Main_Plans
6OA404S48 | 6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | SHIP_NY_44K_100_15.625-12.5-6.1875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OA404S48 | 6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | SHIP_NY_44K_100_15.625-12.5-6.1875 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6 |  | in | FDA,4 | 6 | 6.375 | TM 141 / Visual | Main_Plans
6OA404S48 | 6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | SHIP_NY_44K_100_15.625-12.5-6.1875 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.625 |  | in | FDA,4 | 15.4375 | 15.8125 | TM 141 / Visual | Main_Plans
6OA404S48 | 6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | SHIP_NY_44K_100_15.625-12.5-6.1875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OA404S48 | 6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | SHIP_NY_44K_100_15.625-12.5-6.1875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OA404S48 | 6OA404S48/SHIP_NY_44K_100_15.625-12.5-6.1875 | SHIP_NY_44K_100_15.625-12.5-6.1875 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 12.5625 |  | in | FDA,4 | 12.3125 | 12.6875 | TM 141 / Visual | Main_Plans
6OA540S48 | 6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6OA540S48 | 6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OA540S48 | 6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6OA540S48 | 6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OA540S48 | 6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 12.625 |  | in | FDA,4 | 12.4375 | 12.8125 | TM 141 / Visual | Main_Plans
6OA540S48 | 6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.875 |  | in | FDA,4 | 11.6875 | 12.0625 | TM 141 / Visual | Main_Plans
6OA540S48 | 6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OA540S48 | 6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OA540S48 | 6OA540S48/SHIP_NY_32K_100_11.875-7.75-12.625 | SHIP_NY_32K_100_11.875-7.75-12.625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.75 |  | in | FDA,4 | 7.5625 | 7.9375 | TM 141 / Visual | Main_Plans
6OA632S60 | 6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6OA632S60 | 6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OA632S60 | 6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6OA632S60 | 6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OA632S60 | 6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 14.25 |  | in | FDA,4 | 13.9375 | 14.3125 | TM 141 / Visual | Main_Plans
6OA632S60 | 6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 19.5 |  | in | FDA,4 | 19.1875 | 19.5625 | TM 141 / Visual | Main_Plans
6OA632S60 | 6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OA632S60 | 6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | HSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OA632S60 | 6OA632S60/SHIP_NY_32K_150_19.375-8.25-14.125 | SHIP_NY_32K_150_19.375-8.25-14.125 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 8.25 |  | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual | Main_Plans
6OB123ST | 6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6OB123ST | 6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OB123ST | 6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6OB123ST | 6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OB123ST | 6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 5.875 |  | in | FDA,4 | 5.5625 | 5.9375 | TM 141 / Visual | Main_Plans
6OB123ST | 6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 12.5 |  | in | FDA,4 | 12.1875 | 12.5625 | TM 141 / Visual | Main_Plans
6OB123ST | 6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OB123ST | 6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OB123ST | 6OB123ST/SHIP_NY_32K_100_12.375-10.75-5.75 | SHIP_NY_32K_100_12.375-10.75-5.75 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 10.75 |  | in | FDA,4 | 10.5625 | 10.9375 | TM 141 / Visual | Main_Plans
6OB31400 | 6OB31400/SHIP_32_1AL | SHIP_32_1AL | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OB31400 | 6OB31400/SHIP_32_1AL | SHIP_32_1AL | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6OB31400 | 6OB31400/SHIP_32_1AL | SHIP_32_1AL | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OB31400 | 6OB31400/SHIP_32_1AL | SHIP_32_1AL | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 5 |  | in | FDA,4 | 4.8125 | 5.1875 | TM 141 / Visual | Main_Plans
6OB31400 | 6OB31400/SHIP_32_1AL | SHIP_32_1AL | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 14.625 |  | in | FDA,4 | 14.4375 | 14.8125 | TM 141 / Visual | Main_Plans
6OB31400 | 6OB31400/SHIP_32_1AL | SHIP_32_1AL | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OB31400 | 6OB31400/SHIP_32_1AL | SHIP_32_1AL | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OB31400 | 6OB31400/SHIP_32_1AL | SHIP_32_1AL | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.375 |  | in | FDA,4 | 6.1875 | 6.5625 | TM 141 / Visual | Main_Plans
6OB389PD | 6OB389PD/Ship_NA_3A | Ship_NA_3A | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OB389PD | 6OB389PD/Ship_NA_3A | Ship_NA_3A | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 14.5 |  | in | FDA,4 | 14.3125 | 14.6875 | TM 141 / Visual | Main_Plans
6OB389PD | 6OB389PD/Ship_NA_3A | Ship_NA_3A | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | Layer Pad (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OB389PD | 6OB389PD/Ship_NA_3A | Ship_NA_3A | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.437 |  | in | FDA,4 | 6.0625 | 6.4375 | TM 141 / Visual | Main_Plans
6OC22324 | 6OC22324/SHIP_32_1N | SHIP_32_1N | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OC22324 | 6OC22324/SHIP_32_1N | SHIP_32_1N | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6OC22324 | 6OC22324/SHIP_32_1N | SHIP_32_1N | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OC22324 | 6OC22324/SHIP_32_1N | SHIP_32_1N | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 10.5 |  | in | FDA,4 | 10.1875 | 10.5625 | TM 141 / Visual | Main_Plans
6OC22324 | 6OC22324/SHIP_32_1N | SHIP_32_1N | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.125 |  | in | FDA,4 | 10.9375 | 11.3125 | TM 141 / Visual | Main_Plans
6OC22324 | 6OC22324/SHIP_32_1N | SHIP_32_1N | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OC22324 | 6OC22324/SHIP_32_1N | SHIP_32_1N | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OC22324 | 6OC22324/SHIP_32_1N | SHIP_32_1N | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.375 |  | in | FDA,4 | 6.1875 | 6.5625 | TM 141 / Visual | Main_Plans
6OD01524 | 6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6OD01524 | 6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OD01524 | 6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6OD01524 | 6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OD01524 | 6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 4.75 |  | in | FDA,4 | 4.5 | 4.875 | TM 141 / Visual | Main_Plans
6OD01524 | 6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 8.25 |  | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual | Main_Plans
6OD01524 | 6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OD01524 | 6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OD01524 | 6OD01524/SHIP_NY_32K_100_8.25-6.25-4.6875 | SHIP_NY_32K_100_8.25-6.25-4.6875 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.25 |  | in | FDA,4 | 6.0625 | 6.4375 | TM 141 / Visual | Main_Plans
6OD13500A | 6OD13500A/SHIP_32_1AD | SHIP_32_1AD | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OD13500A | 6OD13500A/SHIP_32_1AD | SHIP_32_1AD | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6OD13500A | 6OD13500A/SHIP_32_1AD | SHIP_32_1AD | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OD13500A | 6OD13500A/SHIP_32_1AD | SHIP_32_1AD | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6.25 |  | in | FDA,4 | 5.9375 | 6.3125 | TM 141 / Visual | Main_Plans
6OD13500A | 6OD13500A/SHIP_32_1AD | SHIP_32_1AD | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 13.25 |  | in | FDA,4 | 13.0625 | 13.4375 | TM 141 / Visual | Main_Plans
6OD13500A | 6OD13500A/SHIP_32_1AD | SHIP_32_1AD | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OD13500A | 6OD13500A/SHIP_32_1AD | SHIP_32_1AD | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OD13500A | 6OD13500A/SHIP_32_1AD | SHIP_32_1AD | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 12 |  | in | FDA,4 | 11.8125 | 12.1875 | TM 141 / Visual | Main_Plans
6OD239DR | 6OD239DR/SHIP_44_1A | SHIP_44_1A | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | Divider | [5 empty] | Shipper Style | Main_Plans
6OD239DR | 6OD239DR/SHIP_44_1A | SHIP_44_1A | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OD239DR | 6OD239DR/SHIP_44_1A | SHIP_44_1A | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | Main_Plans
6OD239DR | 6OD239DR/SHIP_44_1A | SHIP_44_1A | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OD239DR | 6OD239DR/SHIP_44_1A | SHIP_44_1A | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 7.75 |  | in | FDA,4 | 7.4375 | 7.8125 | TM 141 / Visual | Main_Plans
6OD239DR | 6OD239DR/SHIP_44_1A | SHIP_44_1A | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OD239DR | 6OD239DR/SHIP_44_1A | SHIP_44_1A | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 15.375 |  | in | FDA,4 | 15.1875 | 15.5625 | TM 141 / Visual | Main_Plans
6OD36055 | 6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6OD36055 | 6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OD36055 | 6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6OD36055 | 6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OD36055 | 6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 11.375 |  | in | FDA,4 | 11.0625 | 11.4375 | TM 141 / Visual | Main_Plans
6OD36055 | 6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 10.5 |  | in | FDA,4 | 10.4375 | 10.8125 | TM 141 / Visual | Main_Plans
6OD36055 | 6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Does not meet specifications (OOS) | [5 empty] | TM 141 / Visual | Main_Plans
6OD36055 | 6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC w/ shyed flaps (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OD36055 | 6OD36055/SHIP_NY_32K_102_10.625-6.375-11.25 | SHIP_NY_32K_102_10.625-6.375-11.25 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.375 |  | in | FDA,4 | 6.1875 | 6.5625 | TM 141 / Visual | Main_Plans
6OH14324 | 6OH14324/SHIP_32_1Y | SHIP_32_1Y | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OH14324 | 6OH14324/SHIP_32_1Y | SHIP_32_1Y | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6OH14324 | 6OH14324/SHIP_32_1Y | SHIP_32_1Y | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OH14324 | 6OH14324/SHIP_32_1Y | SHIP_32_1Y | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6.25 |  | in | FDA,4 | 6.0625 | 6.4375 | TM 141 / Visual | Main_Plans
6OH14324 | 6OH14324/SHIP_32_1Y | SHIP_32_1Y | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 12.4375 |  | in | FDA,4 | 12.25 | 12.625 | TM 141 / Visual | Main_Plans
6OH14324 | 6OH14324/SHIP_32_1Y | SHIP_32_1Y | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OH14324 | 6OH14324/SHIP_32_1Y | SHIP_32_1Y | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OH14324 | 6OH14324/SHIP_32_1Y | SHIP_32_1Y | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 9 |  | in | FDA,4 | 8.8125 | 9.1875 | TM 141 / Visual | Main_Plans
6OH21112 | 6OH21112/SHIP_32_1F | SHIP_32_1F | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OH21112 | 6OH21112/SHIP_32_1F | SHIP_32_1F | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6OH21112 | 6OH21112/SHIP_32_1F | SHIP_32_1F | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OH21112 | 6OH21112/SHIP_32_1F | SHIP_32_1F | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6.75 |  | in | FDA,4 | 6.5625 | 6.9375 | TM 141 / Visual | Main_Plans
6OH21112 | 6OH21112/SHIP_32_1F | SHIP_32_1F | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 9.75 |  | in | FDA,4 | 9.5625 | 9.9375 | TM 141 / Visual | Main_Plans
6OH21112 | 6OH21112/SHIP_32_1F | SHIP_32_1F | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OH21112 | 6OH21112/SHIP_32_1F | SHIP_32_1F | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OH21112 | 6OH21112/SHIP_32_1F | SHIP_32_1F | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7.5 |  | in | FDA,4 | 7.4375 | 7.8125 | TM 141 / Visual | Main_Plans
6OH45016 | 6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | SHIP_NY_32K_100_8.25-6.875-7.375 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OH45016 | 6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | SHIP_NY_32K_100_8.25-6.875-7.375 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6OH45016 | 6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | SHIP_NY_32K_100_8.25-6.875-7.375 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OH45016 | 6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | SHIP_NY_32K_100_8.25-6.875-7.375 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7.25 |  | in | FDA,4 | 7.1875 | 7.5625 | TM 141 / Visual | Main_Plans
6OH45016 | 6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | SHIP_NY_32K_100_8.25-6.875-7.375 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 8.1562 |  | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual | Main_Plans
6OH45016 | 6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | SHIP_NY_32K_100_8.25-6.875-7.375 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OH45016 | 6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | SHIP_NY_32K_100_8.25-6.875-7.375 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OH45016 | 6OH45016/SHIP_NY_32K_100_8.25-6.875-7.375 | SHIP_NY_32K_100_8.25-6.875-7.375 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 7 |  | in | FDA,4 | 6.6875 | 7.0625 | TM 141 / Visual | Main_Plans
6OK75595 | 6OK75595/SHIP_32_1R | SHIP_32_1R | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OK75595 | 6OK75595/SHIP_32_1R | SHIP_32_1R | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6OK75595 | 6OK75595/SHIP_32_1R | SHIP_32_1R | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OK75595 | 6OK75595/SHIP_32_1R | SHIP_32_1R | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7.75 |  | in | FDA,4 | 7.5 | 7.875 | TM 141 / Visual | Main_Plans
6OK75595 | 6OK75595/SHIP_32_1R | SHIP_32_1R | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.75 |  | in | FDA,4 | 11.5625 | 11.9375 | TM 141 / Visual | Main_Plans
6OK75595 | 6OK75595/SHIP_32_1R | SHIP_32_1R | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OK75595 | 6OK75595/SHIP_32_1R | SHIP_32_1R | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OK75595 | 6OK75595/SHIP_32_1R | SHIP_32_1R | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 9.5 |  | in | FDA,4 | 9.4375 | 9.8125 | TM 141 / Visual | Main_Plans
6OM960CB | 6OM960CB/SHIP_NY_NA_200_2.5-2.5-33.75 | SHIP_NY_NA_200_2.5-2.5-33.75 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OM960CB | 6OM960CB/SHIP_NY_NA_200_2.5-2.5-33.75 | SHIP_NY_NA_200_2.5-2.5-33.75 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 33.625 |  | in | FDA,4 | 33.5625 | 33.9375 | TM 141 / Visual | Main_Plans
6OM960CB | 6OM960CB/SHIP_NY_NA_200_2.5-2.5-33.75 | SHIP_NY_NA_200_2.5-2.5-33.75 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 2.5 |  | in | FDA,4 | 2.3125 | 2.6875 | TM 141 / Visual | Main_Plans
6OM960CB | 6OM960CB/SHIP_NY_NA_200_2.5-2.5-33.75 | SHIP_NY_NA_200_2.5-2.5-33.75 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | ITW STACK MASTER 2600 (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OM960CB | 6OM960CB/SHIP_NY_NA_200_2.5-2.5-33.75 | SHIP_NY_NA_200_2.5-2.5-33.75 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 2.5 |  | in | FDA,4 | 2.3125 | 2.6875 | TM 141 / Visual | Main_Plans
6OP027LP | 6OP027LP/SHIP_32_13A | SHIP_32_13A | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | Layer Pad | [5 empty] | Shipper Style | Main_Plans
6OP027LP | 6OP027LP/SHIP_32_13A | SHIP_32_13A | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OP027LP | 6OP027LP/SHIP_32_13A | SHIP_32_13A | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6OP027LP | 6OP027LP/SHIP_32_13A | SHIP_32_13A | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OP027LP | 6OP027LP/SHIP_32_13A | SHIP_32_13A | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 10.75 |  | in | FDA,4 | 10.5625 | 10.9375 | TM 141 / Visual | Main_Plans
6OP027LP | 6OP027LP/SHIP_32_13A | SHIP_32_13A | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OP027LP | 6OP027LP/SHIP_32_13A | SHIP_32_13A | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 5.5625 |  | in | FDA,4 | 5.375 | 5.75 | TM 141 / Visual | Main_Plans
6OP081LP | 6OP081LP/SHIP_NY_32NA_250_17.250-11.938 | SHIP_NY_32NA_250_17.250-11.938 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6OP081LP | 6OP081LP/SHIP_NY_32NA_250_17.250-11.938 | SHIP_NY_32NA_250_17.250-11.938 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OP081LP | 6OP081LP/SHIP_NY_32NA_250_17.250-11.938 | SHIP_NY_32NA_250_17.250-11.938 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6OP081LP | 6OP081LP/SHIP_NY_32NA_250_17.250-11.938 | SHIP_NY_32NA_250_17.250-11.938 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OP081LP | 6OP081LP/SHIP_NY_32NA_250_17.250-11.938 | SHIP_NY_32NA_250_17.250-11.938 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 17.125 |  | in | FDA,3 | 17.125 | 17.375 | TM 141 / Visual | Main_Plans
6OP081LP | 6OP081LP/SHIP_NY_32NA_250_17.250-11.938 | SHIP_NY_32NA_250_17.250-11.938 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | Inserts (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OP081LP | 6OP081LP/SHIP_NY_32NA_250_17.250-11.938 | SHIP_NY_32NA_250_17.250-11.938 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 11.875 |  | in | FDA,3 | 11.813 | 12.063 | TM 141 / Visual | Main_Plans
6OP34524 | 6OP34524/SHIP_44_2G | SHIP_44_2G | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OP34524 | 6OP34524/SHIP_44_2G | SHIP_44_2G | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | Main_Plans
6OP34524 | 6OP34524/SHIP_44_2G | SHIP_44_2G | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OP34524 | 6OP34524/SHIP_44_2G | SHIP_44_2G | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 8.25 |  | in | FDA,4 | 8.0625 | 8.4375 | TM 141 / Visual | Main_Plans
6OP34524 | 6OP34524/SHIP_44_2G | SHIP_44_2G | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 13.5 |  | in | FDA,4 | 13.3125 | 13.6875 | TM 141 / Visual | Main_Plans
6OP34524 | 6OP34524/SHIP_44_2G | SHIP_44_2G | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OP34524 | 6OP34524/SHIP_44_2G | SHIP_44_2G | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OP34524 | 6OP34524/SHIP_44_2G | SHIP_44_2G | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 11.75 |  | in | FDA,4 | 11.5625 | 11.9375 | TM 141 / Visual | Main_Plans
6OP75510 | 6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 32 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6OP75510 | 6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OP75510 | 6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | Main_Plans
6OP75510 | 6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OP75510 | 6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 10.5625 |  | in | FDA,4 | 10.375 | 10.75 | TM 141 / Visual | Main_Plans
6OP75510 | 6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.375 |  | in | FDA,4 | 15.1875 | 15.5625 | TM 141 / Visual | Main_Plans
6OP75510 | 6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OP75510 | 6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OP75510 | 6OP75510/SHIP_NY_32K_100_15.375-11.5625-10.5625 | SHIP_NY_32K_100_15.375-11.5625-10.5625 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 11.5625 |  | in | FDA,4 | 11.375 | 11.75 | TM 141 / Visual | Main_Plans
6OQ22272 | 6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | SHIP_NY_44K_100_15.188-10.188-15.188 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OQ22272 | 6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | SHIP_NY_44K_100_15.188-10.188-15.188 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | Main_Plans
6OQ22272 | 6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | SHIP_NY_44K_100_15.188-10.188-15.188 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OQ22272 | 6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | SHIP_NY_44K_100_15.188-10.188-15.188 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 15.188 |  | in | FDA,4 | 15.0005 | 15.3755 | TM 141 / Visual | Main_Plans
6OQ22272 | 6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | SHIP_NY_44K_100_15.188-10.188-15.188 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.188 |  | in | FDA,4 | 15.0005 | 15.3755 | TM 141 / Visual | Main_Plans
6OQ22272 | 6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | SHIP_NY_44K_100_15.188-10.188-15.188 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OQ22272 | 6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | SHIP_NY_44K_100_15.188-10.188-15.188 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OQ22272 | 6OQ22272/SHIP_NY_44K_100_15.188-10.188-15.188 | SHIP_NY_44K_100_15.188-10.188-15.188 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 10.188 |  | in | FDA,4 | 10.0005 | 10.3755 | TM 141 / Visual | Main_Plans
6OS12350 | 6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6OS12350 | 6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OS12350 | 6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OS12350 | 6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6OS12350 | 6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OS12350 | 6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 5.938 |  | in | FDA,4 | 5.7505 | 6.1255 | TM 141 / Visual | Main_Plans
6OS12350 | 6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.75 |  | in | FDA,4 | 11.4375 | 11.8125 | TM 141 / Visual | Main_Plans
6OS12350 | 6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OS12350 | 6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OS12350 | 6OS12350/SHIP_NY_32K_100_11.625-6.5-5.938 | SHIP_NY_32K_100_11.625-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.375 |  | in | FDA,4 | 6.3125 | 6.6875 | TM 141 / Visual | Main_Plans
6OS42850 | 6OS42850/SHIP_44_2N | SHIP_44_2N | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OS42850 | 6OS42850/SHIP_44_2N | SHIP_44_2N | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | Main_Plans
6OS42850 | 6OS42850/SHIP_44_2N | SHIP_44_2N | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OS42850 | 6OS42850/SHIP_44_2N | SHIP_44_2N | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 5.625 |  | in | FDA,4 | 5.3125 | 5.6875 | TM 141 / Visual | Main_Plans
6OS42850 | 6OS42850/SHIP_44_2N | SHIP_44_2N | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.375 |  | in | FDA,4 | 15.1875 | 15.5625 | TM 141 / Visual | Main_Plans
6OS42850 | 6OS42850/SHIP_44_2N | SHIP_44_2N | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OS42850 | 6OS42850/SHIP_44_2N | SHIP_44_2N | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OS42850 | 6OS42850/SHIP_44_2N | SHIP_44_2N | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 11.375 |  | in | FDA,4 | 11.1875 | 11.5625 | TM 141 / Visual | Main_Plans
6OS45625 | 6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6OS45625 | 6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OS45625 | 6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OS45625 | 6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6OS45625 | 6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OS45625 | 6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6 |  | in | FDA,4 | 5.7505 | 6.1255 | TM 141 / Visual | Main_Plans
6OS45625 | 6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 6.75 |  | in | FDA,4 | 6.6875 | 7.0625 | TM 141 / Visual | Main_Plans
6OS45625 | 6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OS45625 | 6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OS45625 | 6OS45625/SHIP_NY_32K_100_6.875-6.5-5.938 | SHIP_NY_32K_100_6.875-6.5-5.938 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.5 |  | in | FDA,4 | 6.3125 | 6.6875 | TM 141 / Visual | Main_Plans
6OS46250 | 6OS46250/SHIP_29_1J | SHIP_29_1J | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OS46250 | 6OS46250/SHIP_29_1J | SHIP_29_1J | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 29 |  | Lbs./In. | FDA,0 | 29 |  | TM 141 / Visual | Main_Plans
6OS46250 | 6OS46250/SHIP_29_1J | SHIP_29_1J | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OS46250 | 6OS46250/SHIP_29_1J | SHIP_29_1J | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 11.75 |  | in | FDA,4 | 11.5625 | 11.9375 | TM 141 / Visual | Main_Plans
6OS46250 | 6OS46250/SHIP_29_1J | SHIP_29_1J | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 17.125 |  | in | FDA,4 | 16.9375 | 17.3125 | TM 141 / Visual | Main_Plans
6OS46250 | 6OS46250/SHIP_29_1J | SHIP_29_1J | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OS46250 | 6OS46250/SHIP_29_1J | SHIP_29_1J | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OS46250 | 6OS46250/SHIP_29_1J | SHIP_29_1J | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.6875 |  | in | FDA,4 | 6.5 | 6.875 | TM 141 / Visual | Main_Plans
6OU12066 | 6OU12066/SHIP_40_1M | SHIP_40_1M | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OU12066 | 6OU12066/SHIP_40_1M | SHIP_40_1M | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 40 |  | Lbs./In. | FDA,0 | 40 |  | TM 141 / Visual | Main_Plans
6OU12066 | 6OU12066/SHIP_40_1M | SHIP_40_1M | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OU12066 | 6OU12066/SHIP_40_1M | SHIP_40_1M | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 9 |  | in | FDA,4 | 8.8125 | 9.1875 | TM 141 / Visual | Main_Plans
6OU12066 | 6OU12066/SHIP_40_1M | SHIP_40_1M | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 17.75 |  | in | FDA,4 | 17.625 | 18 | TM 141 / Visual | Main_Plans
6OU12066 | 6OU12066/SHIP_40_1M | SHIP_40_1M | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OU12066 | 6OU12066/SHIP_40_1M | SHIP_40_1M | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OU12066 | 6OU12066/SHIP_40_1M | SHIP_40_1M | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 15.125 |  | in | FDA,4 | 14.9375 | 15.3125 | TM 141 / Visual | Main_Plans
6OU21470 | 6OU21470/SHIP_40_1J | SHIP_40_1J | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OU21470 | 6OU21470/SHIP_40_1J | SHIP_40_1J | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 40 |  | Lbs./In. | FDA,0 | 40 |  | TM 141 / Visual | Main_Plans
6OU21470 | 6OU21470/SHIP_40_1J | SHIP_40_1J | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OU21470 | 6OU21470/SHIP_40_1J | SHIP_40_1J | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 9.4375 |  | in | FDA,4 | 9.0625 | 9.4375 | TM 141 / Visual | Main_Plans
6OU21470 | 6OU21470/SHIP_40_1J | SHIP_40_1J | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.5 |  | in | FDA,4 | 15.1875 | 15.5625 | TM 141 / Visual | Main_Plans
6OU21470 | 6OU21470/SHIP_40_1J | SHIP_40_1J | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OU21470 | 6OU21470/SHIP_40_1J | SHIP_40_1J | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OU21470 | 6OU21470/SHIP_40_1J | SHIP_40_1J | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 11.25 |  | in | FDA,4 | 11.1875 | 11.5625 | TM 141 / Visual | Main_Plans
6OU31450 | 6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6OU31450 | 6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OU31450 | 6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 44 |  | Lbs./In. | FDA,0 | 44 |  | TM 141 / Visual | Main_Plans
6OU31450 | 6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OU31450 | 6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6.125 |  | in | FDA,4 | 5.9375 | 6.3125 | TM 141 / Visual | Main_Plans
6OU31450 | 6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 18.5 |  | in | FDA,4 | 18.3125 | 18.6875 | TM 141 / Visual | Main_Plans
6OU31450 | 6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OU31450 | 6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OU31450 | 6OU31450/SHIP_NY_44K_100_18.5-13.5-6.125 | SHIP_NY_44K_100_18.5-13.5-6.125 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 13.5 |  | in | FDA,4 | 13.3125 | 13.6875 | TM 141 / Visual | Main_Plans
6OU35199 | 6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6OU35199 | 6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OU35199 | 6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6OU35199 | 6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OU35199 | 6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 11.5 |  | in | FDA,4 | 11.25 | 11.5 | TM 141 / Visual | Main_Plans
6OU35199 | 6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 19.75 |  | in | FDA,4 | 19.625 | 19.875 | TM 141 / Visual | Main_Plans
6OU35199 | 6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OU35199 | 6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OU35199 | 6OU35199/SHIP_NY_32K_100_19.75-12-11.375 | SHIP_NY_32K_100_19.75-12-11.375 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 12 |  | in | FDA,4 | 11.875 | 12.125 | TM 141 / Visual | Main_Plans
6OU351LP | 6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | SHIP_NY_32K_NA_19.500X11.750 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | Pad | [5 empty] | Shipper Style | Main_Plans
6OU351LP | 6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | SHIP_NY_32K_NA_19.500X11.750 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6OU351LP | 6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | SHIP_NY_32K_NA_19.500X11.750 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OU351LP | 6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | SHIP_NY_32K_NA_19.500X11.750 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6OU351LP | 6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | SHIP_NY_32K_NA_19.500X11.750 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OU351LP | 6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | SHIP_NY_32K_NA_19.500X11.750 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 19.5 |  | in | FDA,4 | 19.375 | 19.625 | TM 141 / Visual | Main_Plans
6OU351LP | 6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | SHIP_NY_32K_NA_19.500X11.750 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OU351LP | 6OU351LP/SHIP_NY_32K_NA_19.500X11.750 | SHIP_NY_32K_NA_19.500X11.750 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 11.75 |  | in | FDA,4 | 11.625 | 11.875 | TM 141 / Visual | Main_Plans
6OU351TY | 6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | D/C Tray | [5 empty] | Shipper Style | Main_Plans
6OU351TY | 6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6OU351TY | 6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OU351TY | 6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6OU351TY | 6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OU351TY | 6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 3.5 |  | in | FDA,4 | 3.4375 | 3.5625 | TM 141 / Visual | Main_Plans
6OU351TY | 6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 19.062 |  | in | FDA,4 | 19 | 19.125 | TM 141 / Visual | Main_Plans
6OU351TY | 6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OU351TY | 6OU351TY/SHIP_NY_32K_201_19.0625-2.8125-3.500 | SHIP_NY_32K_201_19.0625-2.8125-3.500 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 2.812 |  | in | FDA,4 | 2.75 | 2.875 | TM 141 / Visual | Main_Plans
6OU38510 | 6OU38510/SHIP_32_1BG | SHIP_32_1BG | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OU38510 | 6OU38510/SHIP_32_1BG | SHIP_32_1BG | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6OU38510 | 6OU38510/SHIP_32_1BG | SHIP_32_1BG | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OU38510 | 6OU38510/SHIP_32_1BG | SHIP_32_1BG | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 6 |  | in | FDA,4 | 5.8125 | 6.1875 | TM 141 / Visual | Main_Plans
6OU38510 | 6OU38510/SHIP_32_1BG | SHIP_32_1BG | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 8 |  | in | FDA,4 | 7.8125 | 8.1875 | TM 141 / Visual | Main_Plans
6OU38510 | 6OU38510/SHIP_32_1BG | SHIP_32_1BG | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OU38510 | 6OU38510/SHIP_32_1BG | SHIP_32_1BG | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OU38510 | 6OU38510/SHIP_32_1BG | SHIP_32_1BG | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.125 |  | in | FDA,4 | 5.9375 | 6.3125 | TM 141 / Visual | Main_Plans
6OU55510 | 6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6OU55510 | 6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OU55510 | 6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6OU55510 | 6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OU55510 | 6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 8 |  | in | FDA,4 | 7.8125 | 8.1875 | TM 141 / Visual | Main_Plans
6OU55510 | 6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.75 |  | in | FDA,4 | 11.5625 | 11.9375 | TM 141 / Visual | Main_Plans
6OU55510 | 6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OU55510 | 6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OU55510 | 6OU55510/SHIP_NY_32K_100_11.75-6.25-8 | SHIP_NY_32K_100_11.75-6.25-8 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 6.375 |  | in | FDA,4 | 6.0625 | 6.4375 | TM 141 / Visual | Main_Plans
6OU62620 | 6OU62620/SHIP_NY_32K_100_15-9.5-8 | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6OU62620 | 6OU62620/SHIP_NY_32K_100_15-9.5-8 | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OU62620 | 6OU62620/SHIP_NY_32K_100_15-9.5-8 | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6OU62620 | 6OU62620/SHIP_NY_32K_100_15-9.5-8 | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OU62620 | 6OU62620/SHIP_NY_32K_100_15-9.5-8 | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 8 |  | in | FDA,4 | 7.8125 | 8.1875 | TM 141 / Visual | Main_Plans
6OU62620 | 6OU62620/SHIP_NY_32K_100_15-9.5-8 | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 15.125 |  | in | FDA,4 | 14.8125 | 15.1875 | TM 141 / Visual | Main_Plans
6OU62620 | 6OU62620/SHIP_NY_32K_100_15-9.5-8 | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OU62620 | 6OU62620/SHIP_NY_32K_100_15-9.5-8 | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | RSC (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OU62620 | 6OU62620/SHIP_NY_32K_100_15-9.5-8 | SHIP_NY_32K_100_15-9.5-8 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 9.375 |  | in | FDA,4 | 9.3125 | 9.6875 | TM 141 / Visual | Main_Plans
6OX46500 | 6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size | 125 |  |  | FDA,0 |  |  | TM 141 / Visual | Main_Plans
6OX46500 | 6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OX46500 | 6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6OX46500 | 6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OX46500 | 6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height | 7.5 |  | in | FDA,4 | 7.3125 | 7.6875 | TM 141 / Visual | Main_Plans
6OX46500 | 6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 23.75 |  | in | FDA,4 | 23.5005 | 23.8755 | TM 141 / Visual | Main_Plans
6OX46500 | 6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | Kraft (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OX46500 | 6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | Shipper (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6OX46500 | 6OX46500/SHIP_NY_32K_100_23.688-15.688-7.500 | SHIP_NY_32K_100_23.688-15.688-7.500 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 15.6 |  | in | FDA,4 | 15.5005 | 15.8755 | TM 141 / Visual | Main_Plans
6RQ55172 | 6RQ55172/SHIP_NY_44K_100_15.1875-10.1875-15.1875 | SHIP_NY_44K_100_15.1875-10.1875-15.1875 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Shipper Inspection | Main_Plans
6ZS505PD | 6ZS505PD/SHIP_32_14A | SHIP_32_14A | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | Go - Meets Specifications (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6ZS505PD | 6ZS505PD/SHIP_32_14A | SHIP_32_14A | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) | 32 |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Main_Plans
6ZS505PD | 6ZS505PD/SHIP_32_14A | SHIP_32_14A | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | Accept (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6ZS505PD | 6ZS505PD/SHIP_32_14A | SHIP_32_14A | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length | 11.4375 |  | in | FDA,4 | 11.1875 | 11.5625 | TM 141 / Visual | Main_Plans
6ZS505PD | 6ZS505PD/SHIP_32_14A | SHIP_32_14A | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | Cut Pad - Sheet (Done) | [5 empty] | TM 141 / Visual | Main_Plans
6ZS505PD | 6ZS505PD/SHIP_32_14A | SHIP_32_14A | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width | 5.75 |  | in | FDA,4 | 5.5625 | 5.9375 | TM 141 / Visual | Main_Plans
7A4AA580 | 7A4AA580/Chem_NY_Disodium Cocoamphodipropinate | Chem_NY_Disodium Cocoamphodipropinate | IN-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | COA | Main_Plans
7A4AA580 | 7A4AA580/Chem_NY_Disodium Cocoamphodipropinate | Chem_NY_Disodium Cocoamphodipropinate | IN-Analytical | 447 | pH as is | pH | pH | 9.92 |  |  | FDA,1 | 9 | 10 | TM 200 | Main_Plans
7A4AA580 | 7A4AA580/Chem_NY_Disodium Cocoamphodipropinate | Chem_NY_Disodium Cocoamphodipropinate | IN-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7A4AA580 | 7A4AA580/Chem_NY_Disodium Cocoamphodipropinate | Chem_NY_Disodium Cocoamphodipropinate | IN-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear to Hazy liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4AA580 | 7A4AA580/Chem_NY_Disodium Cocoamphodipropinate | Chem_NY_Disodium Cocoamphodipropinate | IN-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to Yellow (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4AA580 | 7A4AA580/Chem_NY_Disodium Cocoamphodipropinate | Chem_NY_Disodium Cocoamphodipropinate | IN-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4AA580 | 7A4AA580/Chem_NY_Disodium Cocoamphodipropinate | Chem_NY_Disodium Cocoamphodipropinate | IN-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4AL930 | 7A4AL930/Chem_08A | Chem_08A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7A4AL930 | 7A4AL930/Chem_08A | Chem_08A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Pellets (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4AL930 | 7A4AL930/Chem_08A | Chem_08A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White/Grey (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4AL930 | 7A4AL930/Chem_08A | Chem_08A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4AL930 | 7A4AL930/Chem_08A | Chem_08A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4AL930 | 7A4AL930/Chem_08A | Chem_08A | NY-Analytical | 390 | Identification | Identification for Potassium (USP <191>) | Identification for Potassium | Pass (Done) | [5 empty] | TM400/408b/421/422/USP | Main_Plans
7A4AS162 | 7A4AS162/Chem_09A | Chem_09A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Specific Gravity at 25°C | Main_Plans
7A4AS162 | 7A4AS162/Chem_09A | Chem_09A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.2846 |  |  | FDA,3 | 1.249 |  | TM 202 | Main_Plans
7A4AS162 | 7A4AS162/Chem_09A | Chem_09A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7A4AS162 | 7A4AS162/Chem_09A | Chem_09A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, syrupy liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4AS162 | 7A4AS162/Chem_09A | Chem_09A | NY-Analytical | 792 | Chemical Inspection | Color (USP) | Color | Meet USP requirements (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4AS162 | 7A4AS162/Chem_09A | Chem_09A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4AS162 | 7A4AS162/Chem_09A | Chem_09A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4AS162 | 7A4AS162/Chem_09A | Chem_09A | NY-Analytical | 816 | Related Compounds | Individual Impurity | Individual Impurity | 0.01 |  | % | FDA,3 |  | 0.1 | USP <621> | Main_Plans
7A4AS162 | 7A4AS162/Chem_09A | Chem_09A | NY-Analytical | 816 | Related Compounds | Total Impurities | Total Impurities | 0.02 |  | % | FDA,3 |  | 1 | USP <621> | Main_Plans
7A4BL625 | 7A4BL625/Chem_11A | Chem_11A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Identification | Main_Plans
7A4BL625 | 7A4BL625/Chem_11A | Chem_11A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7A4BL625 | 7A4BL625/Chem_11A | Chem_11A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4BL625 | 7A4BL625/Chem_11A | Chem_11A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Light Yellow to Yellow (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4BL625 | 7A4BL625/Chem_11A | Chem_11A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4BL625 | 7A4BL625/Chem_11A | Chem_11A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Lard-Like (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4CD419 | 7A4CD419/Chem_13A | Chem_13A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Identification | Main_Plans
7A4CD419 | 7A4CD419/Chem_13A | Chem_13A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7A4CD419 | 7A4CD419/Chem_13A | Chem_13A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Crystalline Powder (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4CD419 | 7A4CD419/Chem_13A | Chem_13A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White to off White (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4CD419 | 7A4CD419/Chem_13A | Chem_13A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4CD419 | 7A4CD419/Chem_13A | Chem_13A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Faint phenolic (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4FD110 | 7A4FD110/Chem_Sodium Chloride USP/NF | Chem_Sodium Chloride USP/NF | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7A4FD110 | 7A4FD110/Chem_Sodium Chloride USP/NF | Chem_Sodium Chloride USP/NF | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Crystals/crystalline powder (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4FD110 | 7A4FD110/Chem_Sodium Chloride USP/NF | Chem_Sodium Chloride USP/NF | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to white (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4FD110 | 7A4FD110/Chem_Sodium Chloride USP/NF | Chem_Sodium Chloride USP/NF | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4FD110 | 7A4FD110/Chem_Sodium Chloride USP/NF | Chem_Sodium Chloride USP/NF | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4FD110 | 7A4FD110/Chem_Sodium Chloride USP/NF | Chem_Sodium Chloride USP/NF | NY-Analytical | 390 | Identification | Identification for Chloride (USP <191>) | Identification for Chloride | Meets USP (Done) | [5 empty] | TM400/408b/421/422/USP | Main_Plans
7A4FD110 | 7A4FD110/Chem_Sodium Chloride USP/NF | Chem_Sodium Chloride USP/NF | NY-Analytical | 390 | Identification | Identification for Sodium (USP <191>) | Identification for Sodium | Meets USP (Done) | [5 empty] | TM400/408b/421/422/USP | Main_Plans
7A4FD110 | 7A4FD110/Chem_Sodium Chloride USP/NF | Chem_Sodium Chloride USP/NF | NY-Analytical | 802 | Loss on Drying | Loss on Drying | Loss on Drying |  |  | % | FDA,1 |  | 0.5 | USP | Main_Plans
7A4FD110 | 7A4FD110/Chem_Sodium Chloride USP/NF | Chem_Sodium Chloride USP/NF | NY-Analytical | 802 | Loss on Drying | Loss on Drying Final | Loss on Drying Final |  |  | % | FDA,1 |  | 0.5 | USP | Main_Plans
7A4FD110 | 7A4FD110/Chem_Sodium Chloride USP/NF | Chem_Sodium Chloride USP/NF | NY-Analytical | 796 | Assay | Assay on Dried Basis | Assay on Dried Basis | 99.7 |  | % | FDA,1 | 99 | 100.5 | USP <541> | Main_Plans
7A4FD110 | 7A4FD110/Chem_Sodium Chloride USP/NF | Chem_Sodium Chloride USP/NF | NY-Analytical | 796 | Assay | Assay on Dried Basis Final | Assay on Dried Basis Final | 99.7 |  | % | FDA,1 | 99 | 100.5 | USP <541> | Main_Plans
7A4OA347 | 7A4OA347/Chem_NY_Sodium-L(+) Lactate 60% | Chem_NY_Sodium-L(+) Lactate 60% | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | pH | Main_Plans
7A4OA347 | 7A4OA347/Chem_NY_Sodium-L(+) Lactate 60% | Chem_NY_Sodium-L(+) Lactate 60% | NY-Analytical | 398 | pH | pH (1 in 20 solution) | pH (1 in 20 solution) | 6.86 |  |  | FDA,1 | 6.5 | 7.5 | TM 200 | Main_Plans
7A4OA347 | 7A4OA347/Chem_NY_Sodium-L(+) Lactate 60% | Chem_NY_Sodium-L(+) Lactate 60% | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Comparison of scans of sample vs standard (Done) | [5 empty] | TM 411 | Main_Plans
7A4OA347 | 7A4OA347/Chem_NY_Sodium-L(+) Lactate 60% | Chem_NY_Sodium-L(+) Lactate 60% | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Transparent liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OA347 | 7A4OA347/Chem_NY_Sodium-L(+) Lactate 60% | Chem_NY_Sodium-L(+) Lactate 60% | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OA347 | 7A4OA347/Chem_NY_Sodium-L(+) Lactate 60% | Chem_NY_Sodium-L(+) Lactate 60% | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OA347 | 7A4OA347/Chem_NY_Sodium-L(+) Lactate 60% | Chem_NY_Sodium-L(+) Lactate 60% | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless or with a slight Characteristic Odor (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OA709 | 7A4OA709/Chem_NY_Chamomile Extract # 11-19-2 | Chem_NY_Chamomile Extract # 11-19-2 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | pH | Main_Plans
7A4OA709 | 7A4OA709/Chem_NY_Chamomile Extract # 11-19-2 | Chem_NY_Chamomile Extract # 11-19-2 | NY-Analytical | 398 | pH | pH | pH | 4.92 |  |  | FDA,1 | 4 | 6.5 | TM 200 | Main_Plans
7A4OA709 | 7A4OA709/Chem_NY_Chamomile Extract # 11-19-2 | Chem_NY_Chamomile Extract # 11-19-2 | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.0482 |  |  | FDA,3 | 1.043 | 1.053 | TM 202 | Main_Plans
7A4OA709 | 7A4OA709/Chem_NY_Chamomile Extract # 11-19-2 | Chem_NY_Chamomile Extract # 11-19-2 | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.412 |  |  | FDA,3 | 1.409 | 1.419 | TM 204 | Main_Plans
7A4OA709 | 7A4OA709/Chem_NY_Chamomile Extract # 11-19-2 | Chem_NY_Chamomile Extract # 11-19-2 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None- free of foreign particles (Done) | [5 empty] | TM 208 | Main_Plans
7A4OA709 | 7A4OA709/Chem_NY_Chamomile Extract # 11-19-2 | Chem_NY_Chamomile Extract # 11-19-2 | NY-Analytical | 792 | Chemical Inspection | Description (TM 208) | Description | Clear Amber to brown liquid having a characteristic odor (Done) | [5 empty] | TM 208 | Main_Plans
7A4OA911 | 7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | Chem_NY_Fragrance Hugs & Kisses Sheer1 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Refractive Index at 25°C | Main_Plans
7A4OA911 | 7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | Chem_NY_Fragrance Hugs & Kisses Sheer1 | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.4537 |  |  | FDA,4 | 1.4536 | 1.4636 | TM 204 | Main_Plans
7A4OA911 | 7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | Chem_NY_Fragrance Hugs & Kisses Sheer1 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Liquid (Done) | [5 empty] | TM 208 | Main_Plans
7A4OA911 | 7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | Chem_NY_Fragrance Hugs & Kisses Sheer1 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to Slightly Yellow (Done) | [5 empty] | TM 208 | Main_Plans
7A4OA911 | 7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | Chem_NY_Fragrance Hugs & Kisses Sheer1 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Main_Plans
7A4OA911 | 7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | Chem_NY_Fragrance Hugs & Kisses Sheer1 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) | [5 empty] | TM 208 | Main_Plans
7A4OA911 | 7A4OA911/Chem_NY_Fragrance Hugs & Kisses Sheer1 | Chem_NY_Fragrance Hugs & Kisses Sheer1 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) | [5 empty] | TM 411 | Main_Plans
7A4OB223 | 7A4OB223/Chem_189A | Chem_189A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Identification | Main_Plans
7A4OB223 | 7A4OB223/Chem_189A | Chem_189A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7A4OB223 | 7A4OB223/Chem_189A | Chem_189A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Translucent crystals or crystalline powder (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OB223 | 7A4OB223/Chem_189A | Chem_189A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to white (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OB223 | 7A4OB223/Chem_189A | Chem_189A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OB223 | 7A4OB223/Chem_189A | Chem_189A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OB339 | 7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.376 |  |  | FDA,3 | 1.376 | 1.378 | TM 204 | Main_Plans
7A4OB339 | 7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 797 | Nonvolatile Residue | End Weight (g) | End Weight | 92.0885 |  | g | FDA,4 |  |  | TM 226 | Main_Plans
7A4OB339 | 7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 797 | Nonvolatile Residue | Nonvolatile Residue (calc) | Nonvolatile Residue | 0 |  | % | FDA,3 |  | 0.005 | TM 226 | Main_Plans
7A4OB339 | 7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 797 | Nonvolatile Residue | Nonvolatile Residue Final (calc) | Nonvolatile Residue Final | 0 |  | % | FDA,3 |  | 0.005 | TM 226 | Main_Plans
7A4OB339 | 7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 797 | Nonvolatile Residue | Start Weight (g) | Start Weight | 92.0885 |  | g | FDA,4 |  |  | TM 226 | Main_Plans
7A4OB339 | 7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 797 | Nonvolatile Residue | Weight of Residue (g) | Weight of Residue | 0 |  | g | FDA,4 |  |  | TM 226 | Main_Plans
7A4OB339 | 7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 795 | Acidity | Acidity (calc) | Acidity | 0.4 |  | mL | FDA,2 |  | 0.7 | TM 375 | Main_Plans
7A4OB339 | 7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 795 | Acidity | Acidity Final (Calc) | Acidity Final (Calc) | 0.425 |  | mL | FDA,2 |  | 0.7 | TM 375 | Main_Plans
7A4OB339 | 7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 795 | Acidity | Titrant End Volume (mL) | Titrant End Volume | 0.5 |  | mL | FDA,2 |  |  | TM 375 | Main_Plans
7A4OB339 | 7A4OB339/Chem_NY_Isopropyl alcohol 99% USP-Inline | Chem_NY_Isopropyl alcohol 99% USP-Inline | NY-Analytical | 795 | Acidity | Titrant Start Volume (mL) | Titrant Start Volume | 0 |  | mL | FDA,2 |  |  | TM 375 | Main_Plans
7A4OB397 | 7A4OB397/Chem_04A | Chem_04A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Assay | Main_Plans
7A4OB397 | 7A4OB397/Chem_04A | Chem_04A | NY-Analytical | 796 | Assay | Assay | Assay | 55.85 |  | % | FDA,1 | 47.5 | 52.5 | TM 301/434 | Main_Plans
7A4OB397 | 7A4OB397/Chem_04A | Chem_04A | NY-Analytical | 796 | Assay | Assay Final | Assay Final | 55.77 |  | % | FDA,1 | 47.5 | 52.5 | TM 301/434 | Main_Plans
7A4OB397 | 7A4OB397/Chem_04A | Chem_04A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear viscous liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OB397 | 7A4OB397/Chem_04A | Chem_04A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to light yellow (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OB397 | 7A4OB397/Chem_04A | Chem_04A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OB397 | 7A4OB397/Chem_04A | Chem_04A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Mild almond (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OB397 | 7A4OB397/Chem_04A | Chem_04A | NY-Analytical | 390 | Identification | Identification (TM 400) | Identification | Formation of precipitate per test (Done) | [5 empty] | TM400/408b/421/422/USP | Main_Plans
7A4OB690 | 7A4OB690/Chem_01A | Chem_01A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Identification | Main_Plans
7A4OB690 | 7A4OB690/Chem_01A | Chem_01A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7A4OB690 | 7A4OB690/Chem_01A | Chem_01A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, mobile liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OB690 | 7A4OB690/Chem_01A | Chem_01A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OB690 | 7A4OB690/Chem_01A | Chem_01A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OB690 | 7A4OB690/Chem_01A | Chem_01A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Solvent (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OF602 | 7A4OF602/Chem_19A | Chem_19A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.968 |  |  | FDA,3 | 0.964 | 0.968 | TM 202 | Main_Plans
7A4OF602 | 7A4OF602/Chem_19A | Chem_19A | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.458 |  |  | FDA,3 | 1.455 | 1.461 | TM 204 | Main_Plans
7A4OF602 | 7A4OF602/Chem_19A | Chem_19A | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) | [5 empty] | TM 411 | Main_Plans
7A4OF602 | 7A4OF602/Chem_19A | Chem_19A | NY-Analytical | 792 | Chemical Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OF602 | 7A4OF602/Chem_19A | Chem_19A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OG169 | 7A4OG169/Chem_20A | Chem_20A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Identification | Main_Plans
7A4OG169 | 7A4OG169/Chem_20A | Chem_20A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7A4OG169 | 7A4OG169/Chem_20A | Chem_20A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OG169 | 7A4OG169/Chem_20A | Chem_20A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Amber (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OG169 | 7A4OG169/Chem_20A | Chem_20A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OH247 | 7A4OH247/Chem_21A | Chem_21A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7A4OH247 | 7A4OH247/Chem_21A | Chem_21A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | White granular material (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OH247 | 7A4OH247/Chem_21A | Chem_21A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Match reference (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OH247 | 7A4OH247/Chem_21A | Chem_21A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OH247 | 7A4OH247/Chem_21A | Chem_21A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match reference (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OK459 | 7A4OK459/Chem_NY_Salicylic Acid USP | Chem_NY_Salicylic Acid USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7A4OK459 | 7A4OK459/Chem_NY_Salicylic Acid USP | Chem_NY_Salicylic Acid USP | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | White crystalline powder (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7A4OK459 | 7A4OK459/Chem_NY_Salicylic Acid USP | Chem_NY_Salicylic Acid USP | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7A4OK459 | 7A4OK459/Chem_NY_Salicylic Acid USP | Chem_NY_Salicylic Acid USP | NY-Analytical | 832 | Chloride | Chloride | Chloride | <0.014 |  | % | FDA,3 |  | 0.014 | USP | Main_Plans
7A4OK459 | 7A4OK459/Chem_NY_Salicylic Acid USP | Chem_NY_Salicylic Acid USP | NY-Analytical | 833 | Heavy Metals | Heavy Metals | Heavy Metals | <20 |  | ug/g | FDA,0 |  | <20 | USP | Main_Plans
7A4OK459 | 7A4OK459/Chem_NY_Salicylic Acid USP | Chem_NY_Salicylic Acid USP | NY-Analytical | 816 | Related Compounds | Impurity | Impurity | 0 |  | % | FDA,4 |  | 0.05 | USP | Main_Plans
7A4OK459 | 7A4OK459/Chem_NY_Salicylic Acid USP | Chem_NY_Salicylic Acid USP | NY-Analytical | 802 | Loss on Drying | Loss on Drying | Loss on Drying | 0.03 |  | % | FDA,3 |  | 0.5 | USP | Main_Plans
7A4OK459 | 7A4OK459/Chem_NY_Salicylic Acid USP | Chem_NY_Salicylic Acid USP | NY-Analytical | 816 | Related Compounds | Phenol | Phenol | 0 |  | % | FDA,4 |  | 0.01 | USP | Main_Plans
7A4OK459 | 7A4OK459/Chem_NY_Salicylic Acid USP | Chem_NY_Salicylic Acid USP | NY-Analytical | 816 | Related Compounds | Related Compound A | Related Compound A | 0.043 |  | % | FDA,4 |  | 0.1 | USP | Main_Plans
7A4OK459 | 7A4OK459/Chem_NY_Salicylic Acid USP | Chem_NY_Salicylic Acid USP | NY-Analytical | 816 | Related Compounds | Related Compound B | Related Compound B | 0.037 |  | % | FDA,4 |  | 0.05 | USP | Main_Plans
7A4OK459 | 7A4OK459/Chem_NY_Salicylic Acid USP | Chem_NY_Salicylic Acid USP | NY-Analytical | 834 | Residue on Ignition | Residue on Ignition | Residue on Ignition | 0.02 |  | % | FDA,2 |  | 0.05 | USP | Main_Plans
7A4OK459 | 7A4OK459/Chem_NY_Salicylic Acid USP | Chem_NY_Salicylic Acid USP | NY-Analytical | 835 | Sulfate Content | Sulfate Content | Sulfate Content | <0.020 |  | % | FDA,3 |  | 0.02 | USP | Main_Plans
7A4OK459 | 7A4OK459/Chem_NY_Salicylic Acid USP | Chem_NY_Salicylic Acid USP | NY-Analytical | 816 | Related Compounds | Total Impurities | Total Impurities | 0.081 |  | % | FDA,4 |  | 0.2 | USP | Main_Plans
7A4OK459 | 7A4OK459/Chem_NY_Salicylic Acid USP | Chem_NY_Salicylic Acid USP | NY-Analytical | 796 | Assay | Assay | Assay | 99.7 |  | % | FDA,1 | 98 | 102 | USP/NF | Main_Plans
7A4OK459 | 7A4OK459/Chem_NY_Salicylic Acid USP | Chem_NY_Salicylic Acid USP | NY-Analytical | 455 | Identification A | Identification A | Identification A | Passes Test (Done) | [5 empty] | USP/NF | Main_Plans
7A4OK459 | 7A4OK459/Chem_NY_Salicylic Acid USP | Chem_NY_Salicylic Acid USP | NY-Analytical | 456 | Identification B | Identification B | Identification B | Passes Test (Done) | [5 empty] | USP/NF | Main_Plans
7A4OM002 | 7A4OM002/Chem_NY_ Dissolvine GL-47-S | Chem_NY_ Dissolvine GL-47-S | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | pH | Main_Plans
7A4OM002 | 7A4OM002/Chem_NY_ Dissolvine GL-47-S | Chem_NY_ Dissolvine GL-47-S | NY-Analytical | 398 | pH | pH (1% w/v aqueous solution) | pH (1% w/v aqueous solution) | 11.14 |  |  | FDA,1 | 11 | 11.8 | TM 200 | Main_Plans
7A4OM002 | 7A4OM002/Chem_NY_ Dissolvine GL-47-S | Chem_NY_ Dissolvine GL-47-S | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) | [5 empty] | TM 208 | Main_Plans
7A4OM002 | 7A4OM002/Chem_NY_ Dissolvine GL-47-S | Chem_NY_ Dissolvine GL-47-S | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Characteristic (Done) | [5 empty] | TM 208 | Main_Plans
7A4OM002 | 7A4OM002/Chem_NY_ Dissolvine GL-47-S | Chem_NY_ Dissolvine GL-47-S | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Main_Plans
7A4OM002 | 7A4OM002/Chem_NY_ Dissolvine GL-47-S | Chem_NY_ Dissolvine GL-47-S | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) | [5 empty] | TM 208 | Main_Plans
7A4OM002 | 7A4OM002/Chem_NY_ Dissolvine GL-47-S | Chem_NY_ Dissolvine GL-47-S | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) | [5 empty] | TM 411 | Main_Plans
7A4OP065 | 7A4OP065/Chem_NY_ Fragrance Plaisir Gourmand 23 | Chem_NY_ Fragrance Plaisir Gourmand 23 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7A4OP065 | 7A4OP065/Chem_NY_ Fragrance Plaisir Gourmand 23 | Chem_NY_ Fragrance Plaisir Gourmand 23 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) | [5 empty] | TM 208 | Main_Plans
7A4OP065 | 7A4OP065/Chem_NY_ Fragrance Plaisir Gourmand 23 | Chem_NY_ Fragrance Plaisir Gourmand 23 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Pale Yellow to Yellow (Done) | [5 empty] | TM 208 | Main_Plans
7A4OP065 | 7A4OP065/Chem_NY_ Fragrance Plaisir Gourmand 23 | Chem_NY_ Fragrance Plaisir Gourmand 23 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Main_Plans
7A4OP065 | 7A4OP065/Chem_NY_ Fragrance Plaisir Gourmand 23 | Chem_NY_ Fragrance Plaisir Gourmand 23 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) | [5 empty] | TM 208 | Main_Plans
7A4OP065 | 7A4OP065/Chem_NY_ Fragrance Plaisir Gourmand 23 | Chem_NY_ Fragrance Plaisir Gourmand 23 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) | [5 empty] | TM 411 | Main_Plans
7A4OP073 | 7A4OP073/Chem_NY_Fragrance Bali Leia 5 | Chem_NY_Fragrance Bali Leia 5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7A4OP073 | 7A4OP073/Chem_NY_Fragrance Bali Leia 5 | Chem_NY_Fragrance Bali Leia 5 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Liquid (Done) | [5 empty] | TM 208 | Main_Plans
7A4OP073 | 7A4OP073/Chem_NY_Fragrance Bali Leia 5 | Chem_NY_Fragrance Bali Leia 5 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Dark yellow to lemon yellow (Done) | [5 empty] | TM 208 | Main_Plans
7A4OP073 | 7A4OP073/Chem_NY_Fragrance Bali Leia 5 | Chem_NY_Fragrance Bali Leia 5 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Main_Plans
7A4OP073 | 7A4OP073/Chem_NY_Fragrance Bali Leia 5 | Chem_NY_Fragrance Bali Leia 5 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) | [5 empty] | TM 208 | Main_Plans
7A4OP073 | 7A4OP073/Chem_NY_Fragrance Bali Leia 5 | Chem_NY_Fragrance Bali Leia 5 | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) | [5 empty] | TM 411 | Main_Plans
7A4OP081 | 7A4OP081/Chem_NY_Fragrance SC 2705 LF | Chem_NY_Fragrance SC 2705 LF | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7A4OP081 | 7A4OP081/Chem_NY_Fragrance SC 2705 LF | Chem_NY_Fragrance SC 2705 LF | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Liquid (Done) | [5 empty] | TM 208 | Main_Plans
7A4OP081 | 7A4OP081/Chem_NY_Fragrance SC 2705 LF | Chem_NY_Fragrance SC 2705 LF | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to very light yellow (Done) | [5 empty] | TM 208 | Main_Plans
7A4OP081 | 7A4OP081/Chem_NY_Fragrance SC 2705 LF | Chem_NY_Fragrance SC 2705 LF | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Main_Plans
7A4OP081 | 7A4OP081/Chem_NY_Fragrance SC 2705 LF | Chem_NY_Fragrance SC 2705 LF | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) | [5 empty] | TM 208 | Main_Plans
7A4OP081 | 7A4OP081/Chem_NY_Fragrance SC 2705 LF | Chem_NY_Fragrance SC 2705 LF | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) | [5 empty] | TM 411 | Main_Plans
7A4OP093 | 7A4OP093/Chem_NY_Essential Oil 27A Fragrance | Chem_NY_Essential Oil 27A Fragrance | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7A4OP093 | 7A4OP093/Chem_NY_Essential Oil 27A Fragrance | Chem_NY_Essential Oil 27A Fragrance | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Liquid (Done) | [5 empty] | TM 208 | Main_Plans
7A4OP093 | 7A4OP093/Chem_NY_Essential Oil 27A Fragrance | Chem_NY_Essential Oil 27A Fragrance | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Amber (Done) | [5 empty] | TM 208 | Main_Plans
7A4OP093 | 7A4OP093/Chem_NY_Essential Oil 27A Fragrance | Chem_NY_Essential Oil 27A Fragrance | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Main_Plans
7A4OP093 | 7A4OP093/Chem_NY_Essential Oil 27A Fragrance | Chem_NY_Essential Oil 27A Fragrance | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) | [5 empty] | TM 208 | Main_Plans
7A4OP093 | 7A4OP093/Chem_NY_Essential Oil 27A Fragrance | Chem_NY_Essential Oil 27A Fragrance | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) | [5 empty] | TM 411 | Main_Plans
7A4OP113 | 7A4OP113/Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | pH | Main_Plans
7A4OP113 | 7A4OP113/Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | NY-Analytical | 398 | pH | pH | pH | 4.11 |  |  | FDA,1 | 2.5 | 6 | TM 200 | Main_Plans
7A4OP113 | 7A4OP113/Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear to slightly cloudy liquid (Done) | [5 empty] | TM 208 | Main_Plans
7A4OP113 | 7A4OP113/Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) | [5 empty] | TM 208 | Main_Plans
7A4OP113 | 7A4OP113/Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Main_Plans
7A4OP113 | 7A4OP113/Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | Chem_NY_Vitis Vinifera (Grape) Fruit Water_01 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) | [5 empty] | TM 208 | Main_Plans
7A4OP197 | 7A4OP197/Chem_NY_D-Panthenol 75W_01 | Chem_NY_D-Panthenol 75W_01 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | pH | Main_Plans
7A4OP197 | 7A4OP197/Chem_NY_D-Panthenol 75W_01 | Chem_NY_D-Panthenol 75W_01 | NY-Analytical | 398 | pH | pH | pH | 6.96 |  |  | FDA,1 | 6 | 7.5 | TM 200 | Main_Plans
7A4OP197 | 7A4OP197/Chem_NY_D-Panthenol 75W_01 | Chem_NY_D-Panthenol 75W_01 | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.466 |  |  | FDA,3 | 1.466 | 1.475 | TM 204 | Main_Plans
7A4OP197 | 7A4OP197/Chem_NY_D-Panthenol 75W_01 | Chem_NY_D-Panthenol 75W_01 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) | [5 empty] | TM 208 | Main_Plans
7A4OP197 | 7A4OP197/Chem_NY_D-Panthenol 75W_01 | Chem_NY_D-Panthenol 75W_01 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Slightly Yellowish or Colorless (Done) | [5 empty] | TM 208 | Main_Plans
7A4OP197 | 7A4OP197/Chem_NY_D-Panthenol 75W_01 | Chem_NY_D-Panthenol 75W_01 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Main_Plans
7A4OP197 | 7A4OP197/Chem_NY_D-Panthenol 75W_01 | Chem_NY_D-Panthenol 75W_01 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) | [5 empty] | TM 208 | Main_Plans
7A4OP197 | 7A4OP197/Chem_NY_D-Panthenol 75W_01 | Chem_NY_D-Panthenol 75W_01 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) | [5 empty] | TM 411 | Main_Plans
7A4OP205 | 7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | Chem_NY_Cucumis Sativus Fruit Extract | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7A4OP205 | 7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | Chem_NY_Cucumis Sativus Fruit Extract | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Transparent Liquid to slightly Turbid (Done) | [5 empty] | TM 208 | Main_Plans
7A4OP205 | 7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | Chem_NY_Cucumis Sativus Fruit Extract | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to pale yellow (Done) | [5 empty] | TM 208 | Main_Plans
7A4OP205 | 7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | Chem_NY_Cucumis Sativus Fruit Extract | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Main_Plans
7A4OP205 | 7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | Chem_NY_Cucumis Sativus Fruit Extract | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) | [5 empty] | TM 208 | Main_Plans
7A4OP205 | 7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | Chem_NY_Cucumis Sativus Fruit Extract | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) | [5 empty] | TM 411 | Main_Plans
7A4OP205 | 7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | Chem_NY_Cucumis Sativus Fruit Extract | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram) | TAMC (Average per Gram) | 0 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
7A4OP205 | 7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | Chem_NY_Cucumis Sativus Fruit Extract | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
7A4OP205 | 7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | Chem_NY_Cucumis Sativus Fruit Extract | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram) | TMYC (Average per Gram) | 0 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
7A4OP205 | 7A4OP205/Chem_NY_Cucumis Sativus Fruit Extract | Chem_NY_Cucumis Sativus Fruit Extract | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
7A4OP345 | 7A4OP345/Chem_195A | Chem_195A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | pH as is | Main_Plans
7A4OP345 | 7A4OP345/Chem_195A | Chem_195A | NY-Analytical | 447 | pH as is | pH | pH | 3.92 |  |  | FDA,1 | 3.5 | 4.5 | TM 200 | Main_Plans
7A4OP345 | 7A4OP345/Chem_195A | Chem_195A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7A4OP345 | 7A4OP345/Chem_195A | Chem_195A | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram) | TAMC (Average per Gram) | 0 |  | CFU/g | FDA,1 |  |  | TM 610 | Main_Plans
7A4OP345 | 7A4OP345/Chem_195A | Chem_195A | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
7A4OP345 | 7A4OP345/Chem_195A | Chem_195A | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram) | TMYC (Average per Gram) | 0 |  | CFU/g | FDA,1 |  |  | TM 610 | Main_Plans
7A4OP345 | 7A4OP345/Chem_195A | Chem_195A | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
7A4OP345 | 7A4OP345/Chem_195A | Chem_195A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Hazy liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OP345 | 7A4OP345/Chem_195A | Chem_195A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Light green/amber (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OP345 | 7A4OP345/Chem_195A | Chem_195A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OP345 | 7A4OP345/Chem_195A | Chem_195A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OP448 | 7A4OP448/Chem_208A | Chem_208A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Certificate of Analysis | Main_Plans
7A4OP448 | 7A4OP448/Chem_208A | Chem_208A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.9801 |  |  | FDA,4 | 0.9694 | 0.9894 | TM 202 | Main_Plans
7A4OP448 | 7A4OP448/Chem_208A | Chem_208A | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.469 |  |  | FDA,4 | 1.4655 | 1.4855 | TM 204 | Main_Plans
7A4OP448 | 7A4OP448/Chem_208A | Chem_208A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7A4OP448 | 7A4OP448/Chem_208A | Chem_208A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear to pale yellow liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OP448 | 7A4OP448/Chem_208A | Chem_208A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OP448 | 7A4OP448/Chem_208A | Chem_208A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Fresh Bouquet (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OP567 | 7A4OP567/Chem_NY_Surfonic L12-6_Surfactant | Chem_NY_Surfonic L12-6_Surfactant | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | COA | Main_Plans
7A4OP567 | 7A4OP567/Chem_NY_Surfonic L12-6_Surfactant | Chem_NY_Surfonic L12-6_Surfactant | NY-Analytical | 398 | pH | pH (1%) in 10:6 IPA: H2O | pH (1%) in 10:6 IPA: H2O | 6.9 |  |  | FDA,1 | 6.5 | 8 | TM 200 | Main_Plans
7A4OP567 | 7A4OP567/Chem_NY_Surfonic L12-6_Surfactant | Chem_NY_Surfonic L12-6_Surfactant | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear to slightly turbid, free of suspended matter (Done) | [5 empty] | TM 208 | Main_Plans
7A4OP567 | 7A4OP567/Chem_NY_Surfonic L12-6_Surfactant | Chem_NY_Surfonic L12-6_Surfactant | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to Slightly Yellow (Done) | [5 empty] | TM 208 | Main_Plans
7A4OP567 | 7A4OP567/Chem_NY_Surfonic L12-6_Surfactant | Chem_NY_Surfonic L12-6_Surfactant | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Main_Plans
7A4OP567 | 7A4OP567/Chem_NY_Surfonic L12-6_Surfactant | Chem_NY_Surfonic L12-6_Surfactant | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | To match standard (Done) | [5 empty] | TM 208 | Main_Plans
7A4OP601 | 7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Iodine Value | Main_Plans
7A4OP601 | 7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 828 | Iodine Value | Iodine Value | Iodine Value | 2.2 |  | g I/100g | FDA,2 |  | 5 | J&J TM-003029 | Main_Plans
7A4OP601 | 7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 828 | Iodine Value | Iodine Value Final | Iodine Value Final | 2.2 |  | g I/100g | FDA,2 |  | 5 | J&J TM-003029 | Main_Plans
7A4OP601 | 7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Light yellow pasty fluid (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7A4OP601 | 7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7A4OP601 | 7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) | [5 empty] | J&J TMWW78013 | Main_Plans
7A4OP601 | 7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 398 | pH | pH (5% Solution in DI Water:Methanol 1:1) | pH (5% Solution in DI Water:Methanol 1:1) | 7 |  |  | FDA,1 | 6.5 | 7.5 | J&J TMWW78014 | Main_Plans
7A4OP601 | 7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 810 | Saponification Value | Saponification Value | Saponification Value | 54.2 |  | KOH/g | FDA,1 | 47 | 57 | J&J TMWW78016 | Main_Plans
7A4OP601 | 7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 810 | Saponification Value | Saponification Value Final | Saponification Value Final | 53.35 |  | KOH/g | FDA,1 | 47 | 57 | J&J TMWW78016 | Main_Plans
7A4OP601 | 7A4OP601/Chem_NY_Tegosoft LSE 65 K Soft | Chem_NY_Tegosoft LSE 65 K Soft | NY-Analytical | 801 | % Water | Water Content | Water Content | 35.15 |  | % | FDA,2 | 33 | 37 | J&J TMWW78017 | Main_Plans
7A4OQ028 | 7A4OQ028/Chem_NY_Emulgin VL-75 | Chem_NY_Emulgin VL-75 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | pH as is | Main_Plans
7A4OQ028 | 7A4OQ028/Chem_NY_Emulgin VL-75 | Chem_NY_Emulgin VL-75 | NY-Analytical | 447 | pH as is | pH | pH | 5.76 |  |  | FDA,1 | 5 | 6 | TM 200 | Main_Plans
7A4OQ028 | 7A4OQ028/Chem_NY_Emulgin VL-75 | Chem_NY_Emulgin VL-75 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7A4OQ028 | 7A4OQ028/Chem_NY_Emulgin VL-75 | Chem_NY_Emulgin VL-75 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear to slightly Turbid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OQ028 | 7A4OQ028/Chem_NY_Emulgin VL-75 | Chem_NY_Emulgin VL-75 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Yellowish Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OQ028 | 7A4OQ028/Chem_NY_Emulgin VL-75 | Chem_NY_Emulgin VL-75 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OQ028 | 7A4OQ028/Chem_NY_Emulgin VL-75 | Chem_NY_Emulgin VL-75 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OQ084 | 7A4OQ084/Chem_23A | Chem_23A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Identification | Main_Plans
7A4OQ084 | 7A4OQ084/Chem_23A | Chem_23A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7A4OQ084 | 7A4OQ084/Chem_23A | Chem_23A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OQ084 | 7A4OQ084/Chem_23A | Chem_23A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | None or Pale Straw (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OQ084 | 7A4OQ084/Chem_23A | Chem_23A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OQ084 | 7A4OQ084/Chem_23A | Chem_23A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Little or no distinguisable odor (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OQ132 | 7A4OQ132/Chem_25A | Chem_25A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Identification | Main_Plans
7A4OQ132 | 7A4OQ132/Chem_25A | Chem_25A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7A4OQ132 | 7A4OQ132/Chem_25A | Chem_25A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Transparent liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OQ132 | 7A4OQ132/Chem_25A | Chem_25A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to pale yellow (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OQ132 | 7A4OQ132/Chem_25A | Chem_25A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OQ132 | 7A4OQ132/Chem_25A | Chem_25A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Mild (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OQ229 | 7A4OQ229/Chem_Trisodium Citrate Dihydrate USP | Chem_Trisodium Citrate Dihydrate USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | COA | Main_Plans
7A4OQ229 | 7A4OQ229/Chem_Trisodium Citrate Dihydrate USP | Chem_Trisodium Citrate Dihydrate USP | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Fine granular (Done) | [5 empty] | TM 208 | Main_Plans
7A4OQ229 | 7A4OQ229/Chem_Trisodium Citrate Dihydrate USP | Chem_Trisodium Citrate Dihydrate USP | NY-Analytical | 792 | Chemical Inspection | Color | Color | White (Done) | [5 empty] | TM 208 | Main_Plans
7A4OQ229 | 7A4OQ229/Chem_Trisodium Citrate Dihydrate USP | Chem_Trisodium Citrate Dihydrate USP | NY-Analytical | 792 | Chemical Inspection | Contamination | Contamination | None (Done) | [5 empty] | TM 208 | Main_Plans
7A4OQ229 | 7A4OQ229/Chem_Trisodium Citrate Dihydrate USP | Chem_Trisodium Citrate Dihydrate USP | NY-Analytical | 792 | Chemical Inspection | Odor | Odor | Odorless (Done) | [5 empty] | TM 208 | Main_Plans
7A4OQ229 | 7A4OQ229/Chem_Trisodium Citrate Dihydrate USP | Chem_Trisodium Citrate Dihydrate USP | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7A4OQ260 | 7A4OQ260/Chem_29A | Chem_29A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Specific Gravity at 25°C | Main_Plans
7A4OQ260 | 7A4OQ260/Chem_29A | Chem_29A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.9659 |  |  | FDA,3 | 0.954 | 0.974 | TM 202 | Main_Plans
7A4OQ260 | 7A4OQ260/Chem_29A | Chem_29A | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.48 |  |  | FDA,3 | 1.464 | 1.484 | TM 204 | Main_Plans
7A4OQ260 | 7A4OQ260/Chem_29A | Chem_29A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7A4OQ260 | 7A4OQ260/Chem_29A | Chem_29A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OQ260 | 7A4OQ260/Chem_29A | Chem_29A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to pale yellow (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OQ260 | 7A4OQ260/Chem_29A | Chem_29A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OQ260 | 7A4OQ260/Chem_29A | Chem_29A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Powder Rose (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OQ518 | 7A4OQ518/Chem_31A | Chem_31A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | pH | Main_Plans
7A4OQ518 | 7A4OQ518/Chem_31A | Chem_31A | NY-Analytical | 398 | pH | pH (10% Solution) | pH (10% Solution) | 7.57 |  |  | FDA,1 | 6 | 9 | TM 200 | Main_Plans
7A4OQ518 | 7A4OQ518/Chem_31A | Chem_31A | NY-Analytical | 796 | Assay | Assay | Assay | 51.97 |  | % | FDA,1 | 50 | 54 | TM 301/434 | Main_Plans
7A4OQ518 | 7A4OQ518/Chem_31A | Chem_31A | NY-Analytical | 796 | Assay | Assay Final | Assay Final | 51.98 |  | % | FDA,1 | 50 | 54 | TM 301/434 | Main_Plans
7A4OQ518 | 7A4OQ518/Chem_31A | Chem_31A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear viscous liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OQ518 | 7A4OQ518/Chem_31A | Chem_31A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to light straw (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OQ518 | 7A4OQ518/Chem_31A | Chem_31A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OQ518 | 7A4OQ518/Chem_31A | Chem_31A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Pleasant almond (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OQ890 | 7A4OQ890/Chem_34A | Chem_34A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7A4OQ890 | 7A4OQ890/Chem_34A | Chem_34A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Granular Powder (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OQ890 | 7A4OQ890/Chem_34A | Chem_34A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OQ890 | 7A4OQ890/Chem_34A | Chem_34A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OQ890 | 7A4OQ890/Chem_34A | Chem_34A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OS407 | 7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7A4OS407 | 7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 792 | Chemical Inspection | Description | Description | Transparent, colorless, mobile, volatile liquid (Done) | [5 empty] | ORGANOLEPTIC | Main_Plans
7A4OS407 | 7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 795 | Acidity | Acidity (calc) | Acidity | 0.4 |  | mL | FDA,2 |  | 0.7 | USP | Main_Plans
7A4OS407 | 7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 795 | Acidity | Acidity Final (Calc) | Acidity Final (Calc) | 0.45 |  | mL | FDA,2 |  | 0.7 | USP | Main_Plans
7A4OS407 | 7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 796 | Assay | Assay | Assay | 100 |  | % | FDA,1 | 99 |  | USP | Main_Plans
7A4OS407 | 7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 796 | Assay | Assay Final | Assay Final | 99.975 |  | % | FDA,1 | 99 |  | USP | Main_Plans
7A4OS407 | 7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 797 | Nonvolatile Residue | End Weight (g) | End Weight | 90.0835 |  | g | FDA,4 |  |  | USP | Main_Plans
7A4OS407 | 7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 797 | Nonvolatile Residue | Nonvolatile Residue (calc) | Nonvolatile Residue | 0 |  | % | FDA,3 |  | 0.005 | USP | Main_Plans
7A4OS407 | 7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 797 | Nonvolatile Residue | Nonvolatile Residue Final (calc) | Nonvolatile Residue Final | 0.0008 |  | % | FDA,3 |  | 0.005 | USP | Main_Plans
7A4OS407 | 7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.376 |  |  | FDA,3 | 1.376 | 1.378 | USP | Main_Plans
7A4OS407 | 7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 797 | Nonvolatile Residue | Start Weight (g) | Start Weight | 90.7647 |  | g | FDA,4 |  |  | USP | Main_Plans
7A4OS407 | 7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 795 | Acidity | Titrant End Volume (mL) | Titrant End Volume | 0.9 |  | mL | FDA,2 |  |  | USP | Main_Plans
7A4OS407 | 7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 795 | Acidity | Titrant Start Volume (mL) | Titrant Start Volume | 0.5 |  | mL | FDA,2 |  |  | USP | Main_Plans
7A4OS407 | 7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 797 | Nonvolatile Residue | Weight of Residue (g) | Weight of Residue | 0.0004 |  | g | FDA,4 |  |  | USP | Main_Plans
7A4OS407 | 7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR conform to standard (Done) | [5 empty] | USP <197F> | Main_Plans
7A4OS407 | 7A4OS407/Chem_NY_Solumed Isopropyl Alcohol 99% USP | Chem_NY_Solumed Isopropyl Alcohol 99% USP | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.7832 |  |  | FDA,3 | 0.783 | 0.787 | USP/NF <841> | Main_Plans
7A4OS411 | 7A4OS411/Chem_36A | Chem_36A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7A4OS411 | 7A4OS411/Chem_36A | Chem_36A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Free Flowing Powder (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OS411 | 7A4OS411/Chem_36A | Chem_36A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Brownish-Red (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OS411 | 7A4OS411/Chem_36A | Chem_36A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OS411 | 7A4OS411/Chem_36A | Chem_36A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OS411 | 7A4OS411/Chem_36A | Chem_36A | NY-Analytical | 796 | Assay | Assay as is (Without Drying) | Assay as is (Without Drying) | 10.17 |  | % | FDA,1 |  |  | USP | Main_Plans
7A4OS411 | 7A4OS411/Chem_36A | Chem_36A | NY-Analytical | 796 | Assay | Assay as is Final (Without Drying) | Assay as is Final (Without Drying) | 10.15 |  | % | FDA,1 |  |  | USP | Main_Plans
7A4OS411 | 7A4OS411/Chem_36A | Chem_36A | NY-Analytical | 796 | Assay | Assay on Dried Basis | Assay on Dried Basis | 10.59 |  | % | FDA,1 | 9 | 12 | USP | Main_Plans
7A4OS411 | 7A4OS411/Chem_36A | Chem_36A | NY-Analytical | 796 | Assay | Assay on Dried Basis Final | Assay on Dried Basis Final | 10.54 |  | % | FDA,1 | 9 | 12 | USP | Main_Plans
7A4OS411 | 7A4OS411/Chem_36A | Chem_36A | NY-Analytical | 802 | Loss on Drying | Loss on Drying | Loss on Drying | 5.55 |  | % | FDA,1 |  | 8 | USP | Main_Plans
7A4OS411 | 7A4OS411/Chem_36A | Chem_36A | NY-Analytical | 802 | Loss on Drying | Loss on Drying Final | Loss on Drying Final | 5.55 |  | % | FDA,1 |  | 8 | USP | Main_Plans
7A4OU351 | 7A4OU351/Chem_NY_Dehydrated Alcohol USP | Chem_NY_Dehydrated Alcohol USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | COA | Main_Plans
7A4OU351 | 7A4OU351/Chem_NY_Dehydrated Alcohol USP | Chem_NY_Dehydrated Alcohol USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis (ELLI) | Certificate of Analysis | COA Meets Specification (Done) | [5 empty] | COA | Main_Plans
7A4OU351 | 7A4OU351/Chem_NY_Dehydrated Alcohol USP | Chem_NY_Dehydrated Alcohol USP | NY-Analytical | 822 | Specific Gravity @ 15.56°C/ 60° F | Specific Gravity @ 15.56°C | Specific Gravity at 15.56°C | 0.794 |  |  | FDA,4 |  | 0.7962 | TM 202 | Main_Plans
7A4OU351 | 7A4OU351/Chem_NY_Dehydrated Alcohol USP | Chem_NY_Dehydrated Alcohol USP | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) | [5 empty] | TM 208 | Main_Plans
7A4OU351 | 7A4OU351/Chem_NY_Dehydrated Alcohol USP | Chem_NY_Dehydrated Alcohol USP | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Clear to Colorless and Match standard (Done) | [5 empty] | TM 208 | Main_Plans
7A4OU351 | 7A4OU351/Chem_NY_Dehydrated Alcohol USP | Chem_NY_Dehydrated Alcohol USP | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Main_Plans
7A4OU351 | 7A4OU351/Chem_NY_Dehydrated Alcohol USP | Chem_NY_Dehydrated Alcohol USP | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Alcohol (Done) | [5 empty] | TM 208 | Main_Plans
7A4OU351 | 7A4OU351/Chem_NY_Dehydrated Alcohol USP | Chem_NY_Dehydrated Alcohol USP | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Compare to reference Standard (Done) | [5 empty] | TM 411 | Main_Plans
7A4OU555 | 7A4OU555/Chem_NY_Phenylephrine HCL | Chem_NY_Phenylephrine HCL | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7A4OU555 | 7A4OU555/Chem_NY_Phenylephrine HCL | Chem_NY_Phenylephrine HCL | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Crystalline Powder or Crystals (Done) | [5 empty] | TM 208 | Main_Plans
7A4OU555 | 7A4OU555/Chem_NY_Phenylephrine HCL | Chem_NY_Phenylephrine HCL | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White to almost white (Done) | [5 empty] | TM 208 | Main_Plans
7A4OU555 | 7A4OU555/Chem_NY_Phenylephrine HCL | Chem_NY_Phenylephrine HCL | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Main_Plans
7A4OU555 | 7A4OU555/Chem_NY_Phenylephrine HCL | Chem_NY_Phenylephrine HCL | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) | [5 empty] | TM 208 | Main_Plans
7A4OU555 | 7A4OU555/Chem_NY_Phenylephrine HCL | Chem_NY_Phenylephrine HCL | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7A4OU555 | 7A4OU555/Chem_NY_Phenylephrine HCL | Chem_NY_Phenylephrine HCL | NY-Analytical | 812 | Organic Impurites | Any individual unspecified impuritv | Any individual unspecified impuritv |  |  | % | FDA,2 |  | 0.1 | USP | Main_Plans
7A4OU555 | 7A4OU555/Chem_NY_Phenylephrine HCL | Chem_NY_Phenylephrine HCL | NY-Analytical | 796 | Assay | Assay | Assay | 99.4 |  | % | FDA,1 | 98 | 102 | USP | Main_Plans
7A4OU555 | 7A4OU555/Chem_NY_Phenylephrine HCL | Chem_NY_Phenylephrine HCL | NY-Analytical | 802 | Loss on Drying | Loss on Drying | Loss on Drying |  |  | % | FDA,1 |  | 1 | USP | Main_Plans
7A4OU555 | 7A4OU555/Chem_NY_Phenylephrine HCL | Chem_NY_Phenylephrine HCL | NY-Analytical | 802 | Loss on Drying | Loss on Drying Final | Loss on Drying Final |  |  | % | FDA,1 |  | 1 | USP | Main_Plans
7A4OU555 | 7A4OU555/Chem_NY_Phenylephrine HCL | Chem_NY_Phenylephrine HCL | NY-Analytical | 812 | Organic Impurites | Norphenylephrine | Norphenylephrine |  |  | % | FDA,2 |  | 0.1 | USP | Main_Plans
7A4OU555 | 7A4OU555/Chem_NY_Phenylephrine HCL | Chem_NY_Phenylephrine HCL | NY-Analytical | 1043 | Optical Rotation | Optical Rotation | Optical Rotation |  |  | ° | FDA,0 | -47 | -43 | USP | Main_Plans
7A4OU555 | 7A4OU555/Chem_NY_Phenylephrine HCL | Chem_NY_Phenylephrine HCL | NY-Analytical | 812 | Organic Impurites | Phenylephrine Related Compound C | Phenylephrine Related Compound C |  |  | % | FDA,2 |  | 0.1 | USP | Main_Plans
7A4OU555 | 7A4OU555/Chem_NY_Phenylephrine HCL | Chem_NY_Phenylephrine HCL | NY-Analytical | 812 | Organic Impurites | Phenylephrine Related Compound D | Phenylephrine Related Compound D |  |  | % | FDA,2 |  | 0.1 | USP | Main_Plans
7A4OU555 | 7A4OU555/Chem_NY_Phenylephrine HCL | Chem_NY_Phenylephrine HCL | NY-Analytical | 812 | Organic Impurites | Phenylephrine Related Compound E | Phenylephrine Related Compound E |  |  | % | FDA,2 |  | 0.1 | USP | Main_Plans
7A4OU555 | 7A4OU555/Chem_NY_Phenylephrine HCL | Chem_NY_Phenylephrine HCL | NY-Analytical | 834 | Residue on Ignition | Residue on Ignition | Residue on Ignition |  |  | % | FDA,1 |  | 0.2 | USP | Main_Plans
7A4OU555 | 7A4OU555/Chem_NY_Phenylephrine HCL | Chem_NY_Phenylephrine HCL | NY-Analytical | 835 | Sulfate Content | Sulfate | Sulfate |  |  | % | FDA,2 |  | 0.2 | USP | Main_Plans
7A4OU555 | 7A4OU555/Chem_NY_Phenylephrine HCL | Chem_NY_Phenylephrine HCL | NY-Analytical | 812 | Organic Impurites | Total Impurities | Total Impurities |  |  | % | FDA,2 |  | 0.2 | USP | Main_Plans
7A4OX110 | 7A4OX110/Chem_39A | Chem_39A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Identification | Main_Plans
7A4OX110 | 7A4OX110/Chem_39A | Chem_39A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | [6 empty] | TM 411 | Main_Plans
7A4OX110 | 7A4OX110/Chem_39A | Chem_39A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | [6 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OX110 | 7A4OX110/Chem_39A | Chem_39A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | [6 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OX110 | 7A4OX110/Chem_39A | Chem_39A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | [6 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OX110 | 7A4OX110/Chem_39A | Chem_39A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | [6 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7A4OX231 | 7A4OX231/Chem_NY_D&C Yellow 10 | Chem_NY_D&C Yellow 10 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Appearance | Main_Plans
7A4OX231 | 7A4OX231/Chem_NY_D&C Yellow 10 | Chem_NY_D&C Yellow 10 | NY-Analytical | 397 | Appearance | Physical Properties | Physical Properties | Yellow Powder (Done) | [5 empty] | TM 208 | Main_Plans
7A4OX231 | 7A4OX231/Chem_NY_D&C Yellow 10 | Chem_NY_D&C Yellow 10 | NY-Analytical | 390 | Identification | Spectrophotometric | Spectrophotometric | 410.88 |  | λ | FDA,0 | 408 | 416 | TM 3014 | Main_Plans
7A4OX231 | 7A4OX231/Chem_NY_D&C Yellow 10 | Chem_NY_D&C Yellow 10 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram) | TAMC (Average per Gram) | 0 |  | CFU/g | FDA,1 |  |  | TM 610 | Main_Plans
7A4OX231 | 7A4OX231/Chem_NY_D&C Yellow 10 | Chem_NY_D&C Yellow 10 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
7A4OX231 | 7A4OX231/Chem_NY_D&C Yellow 10 | Chem_NY_D&C Yellow 10 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram) | TMYC (Average per Gram) | 0 |  | CFU/g | FDA,1 |  |  | TM 610 | Main_Plans
7A4OX231 | 7A4OX231/Chem_NY_D&C Yellow 10 | Chem_NY_D&C Yellow 10 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
7ACHE61P | 7ACHE61P/Chem_NY_Sodium Benzoate | Chem_NY_Sodium Benzoate | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7ACHE61P | 7ACHE61P/Chem_NY_Sodium Benzoate | Chem_NY_Sodium Benzoate | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Extruded Pellets (Done) | [5 empty] | TM 208 | Main_Plans
7ACHE61P | 7ACHE61P/Chem_NY_Sodium Benzoate | Chem_NY_Sodium Benzoate | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) | [5 empty] | TM 208 | Main_Plans
7ACHE61P | 7ACHE61P/Chem_NY_Sodium Benzoate | Chem_NY_Sodium Benzoate | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Main_Plans
7ACHE61P | 7ACHE61P/Chem_NY_Sodium Benzoate | Chem_NY_Sodium Benzoate | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless (Done) | [5 empty] | TM 208 | Main_Plans
7ACHE61P | 7ACHE61P/Chem_NY_Sodium Benzoate | Chem_NY_Sodium Benzoate | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7B4AA580 | 7B4AA580/Chem_NY_Acetic Acid - USP | Chem_NY_Acetic Acid - USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7B4AA580 | 7B4AA580/Chem_NY_Acetic Acid - USP | Chem_NY_Acetic Acid - USP | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4AA580 | 7B4AA580/Chem_NY_Acetic Acid - USP | Chem_NY_Acetic Acid - USP | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4AA580 | 7B4AA580/Chem_NY_Acetic Acid - USP | Chem_NY_Acetic Acid - USP | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4AA580 | 7B4AA580/Chem_NY_Acetic Acid - USP | Chem_NY_Acetic Acid - USP | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Pungent (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4AL930 | 7B4AL930/Chem_42A | Chem_42A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | pH as is | Main_Plans
7B4AL930 | 7B4AL930/Chem_42A | Chem_42A | NY-Analytical | 447 | pH as is | pH | pH |  |  |  | FDA,1 | 1.1 | 1.5 | TM 200 | Main_Plans
7B4AL930 | 7B4AL930/Chem_42A | Chem_42A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.2444 |  |  | FDA,2 | 1.21 | 1.27 | TM 202 | Main_Plans
7B4AL930 | 7B4AL930/Chem_42A | Chem_42A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7B4AL930 | 7B4AL930/Chem_42A | Chem_42A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4AL930 | 7B4AL930/Chem_42A | Chem_42A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Light Yellow (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4AL930 | 7B4AL930/Chem_42A | Chem_42A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4AL930 | 7B4AL930/Chem_42A | Chem_42A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4FD110 | 7B4FD110/Chem_44A | Chem_44A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Identification | Main_Plans
7B4FD110 | 7B4FD110/Chem_44A | Chem_44A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7B4FD110 | 7B4FD110/Chem_44A | Chem_44A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | White granular powder (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4FD110 | 7B4FD110/Chem_44A | Chem_44A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4FD110 | 7B4FD110/Chem_44A | Chem_44A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4GB598 | 7B4GB598/Chem_45A | Chem_45A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7B4GB598 | 7B4GB598/Chem_45A | Chem_45A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Grain Powder (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4GB598 | 7B4GB598/Chem_45A | Chem_45A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Off-White (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4GB598 | 7B4GB598/Chem_45A | Chem_45A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4GB598 | 7B4GB598/Chem_45A | Chem_45A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OA242 | 7B4OA242/Chem_48A | Chem_48A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Identification | Main_Plans
7B4OA242 | 7B4OA242/Chem_48A | Chem_48A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7B4OA242 | 7B4OA242/Chem_48A | Chem_48A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Crystals/crystalline powder (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OA242 | 7B4OA242/Chem_48A | Chem_48A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to white (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OA242 | 7B4OA242/Chem_48A | Chem_48A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OA242 | 7B4OA242/Chem_48A | Chem_48A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OA347 | 7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | Chem_NY_Calming Waters INCI w/symfresh | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Specific Gravity at 25°C | Main_Plans
7B4OA347 | 7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | Chem_NY_Calming Waters INCI w/symfresh | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.964 |  |  | FDA,3 | 0.959 | 0.969 | TM 202 | Main_Plans
7B4OA347 | 7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | Chem_NY_Calming Waters INCI w/symfresh | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.472 |  |  | FDA,3 | 1.465 | 1.475 | TM 204 | Main_Plans
7B4OA347 | 7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | Chem_NY_Calming Waters INCI w/symfresh | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) | [5 empty] | TM 208 | Main_Plans
7B4OA347 | 7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | Chem_NY_Calming Waters INCI w/symfresh | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to pale yellow (Done) | [5 empty] | TM 208 | Main_Plans
7B4OA347 | 7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | Chem_NY_Calming Waters INCI w/symfresh | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Main_Plans
7B4OA347 | 7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | Chem_NY_Calming Waters INCI w/symfresh | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) | [5 empty] | TM 208 | Main_Plans
7B4OA347 | 7B4OA347/Chem_NY_Calming Waters INCI w/symfresh | Chem_NY_Calming Waters INCI w/symfresh | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) | [5 empty] | TM 411 | Main_Plans
7B4OA709 | 7B4OA709/Chem_NY_Bell fragrance | Chem_NY_Bell fragrance | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Specific Gravity at 25°C | Main_Plans
7B4OA709 | 7B4OA709/Chem_NY_Bell fragrance | Chem_NY_Bell fragrance | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.0639 |  |  | FDA,3 | 1.06 | 1.07 | TM 202 | Main_Plans
7B4OA709 | 7B4OA709/Chem_NY_Bell fragrance | Chem_NY_Bell fragrance | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.468 |  |  | FDA,3 | 1.467 | 1.477 | TM 204 | Main_Plans
7B4OA709 | 7B4OA709/Chem_NY_Bell fragrance | Chem_NY_Bell fragrance | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) | [5 empty] | TM 208 | Main_Plans
7B4OA709 | 7B4OA709/Chem_NY_Bell fragrance | Chem_NY_Bell fragrance | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Pale to light yellow (Done) | [5 empty] | TM 208 | Main_Plans
7B4OA709 | 7B4OA709/Chem_NY_Bell fragrance | Chem_NY_Bell fragrance | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Main_Plans
7B4OA709 | 7B4OA709/Chem_NY_Bell fragrance | Chem_NY_Bell fragrance | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Conforms to Standard (Done) | [5 empty] | TM 208 | Main_Plans
7B4OB690 | 7B4OB690/Chem_50A | Chem_50A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7B4OB690 | 7B4OB690/Chem_50A | Chem_50A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, mobile liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OB690 | 7B4OB690/Chem_50A | Chem_50A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OB690 | 7B4OB690/Chem_50A | Chem_50A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OB690 | 7B4OB690/Chem_50A | Chem_50A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Solvent (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OD232 | 7B4OD232/Chem_52A | Chem_52A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Specific Gravity at 25°C | Main_Plans
7B4OD232 | 7B4OD232/Chem_52A | Chem_52A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.0361 |  |  | FDA,3 | 1.035 | 1.037 | TM 202 | Main_Plans
7B4OD232 | 7B4OD232/Chem_52A | Chem_52A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7B4OD232 | 7B4OD232/Chem_52A | Chem_52A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear viscous liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OD232 | 7B4OD232/Chem_52A | Chem_52A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OD232 | 7B4OD232/Chem_52A | Chem_52A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OD232 | 7B4OD232/Chem_52A | Chem_52A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OP027 | 7B4OP027/Chem_NY_Abil B88183 PH | Chem_NY_Abil B88183 PH | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7B4OP027 | 7B4OP027/Chem_NY_Abil B88183 PH | Chem_NY_Abil B88183 PH | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear to slightly turbid.Colorless to pale yellow,slightly viscous liquid (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7B4OP027 | 7B4OP027/Chem_NY_Abil B88183 PH | Chem_NY_Abil B88183 PH | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7B4OP027 | 7B4OP027/Chem_NY_Abil B88183 PH | Chem_NY_Abil B88183 PH | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) | [5 empty] | J&J TMWW78013 | Main_Plans
7B4OP027 | 7B4OP027/Chem_NY_Abil B88183 PH | Chem_NY_Abil B88183 PH | NY-Microbiology | 509 | J&J_Pathogens | Biggy or SDA | Biggy or SDA | No Growth_SDA | [5 empty] | TM 869 | Main_Plans
7B4OP027 | 7B4OP027/Chem_NY_Abil B88183 PH | Chem_NY_Abil B88183 PH | NY-Microbiology | 509 | J&J_Pathogens | Cetrimide | Cetrimide | No Growth | [5 empty] | TM 869 | Main_Plans
7B4OP027 | 7B4OP027/Chem_NY_Abil B88183 PH | Chem_NY_Abil B88183 PH | NY-Microbiology | 524 | J&J Fungal | Fungal | Fungal | 0 |  | CFU/g | FDA,0 | 0 | 200 | TM 869 | Main_Plans
7B4OP027 | 7B4OP027/Chem_NY_Abil B88183 PH | Chem_NY_Abil B88183 PH | NY-Microbiology | 509 | J&J_Pathogens | MacConkey | MacConkey | No Growth | [5 empty] | TM 869 | Main_Plans
7B4OP027 | 7B4OP027/Chem_NY_Abil B88183 PH | Chem_NY_Abil B88183 PH | NY-Microbiology | 509 | J&J_Pathogens | MSA or VJ | MSA or VJ | No Growth_MSA | [5 empty] | TM 869 | Main_Plans
7B4OP027 | 7B4OP027/Chem_NY_Abil B88183 PH | Chem_NY_Abil B88183 PH | NY-Microbiology | 700 | J&J SDA Controls | Negative Control | Negative Control | No Growth SDA | [5 empty] | TM 869 | Main_Plans
7B4OP027 | 7B4OP027/Chem_NY_Abil B88183 PH | Chem_NY_Abil B88183 PH | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | NegativeControl | No Growth TSB | [5 empty] | TM 869 | Main_Plans
7B4OP027 | 7B4OP027/Chem_NY_Abil B88183 PH | Chem_NY_Abil B88183 PH | NY-Microbiology | 509 | J&J_Pathogens | TSA | TSA | No Growth | [5 empty] | TM 869 | Main_Plans
7B4OP027 | 7B4OP027/Chem_NY_Abil B88183 PH | Chem_NY_Abil B88183 PH | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | TSA JJ | 0 |  | cfu/mL | FDA,0 | 0 | 200 | TM 869 | Main_Plans
7B4OP027 | 7B4OP027/Chem_NY_Abil B88183 PH | Chem_NY_Abil B88183 PH | NY-Microbiology | 517 | J&J_Enrichment | TSB | TSB | Not Turbid | [5 empty] | TM 869 | Main_Plans
7B4OP065 | 7B4OP065/Chem_NY_Chamomile Flower Extract | Chem_NY_Chamomile Flower Extract | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7B4OP065 | 7B4OP065/Chem_NY_Chamomile Flower Extract | Chem_NY_Chamomile Flower Extract | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Liquid (Done) | [5 empty] | TM 208 | Main_Plans
7B4OP065 | 7B4OP065/Chem_NY_Chamomile Flower Extract | Chem_NY_Chamomile Flower Extract | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Greenish Yellow to Yellow Orange (Done) | [5 empty] | TM 208 | Main_Plans
7B4OP065 | 7B4OP065/Chem_NY_Chamomile Flower Extract | Chem_NY_Chamomile Flower Extract | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Main_Plans
7B4OP065 | 7B4OP065/Chem_NY_Chamomile Flower Extract | Chem_NY_Chamomile Flower Extract | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) | [5 empty] | TM 208 | Main_Plans
7B4OP065 | 7B4OP065/Chem_NY_Chamomile Flower Extract | Chem_NY_Chamomile Flower Extract | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) | [5 empty] | TM 411 | Main_Plans
7B4OP065 | 7B4OP065/Chem_NY_Chamomile Flower Extract | Chem_NY_Chamomile Flower Extract | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram) | TAMC (Average per Gram) | 0 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
7B4OP065 | 7B4OP065/Chem_NY_Chamomile Flower Extract | Chem_NY_Chamomile Flower Extract | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
7B4OP065 | 7B4OP065/Chem_NY_Chamomile Flower Extract | Chem_NY_Chamomile Flower Extract | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram) | TMYC (Average per Gram) |  |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
7B4OP065 | 7B4OP065/Chem_NY_Chamomile Flower Extract | Chem_NY_Chamomile Flower Extract | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
7B4OP065 | 7B4OP065/Chem_NY_Chamomile Flower Extract | Chem_NY_Chamomile Flower Extract | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Non-Subculturable Growth (Done) | [5 empty] | WI QAS 076 | Main_Plans
7B4OP093 | 7B4OP093/Chem_NY_PEG-4 | Chem_NY_PEG-4 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | pH | Main_Plans
7B4OP093 | 7B4OP093/Chem_NY_PEG-4 | Chem_NY_PEG-4 | NY-Analytical | 398 | pH | pH (5% aqueous @25 degree C) | pH (5% aqueous @25 degree C) | 6.01 |  |  | FDA,1 | 4.5 | 7.5 | TM 200 | Main_Plans
7B4OP093 | 7B4OP093/Chem_NY_PEG-4 | Chem_NY_PEG-4 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Not More Than Slightly Hazy (Done) | [5 empty] | TM 208 | Main_Plans
7B4OP093 | 7B4OP093/Chem_NY_PEG-4 | Chem_NY_PEG-4 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) | [5 empty] | TM 208 | Main_Plans
7B4OP093 | 7B4OP093/Chem_NY_PEG-4 | Chem_NY_PEG-4 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Main_Plans
7B4OP093 | 7B4OP093/Chem_NY_PEG-4 | Chem_NY_PEG-4 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) | [5 empty] | TM 208 | Main_Plans
7B4OP093 | 7B4OP093/Chem_NY_PEG-4 | Chem_NY_PEG-4 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7B4OP113 | 7B4OP113/Chem_NY_Mentha Piperita Extract_01 | Chem_NY_Mentha Piperita Extract_01 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | pH | Main_Plans
7B4OP113 | 7B4OP113/Chem_NY_Mentha Piperita Extract_01 | Chem_NY_Mentha Piperita Extract_01 | NY-Analytical | 398 | pH | pH | pH | 4.24 |  |  | FDA,1 | 4 | 5 | TM 200 | Main_Plans
7B4OP113 | 7B4OP113/Chem_NY_Mentha Piperita Extract_01 | Chem_NY_Mentha Piperita Extract_01 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear aqueous solution (Done) | [5 empty] | TM 208 | Main_Plans
7B4OP113 | 7B4OP113/Chem_NY_Mentha Piperita Extract_01 | Chem_NY_Mentha Piperita Extract_01 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Amber (Done) | [5 empty] | TM 208 | Main_Plans
7B4OP113 | 7B4OP113/Chem_NY_Mentha Piperita Extract_01 | Chem_NY_Mentha Piperita Extract_01 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Main_Plans
7B4OP113 | 7B4OP113/Chem_NY_Mentha Piperita Extract_01 | Chem_NY_Mentha Piperita Extract_01 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) | [5 empty] | TM 208 | Main_Plans
7B4OP203 | 7B4OP203/Chem_NY_Diazolidinyl Urea – RI-1157 | Chem_NY_Diazolidinyl Urea – RI-1157 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | pH | Main_Plans
7B4OP203 | 7B4OP203/Chem_NY_Diazolidinyl Urea – RI-1157 | Chem_NY_Diazolidinyl Urea – RI-1157 | NY-Analytical | 398 | pH | pH (1% w/v aqueous solution) | pH (1% w/v aqueous solution) | 7.02 |  |  | FDA,1 | 6.5 | 8 | TM 200 | Main_Plans
7B4OP203 | 7B4OP203/Chem_NY_Diazolidinyl Urea – RI-1157 | Chem_NY_Diazolidinyl Urea – RI-1157 | NY-Analytical | 891 | Solutions and Solubility | Solubility (1% w/v aqueous solution) | Solubility (1% w/v aqueous solution) | Clear and Colorless (Done) | [5 empty] | TM 222 | Main_Plans
7B4OP203 | 7B4OP203/Chem_NY_Diazolidinyl Urea – RI-1157 | Chem_NY_Diazolidinyl Urea – RI-1157 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7B4OP203 | 7B4OP203/Chem_NY_Diazolidinyl Urea – RI-1157 | Chem_NY_Diazolidinyl Urea – RI-1157 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Free flowing powder, free of foreign material (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OP203 | 7B4OP203/Chem_NY_Diazolidinyl Urea – RI-1157 | Chem_NY_Diazolidinyl Urea – RI-1157 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OP203 | 7B4OP203/Chem_NY_Diazolidinyl Urea – RI-1157 | Chem_NY_Diazolidinyl Urea – RI-1157 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Similar in character and intensity to standard, slight pungent odor (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OP230 | 7B4OP230/Chem_NY_Cosvat (Elestab CPN Ultra Pure) | Chem_NY_Cosvat (Elestab CPN Ultra Pure) | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7B4OP230 | 7B4OP230/Chem_NY_Cosvat (Elestab CPN Ultra Pure) | Chem_NY_Cosvat (Elestab CPN Ultra Pure) | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | White to off - white powder or crystalline aggregates. (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7B4OP230 | 7B4OP230/Chem_NY_Cosvat (Elestab CPN Ultra Pure) | Chem_NY_Cosvat (Elestab CPN Ultra Pure) | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) | [5 empty] | J&J TMWW78013 | Main_Plans
7B4OP230 | 7B4OP230/Chem_NY_Cosvat (Elestab CPN Ultra Pure) | Chem_NY_Cosvat (Elestab CPN Ultra Pure) | NY-Analytical | 801 | % Water | Water Content | Water Content | 0.2 |  | % | FDA,1 |  | 1 | J&J TMWW78017 | Main_Plans
7B4OP230 | 7B4OP230/Chem_NY_Cosvat (Elestab CPN Ultra Pure) | Chem_NY_Cosvat (Elestab CPN Ultra Pure) | NY-Analytical | 796 | Assay | Assay | Assay | 100.7 |  | % | FDA,1 | 99 |  | TM-002291 | Main_Plans
7B4OP230 | 7B4OP230/Chem_NY_Cosvat (Elestab CPN Ultra Pure) | Chem_NY_Cosvat (Elestab CPN Ultra Pure) | NY-Analytical | 796 | Assay | Assay Final | Assay Final | 100.7 |  | % | FDA,1 | 99 |  | TM-002291 | Main_Plans
7B4OP284 | 7B4OP284/Chem_NY_Sodium Citrate USP | Chem_NY_Sodium Citrate USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Odor | Main_Plans
7B4OP284 | 7B4OP284/Chem_NY_Sodium Citrate USP | Chem_NY_Sodium Citrate USP | NY-Analytical | 396 | Odor | Odor | Odor | Passes Test (Done) | [5 empty] | J&J TM-003188 | Main_Plans
7B4OP284 | 7B4OP284/Chem_NY_Sodium Citrate USP | Chem_NY_Sodium Citrate USP | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Colorless crystals, or as a white crystalline powder (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7B4OP284 | 7B4OP284/Chem_NY_Sodium Citrate USP | Chem_NY_Sodium Citrate USP | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7B4OP284 | 7B4OP284/Chem_NY_Sodium Citrate USP | Chem_NY_Sodium Citrate USP | NY-Analytical | 813 | Identification - USP | Identification A - Sodium | Identification A - Sodium | Passes Test (Done) | [5 empty] | USP | Main_Plans
7B4OP284 | 7B4OP284/Chem_NY_Sodium Citrate USP | Chem_NY_Sodium Citrate USP | NY-Analytical | 813 | Identification - USP | Identification B - Citrate | Identification B - Citrate | Passes Test (Done) | [5 empty] | USP | Main_Plans
7B4OP284 | 7B4OP284/Chem_NY_Sodium Citrate USP | Chem_NY_Sodium Citrate USP | NY-Analytical | 813 | Identification - USP | Identification C | Identification C | Passes Test (Done) | [5 empty] | USP | Main_Plans
7B4OP661 | 7B4OP661/Chem_NY_Pemulen TR-2 | Chem_NY_Pemulen TR-2 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | COA | Main_Plans
7B4OP661 | 7B4OP661/Chem_NY_Pemulen TR-2 | Chem_NY_Pemulen TR-2 | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7B4OP661 | 7B4OP661/Chem_NY_Pemulen TR-2 | Chem_NY_Pemulen TR-2 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Powder (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OP661 | 7B4OP661/Chem_NY_Pemulen TR-2 | Chem_NY_Pemulen TR-2 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OP661 | 7B4OP661/Chem_NY_Pemulen TR-2 | Chem_NY_Pemulen TR-2 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OP661 | 7B4OP661/Chem_NY_Pemulen TR-2 | Chem_NY_Pemulen TR-2 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Slightly acetic (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OQ028 | 7B4OQ028/Chem_NY_Sorbitan Monolaurate | Chem_NY_Sorbitan Monolaurate | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Identification | Main_Plans
7B4OQ028 | 7B4OQ028/Chem_NY_Sorbitan Monolaurate | Chem_NY_Sorbitan Monolaurate | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) | [5 empty] | TM 411 | Main_Plans
7B4OQ028 | 7B4OQ028/Chem_NY_Sorbitan Monolaurate | Chem_NY_Sorbitan Monolaurate | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Match standard (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OQ028 | 7B4OQ028/Chem_NY_Sorbitan Monolaurate | Chem_NY_Sorbitan Monolaurate | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Match standard (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OQ028 | 7B4OQ028/Chem_NY_Sorbitan Monolaurate | Chem_NY_Sorbitan Monolaurate | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OQ028 | 7B4OQ028/Chem_NY_Sorbitan Monolaurate | Chem_NY_Sorbitan Monolaurate | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OQ084 | 7B4OQ084/Chem_60A | Chem_60A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7B4OQ084 | 7B4OQ084/Chem_60A | Chem_60A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Slightly Opaque Viscous Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OQ084 | 7B4OQ084/Chem_60A | Chem_60A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Slightly White (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OQ084 | 7B4OQ084/Chem_60A | Chem_60A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OQ084 | 7B4OQ084/Chem_60A | Chem_60A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Mild (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OQ209 | 7B4OQ209/Chem_61A | Chem_61A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Identification | Main_Plans
7B4OQ209 | 7B4OQ209/Chem_61A | Chem_61A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | [6 empty] | TM 411 | Main_Plans
7B4OQ209 | 7B4OQ209/Chem_61A | Chem_61A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Free flowing granules or powder (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OQ209 | 7B4OQ209/Chem_61A | Chem_61A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OQ209 | 7B4OQ209/Chem_61A | Chem_61A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OQ209 | 7B4OQ209/Chem_61A | Chem_61A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OQ214 | 7B4OQ214/Chem_62A | Chem_62A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Identification | Main_Plans
7B4OQ214 | 7B4OQ214/Chem_62A | Chem_62A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7B4OQ214 | 7B4OQ214/Chem_62A | Chem_62A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Free flowing crystalline powder (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OQ214 | 7B4OQ214/Chem_62A | Chem_62A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OQ214 | 7B4OQ214/Chem_62A | Chem_62A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OQ214 | 7B4OQ214/Chem_62A | Chem_62A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OQ595 | 7B4OQ595/Chem_63A | Chem_63A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Identification | Main_Plans
7B4OQ595 | 7B4OQ595/Chem_63A | Chem_63A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7B4OQ595 | 7B4OQ595/Chem_63A | Chem_63A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Crystalline Powder (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OQ595 | 7B4OQ595/Chem_63A | Chem_63A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White to almost white (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OQ595 | 7B4OQ595/Chem_63A | Chem_63A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OQ595 | 7B4OQ595/Chem_63A | Chem_63A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Faint (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OQ794 | 7B4OQ794/Chem_NY_Disodium EDTA-USP | Chem_NY_Disodium EDTA-USP | NY-Analytical | 841 | Residual Solvents | Residual Solvents | Residual Solvents | Supplier Attestation Letter | [5 empty] | Residual Solvents | Main_Plans
7B4OQ794 | 7B4OQ794/Chem_NY_Disodium EDTA-USP | Chem_NY_Disodium EDTA-USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7B4OQ794 | 7B4OQ794/Chem_NY_Disodium EDTA-USP | Chem_NY_Disodium EDTA-USP | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | White crystalline powder (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7B4OQ794 | 7B4OQ794/Chem_NY_Disodium EDTA-USP | Chem_NY_Disodium EDTA-USP | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7B4OQ794 | 7B4OQ794/Chem_NY_Disodium EDTA-USP | Chem_NY_Disodium EDTA-USP | NY-Analytical | 398 | pH | pH (1 in 20 solution) | pH (1 in 20 solution) | 4.9 |  |  | FDA,1 | 4 | 6 | J&J TMWW78014 | Main_Plans
7B4OQ794 | 7B4OQ794/Chem_NY_Disodium EDTA-USP | Chem_NY_Disodium EDTA-USP | NY-Analytical | 796 | Assay | Assay | Assay | 99.8 |  | % | FDA,1 | 99 | 101 | USP | Main_Plans
7B4OQ794 | 7B4OQ794/Chem_NY_Disodium EDTA-USP | Chem_NY_Disodium EDTA-USP | NY-Analytical | 796 | Assay | Assay Final | Assay Final | 99.8 |  | % | FDA,1 | 99 | 101 | USP | Main_Plans
7B4OQ794 | 7B4OQ794/Chem_NY_Disodium EDTA-USP | Chem_NY_Disodium EDTA-USP | NY-Analytical | 843 | Calcium | Calcium | Calcium | Passes Test (Done) | [5 empty] | USP | Main_Plans
7B4OQ794 | 7B4OQ794/Chem_NY_Disodium EDTA-USP | Chem_NY_Disodium EDTA-USP | NY-Analytical | 833 | Heavy Metals | Heavy Metals | Heavy Metals | <0.005 |  | % | FDA,3 |  | 0.005 | USP | Main_Plans
7B4OQ794 | 7B4OQ794/Chem_NY_Disodium EDTA-USP | Chem_NY_Disodium EDTA-USP | NY-Analytical | 813 | Identification - USP | Identification A | Identification A | Passes Test (Done) | [5 empty] | USP | Main_Plans
7B4OQ794 | 7B4OQ794/Chem_NY_Disodium EDTA-USP | Chem_NY_Disodium EDTA-USP | NY-Analytical | 813 | Identification - USP | Identification B | Identification B | Passes Test (Done) | [5 empty] | USP | Main_Plans
7B4OQ794 | 7B4OQ794/Chem_NY_Disodium EDTA-USP | Chem_NY_Disodium EDTA-USP | NY-Analytical | 813 | Identification - USP | Identification C | Identification C | Passes Test (Done) | [5 empty] | USP | Main_Plans
7B4OQ794 | 7B4OQ794/Chem_NY_Disodium EDTA-USP | Chem_NY_Disodium EDTA-USP | NY-Analytical | 802 | Loss on Drying | Loss on Drying | Loss on Drying | 9.6 |  | % | FDA,1 | 8.7 | 11.4 | USP | Main_Plans
7B4OQ794 | 7B4OQ794/Chem_NY_Disodium EDTA-USP | Chem_NY_Disodium EDTA-USP | NY-Analytical | 802 | Loss on Drying | Loss on Drying Final | Loss on Drying Final | 9.6 |  | % | FDA,1 | 8.7 | 11.4 | USP | Main_Plans
7B4OQ794 | 7B4OQ794/Chem_NY_Disodium EDTA-USP | Chem_NY_Disodium EDTA-USP | NY-Analytical | 842 | Nitrilotriacetic Acid | Nitrilotriacetic Acid | Nitrilotriacetic Acid | 0 |  | % | FDA,1 |  | 0.1 | USP | Main_Plans
7B4OX232 | 7B4OX232/Chem_40A | Chem_40A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Apparent proof | Main_Plans
7B4OX232 | 7B4OX232/Chem_40A | Chem_40A | NY-Analytical | 936 | Apparent proof | Apparent proof | Apparent proof | 189.7 |  |  | FDA,1 | 189.7 | 190.2 | STM 101 | Main_Plans
7B4OX232 | 7B4OX232/Chem_40A | Chem_40A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.8091 |  |  | FDA,4 | 0.8086 | 0.8098 | TM 202 | Main_Plans
7B4OX232 | 7B4OX232/Chem_40A | Chem_40A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Meets USP requirement (Done) | [5 empty] | TM 411 | Main_Plans
7B4OX232 | 7B4OX232/Chem_40A | Chem_40A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, mobile liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OX232 | 7B4OX232/Chem_40A | Chem_40A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OX232 | 7B4OX232/Chem_40A | Chem_40A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7B4OX232 | 7B4OX232/Chem_40A | Chem_40A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Alcohol (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7C4AA580 | 7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | Chem_NY_Fragrance-Grapefruit Blossom &Orange | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Specific Gravity at 25°C | Main_Plans
7C4AA580 | 7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | Chem_NY_Fragrance-Grapefruit Blossom &Orange | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.0662 |  |  | FDA,4 | 1.0561 | 1.0761 | TM 202 | Main_Plans
7C4AA580 | 7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | Chem_NY_Fragrance-Grapefruit Blossom &Orange | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7C4AA580 | 7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | Chem_NY_Fragrance-Grapefruit Blossom &Orange | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7C4AA580 | 7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | Chem_NY_Fragrance-Grapefruit Blossom &Orange | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Pale Yellow – Light Yellow (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7C4AA580 | 7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | Chem_NY_Fragrance-Grapefruit Blossom &Orange | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7C4AA580 | 7C4AA580/Chem_NY_Fragrance-Grapefruit Blossom &Orange | Chem_NY_Fragrance-Grapefruit Blossom &Orange | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7C4FD110 | 7C4FD110/Chem_69A | Chem_69A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7C4FD110 | 7C4FD110/Chem_69A | Chem_69A | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | White or nearly white, hard, brittle material (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7C4FD110 | 7C4FD110/Chem_69A | Chem_69A | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7C4FD110 | 7C4FD110/Chem_69A | Chem_69A | NY-Analytical | 390 | Identification | Identification B | Identification B | Pass (Done) | [5 empty] | USP | Main_Plans
7C4FD110 | 7C4FD110/Chem_69A | Chem_69A | NY-Analytical | 390 | Identification | Identification for Sodium (USP <191>) | Identification for Sodium | Positive response for Sodium (Done) | [5 empty] | USP | Main_Plans
7C4OD232 | 7C4OD232/Chem_NY_Methylparaben NF/USP | Chem_NY_Methylparaben NF/USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Certificate of Analysis | Main_Plans
7C4OD232 | 7C4OD232/Chem_NY_Methylparaben NF/USP | Chem_NY_Methylparaben NF/USP | NY-Analytical | 390 | Identification | IR Identification | IR Identification | [6 empty] | TM 411 | Main_Plans
7C4OD232 | 7C4OD232/Chem_NY_Methylparaben NF/USP | Chem_NY_Methylparaben NF/USP | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Colorless crystals or white crystalline powder (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7C4OD232 | 7C4OD232/Chem_NY_Methylparaben NF/USP | Chem_NY_Methylparaben NF/USP | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | [6 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7C4OD232 | 7C4OD232/Chem_NY_Methylparaben NF/USP | Chem_NY_Methylparaben NF/USP | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | [6 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7C4OD232 | 7C4OD232/Chem_NY_Methylparaben NF/USP | Chem_NY_Methylparaben NF/USP | NY-Analytical | 796 | Assay | Assay | Assay |  |  | % | FDA,1 | 98 | 102 | USP/NF | Main_Plans
7C4OD232 | 7C4OD232/Chem_NY_Methylparaben NF/USP | Chem_NY_Methylparaben NF/USP | NY-Analytical | 796 | Assay | Assay Final | Assay Final |  |  | % | FDA,1 | 98 | 102 | USP/NF | Main_Plans
7C4OG169 | 7C4OG169/Chem_75A | Chem_75A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7C4OG169 | 7C4OG169/Chem_75A | Chem_75A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Crystals (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7C4OG169 | 7C4OG169/Chem_75A | Chem_75A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7C4OG169 | 7C4OG169/Chem_75A | Chem_75A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7C4OG169 | 7C4OG169/Chem_75A | Chem_75A | NY-Analytical | 390 | Identification | Identification for Ammonium (USP <191>) | Identification for Ammonium | Pass (Done) | [5 empty] | TM400/408b/421/422/USP | Main_Plans
7C4OG169 | 7C4OG169/Chem_75A | Chem_75A | NY-Analytical | 390 | Identification | Identification for Chloride (USP <191>) | Identification for Chloride | Pass (Done) | [5 empty] | TM400/408b/421/422/USP | Main_Plans
7C4OP027 | 7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | Chem_NY_Camellia Oleifera Leaf Extract | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7C4OP027 | 7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | Chem_NY_Camellia Oleifera Leaf Extract | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Light yellow to amber liquid (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7C4OP027 | 7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | Chem_NY_Camellia Oleifera Leaf Extract | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7C4OP027 | 7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | Chem_NY_Camellia Oleifera Leaf Extract | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7C4OP027 | 7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | Chem_NY_Camellia Oleifera Leaf Extract | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) | [5 empty] | J&J TMWW78013 | Main_Plans
7C4OP027 | 7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 509 | J&J_Pathogens | Biggy or SDA | Biggy or SDA | No Growth_SDA | [5 empty] | TM 869 | Main_Plans
7C4OP027 | 7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 509 | J&J_Pathogens | Cetrimide | Cetrimide | No Growth | [5 empty] | TM 869 | Main_Plans
7C4OP027 | 7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 524 | J&J Fungal | Fungal | Fungal | 0 |  | CFU/g | FDA,0 | 0 | 200 | TM 869 | Main_Plans
7C4OP027 | 7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 509 | J&J_Pathogens | MacConkey | MacConkey | No Growth | [5 empty] | TM 869 | Main_Plans
7C4OP027 | 7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 509 | J&J_Pathogens | MSA or VJ | MSA or VJ | No Growth_MSA | [5 empty] | TM 869 | Main_Plans
7C4OP027 | 7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 723 | J&J TSB Controls - TSA | Negative Control | Negative Control | No Growth TSB | [5 empty] | TM 869 | Main_Plans
7C4OP027 | 7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | NegativeControl | No Growth TSB | [5 empty] | TM 869 | Main_Plans
7C4OP027 | 7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 509 | J&J_Pathogens | TSA | TSA | No Growth | [5 empty] | TM 869 | Main_Plans
7C4OP027 | 7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | TSA JJ | 0 |  | cfu/mL | FDA,0 | 0 | 500 | TM 869 | Main_Plans
7C4OP027 | 7C4OP027/Chem_NY_Camellia Oleifera Leaf Extract | Chem_NY_Camellia Oleifera Leaf Extract | NY-Microbiology | 517 | J&J_Enrichment | TSB | TSB | Not Turbid | [5 empty] | TM 869 | Main_Plans
7C4OP065 | 7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | Chem_NY_ Sodium Lauroamphoacetate 30% | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | pH | Main_Plans
7C4OP065 | 7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | Chem_NY_ Sodium Lauroamphoacetate 30% | NY-Analytical | 398 | pH | pH (1 in 20 solution) | pH (1 in 20 solution) | 9.5 |  |  | FDA,1 | 9 | 9.5 | TM 200 | Main_Plans
7C4OP065 | 7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | Chem_NY_ Sodium Lauroamphoacetate 30% | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Clean Liquid (Done) | [5 empty] | TM 208 | Main_Plans
7C4OP065 | 7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | Chem_NY_ Sodium Lauroamphoacetate 30% | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to light yellow (Done) | [5 empty] | TM 208 | Main_Plans
7C4OP065 | 7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | Chem_NY_ Sodium Lauroamphoacetate 30% | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Main_Plans
7C4OP065 | 7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | Chem_NY_ Sodium Lauroamphoacetate 30% | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) | [5 empty] | TM 208 | Main_Plans
7C4OP065 | 7C4OP065/Chem_NY_ Sodium Lauroamphoacetate 30% | Chem_NY_ Sodium Lauroamphoacetate 30% | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) | [5 empty] | TM 411 | Main_Plans
7C4OP284 | 7C4OP284/Chem_NY_Grape Fruit Extract | Chem_NY_Grape Fruit Extract | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Identification | Main_Plans
7C4OP284 | 7C4OP284/Chem_NY_Grape Fruit Extract | Chem_NY_Grape Fruit Extract | NY-Analytical | 390 | Identification | Identification (FTIR) | Identification (FTIR) | Matches that of the Reference Standard (Done) | [5 empty] | J&J TMWW78013 | Main_Plans
7C4OP284 | 7C4OP284/Chem_NY_Grape Fruit Extract | Chem_NY_Grape Fruit Extract | NY-Analytical | 792 | Chemical Inspection | Appearance (J&J TMWW78010) | Appearance | Clear to Pale Yellow Liquid (Done) | [5 empty] | JJTM7252 | Main_Plans
7C4OP284 | 7C4OP284/Chem_NY_Grape Fruit Extract | Chem_NY_Grape Fruit Extract | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) | [5 empty] | JJTM7252 | Main_Plans
7C4OQ028 | 7C4OQ028/Chem_NY_Isohexadecane | Chem_NY_Isohexadecane | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Refractive Index at 20°C | Main_Plans
7C4OQ028 | 7C4OQ028/Chem_NY_Isohexadecane | Chem_NY_Isohexadecane | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.439 |  |  | FDA,1 | 1.4 | 1.5 | TM 204 | Main_Plans
7C4OQ028 | 7C4OQ028/Chem_NY_Isohexadecane | Chem_NY_Isohexadecane | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) | [5 empty] | TM 411 | Main_Plans
7C4OQ028 | 7C4OQ028/Chem_NY_Isohexadecane | Chem_NY_Isohexadecane | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7C4OQ028 | 7C4OQ028/Chem_NY_Isohexadecane | Chem_NY_Isohexadecane | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7C4OQ028 | 7C4OQ028/Chem_NY_Isohexadecane | Chem_NY_Isohexadecane | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7C4OQ028 | 7C4OQ028/Chem_NY_Isohexadecane | Chem_NY_Isohexadecane | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7C4OQ595 | 7C4OQ595/Chem_81A | Chem_81A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7C4OQ595 | 7C4OQ595/Chem_81A | Chem_81A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Free flowing granules or powder (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7C4OQ595 | 7C4OQ595/Chem_81A | Chem_81A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7C4OQ595 | 7C4OQ595/Chem_81A | Chem_81A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7C4OQ595 | 7C4OQ595/Chem_81A | Chem_81A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7C4OQ595 | 7C4OQ595/Chem_81A | Chem_81A | NY-Analytical | 390 | Identification | Identification B | Identification B | Pass (Done) | [5 empty] | TM400/408b/421/422/USP | Main_Plans
7C4OQ595 | 7C4OQ595/Chem_81A | Chem_81A | NY-Analytical | 390 | Identification | Identification for Potassium (USP <191>) | Identification for Potassium | Pass (Done) | [5 empty] | TM400/408b/421/422/USP | Main_Plans
7C4OX110 | 7C4OX110/Chem_84A | Chem_84A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Specific Gravity at 25°C | Main_Plans
7C4OX110 | 7C4OX110/Chem_84A | Chem_84A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.8529 |  |  | FDA,3 | 0.845 | 0.858 | TM 202 | Main_Plans
7C4OX110 | 7C4OX110/Chem_84A | Chem_84A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7C4OX110 | 7C4OX110/Chem_84A | Chem_84A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7C4OX110 | 7C4OX110/Chem_84A | Chem_84A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7C4OX110 | 7C4OX110/Chem_84A | Chem_84A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7D4AA580 | 7D4AA580/Chem_NY_Monoethanolamine | Chem_NY_Monoethanolamine | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Identification | Main_Plans
7D4AA580 | 7D4AA580/Chem_NY_Monoethanolamine | Chem_NY_Monoethanolamine | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) | [5 empty] | TM 411 | Main_Plans
7D4AA580 | 7D4AA580/Chem_NY_Monoethanolamine | Chem_NY_Monoethanolamine | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, Colorless, Viscous Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7D4AA580 | 7D4AA580/Chem_NY_Monoethanolamine | Chem_NY_Monoethanolamine | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7D4AA580 | 7D4AA580/Chem_NY_Monoethanolamine | Chem_NY_Monoethanolamine | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7D4AA580 | 7D4AA580/Chem_NY_Monoethanolamine | Chem_NY_Monoethanolamine | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Ammoniacal (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7D4OG169 | 7D4OG169/Chem_88A | Chem_88A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Specific Gravity at 25°C | Main_Plans
7D4OG169 | 7D4OG169/Chem_88A | Chem_88A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.068 |  |  | FDA,4 | 1.0594 | 1.0794 | TM 202 | Main_Plans
7D4OG169 | 7D4OG169/Chem_88A | Chem_88A | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.4755 |  |  | FDA,4 | 1.4695 | 1.4895 | TM 204 | Main_Plans
7D4OG169 | 7D4OG169/Chem_88A | Chem_88A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7D4OG169 | 7D4OG169/Chem_88A | Chem_88A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7D4OG169 | 7D4OG169/Chem_88A | Chem_88A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Pale straw (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7D4OG169 | 7D4OG169/Chem_88A | Chem_88A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7D4OP065 | 7D4OP065/Chem_NY_ Glydant Plus Liquid | Chem_NY_ Glydant Plus Liquid | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | pH | Main_Plans
7D4OP065 | 7D4OP065/Chem_NY_ Glydant Plus Liquid | Chem_NY_ Glydant Plus Liquid | NY-Analytical | 398 | pH | pH | pH | 7.66 |  |  | FDA,1 | 6 | 8 | TM 200 | Main_Plans
7D4OP065 | 7D4OP065/Chem_NY_ Glydant Plus Liquid | Chem_NY_ Glydant Plus Liquid | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) | [5 empty] | TM 208 | Main_Plans
7D4OP065 | 7D4OP065/Chem_NY_ Glydant Plus Liquid | Chem_NY_ Glydant Plus Liquid | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Water-White to light yellow (Done) | [5 empty] | TM 208 | Main_Plans
7D4OP065 | 7D4OP065/Chem_NY_ Glydant Plus Liquid | Chem_NY_ Glydant Plus Liquid | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Main_Plans
7D4OP065 | 7D4OP065/Chem_NY_ Glydant Plus Liquid | Chem_NY_ Glydant Plus Liquid | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) | [5 empty] | TM 208 | Main_Plans
7D4OP065 | 7D4OP065/Chem_NY_ Glydant Plus Liquid | Chem_NY_ Glydant Plus Liquid | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) | [5 empty] | TM 411 | Main_Plans
7D4OP230 | 7D4OP230/Chem_NY_Zemea_01 | Chem_NY_Zemea_01 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Appearance | Main_Plans
7D4OP230 | 7D4OP230/Chem_NY_Zemea_01 | Chem_NY_Zemea_01 | NY-Analytical | 397 | Appearance | Physical Properties | Physical Properties | A Clear, Colorless Liquid (Done) | [5 empty] | TM 208 | Main_Plans
7D4OP230 | 7D4OP230/Chem_NY_Zemea_01 | Chem_NY_Zemea_01 | NY-Analytical | 390 | Identification | Identification A | Identification A | Passes test. (Done) | [5 empty] | USP/NF | Main_Plans
7D4OP230 | 7D4OP230/Chem_NY_Zemea_01 | Chem_NY_Zemea_01 | NY-Analytical | 390 | Identification | Identification B | Identification B | Passes test. (Done) | [5 empty] | USP/NF | Main_Plans
7D4OP601 | 7D4OP601/Chem_91A | Chem_91A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7D4OP601 | 7D4OP601/Chem_91A | Chem_91A | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear, yellow liquid. (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7D4OP601 | 7D4OP601/Chem_91A | Chem_91A | NY-Analytical | 792 | Chemical Inspection | Odor (J&J TM-003188) | Odor | Meets Odor Specification or of Characteristic Odor. [As per specification](Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7D4OP601 | 7D4OP601/Chem_91A | Chem_91A | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7D4OP601 | 7D4OP601/Chem_91A | Chem_91A | NY-Analytical | 831 | Density @ 20° C | Density | Density | 1.1193 |  |  | FDA,3 | 1.111 | 1.131 | J&J TMWW78012 | Main_Plans
7D4OP601 | 7D4OP601/Chem_91A | Chem_91A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) | [5 empty] | J&J TMWW78013 | Main_Plans
7D4OP601 | 7D4OP601/Chem_91A | Chem_91A | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.53 |  |  | FDA,3 | 1.526 | 1.538 | J&J TMWW78015 | Main_Plans
7D4OP601 | 7D4OP601/Chem_91A | Chem_91A | NY-Analytical | 844 | % Benzoic Acid Phenoxyethanol & Dehydroacetic Acid | Benzoic Acid | Benzoic Acid | 12.13 |  | % | FDA,1 | 11 | 13 | TM 882 | Main_Plans
7D4OP601 | 7D4OP601/Chem_91A | Chem_91A | NY-Analytical | 844 | % Benzoic Acid Phenoxyethanol & Dehydroacetic Acid | Dehydroacetic Acid | Dehydroacetic Acid | 6.91 |  | % | FDA,1 | 6.5 | 7.5 | TM 882 | Main_Plans
7D4OP601 | 7D4OP601/Chem_91A | Chem_91A | NY-Analytical | 844 | % Benzoic Acid Phenoxyethanol & Dehydroacetic Acid | Phenoxyethanol | Phenoxyethanol | 76.51 |  | % | FDA,1 | 71 | 77 | TM 882 | Main_Plans
7D4OQ028 | 7D4OQ028/Chem_NY_Dicaprylyl Ether | Chem_NY_Dicaprylyl Ether | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Refractive Index at 20°C | Main_Plans
7D4OQ028 | 7D4OQ028/Chem_NY_Dicaprylyl Ether | Chem_NY_Dicaprylyl Ether | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.431 |  |  | FDA,3 | 1.431 | 1.435 | TM 204 | Main_Plans
7D4OQ028 | 7D4OQ028/Chem_NY_Dicaprylyl Ether | Chem_NY_Dicaprylyl Ether | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) | [5 empty] | TM 411 | Main_Plans
7D4OQ028 | 7D4OQ028/Chem_NY_Dicaprylyl Ether | Chem_NY_Dicaprylyl Ether | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7D4OQ028 | 7D4OQ028/Chem_NY_Dicaprylyl Ether | Chem_NY_Dicaprylyl Ether | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7D4OQ028 | 7D4OQ028/Chem_NY_Dicaprylyl Ether | Chem_NY_Dicaprylyl Ether | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7D4OQ028 | 7D4OQ028/Chem_NY_Dicaprylyl Ether | Chem_NY_Dicaprylyl Ether | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7D4OQ209 | 7D4OQ209/Chem_92A | Chem_92A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Specific Gravity at 25°C | Main_Plans
7D4OQ209 | 7D4OQ209/Chem_92A | Chem_92A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.0512 |  |  | FDA,4 | 1.0399 | 1.0599 | TM 202 | Main_Plans
7D4OQ209 | 7D4OQ209/Chem_92A | Chem_92A | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.511 |  |  | FDA,4 | 1.507 | 1.5272 | TM 204 | Main_Plans
7D4OQ209 | 7D4OQ209/Chem_92A | Chem_92A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7D4OQ209 | 7D4OQ209/Chem_92A | Chem_92A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Pale Yellow Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7D4OQ209 | 7D4OQ209/Chem_92A | Chem_92A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7D4OQ209 | 7D4OQ209/Chem_92A | Chem_92A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Baby Powder (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7D4OQ214 | 7D4OQ214/Chem_93A | Chem_93A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White to Off-White | [5 empty] | Color (TM 208) | Main_Plans
7D4OQ214 | 7D4OQ214/Chem_93A | Chem_93A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | pH | Main_Plans
7D4OQ214 | 7D4OQ214/Chem_93A | Chem_93A | NY-Analytical | 398 | pH | pH (5% Solution) | pH (5% Solution) | 4.72 |  |  | FDA,1 | 4 | 6 | TM 200 | Main_Plans
7D4OQ214 | 7D4OQ214/Chem_93A | Chem_93A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7D4OQ214 | 7D4OQ214/Chem_93A | Chem_93A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Powder (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7D4OQ214 | 7D4OQ214/Chem_93A | Chem_93A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7D4OQ214 | 7D4OQ214/Chem_93A | Chem_93A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7D4OQ967 | 7D4OQ967/Chem_94A | Chem_94A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Identification | Main_Plans
7D4OQ967 | 7D4OQ967/Chem_94A | Chem_94A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7D4OQ967 | 7D4OQ967/Chem_94A | Chem_94A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear viscous liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7D4OQ967 | 7D4OQ967/Chem_94A | Chem_94A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Yellow to amber (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7D4OQ967 | 7D4OQ967/Chem_94A | Chem_94A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7E4OD232 | 7E4OD232/Chem_97A | Chem_97A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Identification | Main_Plans
7E4OD232 | 7E4OD232/Chem_97A | Chem_97A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7E4OD232 | 7E4OD232/Chem_97A | Chem_97A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Granular Solid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7E4OD232 | 7E4OD232/Chem_97A | Chem_97A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7E4OD232 | 7E4OD232/Chem_97A | Chem_97A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7E4OD232 | 7E4OD232/Chem_97A | Chem_97A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7E4OP065 | 7E4OP065/Chem_NY_Benzyl Alcohol | Chem_NY_Benzyl Alcohol | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7E4OP065 | 7E4OP065/Chem_NY_Benzyl Alcohol | Chem_NY_Benzyl Alcohol | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Colorless Liquid (Done) | [5 empty] | TM 208 | Main_Plans
7E4OP065 | 7E4OP065/Chem_NY_Benzyl Alcohol | Chem_NY_Benzyl Alcohol | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) | [5 empty] | TM 208 | Main_Plans
7E4OP065 | 7E4OP065/Chem_NY_Benzyl Alcohol | Chem_NY_Benzyl Alcohol | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Main_Plans
7E4OP065 | 7E4OP065/Chem_NY_Benzyl Alcohol | Chem_NY_Benzyl Alcohol | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) | [5 empty] | TM 208 | Main_Plans
7E4OP065 | 7E4OP065/Chem_NY_Benzyl Alcohol | Chem_NY_Benzyl Alcohol | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) | [5 empty] | TM 411 | Main_Plans
7E4OP203 | 7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | pH | Main_Plans
7E4OP203 | 7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Analytical | 398 | pH | pH (10% w/v aqueous solution) | pH (10% w/v aqueous solution) | 6.93 |  |  | FDA,1 | 6.5 | 7.5 | TM 200 | Main_Plans
7E4OP203 | 7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7E4OP203 | 7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Microbiology | 867 | Celsis_Rapid | Bacteria Count (Obeserved) | Bacteria Count (Obeserved) | 0 |  | cfu | FDA,0 |  | 100 | TM 603 | Main_Plans
7E4OP203 | 7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Microbiology | 867 | Celsis_Rapid | Bacteria Count (Reported) | Bacteria Count (Reported) | 0 |  | CFU/g | FDA,0 |  | 100 | TM 603 | Main_Plans
7E4OP203 | 7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Microbiology | 867 | Celsis_Rapid | Celsis_ Result | Celsis Result | Negative ATP (Done) | [5 empty] | TM 603 | Main_Plans
7E4OP203 | 7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Microbiology | 867 | Celsis_Rapid | Mold Count (Observed) | Mold Count (Observed) | 0 |  | cfu | FDA,0 |  |  | TM 603 | Main_Plans
7E4OP203 | 7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Microbiology | 867 | Celsis_Rapid | Mold Count (Reported) | Mold Count (Reported) | 0 |  | CFU/g | FDA,0 |  |  | TM 603 | Main_Plans
7E4OP203 | 7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear viscous liquid, free of foreign matter. (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7E4OP203 | 7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Analytical | 792 | Chemical Inspection | Color (TM 230) | Color | 2 |  |  | FDA,0 |  | 6 | TM208/220/230/USP/Kao LM | Main_Plans
7E4OP203 | 7E4OP203/Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | Chem_NY_Cocamidopropyl Dimonium Chloride Phosphate RI-1059 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Similar in character and intensity to standard. Herbal Odor (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7E4OP230 | 7E4OP230/Chem_NY_Lexfeel Natural | Chem_NY_Lexfeel Natural | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7E4OP230 | 7E4OP230/Chem_NY_Lexfeel Natural | Chem_NY_Lexfeel Natural | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear colorless to pale yellow liquid. (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7E4OP230 | 7E4OP230/Chem_NY_Lexfeel Natural | Chem_NY_Lexfeel Natural | NY-Analytical | 792 | Chemical Inspection | Colour (J&J TM7377) | Colour | 2 |  |  | FDA,0 |  | 50 | J&J TMWW78010/J&J TM7252 | Main_Plans
7E4OP230 | 7E4OP230/Chem_NY_Lexfeel Natural | Chem_NY_Lexfeel Natural | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7E4OP230 | 7E4OP230/Chem_NY_Lexfeel Natural | Chem_NY_Lexfeel Natural | NY-Analytical | 826 | Acid Value | Acid Value | Acid Value | 0.03 |  | mg KOH/g | FDA,1 |  | 0.1 | J&J TMWW78011 | Main_Plans
7E4OP230 | 7E4OP230/Chem_NY_Lexfeel Natural | Chem_NY_Lexfeel Natural | NY-Analytical | 826 | Acid Value | Acid Value Final | Acid Value Final | 0.03 |  | mg KOH/g | FDA,1 |  | 0.1 | J&J TMWW78011 | Main_Plans
7E4OP230 | 7E4OP230/Chem_NY_Lexfeel Natural | Chem_NY_Lexfeel Natural | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) | [5 empty] | J&J TMWW78013 | Main_Plans
7E4OP230 | 7E4OP230/Chem_NY_Lexfeel Natural | Chem_NY_Lexfeel Natural | NY-Analytical | 801 | % Water | Moisture Content | Moisture Content | 0.04 |  | % | FDA,1 |  | 0.1 | J&J TMWW78017 | Main_Plans
7E4OP230 | 7E4OP230/Chem_NY_Lexfeel Natural | Chem_NY_Lexfeel Natural | NY-Microbiology | 524 | J&J Fungal | Fungal | Fungal | 0 |  | CFU/g | FDA,0 | 0 | 200 | TM 869 | Main_Plans
7E4OP230 | 7E4OP230/Chem_NY_Lexfeel Natural | Chem_NY_Lexfeel Natural | NY-Microbiology | 701 | J&J TSA Controls | Negative Control | Negative Control | No Growth TSA | [5 empty] | TM 869 | Main_Plans
7E4OP230 | 7E4OP230/Chem_NY_Lexfeel Natural | Chem_NY_Lexfeel Natural | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | NegativeControl | No Growth TSB | [5 empty] | TM 869 | Main_Plans
7E4OP230 | 7E4OP230/Chem_NY_Lexfeel Natural | Chem_NY_Lexfeel Natural | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | TSA JJ | 0 |  | cfu/mL | FDA,0 | 0 | 2000 | TM 869 | Main_Plans
7E4OP601 | 7E4OP601/Chem_102A | Chem_102A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Assay | Main_Plans
7E4OP601 | 7E4OP601/Chem_102A | Chem_102A | NY-Analytical | 796 | Assay | Assay | Assay | 10.16 |  | % | FDA,2 | 9 | 11 | J&J TM-002528 | Main_Plans
7E4OP601 | 7E4OP601/Chem_102A | Chem_102A | NY-Analytical | 796 | Assay | Assay Final | Assay Final | 10.16 |  | % | FDA,1 | 9 | 11 | J&J TM-002528 | Main_Plans
7E4OP601 | 7E4OP601/Chem_102A | Chem_102A | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear, yellow liquid. (Done) | [5 empty] | J&J TMWW78010 | Main_Plans
7E4OP601 | 7E4OP601/Chem_102A | Chem_102A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) | [5 empty] | J&J TMWW78013 | Main_Plans
7E4OP601 | 7E4OP601/Chem_102A | Chem_102A | NY-Analytical | 801 | % Water | Water Content | Water Content | 0.4 |  | % | FDA,2 |  | 1 | J&J TMWW78017 | Main_Plans
7E4OP806 | 7E4OP806/Chem_58B | Chem_58B | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7E4OP806 | 7E4OP806/Chem_58B | Chem_58B | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear colorless liquid free from suspended materials (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7E4OP806 | 7E4OP806/Chem_58B | Chem_58B | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7E4OP806 | 7E4OP806/Chem_58B | Chem_58B | NY-Analytical | 831 | Density @ 20° C | Density | Density | 0.9221 |  |  | FDA,3 | 0.92 | 0.923 | J&J TMWW78012 | Main_Plans
7E4OP806 | 7E4OP806/Chem_58B | Chem_58B | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) | [5 empty] | J&J TMWW78013 | Main_Plans
7E4OP806 | 7E4OP806/Chem_58B | Chem_58B | NY-Analytical | 801 | % Water | Moisture Content (10 grams of sample) | Moisture Content (10 grams of sample) | 0.06 |  | % | FDA,3 |  | 0.1 | J&J TMWW78017 | Main_Plans
7E4OQ028 | 7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | Chem_NY_Ethylhexyl Hydroxystearate | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Specific Gravity at 25°C | Main_Plans
7E4OQ028 | 7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | Chem_NY_Ethylhexyl Hydroxystearate | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.8964 |  |  | FDA,4 | 0.892 | 0.898 | TM 202 | Main_Plans
7E4OQ028 | 7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | Chem_NY_Ethylhexyl Hydroxystearate | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.4555 |  |  | FDA,4 | 1.4555 | 1.458 | TM 204 | Main_Plans
7E4OQ028 | 7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | Chem_NY_Ethylhexyl Hydroxystearate | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) | [5 empty] | TM 411 | Main_Plans
7E4OQ028 | 7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | Chem_NY_Ethylhexyl Hydroxystearate | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear to slightly hazy liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7E4OQ028 | 7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | Chem_NY_Ethylhexyl Hydroxystearate | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Match standard (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7E4OQ028 | 7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | Chem_NY_Ethylhexyl Hydroxystearate | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7E4OQ028 | 7E4OQ028/Chem_NY_Ethylhexyl Hydroxystearate | Chem_NY_Ethylhexyl Hydroxystearate | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Slightly Characteristic (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7F4BP869 | 7F4BP869/Chem_NY_Citric Acid Anhydrous - USP | Chem_NY_Citric Acid Anhydrous - USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | COA | Main_Plans
7F4BP869 | 7F4BP869/Chem_NY_Citric Acid Anhydrous - USP | Chem_NY_Citric Acid Anhydrous - USP | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Crystals/crystalline powder (Done) | [5 empty] | TM 208 | Main_Plans
7F4BP869 | 7F4BP869/Chem_NY_Citric Acid Anhydrous - USP | Chem_NY_Citric Acid Anhydrous - USP | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to white (Done) | [5 empty] | TM 208 | Main_Plans
7F4BP869 | 7F4BP869/Chem_NY_Citric Acid Anhydrous - USP | Chem_NY_Citric Acid Anhydrous - USP | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Main_Plans
7F4BP869 | 7F4BP869/Chem_NY_Citric Acid Anhydrous - USP | Chem_NY_Citric Acid Anhydrous - USP | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless (Done) | [5 empty] | TM 208 | Main_Plans
7F4BP869 | 7F4BP869/Chem_NY_Citric Acid Anhydrous - USP | Chem_NY_Citric Acid Anhydrous - USP | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7F4OA160 | 7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | Chem_NY_PEG-40 Hydrogenated Castor Oil | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | pH | Main_Plans
7F4OA160 | 7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | Chem_NY_PEG-40 Hydrogenated Castor Oil | NY-Analytical | 398 | pH | pH (10% in water) | pH (10% in water) | 6.99 |  |  | FDA,1 | 6 | 7 | TM 200 | Main_Plans
7F4OA160 | 7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | Chem_NY_PEG-40 Hydrogenated Castor Oil | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7F4OA160 | 7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | Chem_NY_PEG-40 Hydrogenated Castor Oil | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear viscous liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7F4OA160 | 7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | Chem_NY_PEG-40 Hydrogenated Castor Oil | NY-Analytical | 792 | Chemical Inspection | Color / Appearance (TM 208) | Color / Appearance | White to yellowish soft or flowing paste. Close Match to Standard (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7F4OA160 | 7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | Chem_NY_PEG-40 Hydrogenated Castor Oil | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7F4OA160 | 7F4OA160/Chem_NY_PEG-40 Hydrogenated Castor Oil | Chem_NY_PEG-40 Hydrogenated Castor Oil | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Close match to standard (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7F4OD110 | 7F4OD110/Chem_106A | Chem_106A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | pH | Main_Plans
7F4OD110 | 7F4OD110/Chem_106A | Chem_106A | NY-Analytical | 398 | pH | pH (1% Solution) | pH (1% Solution) | 10.6 |  |  | FDA,1 | 10.5 | 11.5 | TM 200 | Main_Plans
7F4OD110 | 7F4OD110/Chem_106A | Chem_106A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7F4OD110 | 7F4OD110/Chem_106A | Chem_106A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Crystalline Powder (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7F4OD110 | 7F4OD110/Chem_106A | Chem_106A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7F4OD110 | 7F4OD110/Chem_106A | Chem_106A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7F4OD110 | 7F4OD110/Chem_106A | Chem_106A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7F4OP027 | 7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | Chem_NY_Arlasilk Phospholipid PTC | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | color (Gardner) | Main_Plans
7F4OP027 | 7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | Chem_NY_Arlasilk Phospholipid PTC | NY-Analytical | 944 | color (Gardner) | color (Gardner) | color (Gardner) | 3 |  |  | FDA,0 |  | 4 | J&J TM701 | Main_Plans
7F4OP027 | 7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | Chem_NY_Arlasilk Phospholipid PTC | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear amber liquid. (Done) | [5 empty] | J&J TMWW78010 | Main_Plans
7F4OP027 | 7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | Chem_NY_Arlasilk Phospholipid PTC | NY-Analytical | 792 | Chemical Inspection | Odor (J&J TM-003188) | Odor | Passes Test (Done) | [5 empty] | J&J TMWW78010 | Main_Plans
7F4OP027 | 7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | Chem_NY_Arlasilk Phospholipid PTC | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) | [5 empty] | J&J TMWW78010 | Main_Plans
7F4OP027 | 7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | Chem_NY_Arlasilk Phospholipid PTC | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) | [5 empty] | J&J TMWW78013 | Main_Plans
7F4OP027 | 7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 509 | J&J_Pathogens | Biggy or SDA | Biggy or SDA | No Growth_SDA | [5 empty] | TM 869 | Main_Plans
7F4OP027 | 7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 509 | J&J_Pathogens | Cetrimide | Cetrimide | No Growth | [5 empty] | TM 869 | Main_Plans
7F4OP027 | 7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 524 | J&J Fungal | Fungal | Fungal | 0 |  | CFU/g | FDA,0 | 0 | 200 | TM 869 | Main_Plans
7F4OP027 | 7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 509 | J&J_Pathogens | MacConkey | MacConkey | No Growth | [5 empty] | TM 869 | Main_Plans
7F4OP027 | 7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 509 | J&J_Pathogens | MSA or VJ | MSA or VJ | No Growth_MSA | [5 empty] | TM 869 | Main_Plans
7F4OP027 | 7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 700 | J&J SDA Controls | Negative Control | Negative Control | No Growth SDA | [5 empty] | TM 869 | Main_Plans
7F4OP027 | 7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | NegativeControl | No Growth TSB | [5 empty] | TM 869 | Main_Plans
7F4OP027 | 7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 509 | J&J_Pathogens | TSA | TSA | No Growth | [5 empty] | TM 869 | Main_Plans
7F4OP027 | 7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | TSA JJ | 0 |  | cfu/mL | FDA,0 | 0 | 2000 | TM 869 | Main_Plans
7F4OP027 | 7F4OP027/Chem_NY_Arlasilk Phospholipid PTC | Chem_NY_Arlasilk Phospholipid PTC | NY-Microbiology | 517 | J&J_Enrichment | TSB | TSB | Turbid | [5 empty] | TM 869 | Main_Plans
7F4OP601 | 7F4OP601/Chem_109A | Chem_109A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Odor | Main_Plans
7F4OP601 | 7F4OP601/Chem_109A | Chem_109A | NY-Analytical | 396 | Odor | Odor | Odor | Passes Test (Done) | [5 empty] | J&J TM-003188 | Main_Plans
7F4OP601 | 7F4OP601/Chem_109A | Chem_109A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear colorless liquid with oily consistency (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7F4OP601 | 7F4OP601/Chem_109A | Chem_109A | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7F4OP601 | 7F4OP601/Chem_109A | Chem_109A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) | [5 empty] | J&J TMWW78013 | Main_Plans
7F4OP806 | 7F4OP806/Chem_NY_Phenoxyethanol_7F4OP806 | Chem_NY_Phenoxyethanol_7F4OP806 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7F4OP806 | 7F4OP806/Chem_NY_Phenoxyethanol_7F4OP806 | Chem_NY_Phenoxyethanol_7F4OP806 | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear, thin colorless liquid (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7F4OP806 | 7F4OP806/Chem_NY_Phenoxyethanol_7F4OP806 | Chem_NY_Phenoxyethanol_7F4OP806 | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7F4OP806 | 7F4OP806/Chem_NY_Phenoxyethanol_7F4OP806 | Chem_NY_Phenoxyethanol_7F4OP806 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) | [5 empty] | J&J TMWW78013 | Main_Plans
7F4OP806 | 7F4OP806/Chem_NY_Phenoxyethanol_7F4OP806 | Chem_NY_Phenoxyethanol_7F4OP806 | NY-Analytical | 905 | Phenol | Phenol | Phenol | 21.5 |  | ppm | FDA,0 |  | 25 | USP/NF | Main_Plans
7F4OP806 | 7F4OP806/Chem_NY_Phenoxyethanol_7F4OP806 | Chem_NY_Phenoxyethanol_7F4OP806 | NY-Analytical | 986 | Phenoxyethanol Content | Phenoxyethanol Assay | Phenoxyethanol Assay | 100.9 |  | % | FDA,1 | 98 | 102 | USP/NF | Main_Plans
7F4OP806 | 7F4OP806/Chem_NY_Phenoxyethanol_7F4OP806 | Chem_NY_Phenoxyethanol_7F4OP806 | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C | 1.1095 |  |  | FDA,3 | 1.105 | 1.11 | USP/NF | Main_Plans
7F4OQ028 | 7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | Chem_NY_Octyldodecyl Stearoyl Stearate | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Specific Gravity at 25°C | Main_Plans
7F4OQ028 | 7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | Chem_NY_Octyldodecyl Stearoyl Stearate | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.8747 |  |  | FDA,4 | 0.865 | 0.885 | TM 202 | Main_Plans
7F4OQ028 | 7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | Chem_NY_Octyldodecyl Stearoyl Stearate | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.4585 |  |  | FDA,4 | 1.455 | 1.46 | TM 204 | Main_Plans
7F4OQ028 | 7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | Chem_NY_Octyldodecyl Stearoyl Stearate | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) | [5 empty] | TM 411 | Main_Plans
7F4OQ028 | 7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | Chem_NY_Octyldodecyl Stearoyl Stearate | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7F4OQ028 | 7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | Chem_NY_Octyldodecyl Stearoyl Stearate | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Light Yellow Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7F4OQ028 | 7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | Chem_NY_Octyldodecyl Stearoyl Stearate | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7F4OQ028 | 7F4OQ028/Chem_NY_Octyldodecyl Stearoyl Stearate | Chem_NY_Octyldodecyl Stearoyl Stearate | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Mild Characteristic (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7G4OP027 | 7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | Chem_NY_SD Alcohol 40-B 200 Proof_ADM | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Certificate of Analysis | Main_Plans
7G4OP027 | 7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | Chem_NY_SD Alcohol 40-B 200 Proof_ADM | NY-Analytical | 824 | Identity of Denatonium Benzoate | Identification | Identification | Passes test with 5cm Quartz (Done) | [5 empty] | J&J TM4028 | Main_Plans
7G4OP027 | 7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | Chem_NY_SD Alcohol 40-B 200 Proof_ADM | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear colorless liquid. (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7G4OP027 | 7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | Chem_NY_SD Alcohol 40-B 200 Proof_ADM | NY-Analytical | 792 | Chemical Inspection | Odor (J&J TM-003188) | Odor | Passes Test (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7G4OP027 | 7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | Chem_NY_SD Alcohol 40-B 200 Proof_ADM | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7G4OP027 | 7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | Chem_NY_SD Alcohol 40-B 200 Proof_ADM | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C |  |  |  | FDA,4 | 0.7905 | 0.792 | J&J TMWW78012 | Main_Plans
7G4OP027 | 7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | Chem_NY_SD Alcohol 40-B 200 Proof_ADM | NY-Analytical | 822 | Specific Gravity @ 15.56°C/ 60° F | Specific Gravity | Specific Gravity at 15.56°C | 0.7941 |  |  | FDA,4 |  | 0.7952 | J&J TMWW78012 | Main_Plans
7G4OP027 | 7G4OP027/Chem_NY_SD Alcohol 40-B 200 Proof_ADM | Chem_NY_SD Alcohol 40-B 200 Proof_ADM | NY-Analytical | 390 | Identification | IR Identification | IR Identification | [6 empty] | J&J TMWW78013 | Main_Plans
7G4OP308 | 7G4OP308/Chem_111A | Chem_111A | NY-Analytical | 792 | Chemical Inspection | Odor | Odor | Matches | [5 empty] | TM 208 | Main_Plans
7G4OP308 | 7G4OP308/Chem_111A | Chem_111A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7G4OP308 | 7G4OP308/Chem_111A | Chem_111A | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Colorless to Light Yellow Liquid, Free of Particulate Matter (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7G4OP308 | 7G4OP308/Chem_111A | Chem_111A | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7G4OP308 | 7G4OP308/Chem_111A | Chem_111A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.9464 |  |  | FDA,3 | 0.942 | 0.952 | J&J TMWW78012 | Main_Plans
7G4OP308 | 7G4OP308/Chem_111A | Chem_111A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) | [5 empty] | J&J TMWW78013 | Main_Plans
7G4OP308 | 7G4OP308/Chem_111A | Chem_111A | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.4632 |  |  | FDA,3 | 1.463 | 1.473 | J&J TMWW78015 | Main_Plans
7H4OP308 | 7H4OP308/Chem_NY_Feverfew CMP-02A | Chem_NY_Feverfew CMP-02A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7H4OP308 | 7H4OP308/Chem_NY_Feverfew CMP-02A | Chem_NY_Feverfew CMP-02A | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Red Brown Liquid (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7H4OP308 | 7H4OP308/Chem_NY_Feverfew CMP-02A | Chem_NY_Feverfew CMP-02A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7H4OP308 | 7H4OP308/Chem_NY_Feverfew CMP-02A | Chem_NY_Feverfew CMP-02A | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7H4OP308 | 7H4OP308/Chem_NY_Feverfew CMP-02A | Chem_NY_Feverfew CMP-02A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) | [5 empty] | J&J TMWW78013 | Main_Plans
7H4OP308 | 7H4OP308/Chem_NY_Feverfew CMP-02A | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 509 | J&J_Pathogens | Biggy or SDA | Biggy or SDA | No Growth_SDA | [5 empty] | TM 869 | Main_Plans
7H4OP308 | 7H4OP308/Chem_NY_Feverfew CMP-02A | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 509 | J&J_Pathogens | Cetrimide | Cetrimide | No Growth | [5 empty] | TM 869 | Main_Plans
7H4OP308 | 7H4OP308/Chem_NY_Feverfew CMP-02A | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 524 | J&J Fungal | Fungal | Fungal | 0 |  | CFU/g | FDA,0 | 0 | 200 | TM 869 | Main_Plans
7H4OP308 | 7H4OP308/Chem_NY_Feverfew CMP-02A | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 509 | J&J_Pathogens | MacConkey | MacConkey | No Growth | [5 empty] | TM 869 | Main_Plans
7H4OP308 | 7H4OP308/Chem_NY_Feverfew CMP-02A | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 509 | J&J_Pathogens | MSA or VJ | MSA or VJ | No Growth_MSA | [5 empty] | TM 869 | Main_Plans
7H4OP308 | 7H4OP308/Chem_NY_Feverfew CMP-02A | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 700 | J&J SDA Controls | Negative Control | Negative Control | No Growth SDA | [5 empty] | TM 869 | Main_Plans
7H4OP308 | 7H4OP308/Chem_NY_Feverfew CMP-02A | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | NegativeControl | No Growth TSB | [5 empty] | TM 869 | Main_Plans
7H4OP308 | 7H4OP308/Chem_NY_Feverfew CMP-02A | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 509 | J&J_Pathogens | TSA | TSA | No Growth | [5 empty] | TM 869 | Main_Plans
7H4OP308 | 7H4OP308/Chem_NY_Feverfew CMP-02A | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | TSA JJ | 0 |  | cfu/mL | FDA,0 | 0 | 200 | TM 869 | Main_Plans
7H4OP308 | 7H4OP308/Chem_NY_Feverfew CMP-02A | Chem_NY_Feverfew CMP-02A | NY-Microbiology | 517 | J&J_Enrichment | TSB | TSB | Not Turbid | [5 empty] | TM 869 | Main_Plans
7H4OP601 | 7H4OP601/Chem_NY_DUB PTO | Chem_NY_DUB PTO | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7H4OP601 | 7H4OP601/Chem_NY_DUB PTO | Chem_NY_DUB PTO | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Colorless to pale yellow liquid (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7H4OP601 | 7H4OP601/Chem_NY_DUB PTO | Chem_NY_DUB PTO | NY-Analytical | 792 | Chemical Inspection | Colour (J&J TM7377) | Colour | 3 |  |  | FDA,0 |  | 60 | J&J TMWW78010/J&J TM7252 | Main_Plans
7H4OP601 | 7H4OP601/Chem_NY_DUB PTO | Chem_NY_DUB PTO | NY-Analytical | 792 | Chemical Inspection | Odor (J&J TM-003188) | Odor | Passes Test (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7H4OP601 | 7H4OP601/Chem_NY_DUB PTO | Chem_NY_DUB PTO | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7H4OP601 | 7H4OP601/Chem_NY_DUB PTO | Chem_NY_DUB PTO | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) | [5 empty] | J&J TMWW78013 | Main_Plans
7H4OP601 | 7H4OP601/Chem_NY_DUB PTO | Chem_NY_DUB PTO | NY-Microbiology | 524 | J&J Fungal | Fungal | Fungal |  |  | CFU/g | FDA,0 | 0 | 200 | TM 869 | Main_Plans
7H4OP601 | 7H4OP601/Chem_NY_DUB PTO | Chem_NY_DUB PTO | NY-Microbiology | 701 | J&J TSA Controls | Negative Control | Negative Control | [6 empty] | TM 869 | Main_Plans
7H4OP601 | 7H4OP601/Chem_NY_DUB PTO | Chem_NY_DUB PTO | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | NegativeControl | No Growth TSB | [5 empty] | TM 869 | Main_Plans
7H4OP601 | 7H4OP601/Chem_NY_DUB PTO | Chem_NY_DUB PTO | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | TSA JJ | 0 |  | cfu/mL | FDA,0 | 0 | 2000 | TM 869 | Main_Plans
7H4OP806 | 7H4OP806/Chem_NY_Kelp & Sea Salt Fragrance | Chem_NY_Kelp & Sea Salt Fragrance | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | GC Profile | Main_Plans
7H4OP806 | 7H4OP806/Chem_NY_Kelp & Sea Salt Fragrance | Chem_NY_Kelp & Sea Salt Fragrance | NY-Analytical | 839 | GC Profile | GC Profile | GC Profile | To match standard. Not required when performing IR test. (Done) | [5 empty] | J&J TM-003030 | Main_Plans
7H4OP806 | 7H4OP806/Chem_NY_Kelp & Sea Salt Fragrance | Chem_NY_Kelp & Sea Salt Fragrance | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear yellow to dark yellow liquid (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7H4OP806 | 7H4OP806/Chem_NY_Kelp & Sea Salt Fragrance | Chem_NY_Kelp & Sea Salt Fragrance | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7H4OP806 | 7H4OP806/Chem_NY_Kelp & Sea Salt Fragrance | Chem_NY_Kelp & Sea Salt Fragrance | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C | 0.9621 |  |  | FDA,3 | 0.957 | 0.967 | J&J TMWW78012 | Main_Plans
7H4OP806 | 7H4OP806/Chem_NY_Kelp & Sea Salt Fragrance | Chem_NY_Kelp & Sea Salt Fragrance | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) | [5 empty] | J&J TMWW78013 | Main_Plans
7H4OP806 | 7H4OP806/Chem_NY_Kelp & Sea Salt Fragrance | Chem_NY_Kelp & Sea Salt Fragrance | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.458 |  |  | FDA,4 | 1.456 | 1.466 | J&J TMWW78015 | Main_Plans
7J4OP485 | 7J4OP485/Chem_NY_Butylene Glycol_7J4OP485 | Chem_NY_Butylene Glycol_7J4OP485 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7J4OP485 | 7J4OP485/Chem_NY_Butylene Glycol_7J4OP485 | Chem_NY_Butylene Glycol_7J4OP485 | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Colorless viscous liquid (Done) | [5 empty] | J&J TM-003188 | Main_Plans
7J4OP485 | 7J4OP485/Chem_NY_Butylene Glycol_7J4OP485 | Chem_NY_Butylene Glycol_7J4OP485 | NY-Analytical | 792 | Chemical Inspection | Colour (J&J TMP-552) | Colour | 1 |  |  | FDA,0 |  | 10 | J&J TM-003188 | Main_Plans
7J4OP485 | 7J4OP485/Chem_NY_Butylene Glycol_7J4OP485 | Chem_NY_Butylene Glycol_7J4OP485 | NY-Analytical | 792 | Chemical Inspection | Odor (J&J TM-003188) | Odor | Passes Test (Done) | [5 empty] | J&J TM-003188 | Main_Plans
7J4OP485 | 7J4OP485/Chem_NY_Butylene Glycol_7J4OP485 | Chem_NY_Butylene Glycol_7J4OP485 | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) | [5 empty] | J&J TM-003188 | Main_Plans
7J4OP485 | 7J4OP485/Chem_NY_Butylene Glycol_7J4OP485 | Chem_NY_Butylene Glycol_7J4OP485 | NY-Analytical | 795 | Acidity | Acidity as Acetic Acid(10ml burette 0.05ml subdiv) | Acidity as Acetic Acid(10ml burette 0.05ml subdiv) | 0.002 |  | % | FDA,3 |  | 0.005 | J&J TM-003589 | Main_Plans
7J4OP485 | 7J4OP485/Chem_NY_Butylene Glycol_7J4OP485 | Chem_NY_Butylene Glycol_7J4OP485 | NY-Analytical | 796 | Assay | Assay | Assay | 99.82 |  | % | FDA,2 | 99.5 |  | J&J TM4075 | Main_Plans
7J4OP485 | 7J4OP485/Chem_NY_Butylene Glycol_7J4OP485 | Chem_NY_Butylene Glycol_7J4OP485 | NY-Analytical | 796 | Assay | Assay Final | Assay Final | 99.82 |  | % | FDA,2 | 99.5 |  | J&J TM4075 | Main_Plans
7J4OP485 | 7J4OP485/Chem_NY_Butylene Glycol_7J4OP485 | Chem_NY_Butylene Glycol_7J4OP485 | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C | 1.006 |  |  | FDA,3 | 1.004 | 1.006 | J&J TMWW78012 | Main_Plans
7J4OP485 | 7J4OP485/Chem_NY_Butylene Glycol_7J4OP485 | Chem_NY_Butylene Glycol_7J4OP485 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) | [5 empty] | J&J TMWW78013 | Main_Plans
7J4OP485 | 7J4OP485/Chem_NY_Butylene Glycol_7J4OP485 | Chem_NY_Butylene Glycol_7J4OP485 | NY-Analytical | 801 | % Water | Water Content | Water Content | 0.02 |  | % | FDA,1 |  | 0.5 | J&J TMWW78017 | Main_Plans
7K4OP027 | 7K4OP027/Chem_116A | Chem_116A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7K4OP027 | 7K4OP027/Chem_116A | Chem_116A | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Pale yellow to yellow liquid (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7K4OP027 | 7K4OP027/Chem_116A | Chem_116A | NY-Analytical | 792 | Chemical Inspection | Odor (CoA) | Odor (CoA) | Certified on CoA as Meets Odor Specification or of Characteristic Odor. (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7K4OP027 | 7K4OP027/Chem_116A | Chem_116A | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7K4OP027 | 7K4OP027/Chem_116A | Chem_116A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) | [5 empty] | J&J TMWW78013 | Main_Plans
7L4OP203 | 7L4OP203/Chem_NY_Glycerin - RI-0420 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 801 | % Water | Water | Water | 0.09 |  | % | FDA,1 |  | 5 | APAM 6.018 | Main_Plans
7L4OP203 | 7L4OP203/Chem_NY_Glycerin - RI-0420 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7L4OP203 | 7L4OP203/Chem_NY_Glycerin - RI-0420 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, syrupy liquid, free of foreign matter. (Done) | [5 empty] | See Analyte for Method # | Main_Plans
7L4OP203 | 7L4OP203/Chem_NY_Glycerin - RI-0420 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) | [5 empty] | See Analyte for Method # | Main_Plans
7L4OP203 | 7L4OP203/Chem_NY_Glycerin - RI-0420 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Similar in character and intensity to standard, practically odorless (Done) | [5 empty] | See Analyte for Method # | Main_Plans
7L4OP203 | 7L4OP203/Chem_NY_Glycerin - RI-0420 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.273 |  |  | FDA,3 | 1.249 |  | TM 202 | Main_Plans
7L4OP203 | 7L4OP203/Chem_NY_Glycerin - RI-0420 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7L4OP203 | 7L4OP203/Chem_NY_Glycerin - RI-0420 | Chem_NY_Glycerin - RI-0420 | NY-Microbiology | 867 | Celsis_Rapid | Bacteria Count (Obeserved) | Bacteria Count (Obeserved) | 0 |  | cfu | FDA,0 |  | 100 | TM 603 | Main_Plans
7L4OP203 | 7L4OP203/Chem_NY_Glycerin - RI-0420 | Chem_NY_Glycerin - RI-0420 | NY-Microbiology | 867 | Celsis_Rapid | Bacteria Count (Reported) | Bacteria Count (Reported) | 0 |  | CFU/g | FDA,0 |  | 100 | TM 603 | Main_Plans
7L4OP203 | 7L4OP203/Chem_NY_Glycerin - RI-0420 | Chem_NY_Glycerin - RI-0420 | NY-Microbiology | 867 | Celsis_Rapid | Celsis_ Result | Celsis Result | Negative ATP (Done) | [5 empty] | TM 603 | Main_Plans
7L4OP203 | 7L4OP203/Chem_NY_Glycerin - RI-0420 | Chem_NY_Glycerin - RI-0420 | NY-Microbiology | 867 | Celsis_Rapid | Mold Count (Observed) | Mold Count (Observed) | 0 |  | cfu | FDA,0 |  |  | TM 603 | Main_Plans
7L4OP203 | 7L4OP203/Chem_NY_Glycerin - RI-0420 | Chem_NY_Glycerin - RI-0420 | NY-Microbiology | 867 | Celsis_Rapid | Mold Count (Reported) | Mold Count (Reported) | 0 |  | CFU/g | FDA,0 |  |  | TM 603 | Main_Plans
7L4OP203 | 7L4OP203/Chem_NY_Glycerin - RI-0420 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 825 | Limit of Ethylene Glycol and Diethylene Glycol | Diethylene Glycol | Diethylene Glycol | 0 |  | % | FDA,1 |  | 0.1 | USP | Main_Plans
7L4OP203 | 7L4OP203/Chem_NY_Glycerin - RI-0420 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 825 | Limit of Ethylene Glycol and Diethylene Glycol | Ethylene Glycol | Ethylene Glycol | 0 |  | % | FDA,1 |  | 0.1 | USP | Main_Plans
7L4OP203 | 7L4OP203/Chem_NY_Glycerin - RI-0420 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 825 | Limit of Ethylene Glycol and Diethylene Glycol | Individual Impurity | Individual Impurity | 0.01 |  | % | FDA,1 |  | 0.1 | USP | Main_Plans
7L4OP203 | 7L4OP203/Chem_NY_Glycerin - RI-0420 | Chem_NY_Glycerin - RI-0420 | NY-Analytical | 825 | Limit of Ethylene Glycol and Diethylene Glycol | Total Impurities | Total Impurities | 0.02 |  | % | FDA,0 |  | 1 | USP | Main_Plans
7L4OP485 | 7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | Chem_NY_Polysorbate 20_7L4OP485 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Odor | Main_Plans
7L4OP485 | 7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | Chem_NY_Polysorbate 20_7L4OP485 | NY-Analytical | 396 | Odor | Odor | Odor | Passes Test (Done) | [5 empty] | J&J TM-003188 | Main_Plans
7L4OP485 | 7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | Chem_NY_Polysorbate 20_7L4OP485 | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | A pale yellow to an amber colored liquid, free of foreign matter. (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7L4OP485 | 7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | Chem_NY_Polysorbate 20_7L4OP485 | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7L4OP485 | 7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | Chem_NY_Polysorbate 20_7L4OP485 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) | [5 empty] | J&J TMWW78013 | Main_Plans
7L4OP485 | 7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | Chem_NY_Polysorbate 20_7L4OP485 | NY-Microbiology | 524 | J&J Fungal | Fungal | Fungal | 0 |  | CFU/g | FDA,0 | 0 | 200 | TM 869 | Main_Plans
7L4OP485 | 7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | Chem_NY_Polysorbate 20_7L4OP485 | NY-Microbiology | 723 | J&J TSB Controls - TSA | Negative Control | Negative Control | No Growth TSB | [5 empty] | TM 869 | Main_Plans
7L4OP485 | 7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | Chem_NY_Polysorbate 20_7L4OP485 | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | NegativeControl | No Growth TSB | [5 empty] | TM 869 | Main_Plans
7L4OP485 | 7L4OP485/Chem_NY_Polysorbate 20_7L4OP485 | Chem_NY_Polysorbate 20_7L4OP485 | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | TSA JJ | 0 |  | CFU/g | FDA,0 | 0 | 2000 | TM 869 | Main_Plans
7M4OP027 | 7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7M4OP027 | 7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear colorless liquid. (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7M4OP027 | 7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7M4OP027 | 7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 796 | Assay | Assay | Assay | 48.2 |  | % | FDA,1 | 47.5 | 52.5 | USP/NF | Main_Plans
7M4OP027 | 7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 796 | Assay | Assay Final | Assay Final | 48.2 |  | % | FDA,1 | 47.5 | 52.5 | USP/NF | Main_Plans
7M4OP027 | 7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 846 | Ratio of Alkyl Compounds | C12 + C14 Content of Total | C12 + C14 Content of Total | 95.7 |  | % | FDA,1 | 70 |  | USP/NF | Main_Plans
7M4OP027 | 7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 846 | Ratio of Alkyl Compounds | C12 Content of Total | C12 Content of Total | 56.6 |  | % | FDA,1 | 40 |  | USP/NF | Main_Plans
7M4OP027 | 7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 846 | Ratio of Alkyl Compounds | C14 Content of Total | C14 Content of Total | 39.1 |  | % | FDA,1 | 20 |  | USP/NF | Main_Plans
7M4OP027 | 7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 390 | Identification | Identification A | Identification A | Passes test. (Done) | [5 empty] | USP/NF | Main_Plans
7M4OP027 | 7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 390 | Identification | Identification B | Identification B | Passes test. (Done) | [5 empty] | USP/NF | Main_Plans
7M4OP027 | 7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 390 | Identification | Identification C | Identification C | Passes test. (Done) | [5 empty] | USP/NF | Main_Plans
7M4OP027 | 7M4OP027/Chem_NY_Benzalkonium Chloride 50% -USP | Chem_NY_Benzalkonium Chloride 50% -USP | NY-Analytical | 390 | Identification | Identification D | Identification D | Passes test. (Done) | [5 empty] | USP/NF | Main_Plans
7O4AA580 | 7O4AA580/Chem_NY_Ethylene Glycol Monohexyl Ether | Chem_NY_Ethylene Glycol Monohexyl Ether | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Identification | Main_Plans
7O4AA580 | 7O4AA580/Chem_NY_Ethylene Glycol Monohexyl Ether | Chem_NY_Ethylene Glycol Monohexyl Ether | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) | [5 empty] | TM 411 | Main_Plans
7O4AA580 | 7O4AA580/Chem_NY_Ethylene Glycol Monohexyl Ether | Chem_NY_Ethylene Glycol Monohexyl Ether | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, Colorless, Viscous Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4AA580 | 7O4AA580/Chem_NY_Ethylene Glycol Monohexyl Ether | Chem_NY_Ethylene Glycol Monohexyl Ether | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4AA580 | 7O4AA580/Chem_NY_Ethylene Glycol Monohexyl Ether | Chem_NY_Ethylene Glycol Monohexyl Ether | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4AA580 | 7O4AA580/Chem_NY_Ethylene Glycol Monohexyl Ether | Chem_NY_Ethylene Glycol Monohexyl Ether | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Pungent (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4BJ131 | 7O4BJ131/Chem_NY_Acticide CBM-2 | Chem_NY_Acticide CBM-2 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Identification | Main_Plans
7O4BJ131 | 7O4BJ131/Chem_NY_Acticide CBM-2 | Chem_NY_Acticide CBM-2 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7O4BJ131 | 7O4BJ131/Chem_NY_Acticide CBM-2 | Chem_NY_Acticide CBM-2 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Off-white - greenish dispersion (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4BJ131 | 7O4BJ131/Chem_NY_Acticide CBM-2 | Chem_NY_Acticide CBM-2 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4BJ131 | 7O4BJ131/Chem_NY_Acticide CBM-2 | Chem_NY_Acticide CBM-2 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4BJ131 | 7O4BJ131/Chem_NY_Acticide CBM-2 | Chem_NY_Acticide CBM-2 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4BM301 | 7O4BM301/Chem_126A | Chem_126A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Identification | Main_Plans
7O4BM301 | 7O4BM301/Chem_126A | Chem_126A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7O4BM301 | 7O4BM301/Chem_126A | Chem_126A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4BM301 | 7O4BM301/Chem_126A | Chem_126A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to Amber (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4BM301 | 7O4BM301/Chem_126A | Chem_126A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4BM301 | 7O4BM301/Chem_126A | Chem_126A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4DH247 | 7O4DH247/Chem_201A | Chem_201A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Specific Gravity at 25°C | Main_Plans
7O4DH247 | 7O4DH247/Chem_201A | Chem_201A | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C | 1.2883 |  |  | FDA,3 | 1.27 | 1.31 | TM 202 | Main_Plans
7O4DH247 | 7O4DH247/Chem_201A | Chem_201A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear solution, free from solids (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4DH247 | 7O4DH247/Chem_201A | Chem_201A | NY-Analytical | 792 | Chemical Inspection | Color (TM 230) | Color | 2 |  |  | FDA,1 |  | 2 | TM208/220/230/USP/Kao LM | Main_Plans
7O4DH247 | 7O4DH247/Chem_201A | Chem_201A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4DH247 | 7O4DH247/Chem_201A | Chem_201A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Bleach/chlorine (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4DX220 | 7O4DX220/Chem_NY_Neutralizing Agent | Chem_NY_Neutralizing Agent | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Specific Gravity at 25°C | Main_Plans
7O4DX220 | 7O4DX220/Chem_NY_Neutralizing Agent | Chem_NY_Neutralizing Agent | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.0999 |  |  | FDA,4 | 1.088 | 1.118 | TM 202 | Main_Plans
7O4DX220 | 7O4DX220/Chem_NY_Neutralizing Agent | Chem_NY_Neutralizing Agent | NY-Analytical | 451 | Refractive Index | Refractive Index | Refractive Index at 25°C | 1.4685 |  |  | FDA,4 | 1.46 | 1.48 | TM 204 | Main_Plans
7O4DX220 | 7O4DX220/Chem_NY_Neutralizing Agent | Chem_NY_Neutralizing Agent | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7O4DX220 | 7O4DX220/Chem_NY_Neutralizing Agent | Chem_NY_Neutralizing Agent | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4DX220 | 7O4DX220/Chem_NY_Neutralizing Agent | Chem_NY_Neutralizing Agent | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Light Yellow Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4DX220 | 7O4DX220/Chem_NY_Neutralizing Agent | Chem_NY_Neutralizing Agent | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4DX220 | 7O4DX220/Chem_NY_Neutralizing Agent | Chem_NY_Neutralizing Agent | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic – comparable to standard (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OA234 | 7O4OA234/Chem_131A | Chem_131A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | pH as is | Main_Plans
7O4OA234 | 7O4OA234/Chem_131A | Chem_131A | NY-Analytical | 447 | pH as is | pH | pH | 3.22 |  |  | FDA,2 | 3 | 3.3 | TM 200 | Main_Plans
7O4OA234 | 7O4OA234/Chem_131A | Chem_131A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7O4OA234 | 7O4OA234/Chem_131A | Chem_131A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Opaque, Thin Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OA234 | 7O4OA234/Chem_131A | Chem_131A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OA234 | 7O4OA234/Chem_131A | Chem_131A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OA234 | 7O4OA234/Chem_131A | Chem_131A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OA347 | 7O4OA347/Chem_NY_Lauryl Glucoside | Chem_NY_Lauryl Glucoside | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7O4OA347 | 7O4OA347/Chem_NY_Lauryl Glucoside | Chem_NY_Lauryl Glucoside | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Match standard (Done) | [5 empty] | TM 208 | Main_Plans
7O4OA347 | 7O4OA347/Chem_NY_Lauryl Glucoside | Chem_NY_Lauryl Glucoside | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Match standard (Done) | [5 empty] | TM 208 | Main_Plans
7O4OA347 | 7O4OA347/Chem_NY_Lauryl Glucoside | Chem_NY_Lauryl Glucoside | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Main_Plans
7O4OA347 | 7O4OA347/Chem_NY_Lauryl Glucoside | Chem_NY_Lauryl Glucoside | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) | [5 empty] | TM 208 | Main_Plans
7O4OA347 | 7O4OA347/Chem_NY_Lauryl Glucoside | Chem_NY_Lauryl Glucoside | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Comparison of scans of sample vs standard (Done) | [5 empty] | TM 411 | Main_Plans
7O4OA401 | 7O4OA401/Chem_132A | Chem_132A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Specific Gravity at 25°C | Main_Plans
7O4OA401 | 7O4OA401/Chem_132A | Chem_132A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.0225 |  |  | FDA,4 | 1.0115 | 1.0315 | TM 202 | Main_Plans
7O4OA401 | 7O4OA401/Chem_132A | Chem_132A | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.485 |  |  | FDA,3 | 1.4794 | 1.4994 | TM 204 | Main_Plans
7O4OA401 | 7O4OA401/Chem_132A | Chem_132A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7O4OA401 | 7O4OA401/Chem_132A | Chem_132A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OA401 | 7O4OA401/Chem_132A | Chem_132A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to light yellow (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OA401 | 7O4OA401/Chem_132A | Chem_132A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OA401 | 7O4OA401/Chem_132A | Chem_132A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Fresh Floral (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OA490 | 7O4OA490/Chem_133A | Chem_133A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Specific Gravity at 25°C | Main_Plans
7O4OA490 | 7O4OA490/Chem_133A | Chem_133A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.9182 |  |  | FDA,3 | 0.909 | 0.929 | TM 202 | Main_Plans
7O4OA490 | 7O4OA490/Chem_133A | Chem_133A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7O4OA490 | 7O4OA490/Chem_133A | Chem_133A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OA490 | 7O4OA490/Chem_133A | Chem_133A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OA490 | 7O4OA490/Chem_133A | Chem_133A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OA490 | 7O4OA490/Chem_133A | Chem_133A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Ether (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OA580 | 7O4OA580/Chem_NY_Glycacil SG | Chem_NY_Glycacil SG | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Identification | Main_Plans
7O4OA580 | 7O4OA580/Chem_NY_Glycacil SG | Chem_NY_Glycacil SG | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7O4OA580 | 7O4OA580/Chem_NY_Glycacil SG | Chem_NY_Glycacil SG | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OA580 | 7O4OA580/Chem_NY_Glycacil SG | Chem_NY_Glycacil SG | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to Amber (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OA580 | 7O4OA580/Chem_NY_Glycacil SG | Chem_NY_Glycacil SG | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OA580 | 7O4OA580/Chem_NY_Glycacil SG | Chem_NY_Glycacil SG | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic Match Standard (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OA634 | 7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | Chem_NY_Water Soluble Flowering Herbs Fragrance | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Specific Gravity at 25°C | Main_Plans
7O4OA634 | 7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | Chem_NY_Water Soluble Flowering Herbs Fragrance | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.0629 |  |  | FDA,4 | 1.0531 | 1.0831 | TM 202 | Main_Plans
7O4OA634 | 7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | Chem_NY_Water Soluble Flowering Herbs Fragrance | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.47 |  |  | FDA,4 | 1.4655 | 1.4855 | TM 204 | Main_Plans
7O4OA634 | 7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | Chem_NY_Water Soluble Flowering Herbs Fragrance | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7O4OA634 | 7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | Chem_NY_Water Soluble Flowering Herbs Fragrance | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OA634 | 7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | Chem_NY_Water Soluble Flowering Herbs Fragrance | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Pale Yellow to Straw (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OA634 | 7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | Chem_NY_Water Soluble Flowering Herbs Fragrance | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OA634 | 7O4OA634/Chem_NY_Water Soluble Flowering Herbs Fragrance | Chem_NY_Water Soluble Flowering Herbs Fragrance | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Floral (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OA709 | 7O4OA709/Chem_NY_Cucumber Extract # 11-19-3 | Chem_NY_Cucumber Extract # 11-19-3 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [6 empty] | Specific Gravity at 25°C | Main_Plans
7O4OA709 | 7O4OA709/Chem_NY_Cucumber Extract # 11-19-3 | Chem_NY_Cucumber Extract # 11-19-3 | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C |  |  |  | FDA,3 | 1.034 | 1.044 | TM 202 | Main_Plans
7O4OA709 | 7O4OA709/Chem_NY_Cucumber Extract # 11-19-3 | Chem_NY_Cucumber Extract # 11-19-3 | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.425 |  |  | FDA,3 | 1.42 | 1.43 | TM 204 | Main_Plans
7O4OA709 | 7O4OA709/Chem_NY_Cucumber Extract # 11-19-3 | Chem_NY_Cucumber Extract # 11-19-3 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Main_Plans
7O4OA709 | 7O4OA709/Chem_NY_Cucumber Extract # 11-19-3 | Chem_NY_Cucumber Extract # 11-19-3 | NY-Analytical | 792 | Chemical Inspection | Description (TM 208) | Description | Clear colorless liquid having a characteristic odor (Done) | [5 empty] | TM 208 | Main_Plans
7O4OA911 | 7O4OA911/Chem_NY_Zinc Sulfate Heptahydrate | Chem_NY_Zinc Sulfate Heptahydrate | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7O4OA911 | 7O4OA911/Chem_NY_Zinc Sulfate Heptahydrate | Chem_NY_Zinc Sulfate Heptahydrate | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Crystals/crystalline powder (Done) | [5 empty] | TM 208 | Main_Plans
7O4OA911 | 7O4OA911/Chem_NY_Zinc Sulfate Heptahydrate | Chem_NY_Zinc Sulfate Heptahydrate | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to white (Done) | [5 empty] | TM 208 | Main_Plans
7O4OA911 | 7O4OA911/Chem_NY_Zinc Sulfate Heptahydrate | Chem_NY_Zinc Sulfate Heptahydrate | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Main_Plans
7O4OA911 | 7O4OA911/Chem_NY_Zinc Sulfate Heptahydrate | Chem_NY_Zinc Sulfate Heptahydrate | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) | [5 empty] | TM 208 | Main_Plans
7O4OA911 | 7O4OA911/Chem_NY_Zinc Sulfate Heptahydrate | Chem_NY_Zinc Sulfate Heptahydrate | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) | [5 empty] | TM 411 | Main_Plans
7O4OD351 | 7O4OD351/Chem_NY_Benzalkonium Chloride 50% | Chem_NY_Benzalkonium Chloride 50% | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [6 empty] | Appearance | Main_Plans
7O4OD351 | 7O4OD351/Chem_NY_Benzalkonium Chloride 50% | Chem_NY_Benzalkonium Chloride 50% | NY-Analytical | 397 | Appearance | Appearance | Appearance | Clear liquid, free from foreign material (FFFM) (Done) | [5 empty] | TM 208 | Main_Plans
7O4OD351 | 7O4OD351/Chem_NY_Benzalkonium Chloride 50% | Chem_NY_Benzalkonium Chloride 50% | NY-Analytical | 938 | Ethanol | Ethanol | Ethanol |  |  | % v/v | FDA,2 | 5.9 | 6.5 | TM 323 | Main_Plans
7O4OD351 | 7O4OD351/Chem_NY_Benzalkonium Chloride 50% | Chem_NY_Benzalkonium Chloride 50% | NY-Analytical | 938 | Ethanol | Ethanol Final | Ethanol Final |  |  | % v/v | FDA,2 | 5.9 | 6.5 | TM 323 | Main_Plans
7O4OD351 | 7O4OD351/Chem_NY_Benzalkonium Chloride 50% | Chem_NY_Benzalkonium Chloride 50% | NY-Analytical | 455 | Identification A | Identification A - Chloride | Identification A - Chloride | [6 empty] | TM 4001 | Main_Plans
7O4OD351 | 7O4OD351/Chem_NY_Benzalkonium Chloride 50% | Chem_NY_Benzalkonium Chloride 50% | NY-Analytical | 456 | Identification B | Identification B - Benzalkonium | Identification B - Benzalkonium | [6 empty] | TM 4001 | Main_Plans
7O4OD351 | 7O4OD351/Chem_NY_Benzalkonium Chloride 50% | Chem_NY_Benzalkonium Chloride 50% | NY-Analytical | 1064 | Total % ADBAC | % BZK and Dual-Quat | % BZK and Dual-Quat |  |  | % | FDA,1 | 47.5 | 52.5 | TM 434 | Main_Plans
7O4OD363 | 7O4OD363/Chem_137A | Chem_137A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | pH as is | Main_Plans
7O4OD363 | 7O4OD363/Chem_137A | Chem_137A | NY-Analytical | 447 | pH as is | pH | pH | 5.06 |  |  | FDA,1 | 4.5 | 5.5 | TM 200 | Main_Plans
7O4OD363 | 7O4OD363/Chem_137A | Chem_137A | NY-Analytical | 806 | Total Solids | Total Solids (2 g for 3 hr @ 105°C) | Total Solids (2 g for 3 hr @ 105°C) | 43.5 |  | % | FDA,1 | 41 | 45 | TM 206 | Main_Plans
7O4OD363 | 7O4OD363/Chem_137A | Chem_137A | NY-Analytical | 806 | Total Solids | Total Solids Final (2 g for 3 hr @ 105°C) | Total Solids Final (2 g for 3 hr @ 105°C) | 43.45 |  | % | FDA,1 | 41 | 45 | TM 206 | Main_Plans
7O4OD363 | 7O4OD363/Chem_137A | Chem_137A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7O4OD363 | 7O4OD363/Chem_137A | Chem_137A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OD363 | 7O4OD363/Chem_137A | Chem_137A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to pale yellow (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OD419 | 7O4OD419/Chem_138A | Chem_138A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA does not meet specifications (OOS) | [5 empty] | Identification | Main_Plans
7O4OD419 | 7O4OD419/Chem_138A | Chem_138A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7O4OD419 | 7O4OD419/Chem_138A | Chem_138A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Semi Solid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OD419 | 7O4OD419/Chem_138A | Chem_138A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White to light yellow (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OD419 | 7O4OD419/Chem_138A | Chem_138A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OD419 | 7O4OD419/Chem_138A | Chem_138A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Fatty (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OD748 | 7O4OD748/Chem_139A | Chem_139A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Identification | Main_Plans
7O4OD748 | 7O4OD748/Chem_139A | Chem_139A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7O4OD748 | 7O4OD748/Chem_139A | Chem_139A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | White solid crystals (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OD748 | 7O4OD748/Chem_139A | Chem_139A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White Crystals (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OD748 | 7O4OD748/Chem_139A | Chem_139A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OD748 | 7O4OD748/Chem_139A | Chem_139A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | To match standard (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OF575 | 7O4OF575/Che_NY_Plantaren 2000N - UP | Che_NY_Plantaren 2000N - UP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Identification | Main_Plans
7O4OF575 | 7O4OF575/Che_NY_Plantaren 2000N - UP | Che_NY_Plantaren 2000N - UP | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7O4OF575 | 7O4OF575/Che_NY_Plantaren 2000N - UP | Che_NY_Plantaren 2000N - UP | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Hazy, Pale Yellow Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OF575 | 7O4OF575/Che_NY_Plantaren 2000N - UP | Che_NY_Plantaren 2000N - UP | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OF602 | 7O4OF602/Chem_141A | Chem_141A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Identification | Main_Plans
7O4OF602 | 7O4OF602/Chem_141A | Chem_141A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7O4OF602 | 7O4OF602/Chem_141A | Chem_141A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OG169 | 7O4OG169/Chem_142A | Chem_142A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Specific Gravity at 25°C | Main_Plans
7O4OG169 | 7O4OG169/Chem_142A | Chem_142A | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C | 1.08 |  |  | FDA,3 | 1.08 | 1.09 | TM 202 | Main_Plans
7O4OG169 | 7O4OG169/Chem_142A | Chem_142A | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.435 |  |  | FDA,3 | 1.435 | 1.437 | TM 204 | Main_Plans
7O4OG169 | 7O4OG169/Chem_142A | Chem_142A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7O4OG169 | 7O4OG169/Chem_142A | Chem_142A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OG169 | 7O4OG169/Chem_142A | Chem_142A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OG169 | 7O4OG169/Chem_142A | Chem_142A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OH247 | 7O4OH247/Chem_143A | Chem_143A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | pH | Main_Plans
7O4OH247 | 7O4OH247/Chem_143A | Chem_143A | NY-Analytical | 398 | pH | pH (1% w/v aqueous solution) | pH (1% w/v aqueous solution) | 2 |  |  | FDA,1 |  | 2 | TM 200 | Main_Plans
7O4OH247 | 7O4OH247/Chem_143A | Chem_143A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7O4OH247 | 7O4OH247/Chem_143A | Chem_143A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear viscous liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OH247 | 7O4OH247/Chem_143A | Chem_143A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to dark amber (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OH247 | 7O4OH247/Chem_143A | Chem_143A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OJ123 | 7O4OJ123/Chem_NY_NET-KOB | Chem_NY_NET-KOB | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Identification | Main_Plans
7O4OJ123 | 7O4OJ123/Chem_NY_NET-KOB | Chem_NY_NET-KOB | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7O4OJ123 | 7O4OJ123/Chem_NY_NET-KOB | Chem_NY_NET-KOB | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Viscous Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OJ123 | 7O4OJ123/Chem_NY_NET-KOB | Chem_NY_NET-KOB | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to pale yellow (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OJ123 | 7O4OJ123/Chem_NY_NET-KOB | Chem_NY_NET-KOB | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OJ123 | 7O4OJ123/Chem_NY_NET-KOB | Chem_NY_NET-KOB | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristic (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OM165 | 7O4OM165/Chem_146A | Chem_146A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Identification | Main_Plans
7O4OM165 | 7O4OM165/Chem_146A | Chem_146A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7O4OM165 | 7O4OM165/Chem_146A | Chem_146A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, slightly oily viscous liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OM165 | 7O4OM165/Chem_146A | Chem_146A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Light Yellow (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OM165 | 7O4OM165/Chem_146A | Chem_146A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OM165 | 7O4OM165/Chem_146A | Chem_146A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | To match standard (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OM666 | 7O4OM666/Chem_NY_Sodium Benzoate EDF NF/FCC | Chem_NY_Sodium Benzoate EDF NF/FCC | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7O4OM666 | 7O4OM666/Chem_NY_Sodium Benzoate EDF NF/FCC | Chem_NY_Sodium Benzoate EDF NF/FCC | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Extruded Pellets (Done) | [5 empty] | TM 208 | Main_Plans
7O4OM666 | 7O4OM666/Chem_NY_Sodium Benzoate EDF NF/FCC | Chem_NY_Sodium Benzoate EDF NF/FCC | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) | [5 empty] | TM 208 | Main_Plans
7O4OM666 | 7O4OM666/Chem_NY_Sodium Benzoate EDF NF/FCC | Chem_NY_Sodium Benzoate EDF NF/FCC | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Main_Plans
7O4OM666 | 7O4OM666/Chem_NY_Sodium Benzoate EDF NF/FCC | Chem_NY_Sodium Benzoate EDF NF/FCC | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Odorless (Done) | [5 empty] | TM 208 | Main_Plans
7O4OM666 | 7O4OM666/Chem_NY_Sodium Benzoate EDF NF/FCC | Chem_NY_Sodium Benzoate EDF NF/FCC | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7O4OP065 | 7O4OP065/Chem_NY_Actipone White Tea | Chem_NY_Actipone White Tea | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Density @ 20° C | Main_Plans
7O4OP065 | 7O4OP065/Chem_NY_Actipone White Tea | Chem_NY_Actipone White Tea | NY-Analytical | 831 | Density @ 20° C | Density | Density | 1.1173 |  |  | FDA,3 | 1.112 | 1.122 | TM 202 | Main_Plans
7O4OP065 | 7O4OP065/Chem_NY_Actipone White Tea | Chem_NY_Actipone White Tea | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.392 |  |  | FDA,3 | 1.388 | 1.398 | TM 204 | Main_Plans
7O4OP065 | 7O4OP065/Chem_NY_Actipone White Tea | Chem_NY_Actipone White Tea | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear to slightly cloudy liquid (Done) | [5 empty] | TM 208 | Main_Plans
7O4OP065 | 7O4OP065/Chem_NY_Actipone White Tea | Chem_NY_Actipone White Tea | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Yellow to Dark Brown (Done) | [5 empty] | TM 208 | Main_Plans
7O4OP065 | 7O4OP065/Chem_NY_Actipone White Tea | Chem_NY_Actipone White Tea | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Main_Plans
7O4OP065 | 7O4OP065/Chem_NY_Actipone White Tea | Chem_NY_Actipone White Tea | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) | [5 empty] | TM 208 | Main_Plans
7O4OP065 | 7O4OP065/Chem_NY_Actipone White Tea | Chem_NY_Actipone White Tea | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) | [5 empty] | TM 411 | Main_Plans
7O4OP065 | 7O4OP065/Chem_NY_Actipone White Tea | Chem_NY_Actipone White Tea | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram) | TAMC (Average per Gram) | 0 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
7O4OP065 | 7O4OP065/Chem_NY_Actipone White Tea | Chem_NY_Actipone White Tea | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
7O4OP065 | 7O4OP065/Chem_NY_Actipone White Tea | Chem_NY_Actipone White Tea | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram) | TMYC (Average per Gram) |  |  | CFU/g | FDA,1 |  | 100 | TM 610 | Main_Plans
7O4OP065 | 7O4OP065/Chem_NY_Actipone White Tea | Chem_NY_Actipone White Tea | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  |  | cfu | FDA,0 |  |  | TM 610 | Main_Plans
7O4OP065 | 7O4OP065/Chem_NY_Actipone White Tea | Chem_NY_Actipone White Tea | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Negative Rods (OOS) | [5 empty] | WI QAS 076 | Main_Plans
7O4OP075 | 7O4OP075/Chem_NY_Magique Lift 02LF | Chem_NY_Magique Lift 02LF | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA does not meet specifications (OOS) | [5 empty] | Specific Gravity at 25°C | Main_Plans
7O4OP075 | 7O4OP075/Chem_NY_Magique Lift 02LF | Chem_NY_Magique Lift 02LF | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.9516 |  |  | FDA,3 | 0.938 | 0.948 | TM 202 | Main_Plans
7O4OP075 | 7O4OP075/Chem_NY_Magique Lift 02LF | Chem_NY_Magique Lift 02LF | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.48 |  |  | FDA,3 | 1.474 | 1.484 | TM 204 | Main_Plans
7O4OP075 | 7O4OP075/Chem_NY_Magique Lift 02LF | Chem_NY_Magique Lift 02LF | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7O4OP075 | 7O4OP075/Chem_NY_Magique Lift 02LF | Chem_NY_Magique Lift 02LF | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OP075 | 7O4OP075/Chem_NY_Magique Lift 02LF | Chem_NY_Magique Lift 02LF | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless to very light yellow (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OP075 | 7O4OP075/Chem_NY_Magique Lift 02LF | Chem_NY_Magique Lift 02LF | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OP075 | 7O4OP075/Chem_NY_Magique Lift 02LF | Chem_NY_Magique Lift 02LF | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OP093 | 7O4OP093/Chem_NY_PEG-11 | Chem_NY_PEG-11 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | COA | Main_Plans
7O4OP093 | 7O4OP093/Chem_NY_PEG-11 | Chem_NY_PEG-11 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, Light liquid (Done) | [5 empty] | TM 208 | Main_Plans
7O4OP093 | 7O4OP093/Chem_NY_PEG-11 | Chem_NY_PEG-11 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Match standard (Done) | [5 empty] | TM 208 | Main_Plans
7O4OP093 | 7O4OP093/Chem_NY_PEG-11 | Chem_NY_PEG-11 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Main_Plans
7O4OP093 | 7O4OP093/Chem_NY_PEG-11 | Chem_NY_PEG-11 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) | [5 empty] | TM 208 | Main_Plans
7O4OP093 | 7O4OP093/Chem_NY_PEG-11 | Chem_NY_PEG-11 | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) | [5 empty] | TM 411 | Main_Plans
7O4OP113 | 7O4OP113/Chem_NY_Robertet The Matcha I | Chem_NY_Robertet The Matcha I | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Specific Gravity at 25°C | Main_Plans
7O4OP113 | 7O4OP113/Chem_NY_Robertet The Matcha I | Chem_NY_Robertet The Matcha I | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C | 0.9652 |  |  | FDA,3 | 0.96 | 0.97 | TM 202 | Main_Plans
7O4OP113 | 7O4OP113/Chem_NY_Robertet The Matcha I | Chem_NY_Robertet The Matcha I | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.47 |  |  | FDA,3 | 1.469 | 1.479 | TM 204 | Main_Plans
7O4OP113 | 7O4OP113/Chem_NY_Robertet The Matcha I | Chem_NY_Robertet The Matcha I | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Oily liquid (Done) | [5 empty] | TM 208 | Main_Plans
7O4OP113 | 7O4OP113/Chem_NY_Robertet The Matcha I | Chem_NY_Robertet The Matcha I | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Pale Yellow – Light Yellow (Done) | [5 empty] | TM 208 | Main_Plans
7O4OP113 | 7O4OP113/Chem_NY_Robertet The Matcha I | Chem_NY_Robertet The Matcha I | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Main_Plans
7O4OP113 | 7O4OP113/Chem_NY_Robertet The Matcha I | Chem_NY_Robertet The Matcha I | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) | [5 empty] | TM 208 | Main_Plans
7O4OP113 | 7O4OP113/Chem_NY_Robertet The Matcha I | Chem_NY_Robertet The Matcha I | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) | [5 empty] | TM 411 | Main_Plans
7O4OP197 | 7O4OP197/Chem_NY_Emulgade Smart/MB_01 | Chem_NY_Emulgade Smart/MB_01 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | pH | Main_Plans
7O4OP197 | 7O4OP197/Chem_NY_Emulgade Smart/MB_01 | Chem_NY_Emulgade Smart/MB_01 | NY-Analytical | 398 | pH | pH | pH | 3 |  |  | FDA,1 | 3 | 3.5 | TM 200 | Main_Plans
7O4OP197 | 7O4OP197/Chem_NY_Emulgade Smart/MB_01 | Chem_NY_Emulgade Smart/MB_01 | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Match standard (Done) | [5 empty] | TM 208 | Main_Plans
7O4OP197 | 7O4OP197/Chem_NY_Emulgade Smart/MB_01 | Chem_NY_Emulgade Smart/MB_01 | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White Liquid (Done) | [5 empty] | TM 208 | Main_Plans
7O4OP197 | 7O4OP197/Chem_NY_Emulgade Smart/MB_01 | Chem_NY_Emulgade Smart/MB_01 | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Main_Plans
7O4OP197 | 7O4OP197/Chem_NY_Emulgade Smart/MB_01 | Chem_NY_Emulgade Smart/MB_01 | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) | [5 empty] | TM 208 | Main_Plans
7O4OP197 | 7O4OP197/Chem_NY_Emulgade Smart/MB_01 | Chem_NY_Emulgade Smart/MB_01 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) | [5 empty] | TM 411 | Main_Plans
7O4OP230 | 7O4OP230/Chem_NY_Hydromanil | Chem_NY_Hydromanil | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7O4OP230 | 7O4OP230/Chem_NY_Hydromanil | Chem_NY_Hydromanil | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | An opalescent, pale yellow to pale brown viscous liquid. (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7O4OP230 | 7O4OP230/Chem_NY_Hydromanil | Chem_NY_Hydromanil | NY-Analytical | 792 | Chemical Inspection | Odor (J&J TM-003188) | Odor | Passes Test (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7O4OP230 | 7O4OP230/Chem_NY_Hydromanil | Chem_NY_Hydromanil | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7O4OP230 | 7O4OP230/Chem_NY_Hydromanil | Chem_NY_Hydromanil | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C | 1.1158 |  |  | FDA,3 | 1.1 | 1.14 | J&J TMWW78012 | Main_Plans
7O4OP230 | 7O4OP230/Chem_NY_Hydromanil | Chem_NY_Hydromanil | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) | [5 empty] | J&J TMWW78013 | Main_Plans
7O4OP230 | 7O4OP230/Chem_NY_Hydromanil | Chem_NY_Hydromanil | NY-Microbiology | 509 | J&J_Pathogens | Biggy or SDA | Biggy or SDA | No Growth_SDA | [5 empty] | TM 869 | Main_Plans
7O4OP230 | 7O4OP230/Chem_NY_Hydromanil | Chem_NY_Hydromanil | NY-Microbiology | 509 | J&J_Pathogens | Cetrimide | Cetrimide | No Growth | [5 empty] | TM 869 | Main_Plans
7O4OP230 | 7O4OP230/Chem_NY_Hydromanil | Chem_NY_Hydromanil | NY-Microbiology | 524 | J&J Fungal | Fungal | Fungal | 0 |  | CFU/g | FDA,0 | 0 | 200 | TM 869 | Main_Plans
7O4OP230 | 7O4OP230/Chem_NY_Hydromanil | Chem_NY_Hydromanil | NY-Microbiology | 509 | J&J_Pathogens | MacConkey | MacConkey | No Growth | [5 empty] | TM 869 | Main_Plans
7O4OP230 | 7O4OP230/Chem_NY_Hydromanil | Chem_NY_Hydromanil | NY-Microbiology | 509 | J&J_Pathogens | MSA or VJ | MSA or VJ | No Growth_MSA | [5 empty] | TM 869 | Main_Plans
7O4OP230 | 7O4OP230/Chem_NY_Hydromanil | Chem_NY_Hydromanil | NY-Microbiology | 700 | J&J SDA Controls | Negative Control | Negative Control | No Growth SDA | [5 empty] | TM 869 | Main_Plans
7O4OP230 | 7O4OP230/Chem_NY_Hydromanil | Chem_NY_Hydromanil | NY-Microbiology | 724 | J&J TSB Controls - SAB | NegativeControl | NegativeControl | No Growth TSB | [5 empty] | TM 869 | Main_Plans
7O4OP230 | 7O4OP230/Chem_NY_Hydromanil | Chem_NY_Hydromanil | NY-Microbiology | 509 | J&J_Pathogens | TSA | TSA | No Growth | [5 empty] | TM 869 | Main_Plans
7O4OP230 | 7O4OP230/Chem_NY_Hydromanil | Chem_NY_Hydromanil | NY-Microbiology | 537 | J&J_Bacteria | TSA JJ | TSA JJ | 0 |  | cfu/mL | FDA,0 | 0 | 2000 | TM 869 | Main_Plans
7O4OP230 | 7O4OP230/Chem_NY_Hydromanil | Chem_NY_Hydromanil | NY-Microbiology | 517 | J&J_Enrichment | TSB | TSB | Not Turbid | [5 empty] | TM 869 | Main_Plans
7O4OP230 | 7O4OP230/Chem_NY_Hydromanil | Chem_NY_Hydromanil | NY-Analytical | 447 | pH as is | pH | pH | 4.87 |  |  | FDA,1 | 4.5 | 5.5 | TM WW78014 | Main_Plans
7O4OP230 | 7O4OP230/Chem_NY_Hydromanil | Chem_NY_Hydromanil | NY-Analytical | 917 | Total Sugar | Total Sugar | Total Sugar | 8.2 |  | % | FDA,1 | 6.5 | 9 | TM-002568 | Main_Plans
7O4OP230 | 7O4OP230/Chem_NY_Hydromanil | Chem_NY_Hydromanil | NY-Analytical | 916 | Reducing Sugars | Reducing Sugars | Reducing Sugars | 3.3 |  | % | FDA,1 | 2.5 | 3.6 | TM-002569 | Main_Plans
7O4OP253 | 7O4OP253/Chem_NY_BTC 1010 | Chem_NY_BTC 1010 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | pH | Main_Plans
7O4OP253 | 7O4OP253/Chem_NY_BTC 1010 | Chem_NY_BTC 1010 | NY-Analytical | 398 | pH | pH (10% in water) | pH (10% in water) | 7.19 |  |  | FDA,1 | 6 | 9 | TM 200 | Main_Plans
7O4OP253 | 7O4OP253/Chem_NY_BTC 1010 | Chem_NY_BTC 1010 | NY-Analytical | 796 | Assay | Assay | Assay | 50.8 |  | % | FDA,0 | 50 | 55 | TM 301/434 | Main_Plans
7O4OP253 | 7O4OP253/Chem_NY_BTC 1010 | Chem_NY_BTC 1010 | NY-Analytical | 796 | Assay | Assay Final | Assay Final | 50.85 |  | % | FDA,0 | 50 | 55 | TM 301/434 | Main_Plans
7O4OP284 | 7O4OP284/Chem_155A | Chem_155A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | GC Profile | Main_Plans
7O4OP284 | 7O4OP284/Chem_155A | Chem_155A | NY-Analytical | 839 | GC Profile | GC Profile | GC Profile | To match standard. Not required when performing IR test. (Done) | [5 empty] | J&J TM-001862 | Main_Plans
7O4OP284 | 7O4OP284/Chem_155A | Chem_155A | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear, yellow liquid. (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7O4OP284 | 7O4OP284/Chem_155A | Chem_155A | NY-Analytical | 792 | Chemical Inspection | Odor (J&J TM-003188) | Odor | Supplier certifies on C of A that odor matches standard (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7O4OP284 | 7O4OP284/Chem_155A | Chem_155A | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7O4OP284 | 7O4OP284/Chem_155A | Chem_155A | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C | 0.9722 |  |  | FDA,3 | 0.963 | 0.973 | J&J TMWW78012 | Main_Plans
7O4OP284 | 7O4OP284/Chem_155A | Chem_155A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) | [5 empty] | J&J TMWW78013 | Main_Plans
7O4OP284 | 7O4OP284/Chem_155A | Chem_155A | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.442 |  |  | FDA,4 | 1.442 | 1.452 | J&J TMWW78015 | Main_Plans
7O4OP345 | 7O4OP345/Chem_157A | Chem_157A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Identification | Main_Plans
7O4OP345 | 7O4OP345/Chem_157A | Chem_157A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7O4OP345 | 7O4OP345/Chem_157A | Chem_157A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OP345 | 7O4OP345/Chem_157A | Chem_157A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Yellow (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OP345 | 7O4OP345/Chem_157A | Chem_157A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OP345 | 7O4OP345/Chem_157A | Chem_157A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Phenoxyethanol-like (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OP567 | 7O4OP567/Chem_NY_BTC 1210-80% | Chem_NY_BTC 1210-80% | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | COA | Main_Plans
7O4OP567 | 7O4OP567/Chem_NY_BTC 1210-80% | Chem_NY_BTC 1210-80% | NY-Analytical | 398 | pH | pH (10% in water) | pH (10% in water) | 7.79 |  |  | FDA,1 | 6 | 8.5 | TM 200 | Main_Plans
7O4OP567 | 7O4OP567/Chem_NY_BTC 1210-80% | Chem_NY_BTC 1210-80% | NY-Analytical | 397 | Appearance | Appearance | Appearance | Colorless to slight yellow (Done) | [5 empty] | TM 208 | Main_Plans
7O4OP567 | 7O4OP567/Chem_NY_BTC 1210-80% | Chem_NY_BTC 1210-80% | NY-Analytical | 1015 | Active (BTC 1210-80%) | Active (BTC 1210-80%) | Active (BTC 1210-80%) | 80.9 |  | % | FDA,1 | 80 | 82 | TM 3005 | Main_Plans
7O4OP601 | 7O4OP601/Chem_107A | Chem_107A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | GC Profile | Main_Plans
7O4OP601 | 7O4OP601/Chem_107A | Chem_107A | NY-Analytical | 839 | GC Profile | GC Profile | GC Profile | To match standard. Not required when performing IR test. (Done) | [5 empty] | J&J TM-003030 | Main_Plans
7O4OP601 | 7O4OP601/Chem_107A | Chem_107A | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Colorless to pale yellow clear liquid. (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7O4OP601 | 7O4OP601/Chem_107A | Chem_107A | NY-Analytical | 792 | Chemical Inspection | Odor (J&J TM-003188) | Odor | Meets Odor Specification or of Characteristic Odor. [As per specification](Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7O4OP601 | 7O4OP601/Chem_107A | Chem_107A | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7O4OP601 | 7O4OP601/Chem_107A | Chem_107A | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C | 0.9921 |  |  | FDA,3 | 0.986 | 0.996 | J&J TMWW78012 | Main_Plans
7O4OP601 | 7O4OP601/Chem_107A | Chem_107A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) | [5 empty] | J&J TMWW78013 | Main_Plans
7O4OP601 | 7O4OP601/Chem_107A | Chem_107A | NY-Analytical | 798 | Refractive Index at 20°C | Refractive Index at 20°C | Refractive Index at 20°C | 1.452 |  |  | FDA,4 | 1.452 | 1.462 | J&J TMWW78015 | Main_Plans
7O4OP638 | 7O4OP638/Chem_192A | Chem_192A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7O4OP638 | 7O4OP638/Chem_192A | Chem_192A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, colorless to light yellow liquid / White solid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OP705 | 7O4OP705/Chem_NY_Cetaphil Concentrate | Chem_NY_Cetaphil Concentrate | NY-Analytical | 1051 | Temptale Reading | Temptale Reading | Temptale Reading |  |  | ºC | FDA,0 | 5 | 40 | Temptale Reading | Main_Plans
7O4OP705 | 7O4OP705/Chem_NY_Cetaphil Concentrate | Chem_NY_Cetaphil Concentrate | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | pH as is | Main_Plans
7O4OP705 | 7O4OP705/Chem_NY_Cetaphil Concentrate | Chem_NY_Cetaphil Concentrate | NY-Analytical | 447 | pH as is | pH | pH | 7 |  |  | FDA,1 | 5.5 | 7 | TM 200 | Main_Plans
7O4OP705 | 7O4OP705/Chem_NY_Cetaphil Concentrate | Chem_NY_Cetaphil Concentrate | NY-Analytical | 747 | Viscosity | Viscosity (Spindle 3 @ 60RPM in 300 ml beaker) | Viscosity (Spindle 3 @ 60RPM in 300 ml beaker) | 915 |  | cps | FDA,0 | 400 | 1600 | TM 203 | Main_Plans
7O4OP705 | 7O4OP705/Chem_NY_Cetaphil Concentrate | Chem_NY_Cetaphil Concentrate | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Main_Plans
7O4OP705 | 7O4OP705/Chem_NY_Cetaphil Concentrate | Chem_NY_Cetaphil Concentrate | NY-Analytical | 792 | Chemical Inspection | Description (TM 208) | Description | Viscous, slightly translucent, offwhite liquid: cetyl alcohol odor (Done) | [5 empty] | TM 208 | Main_Plans
7O4OP806 | 7O4OP806/Chem_NY_Carbopol 2020 | Chem_NY_Carbopol 2020 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Loss on Drying | Main_Plans
7O4OP806 | 7O4OP806/Chem_NY_Carbopol 2020 | Chem_NY_Carbopol 2020 | NY-Analytical | 802 | Loss on Drying | Loss on Drying (80C Vacuum 1 hour) | Loss on Drying (80C Vacuum 1 hour) | 0.49 |  | % | FDA,2 |  | 2 | J&J TM7221 | Main_Plans
7O4OP806 | 7O4OP806/Chem_NY_Carbopol 2020 | Chem_NY_Carbopol 2020 | NY-Analytical | 802 | Loss on Drying | Loss on Drying (80C Vacuum 1 hour) Final | Loss on Drying (80C Vacuum 1 hour) Final | 0.49 |  | % | FDA,2 |  | 2 | J&J TM7221 | Main_Plans
7O4OP806 | 7O4OP806/Chem_NY_Carbopol 2020 | Chem_NY_Carbopol 2020 | NY-Analytical | 747 | Viscosity | Viscosity | Viscosity | 48667 |  | cps | FDA,0 | 47000 | 77000 | J&J TM7423 | Main_Plans
7O4OP806 | 7O4OP806/Chem_NY_Carbopol 2020 | Chem_NY_Carbopol 2020 | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Fluffy white powder free of lumps (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7O4OP806 | 7O4OP806/Chem_NY_Carbopol 2020 | Chem_NY_Carbopol 2020 | NY-Analytical | 748 | Visual Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7O4OP806 | 7O4OP806/Chem_NY_Carbopol 2020 | Chem_NY_Carbopol 2020 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) | [5 empty] | J&J TMWW78013 | Main_Plans
7O4OQ028 | 7O4OQ028/Chem_NY_Hexylene Glycol | Chem_NY_Hexylene Glycol | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Certificate of Analysis | Main_Plans
7O4OQ028 | 7O4OQ028/Chem_NY_Hexylene Glycol | Chem_NY_Hexylene Glycol | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C |  |  |  | FDA,4 | 0.92 | 0.923 | TM 202 | Main_Plans
7O4OQ028 | 7O4OQ028/Chem_NY_Hexylene Glycol | Chem_NY_Hexylene Glycol | NY-Analytical | 390 | Identification | IR Identification | IR Identification | To match standard. (Done) | [5 empty] | TM 411 | Main_Plans
7O4OQ028 | 7O4OQ028/Chem_NY_Hexylene Glycol | Chem_NY_Hexylene Glycol | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OQ028 | 7O4OQ028/Chem_NY_Hexylene Glycol | Chem_NY_Hexylene Glycol | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OQ028 | 7O4OQ028/Chem_NY_Hexylene Glycol | Chem_NY_Hexylene Glycol | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OQ028 | 7O4OQ028/Chem_NY_Hexylene Glycol | Chem_NY_Hexylene Glycol | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OQ243 | 7O4OQ243/Chem_165A | Chem_165A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Identification | Main_Plans
7O4OQ243 | 7O4OQ243/Chem_165A | Chem_165A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7O4OQ243 | 7O4OQ243/Chem_165A | Chem_165A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Fine, free-flowing powder (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OQ243 | 7O4OQ243/Chem_165A | Chem_165A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | White (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OQ243 | 7O4OQ243/Chem_165A | Chem_165A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OQ243 | 7O4OQ243/Chem_165A | Chem_165A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Characteristically mild (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OQ342 | 7O4OQ342/Chem_166A | Chem_166A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Specific Gravity at 25°C | Main_Plans
7O4OQ342 | 7O4OQ342/Chem_166A | Chem_166A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.9558 |  |  | FDA,4 | 0.9455 | 0.9655 | TM 202 | Main_Plans
7O4OQ342 | 7O4OQ342/Chem_166A | Chem_166A | NY-Analytical | 451 | Refractive Index at 25°C | Refractive Index | Refractive Index at 25°C | 1.458 |  |  | FDA,4 | 1.4505 | 1.4705 | TM 204 | Main_Plans
7O4OQ342 | 7O4OQ342/Chem_166A | Chem_166A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear, mobile liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OQ342 | 7O4OQ342/Chem_166A | Chem_166A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Very pale yellow (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OQ342 | 7O4OQ342/Chem_166A | Chem_166A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OQ342 | 7O4OQ342/Chem_166A | Chem_166A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Citrus Fresh (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OQ437 | 7O4OQ437/Chem_168A | Chem_168A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Specific Gravity at 25°C | Main_Plans
7O4OQ437 | 7O4OQ437/Chem_168A | Chem_168A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 0.944 |  |  | FDA,4 | 0.938 | 0.9465 | TM 202 | Main_Plans
7O4OQ437 | 7O4OQ437/Chem_168A | Chem_168A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7O4OQ437 | 7O4OQ437/Chem_168A | Chem_168A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Mobile Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OQ437 | 7O4OQ437/Chem_168A | Chem_168A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OQ437 | 7O4OQ437/Chem_168A | Chem_168A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OQ437 | 7O4OQ437/Chem_168A | Chem_168A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Amine (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OQ890 | 7O4OQ890/Chem_170A | Chem_170A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Identification | Main_Plans
7O4OQ890 | 7O4OQ890/Chem_170A | Chem_170A | NY-Analytical | 390 | Identification | IR Identification | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7O4OQ890 | 7O4OQ890/Chem_170A | Chem_170A | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear Liquid (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OQ890 | 7O4OQ890/Chem_170A | Chem_170A | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OQ890 | 7O4OQ890/Chem_170A | Chem_170A | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OQ890 | 7O4OQ890/Chem_170A | Chem_170A | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Ether-like (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 792 | Chemical Inspection | Description | Description | An almost colorless or paleyellowish liquid (Done) | [5 empty] | TM 208 | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Any individual unspecified impuritv | Any individual unspecified impuritv | 0.077 |  | % | FDA,1 |  | 0.1 | USP | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Ch!orhexidine oxazinone analog | Ch!orhexidine oxazinone analog | 0.042 |  | % | FDA,1 |  | 0.2 | USP | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Chlorhexidine | Chlorhexidine | 93.338 |  | % | FDA,1 |  |  | USP | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Chlorhexidine urea | Chlorhexidine urea | 0.085 |  | % | FDA,1 |  | 0.2 | USP | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Chlorhexidine amine | Chlorhexidine amine | 0.047 |  | % | FDA,1 |  | 0.3 | USP | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Chlorhexidine dimer | Chlorhexidine dimer | 0.09 |  | % | FDA,1 |  | 0.5 | USP | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Chlorhexidine glucityl triazine | Chlorhexidine glucityl triazine | 0.083 |  | % | FDA,1 |  | 0.4 | USP | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Chlorhexidine guanidine | Chlorhexidine guanidine | 0.083 |  | % | FDA,1 |  | 0.3 | USP | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Chlorhexidine nitrile | Chlorhexidine nitrile | 0.114 |  | % | FDA,1 |  | 0.4 | USP | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 390 | Identification | Identification A (USP <197k>) | Identification A | Conforms to USP (Done) | [5 empty] | USP | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 390 | Identification | Identification B (USP <201>) | Identification B | Conforms to USP (Done) | [5 empty] | USP | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | o-Chlorhexidine and specified unidentified impurit | o-Chlorhexidine and specified unidentified impurit | 0.164 |  | % | FDA,1 |  | 0.4 | USP | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Oxochlorhexidine | Oxochlorhexidine | 0.288 |  | % | FDA,1 |  | 0.4 | USP | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 812 | Organic Impurites | p-Chloroaniline | p-Chloroaniline | 0 |  | ug/mL | FDA,0 |  | 100 | USP | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 812 | Organic Impurites | p-Chloroaniline Final | p-Chloroaniline Final | 0 |  | ug/mL | FDA,0 |  | 100 | USP | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | p-Chlorophenvl urea | p-Chlorophenvl urea | 0 |  | % | FDA,1 |  | 0.2 | USP | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Any individual unspecified impuritv | RRT of Any individual unspecified impuritv | 1.13 |  |  | FDA,2 |  |  | USP | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Ch!orhexidine oxazinone analog | RRT of Ch!orhexidine oxazinone analog | 0.16 |  |  | FDA,2 |  |  | USP | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Chlorhexidine | RRT of Chlorhexidine | 1.1 |  |  | FDA,2 |  |  | USP | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Chlorhexidine urea | RRT of Chlorhexidine urea | 0.32 |  |  | FDA,2 |  |  | USP | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Chlorhexidine amine | RRT of Chlorhexidine amine | 0.21 |  |  | FDA,2 |  |  | USP | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Chlorhexidine dimer | RRT of Chlorhexidine dimer | 0.87 |  |  | FDA,2 |  |  | USP | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Chlorhexidine glucityl triazine | RRT of Chlorhexidine glucityl triazine | 0.95 |  |  | FDA,2 |  |  | USP | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Chlorhexidine guanidine | RRT of Chlorhexidine guanidine | 0.31 |  |  | FDA,2 |  |  | USP | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Chlorhexidine nitrile | RRT of Chlorhexidine nitrile | 0.52 |  |  | FDA,2 |  |  | USP | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of o-Chlorhexidine and specified unidentified | RRT of o-Chlorhexidine and specified unidentified | 0.9 |  |  | FDA,2 |  |  | USP | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Oxochlorhexidine | RRT of Oxochlorhexidine | 1.43 |  |  | FDA,2 |  |  | USP | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of p-Chlorophenvl urea | RRT of p-Chlorophenvl urea | 0 |  |  | FDA,2 |  |  | USP | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | RRT of Specified unidentified impuritv 1 | RRT of Specified unidentified impuritv 1 | 0.17 |  |  | FDA,2 |  |  | USP | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Specified unidentified impuritv 1 | Specified unidentified impuritv 1 | 0.011 |  | % | FDA,1 |  | 0.2 | USP | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 975 | Organic Impurities Part II | Total Impurities | Total Impurities | 0.983 |  | % | FDA,1 |  | 3 | USP | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 398 | pH | pH (Diluted 1 in 20) | pH (Diluted 1 in 20) | 5.55 |  |  | FDA,1 | 5.5 | 7 | USP <791> | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 796 | Assay | Assay | Assay | 20.1 |  | % | FDA,1 | 19 | 21 | USP-HPLC | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 796 | Assay | Assay Final | Assay Final | 20.1 |  | % | FDA,1 | 19 | 21 | USP-HPLC | Main_Plans
7O4OS407 | 7O4OS407/Chem_173A | Chem_173A | NY-Analytical | 402 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 25°C | 1.0648 |  |  | FDA,2 | 1.06 | 1.07 | USP/NF <841> | Main_Plans
7O4OU181 | 7O4OU181/Chem_NY_Isopropanol 70% | Chem_NY_Isopropanol 70% | IN-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | COA | Main_Plans
7O4OU181 | 7O4OU181/Chem_NY_Isopropanol 70% | Chem_NY_Isopropanol 70% | IN-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear (Done) | [5 empty] | TM 208 | Main_Plans
7O4OU181 | 7O4OU181/Chem_NY_Isopropanol 70% | Chem_NY_Isopropanol 70% | IN-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Colorless (Done) | [5 empty] | TM 208 | Main_Plans
7O4OU181 | 7O4OU181/Chem_NY_Isopropanol 70% | Chem_NY_Isopropanol 70% | IN-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Main_Plans
7O4OU181 | 7O4OU181/Chem_NY_Isopropanol 70% | Chem_NY_Isopropanol 70% | IN-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) | [5 empty] | TM 208 | Main_Plans
7O4OU181 | 7O4OU181/Chem_NY_Isopropanol 70% | Chem_NY_Isopropanol 70% | IN-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Matches that of the Reference Standard (Done) | [5 empty] | TM 411 | Main_Plans
7O4OU225 | 7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | pH | Main_Plans
7O4OU225 | 7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 398 | pH | pH (10% Solution) | pH (10% Solution) | 11.8 |  |  | FDA,1 | 10.5 | 13 | TM 200 | Main_Plans
7O4OU225 | 7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 807 | % NaOCl | % NaOCl by weight (calc) | % NaOCl by weight | 9.9316 |  | % | FDA,2 | 11.5 | 14 | TM 366 | Main_Plans
7O4OU225 | 7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 807 | % NaOCl | % NaOCl by weight Final (calc) | % NaOCl by weight Final | 10.0533 |  | % | FDA,2 | 11.5 | 14 | TM 366 | Main_Plans
7O4OU225 | 7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 807 | % NaOCl | Sample Weight (g) | Sample Weight | 1.0023 |  | g | FDA,4 |  |  | TM 366 | Main_Plans
7O4OU225 | 7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 807 | % NaOCl | Titrant End Volume (mL) | Titrant End Volume | 27.5 |  | mL | FDA,1 |  |  | TM 366 | Main_Plans
7O4OU225 | 7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 807 | % NaOCl | Titrant Start Volume (mL) | Titrant Start Volume | 0 |  | mL | FDA,1 |  |  | TM 366 | Main_Plans
7O4OU225 | 7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 807 | % NaOCl | Total Volume of Titrant (mL) | Total Volume of Titrant | 27.5 |  | mL | FDA,1 |  |  | TM 366 | Main_Plans
7O4OU225 | 7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Clear liquid with a chlorine odor (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OU225 | 7O4OU225/Chem_NY_Sodium Hypochlorite 12.5% | Chem_NY_Sodium Hypochlorite 12.5% | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | straw-yellow (Done) | [5 empty] | TM208/220/230/USP/Kao LM | Main_Plans
7O4OU555 | 7O4OU555/Chem_NY_Lidocaine Base USP | Chem_NY_Lidocaine Base USP | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7O4OU555 | 7O4OU555/Chem_NY_Lidocaine Base USP | Chem_NY_Lidocaine Base USP | NY-Analytical | 792 | Chemical Inspection | Appearance (TM 208) | Appearance | Match standard (Done) | [5 empty] | TM 208 | Main_Plans
7O4OU555 | 7O4OU555/Chem_NY_Lidocaine Base USP | Chem_NY_Lidocaine Base USP | NY-Analytical | 792 | Chemical Inspection | Color (TM 208) | Color | Match standard (Done) | [5 empty] | TM 208 | Main_Plans
7O4OU555 | 7O4OU555/Chem_NY_Lidocaine Base USP | Chem_NY_Lidocaine Base USP | NY-Analytical | 792 | Chemical Inspection | Contamination (TM 208) | Contamination | None (Done) | [5 empty] | TM 208 | Main_Plans
7O4OU555 | 7O4OU555/Chem_NY_Lidocaine Base USP | Chem_NY_Lidocaine Base USP | NY-Analytical | 792 | Chemical Inspection | Odor (TM 208) | Odor | Match standard (Done) | [5 empty] | TM 208 | Main_Plans
7O4OU555 | 7O4OU555/Chem_NY_Lidocaine Base USP | Chem_NY_Lidocaine Base USP | NY-Analytical | 390 | Identification | IR Identification (TM 411) | IR Identification | Close Match to Standard Sample or Spectrum (Done) | [5 empty] | TM 411 | Main_Plans
7O4OU555 | 7O4OU555/Chem_NY_Lidocaine Base USP | Chem_NY_Lidocaine Base USP | NY-Analytical | 796 | Assay | Assay | Assay | 100.4 |  | % | FDA,1 | 97.5 | 102.5 | USP | Main_Plans
7O4OU555 | 7O4OU555/Chem_NY_Lidocaine Base USP | Chem_NY_Lidocaine Base USP | NY-Analytical | 832 | Chloride | Chloride | Chloride |  |  | % | FDA,4 |  | 0.0035 | USP | Main_Plans
7O4OU555 | 7O4OU555/Chem_NY_Lidocaine Base USP | Chem_NY_Lidocaine Base USP | NY-Analytical | 833 | Heavy Metals | Heavy Metals | Heavy Metals |  |  | ug/g | FDA,3 |  | 0.002 | USP | Main_Plans
7O4OU555 | 7O4OU555/Chem_NY_Lidocaine Base USP | Chem_NY_Lidocaine Base USP | NY-Analytical | 809 | Melting Range | Melting Range - End | Melting Range - End |  |  | C | FDA,1 | 66 | 69 | USP | Main_Plans
7O4OU555 | 7O4OU555/Chem_NY_Lidocaine Base USP | Chem_NY_Lidocaine Base USP | NY-Analytical | 809 | Melting Range | Melting Range - Start | Melting Range - Start |  |  | C | FDA,1 | 66 | 69 | USP | Main_Plans
7O4OU555 | 7O4OU555/Chem_NY_Lidocaine Base USP | Chem_NY_Lidocaine Base USP | NY-Analytical | 834 | Residue on Ignition | Residue on Ignition | Residue on Ignition |  |  | % | FDA,1 |  | 0.1 | USP | Main_Plans
7O4OU555 | 7O4OU555/Chem_NY_Lidocaine Base USP | Chem_NY_Lidocaine Base USP | NY-Analytical | 835 | Sulfate Content | Sulfate | Sulfate |  |  | % | FDA,2 |  | 0.1 | USP | Main_Plans
7O4OX231 | 7O4OX231/Chem_NY_FD&C Red 40 | Chem_NY_FD&C Red 40 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Appearance | Main_Plans
7O4OX231 | 7O4OX231/Chem_NY_FD&C Red 40 | Chem_NY_FD&C Red 40 | NY-Analytical | 397 | Appearance | Physical Properties | Physical Properties | Brown Powder with Purple/red tint Powder (Done) | [5 empty] | TM 208 | Main_Plans
7O4OX231 | 7O4OX231/Chem_NY_FD&C Red 40 | Chem_NY_FD&C Red 40 | NY-Analytical | 1061 | Total Color | Spectrophotometric | Spectrophotometric | 89.4 |  | % | FDA,0 | 85 |  | TM 3014 | Main_Plans
7R4OP485 | 7R4OP485/Chem_NY_Phenoxyethanol_7R4OP485 | Chem_NY_Phenoxyethanol_7R4OP485 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Chemical Inspection | Main_Plans
7R4OP485 | 7R4OP485/Chem_NY_Phenoxyethanol_7R4OP485 | Chem_NY_Phenoxyethanol_7R4OP485 | NY-Analytical | 792 | Chemical Inspection | Appearance | Appearance | Clear, thin colorless liquid (Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7R4OP485 | 7R4OP485/Chem_NY_Phenoxyethanol_7R4OP485 | Chem_NY_Phenoxyethanol_7R4OP485 | NY-Analytical | 792 | Chemical Inspection | Packaging Damage or Contamination | Packaging Damage or Contamination | Packaging is suitable to protect from contamination [As per specification](Done) | [5 empty] | J&J TMWW78010/J&J TM7252 | Main_Plans
7R4OP485 | 7R4OP485/Chem_NY_Phenoxyethanol_7R4OP485 | Chem_NY_Phenoxyethanol_7R4OP485 | NY-Analytical | 390 | Identification | IR Identification | IR Identification | IR absorption bands are similar to the standard [As per specification](Done) | [5 empty] | J&J TMWW78013 | Main_Plans
7R4OP485 | 7R4OP485/Chem_NY_Phenoxyethanol_7R4OP485 | Chem_NY_Phenoxyethanol_7R4OP485 | NY-Analytical | 892 | Limit of Phenol | Phenol | Phenol | 18 |  | ppm | FDA,0 |  | 25 | USP/NF | Main_Plans
7R4OP485 | 7R4OP485/Chem_NY_Phenoxyethanol_7R4OP485 | Chem_NY_Phenoxyethanol_7R4OP485 | NY-Analytical | 849 | Purity | Purity | Purity | 99.4 |  | % | FDA,1 | 98 | 102 | USP/NF | Main_Plans
7R4OP485 | 7R4OP485/Chem_NY_Phenoxyethanol_7R4OP485 | Chem_NY_Phenoxyethanol_7R4OP485 | NY-Analytical | 432 | Specific Gravity at 25°C | Specific Gravity at 25°C | Specific Gravity at 20°C | 1.1077 |  |  | FDA,3 | 1.105 | 1.11 | USP/NF | Main_Plans
8AA70900 | 8AA70900/Label_NY_NP9990_2.5-4.625 | Label_NY_NP9990_2.5-4.625 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Label Inspection | Main_Plans
8AA70900 | 8AA70900/Label_NY_NP9990_2.5-4.625 | Label_NY_NP9990_2.5-4.625 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AA70900 | 8AA70900/Label_NY_NP9990_2.5-4.625 | Label_NY_NP9990_2.5-4.625 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AA70900 | 8AA70900/Label_NY_NP9990_2.5-4.625 | Label_NY_NP9990_2.5-4.625 | NY-Analytical | 770 | Label Inspection | Length | Length | 2.5 |  | in | FDA,4 | 2.4375 | 2.5625 | TM 143 / Visual | Main_Plans
8AA70900 | 8AA70900/Label_NY_NP9990_2.5-4.625 | Label_NY_NP9990_2.5-4.625 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AA70900 | 8AA70900/Label_NY_NP9990_2.5-4.625 | Label_NY_NP9990_2.5-4.625 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AA70900 | 8AA70900/Label_NY_NP9990_2.5-4.625 | Label_NY_NP9990_2.5-4.625 | NY-Analytical | 770 | Label Inspection | Width | Width | 4.625 |  | in | FDA,4 | 4.5625 | 4.6875 | TM 143 / Visual | Main_Plans
8AP02700RF | 8AP02700RF/Label_NY_NP9815_73-85 | Label_NY_NP9815_73-85 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Label Inspection | Main_Plans
8AP02700RF | 8AP02700RF/Label_NY_NP9815_73-85 | Label_NY_NP9815_73-85 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AP02700RF | 8AP02700RF/Label_NY_NP9815_73-85 | Label_NY_NP9815_73-85 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AP02700RF | 8AP02700RF/Label_NY_NP9815_73-85 | Label_NY_NP9815_73-85 | NY-Analytical | 770 | Label Inspection | Length | Length | 73 |  | mm | FDA,4 | 71.4125 | 74.5875 | TM 143 / Visual | Main_Plans
8AP02700RF | 8AP02700RF/Label_NY_NP9815_73-85 | Label_NY_NP9815_73-85 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AP02700RF | 8AP02700RF/Label_NY_NP9815_73-85 | Label_NY_NP9815_73-85 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AP02700RF | 8AP02700RF/Label_NY_NP9815_73-85 | Label_NY_NP9815_73-85 | NY-Analytical | 770 | Label Inspection | Width | Width | 85 |  | mm | FDA,4 | 83.4125 | 86.5875 | TM 143 / Visual | Main_Plans
8AP034RSL | 8AP034RSL/Label_NY_NP9800_3.375-2.3437 | Label_NY_NP9800_3.375-2.3437 | IN-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AP034RSL | 8AP034RSL/Label_NY_NP9800_3.375-2.3437 | Label_NY_NP9800_3.375-2.3437 | IN-Analytical | 770 | Label Inspection | Certificate of Conformance | Certificate of Conformance | CofC meets specifications (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AP034RSL | 8AP034RSL/Label_NY_NP9800_3.375-2.3437 | Label_NY_NP9800_3.375-2.3437 | IN-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AP034RSL | 8AP034RSL/Label_NY_NP9800_3.375-2.3437 | Label_NY_NP9800_3.375-2.3437 | IN-Analytical | 770 | Label Inspection | Length | Length | 3.375 |  | in | FDA,4 | 3.3125 | 3.4375 | TM 143 / Visual | Main_Plans
8AP034RSL | 8AP034RSL/Label_NY_NP9800_3.375-2.3437 | Label_NY_NP9800_3.375-2.3437 | IN-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AP034RSL | 8AP034RSL/Label_NY_NP9800_3.375-2.3437 | Label_NY_NP9800_3.375-2.3437 | IN-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AP034RSL | 8AP034RSL/Label_NY_NP9800_3.375-2.3437 | Label_NY_NP9800_3.375-2.3437 | IN-Analytical | 770 | Label Inspection | Width | Width | 2.3437 |  | in | FDA,4 | 2.2812 | 2.4062 | TM 143 / Visual | Main_Plans
8AP13800 | 8AP13800/Label_NY_NP7200_6.25-4 | Label_NY_NP7200_6.25-4 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Label Inspection | Main_Plans
8AP13800 | 8AP13800/Label_NY_NP7200_6.25-4 | Label_NY_NP7200_6.25-4 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AP13800 | 8AP13800/Label_NY_NP7200_6.25-4 | Label_NY_NP7200_6.25-4 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AP13800 | 8AP13800/Label_NY_NP7200_6.25-4 | Label_NY_NP7200_6.25-4 | NY-Analytical | 770 | Label Inspection | Length | Length | 6.1875 |  | in | FDA,4 | 6.1875 | 6.3125 | TM 143 / Visual | Main_Plans
8AP13800 | 8AP13800/Label_NY_NP7200_6.25-4 | Label_NY_NP7200_6.25-4 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AP13800 | 8AP13800/Label_NY_NP7200_6.25-4 | Label_NY_NP7200_6.25-4 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AP13800 | 8AP13800/Label_NY_NP7200_6.25-4 | Label_NY_NP7200_6.25-4 | NY-Analytical | 770 | Label Inspection | Width | Width | 3.9375 |  | in | FDA,4 | 3.9375 | 4.0625 | TM 143 / Visual | Main_Plans
8AP21100A | 8AP21100A/Label_NY_NP9890_85-73 | Label_NY_NP9890_85-73 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Label Inspection | Main_Plans
8AP21100A | 8AP21100A/Label_NY_NP9890_85-73 | Label_NY_NP9890_85-73 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AP21100A | 8AP21100A/Label_NY_NP9890_85-73 | Label_NY_NP9890_85-73 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AP21100A | 8AP21100A/Label_NY_NP9890_85-73 | Label_NY_NP9890_85-73 | NY-Analytical | 770 | Label Inspection | Length | Length | 85 |  | mm | FDA,4 | 83.4125 | 86.5875 | TM 143 / Visual | Main_Plans
8AP21100A | 8AP21100A/Label_NY_NP9890_85-73 | Label_NY_NP9890_85-73 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AP21100A | 8AP21100A/Label_NY_NP9890_85-73 | Label_NY_NP9890_85-73 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AP21100A | 8AP21100A/Label_NY_NP9890_85-73 | Label_NY_NP9890_85-73 | NY-Analytical | 770 | Label Inspection | Width | Width | 73 |  | mm | FDA,4 | 71.4125 | 74.5875 | TM 143 / Visual | Main_Plans
8AP22800 | 8AP22800/Label_NY_NP7200_5.8125-6.5 | Label_NY_NP7200_5.8125-6.5 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | [6 empty] | Label Inspection | Main_Plans
8AP22800 | 8AP22800/Label_NY_NP7200_5.8125-6.5 | Label_NY_NP7200_5.8125-6.5 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AP22800 | 8AP22800/Label_NY_NP7200_5.8125-6.5 | Label_NY_NP7200_5.8125-6.5 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AP22800 | 8AP22800/Label_NY_NP7200_5.8125-6.5 | Label_NY_NP7200_5.8125-6.5 | NY-Analytical | 770 | Label Inspection | Length | Length | 5.75 |  | in | FDA,4 | 5.75 | 5.875 | TM 143 / Visual | Main_Plans
8AP22800 | 8AP22800/Label_NY_NP7200_5.8125-6.5 | Label_NY_NP7200_5.8125-6.5 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AP22800 | 8AP22800/Label_NY_NP7200_5.8125-6.5 | Label_NY_NP7200_5.8125-6.5 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AP22800 | 8AP22800/Label_NY_NP7200_5.8125-6.5 | Label_NY_NP7200_5.8125-6.5 | NY-Analytical | 770 | Label Inspection | Width | Width | 6.5 |  | in | FDA,4 | 6.4375 | 6.5625 | TM 143 / Visual | Main_Plans
8AP32800 | 8AP32800/Label_NY_NP98372_73-85 | Label_NY_NP98372_73-85 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Label Inspection | Main_Plans
8AP32800 | 8AP32800/Label_NY_NP98372_73-85 | Label_NY_NP98372_73-85 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AP32800 | 8AP32800/Label_NY_NP98372_73-85 | Label_NY_NP98372_73-85 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AP32800 | 8AP32800/Label_NY_NP98372_73-85 | Label_NY_NP98372_73-85 | NY-Analytical | 770 | Label Inspection | Length | Length | 73 |  | mm | FDA,4 | 71.4125 | 74.5875 | TM 143 / Visual | Main_Plans
8AP32800 | 8AP32800/Label_NY_NP98372_73-85 | Label_NY_NP98372_73-85 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AP32800 | 8AP32800/Label_NY_NP98372_73-85 | Label_NY_NP98372_73-85 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AP32800 | 8AP32800/Label_NY_NP98372_73-85 | Label_NY_NP98372_73-85 | NY-Analytical | 770 | Label Inspection | Width | Width | 85 |  | mm | FDA,4 | 83.4125 | 86.5875 | TM 143 / Visual | Main_Plans
8AP60100RF | 8AP60100RF/Label_NY_NP9890_73-85 | Label_NY_NP9890_73-85 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Label Inspection | Main_Plans
8AP60100RF | 8AP60100RF/Label_NY_NP9890_73-85 | Label_NY_NP9890_73-85 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AP60100RF | 8AP60100RF/Label_NY_NP9890_73-85 | Label_NY_NP9890_73-85 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AP60100RF | 8AP60100RF/Label_NY_NP9890_73-85 | Label_NY_NP9890_73-85 | NY-Analytical | 770 | Label Inspection | Length | Length | 72 |  | mm | FDA,4 | 71.4125 | 74.5875 | TM 143 / Visual | Main_Plans
8AP60100RF | 8AP60100RF/Label_NY_NP9890_73-85 | Label_NY_NP9890_73-85 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AP60100RF | 8AP60100RF/Label_NY_NP9890_73-85 | Label_NY_NP9890_73-85 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AP60100RF | 8AP60100RF/Label_NY_NP9890_73-85 | Label_NY_NP9890_73-85 | NY-Analytical | 770 | Label Inspection | Width | Width | 85 |  | mm | FDA,4 | 83.4125 | 86.5875 | TM 143 / Visual | Main_Plans
8AU120CK | 8AU120CK/Label_NY_NP7030_6-4 | Label_NY_NP7030_6-4 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Label Inspection | Main_Plans
8AU120CK | 8AU120CK/Label_NY_NP7030_6-4 | Label_NY_NP7030_6-4 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AU120CK | 8AU120CK/Label_NY_NP7030_6-4 | Label_NY_NP7030_6-4 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AU120CK | 8AU120CK/Label_NY_NP7030_6-4 | Label_NY_NP7030_6-4 | NY-Analytical | 770 | Label Inspection | Length | Length | 6 |  | in | FDA,4 | 5.9375 | 6.0625 | TM 143 / Visual | Main_Plans
8AU120CK | 8AU120CK/Label_NY_NP7030_6-4 | Label_NY_NP7030_6-4 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AU120CK | 8AU120CK/Label_NY_NP7030_6-4 | Label_NY_NP7030_6-4 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8AU120CK | 8AU120CK/Label_NY_NP7030_6-4 | Label_NY_NP7030_6-4 | NY-Analytical | 770 | Label Inspection | Width | Width | 4 |  | in | FDA,4 | 3.9375 | 4.0625 | TM 143 / Visual | Main_Plans
8BB196CK | 8BB196CK/Label_NY_NP7030_6-5 | Label_NY_NP7030_6-5 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Label Inspection | Main_Plans
8BB196CK | 8BB196CK/Label_NY_NP7030_6-5 | Label_NY_NP7030_6-5 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8BB196CK | 8BB196CK/Label_NY_NP7030_6-5 | Label_NY_NP7030_6-5 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8BB196CK | 8BB196CK/Label_NY_NP7030_6-5 | Label_NY_NP7030_6-5 | NY-Analytical | 770 | Label Inspection | Length | Length | 5.99 |  | in | FDA,4 | 5.9375 | 6.0625 | TM 143 / Visual | Main_Plans
8BB196CK | 8BB196CK/Label_NY_NP7030_6-5 | Label_NY_NP7030_6-5 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8BB196CK | 8BB196CK/Label_NY_NP7030_6-5 | Label_NY_NP7030_6-5 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8BB196CK | 8BB196CK/Label_NY_NP7030_6-5 | Label_NY_NP7030_6-5 | NY-Analytical | 770 | Label Inspection | Width | Width | 5.01 |  | in | FDA,4 | 4.9375 | 5.0625 | TM 143 / Visual | Main_Plans
8BB512CL | 8BB512CL/Label_NY_NP7030_6-4 | Label_NY_NP7030_6-4 | NY-Analytical | 770 | Label Inspection | AQL Sample Size | AQL Sample Size | 50 | [5 empty] | TM 143 / Visual | Main_Plans
8BB512CL | 8BB512CL/Label_NY_NP7030_6-4 | Label_NY_NP7030_6-4 | NY-Analytical | 770 | Label Inspection | Artwork_Copy (Color & Text) | Artwork / Copy | Sample(s) all meet specifications (Done) |  | in |  |  |  | TM 143 / Visual | Main_Plans
8BP12801 | 8BP12801/Label_NY_NP7200_4.5-4 | Label_NY_NP7200_4.5-4 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Label Inspection | Main_Plans
8BP12801 | 8BP12801/Label_NY_NP7200_4.5-4 | Label_NY_NP7200_4.5-4 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8BP12801 | 8BP12801/Label_NY_NP7200_4.5-4 | Label_NY_NP7200_4.5-4 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8BP12801 | 8BP12801/Label_NY_NP7200_4.5-4 | Label_NY_NP7200_4.5-4 | NY-Analytical | 770 | Label Inspection | Length | Length | 4.48 |  | in | FDA,4 | 4.4375 | 4.5625 | TM 143 / Visual | Main_Plans
8BP12801 | 8BP12801/Label_NY_NP7200_4.5-4 | Label_NY_NP7200_4.5-4 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8BP12801 | 8BP12801/Label_NY_NP7200_4.5-4 | Label_NY_NP7200_4.5-4 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8BP12801 | 8BP12801/Label_NY_NP7200_4.5-4 | Label_NY_NP7200_4.5-4 | NY-Analytical | 770 | Label Inspection | Width | Width | 3.98 |  | in | FDA,4 | 3.9375 | 4.0625 | TM 143 / Visual | Main_Plans
8BP24500 | 8BP24500/Label_NY_NP7200_4-4.5 | Label_NY_NP7200_4-4.5 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Label Inspection | Main_Plans
8BP24500 | 8BP24500/Label_NY_NP7200_4-4.5 | Label_NY_NP7200_4-4.5 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8BP24500 | 8BP24500/Label_NY_NP7200_4-4.5 | Label_NY_NP7200_4-4.5 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8BP24500 | 8BP24500/Label_NY_NP7200_4-4.5 | Label_NY_NP7200_4-4.5 | NY-Analytical | 770 | Label Inspection | Length | Length | 4 |  | in | FDA,4 | 3.9375 | 4.0625 | TM 143 / Visual | Main_Plans
8BP24500 | 8BP24500/Label_NY_NP7200_4-4.5 | Label_NY_NP7200_4-4.5 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8BP24500 | 8BP24500/Label_NY_NP7200_4-4.5 | Label_NY_NP7200_4-4.5 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8BP24500 | 8BP24500/Label_NY_NP7200_4-4.5 | Label_NY_NP7200_4-4.5 | NY-Analytical | 770 | Label Inspection | Width | Width | 4.5 |  | in | FDA,4 | 4.4375 | 4.5625 | TM 143 / Visual | Main_Plans
8BQ76500 | 8BQ76500/Label_NY_NP7200_6-3.75 | Label_NY_NP7200_6-3.75 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Label Inspection | Main_Plans
8BQ76500 | 8BQ76500/Label_NY_NP7200_6-3.75 | Label_NY_NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8BQ76500 | 8BQ76500/Label_NY_NP7200_6-3.75 | Label_NY_NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8BQ76500 | 8BQ76500/Label_NY_NP7200_6-3.75 | Label_NY_NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Length | Length | 6 |  | in | FDA,4 | 5.9375 | 6.0625 | TM 143 / Visual | Main_Plans
8BQ76500 | 8BQ76500/Label_NY_NP7200_6-3.75 | Label_NY_NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8BQ76500 | 8BQ76500/Label_NY_NP7200_6-3.75 | Label_NY_NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8BQ76500 | 8BQ76500/Label_NY_NP7200_6-3.75 | Label_NY_NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Width | Width | 3.6875 |  | in | FDA,4 | 3.6875 | 3.8125 | TM 143 / Visual | Main_Plans
8BQ76501 | 8BQ76501/Label_NY_NP7200_3.75-6 | Label_NY_NP7200_3.75-6 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Label Inspection | Main_Plans
8BQ76501 | 8BQ76501/Label_NY_NP7200_3.75-6 | Label_NY_NP7200_3.75-6 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8BQ76501 | 8BQ76501/Label_NY_NP7200_3.75-6 | Label_NY_NP7200_3.75-6 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8BQ76501 | 8BQ76501/Label_NY_NP7200_3.75-6 | Label_NY_NP7200_3.75-6 | NY-Analytical | 770 | Label Inspection | Length | Length | 3.6875 |  | in | FDA,4 | 3.6875 | 3.8125 | TM 143 / Visual | Main_Plans
8BQ76501 | 8BQ76501/Label_NY_NP7200_3.75-6 | Label_NY_NP7200_3.75-6 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8BQ76501 | 8BQ76501/Label_NY_NP7200_3.75-6 | Label_NY_NP7200_3.75-6 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8BQ76501 | 8BQ76501/Label_NY_NP7200_3.75-6 | Label_NY_NP7200_3.75-6 | NY-Analytical | 770 | Label Inspection | Width | Width | 6 |  | in | FDA,4 | 5.9375 | 6.0625 | TM 143 / Visual | Main_Plans
8CH211CK | 8CH211CK/Label_A12 | Label_A12 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Label Inspection | Main_Plans
8CH211CK | 8CH211CK/Label_A12 | Label_A12 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8CH211CK | 8CH211CK/Label_A12 | Label_A12 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8CH211CK | 8CH211CK/Label_A12 | Label_A12 | NY-Analytical | 770 | Label Inspection | Length | Length | 9 |  | in | FDA,4 | 8.9375 | 9.0625 | TM 143 / Visual | Main_Plans
8CH211CK | 8CH211CK/Label_A12 | Label_A12 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8CH211CK | 8CH211CK/Label_A12 | Label_A12 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8CH211CK | 8CH211CK/Label_A12 | Label_A12 | NY-Analytical | 770 | Label Inspection | Width | Width | 2.625 |  | in | FDA,4 | 2.5625 | 2.6875 | TM 143 / Visual | Main_Plans
8CP13400 | 8CP13400/Label_NY_NP7100_4.5-4 | Label_NY_NP7100_4.5-4 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Label Inspection | Main_Plans
8CP13400 | 8CP13400/Label_NY_NP7100_4.5-4 | Label_NY_NP7100_4.5-4 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8CP13400 | 8CP13400/Label_NY_NP7100_4.5-4 | Label_NY_NP7100_4.5-4 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8CP13400 | 8CP13400/Label_NY_NP7100_4.5-4 | Label_NY_NP7100_4.5-4 | NY-Analytical | 770 | Label Inspection | Length | Length | 4.5 |  | in | FDA,4 | 4.4375 | 4.5625 | TM 143 / Visual | Main_Plans
8CP13400 | 8CP13400/Label_NY_NP7100_4.5-4 | Label_NY_NP7100_4.5-4 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8CP13400 | 8CP13400/Label_NY_NP7100_4.5-4 | Label_NY_NP7100_4.5-4 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8CP13400 | 8CP13400/Label_NY_NP7100_4.5-4 | Label_NY_NP7100_4.5-4 | NY-Analytical | 770 | Label Inspection | Width | Width | 4 |  | in | FDA,4 | 3.9375 | 4.0625 | TM 143 / Visual | Main_Plans
8CP15900 | 8CP15900/Label_NY_NP7100/NP7200_6-3.75 | Label_NY_NP7100/NP7200_6-3.75 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Label Inspection | Main_Plans
8CP15900 | 8CP15900/Label_NY_NP7100/NP7200_6-3.75 | Label_NY_NP7100/NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8CP15900 | 8CP15900/Label_NY_NP7100/NP7200_6-3.75 | Label_NY_NP7100/NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8CP15900 | 8CP15900/Label_NY_NP7100/NP7200_6-3.75 | Label_NY_NP7100/NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Length | Length | 5.99 |  | in | FDA,4 | 5.9375 | 6.0625 | TM 143 / Visual | Main_Plans
8CP15900 | 8CP15900/Label_NY_NP7100/NP7200_6-3.75 | Label_NY_NP7100/NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8CP15900 | 8CP15900/Label_NY_NP7100/NP7200_6-3.75 | Label_NY_NP7100/NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8CP15900 | 8CP15900/Label_NY_NP7100/NP7200_6-3.75 | Label_NY_NP7100/NP7200_6-3.75 | NY-Analytical | 770 | Label Inspection | Width | Width | 3.705 |  | in | FDA,4 | 3.6875 | 3.8125 | TM 143 / Visual | Main_Plans
8DP92000 | 8DP92000/Label_NY_NP7200_6.5-5.8125 | Label_NY_NP7200_6.5-5.8125 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | [6 empty] | Label Inspection | Main_Plans
8DP92000 | 8DP92000/Label_NY_NP7200_6.5-5.8125 | Label_NY_NP7200_6.5-5.8125 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8DP92000 | 8DP92000/Label_NY_NP7200_6.5-5.8125 | Label_NY_NP7200_6.5-5.8125 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8DP92000 | 8DP92000/Label_NY_NP7200_6.5-5.8125 | Label_NY_NP7200_6.5-5.8125 | NY-Analytical | 770 | Label Inspection | Length | Length | 6.5 |  | in | FDA,4 | 6.4375 | 6.5625 | TM 143 / Visual | Main_Plans
8DP92000 | 8DP92000/Label_NY_NP7200_6.5-5.8125 | Label_NY_NP7200_6.5-5.8125 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8DP92000 | 8DP92000/Label_NY_NP7200_6.5-5.8125 | Label_NY_NP7200_6.5-5.8125 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8DP92000 | 8DP92000/Label_NY_NP7200_6.5-5.8125 | Label_NY_NP7200_6.5-5.8125 | NY-Analytical | 770 | Label Inspection | Width | Width | 5.8125 |  | in | FDA,4 | 5.75 | 5.875 | TM 143 / Visual | Main_Plans
8DQ85400 | 8DQ85400/Label_A41 | Label_A41 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Label Inspection | Main_Plans
8DQ85400 | 8DQ85400/Label_A41 | Label_A41 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8DQ85400 | 8DQ85400/Label_A41 | Label_A41 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8DQ85400 | 8DQ85400/Label_A41 | Label_A41 | NY-Analytical | 770 | Label Inspection | Length | Length | 5.5625 |  | in | FDA,4 | 5.5 | 5.625 | TM 143 / Visual | Main_Plans
8DQ85400 | 8DQ85400/Label_A41 | Label_A41 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8DQ85400 | 8DQ85400/Label_A41 | Label_A41 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8DQ85400 | 8DQ85400/Label_A41 | Label_A41 | NY-Analytical | 770 | Label Inspection | Width | Width | 6.25 |  | in | FDA,4 | 6.1875 | 6.3125 | TM 143 / Visual | Main_Plans
8DQ85900 | 8DQ85900/Label_NY_NP7200_5.5625-6.25 | Label_NY_NP7200_5.5625-6.25 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Label Inspection | Main_Plans
8DQ85900 | 8DQ85900/Label_NY_NP7200_5.5625-6.25 | Label_NY_NP7200_5.5625-6.25 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8DQ85900 | 8DQ85900/Label_NY_NP7200_5.5625-6.25 | Label_NY_NP7200_5.5625-6.25 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8DQ85900 | 8DQ85900/Label_NY_NP7200_5.5625-6.25 | Label_NY_NP7200_5.5625-6.25 | NY-Analytical | 770 | Label Inspection | Length | Length | 5.5625 |  | in | FDA,4 | 5.5 | 5.625 | TM 143 / Visual | Main_Plans
8DQ85900 | 8DQ85900/Label_NY_NP7200_5.5625-6.25 | Label_NY_NP7200_5.5625-6.25 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8DQ85900 | 8DQ85900/Label_NY_NP7200_5.5625-6.25 | Label_NY_NP7200_5.5625-6.25 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8DQ85900 | 8DQ85900/Label_NY_NP7200_5.5625-6.25 | Label_NY_NP7200_5.5625-6.25 | NY-Analytical | 770 | Label Inspection | Width | Width | 6.25 |  | in | FDA,4 | 6.1875 | 6.3125 | TM 143 / Visual | Main_Plans
8DQ914SL | 8DQ914SL/Label_NY_NP7030_9-7 | Label_NY_NP7030_9-7 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | [6 empty] | Label Inspection | Main_Plans
8DQ914SL | 8DQ914SL/Label_NY_NP7030_9-7 | Label_NY_NP7030_9-7 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8DQ914SL | 8DQ914SL/Label_NY_NP7030_9-7 | Label_NY_NP7030_9-7 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8DQ914SL | 8DQ914SL/Label_NY_NP7030_9-7 | Label_NY_NP7030_9-7 | NY-Analytical | 770 | Label Inspection | Length | Length | 9 |  | in | FDA,4 | 8.9375 | 9.0625 | TM 143 / Visual | Main_Plans
8DQ914SL | 8DQ914SL/Label_NY_NP7030_9-7 | Label_NY_NP7030_9-7 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8DQ914SL | 8DQ914SL/Label_NY_NP7030_9-7 | Label_NY_NP7030_9-7 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8DQ914SL | 8DQ914SL/Label_NY_NP7030_9-7 | Label_NY_NP7030_9-7 | NY-Analytical | 770 | Label Inspection | Width | Width | 7 |  | in | FDA,4 | 6.9375 | 7.0625 | TM 143 / Visual | Main_Plans
8GQ772SK | 8GQ772SK/Label_B03 | Label_B03 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Label Inspection | Main_Plans
8GQ772SK | 8GQ772SK/Label_B03 | Label_B03 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8GQ772SK | 8GQ772SK/Label_B03 | Label_B03 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8GQ772SK | 8GQ772SK/Label_B03 | Label_B03 | NY-Analytical | 770 | Label Inspection | Diameter | Diameter | 1.875 |  | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | Main_Plans
8GQ772SK | 8GQ772SK/Label_B03 | Label_B03 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8GQ772SK | 8GQ772SK/Label_B03 | Label_B03 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OH217UP | 8OH217UP/Label_A25 | Label_A25 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Label Inspection | Main_Plans
8OH217UP | 8OH217UP/Label_A25 | Label_A25 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OH217UP | 8OH217UP/Label_A25 | Label_A25 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OH217UP | 8OH217UP/Label_A25 | Label_A25 | NY-Analytical | 770 | Label Inspection | Length | Length | 10 |  | in | FDA,4 | 9.9375 | 10.0625 | TM 143 / Visual | Main_Plans
8OH217UP | 8OH217UP/Label_A25 | Label_A25 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OH217UP | 8OH217UP/Label_A25 | Label_A25 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OH217UP | 8OH217UP/Label_A25 | Label_A25 | NY-Analytical | 770 | Label Inspection | Width | Width | 4.0625 |  | in | FDA,4 | 3.9375 | 4.0625 | TM 143 / Visual | Main_Plans
8OJ22750 | 8OJ22750/Label_A11 | Label_A11 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Label Inspection | Main_Plans
8OJ22750 | 8OJ22750/Label_A11 | Label_A11 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OJ22750 | 8OJ22750/Label_A11 | Label_A11 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OJ22750 | 8OJ22750/Label_A11 | Label_A11 | NY-Analytical | 770 | Label Inspection | Length | Length | 5.35 |  | in | FDA,4 | 5.3125 | 5.4375 | TM 143 / Visual | Main_Plans
8OJ22750 | 8OJ22750/Label_A11 | Label_A11 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OJ22750 | 8OJ22750/Label_A11 | Label_A11 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OJ22750 | 8OJ22750/Label_A11 | Label_A11 | NY-Analytical | 770 | Label Inspection | Width | Width | 2.61 |  | in | FDA,4 | 2.5625 | 2.6875 | TM 143 / Visual | Main_Plans
8OP03700 | 8OP03700/Label_NY_NP9840_87-75 | Label_NY_NP9840_87-75 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Label Inspection | Main_Plans
8OP03700 | 8OP03700/Label_NY_NP9840_87-75 | Label_NY_NP9840_87-75 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP03700 | 8OP03700/Label_NY_NP9840_87-75 | Label_NY_NP9840_87-75 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP03700 | 8OP03700/Label_NY_NP9840_87-75 | Label_NY_NP9840_87-75 | NY-Analytical | 770 | Label Inspection | Length | Length | 86.97 |  | mm | FDA,4 | 85.4125 | 88.5875 | TM 143 / Visual | Main_Plans
8OP03700 | 8OP03700/Label_NY_NP9840_87-75 | Label_NY_NP9840_87-75 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP03700 | 8OP03700/Label_NY_NP9840_87-75 | Label_NY_NP9840_87-75 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP03700 | 8OP03700/Label_NY_NP9840_87-75 | Label_NY_NP9840_87-75 | NY-Analytical | 770 | Label Inspection | Width | Width | 75.06 |  | mm | FDA,4 | 73.4125 | 76.5875 | TM 143 / Visual | Main_Plans
8OP065RSL | 8OP065RSL/Label_NY_NP9995_2.625-3.375 | Label_NY_NP9995_2.625-3.375 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Label Inspection | Main_Plans
8OP065RSL | 8OP065RSL/Label_NY_NP9995_2.625-3.375 | Label_NY_NP9995_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP065RSL | 8OP065RSL/Label_NY_NP9995_2.625-3.375 | Label_NY_NP9995_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP065RSL | 8OP065RSL/Label_NY_NP9995_2.625-3.375 | Label_NY_NP9995_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Length | Length | 2.625 |  | in | FDA,4 | 2.5625 | 2.6875 | TM 143 / Visual | Main_Plans
8OP065RSL | 8OP065RSL/Label_NY_NP9995_2.625-3.375 | Label_NY_NP9995_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP065RSL | 8OP065RSL/Label_NY_NP9995_2.625-3.375 | Label_NY_NP9995_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP065RSL | 8OP065RSL/Label_NY_NP9995_2.625-3.375 | Label_NY_NP9995_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Width | Width | 3.375 |  | in | FDA,4 | 3.3125 | 3.4375 | TM 143 / Visual | Main_Plans
8OP08700 | 8OP08700/Label_NY_NP9800_3.375-2.625 | Label_NY_NP9800_3.375-2.625 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Label Inspection | Main_Plans
8OP08700 | 8OP08700/Label_NY_NP9800_3.375-2.625 | Label_NY_NP9800_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP08700 | 8OP08700/Label_NY_NP9800_3.375-2.625 | Label_NY_NP9800_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP08700 | 8OP08700/Label_NY_NP9800_3.375-2.625 | Label_NY_NP9800_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Length | Length | 3.375 |  | in | FDA,4 | 3.3125 | 3.4375 | TM 143 / Visual | Main_Plans
8OP08700 | 8OP08700/Label_NY_NP9800_3.375-2.625 | Label_NY_NP9800_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP08700 | 8OP08700/Label_NY_NP9800_3.375-2.625 | Label_NY_NP9800_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP08700 | 8OP08700/Label_NY_NP9800_3.375-2.625 | Label_NY_NP9800_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Width | Width | 2.5625 |  | in | FDA,4 | 2.5625 | 2.6875 | TM 143 / Visual | Main_Plans
8OP09100 | 8OP09100/Label_NY_NP9991_3.375-2.625 | Label_NY_NP9991_3.375-2.625 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Label Inspection | Main_Plans
8OP09100 | 8OP09100/Label_NY_NP9991_3.375-2.625 | Label_NY_NP9991_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP09100 | 8OP09100/Label_NY_NP9991_3.375-2.625 | Label_NY_NP9991_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP09100 | 8OP09100/Label_NY_NP9991_3.375-2.625 | Label_NY_NP9991_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Length | Length | 3.375 |  | in | FDA,4 | 3.3125 | 3.4375 | TM 143 / Visual | Main_Plans
8OP09100 | 8OP09100/Label_NY_NP9991_3.375-2.625 | Label_NY_NP9991_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP09100 | 8OP09100/Label_NY_NP9991_3.375-2.625 | Label_NY_NP9991_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP09100 | 8OP09100/Label_NY_NP9991_3.375-2.625 | Label_NY_NP9991_3.375-2.625 | NY-Analytical | 770 | Label Inspection | Width | Width | 2.625 |  | in | FDA,4 | 2.5625 | 2.6875 | TM 143 / Visual | Main_Plans
8OP09300 | 8OP09300/Label_NY_NP7110_2.55-3 | Label_NY_NP7110_2.55-3 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP09300 | 8OP09300/Label_NY_NP7110_2.55-3 | Label_NY_NP7110_2.55-3 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP09300 | 8OP09300/Label_NY_NP7110_2.55-3 | Label_NY_NP7110_2.55-3 | NY-Analytical | 770 | Label Inspection | Length | Length | 2.5 |  | in | FDA,4 | 2.4875 | 2.6125 | TM 143 / Visual | Main_Plans
8OP09300 | 8OP09300/Label_NY_NP7110_2.55-3 | Label_NY_NP7110_2.55-3 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP09300 | 8OP09300/Label_NY_NP7110_2.55-3 | Label_NY_NP7110_2.55-3 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP09300 | 8OP09300/Label_NY_NP7110_2.55-3 | Label_NY_NP7110_2.55-3 | NY-Analytical | 770 | Label Inspection | Width | Width | 3 |  | in | FDA,4 | 2.9375 | 3.0625 | TM 143 / Visual | Main_Plans
8OP09300 | 8OP09300/Label_NY_NP7110_2.55-3 | Label_NY_NP7110_2.55-3 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Visual | Main_Plans
8OP09301 | 8OP09301/Label_NY_NP7110_2.55-5 | Label_NY_NP7110_2.55-5 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP09301 | 8OP09301/Label_NY_NP7110_2.55-5 | Label_NY_NP7110_2.55-5 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP09301 | 8OP09301/Label_NY_NP7110_2.55-5 | Label_NY_NP7110_2.55-5 | NY-Analytical | 770 | Label Inspection | Length | Length | 2.5313 |  | in | FDA,4 | 2.4875 | 2.6125 | TM 143 / Visual | Main_Plans
8OP09301 | 8OP09301/Label_NY_NP7110_2.55-5 | Label_NY_NP7110_2.55-5 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP09301 | 8OP09301/Label_NY_NP7110_2.55-5 | Label_NY_NP7110_2.55-5 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP09301 | 8OP09301/Label_NY_NP7110_2.55-5 | Label_NY_NP7110_2.55-5 | NY-Analytical | 770 | Label Inspection | Width | Width | 4.9687 |  | in | FDA,4 | 4.9375 | 5.0625 | TM 143 / Visual | Main_Plans
8OP09301 | 8OP09301/Label_NY_NP7110_2.55-5 | Label_NY_NP7110_2.55-5 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Visual | Main_Plans
8OP227RSL | 8OP227RSL/Label_NY_NP9997_2.625-3.375 | Label_NY_NP9997_2.625-3.375 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | [6 empty] | Label Inspection | Main_Plans
8OP227RSL | 8OP227RSL/Label_NY_NP9997_2.625-3.375 | Label_NY_NP9997_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP227RSL | 8OP227RSL/Label_NY_NP9997_2.625-3.375 | Label_NY_NP9997_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP227RSL | 8OP227RSL/Label_NY_NP9997_2.625-3.375 | Label_NY_NP9997_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Length | Length | 2.63 |  | in | FDA,4 | 2.5625 | 2.6875 | TM 143 / Visual | Main_Plans
8OP227RSL | 8OP227RSL/Label_NY_NP9997_2.625-3.375 | Label_NY_NP9997_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP227RSL | 8OP227RSL/Label_NY_NP9997_2.625-3.375 | Label_NY_NP9997_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP227RSL | 8OP227RSL/Label_NY_NP9997_2.625-3.375 | Label_NY_NP9997_2.625-3.375 | NY-Analytical | 770 | Label Inspection | Width | Width | 3.37 |  | in | FDA,4 | 3.3125 | 3.4375 | TM 143 / Visual | Main_Plans
8OP22900RF | 8OP22900RF/Label_NY_NP98371_78-61 | Label_NY_NP98371_78-61 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Label Inspection | Main_Plans
8OP22900RF | 8OP22900RF/Label_NY_NP98371_78-61 | Label_NY_NP98371_78-61 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP22900RF | 8OP22900RF/Label_NY_NP98371_78-61 | Label_NY_NP98371_78-61 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP22900RF | 8OP22900RF/Label_NY_NP98371_78-61 | Label_NY_NP98371_78-61 | NY-Analytical | 770 | Label Inspection | Length | Length | 78.75 |  | mm | FDA,4 | 76.4125 | 79.5875 | TM 143 / Visual | Main_Plans
8OP22900RF | 8OP22900RF/Label_NY_NP98371_78-61 | Label_NY_NP98371_78-61 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP22900RF | 8OP22900RF/Label_NY_NP98371_78-61 | Label_NY_NP98371_78-61 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP22900RF | 8OP22900RF/Label_NY_NP98371_78-61 | Label_NY_NP98371_78-61 | NY-Analytical | 770 | Label Inspection | Width | Width | 61.36 |  | mm | FDA,4 | 59.4125 | 62.5875 | TM 143 / Visual | Main_Plans
8OP260RSL | 8OP260RSL/Label_NY_NP7700_3.375-2.3437 | Label_NY_NP7700_3.375-2.3437 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Label Inspection | Main_Plans
8OP260RSL | 8OP260RSL/Label_NY_NP7700_3.375-2.3437 | Label_NY_NP7700_3.375-2.3437 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP260RSL | 8OP260RSL/Label_NY_NP7700_3.375-2.3437 | Label_NY_NP7700_3.375-2.3437 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP260RSL | 8OP260RSL/Label_NY_NP7700_3.375-2.3437 | Label_NY_NP7700_3.375-2.3437 | NY-Analytical | 770 | Label Inspection | Length | Length | 3.3125 |  | in | FDA,4 | 3.3125 | 3.4375 | TM 143 / Visual | Main_Plans
8OP260RSL | 8OP260RSL/Label_NY_NP7700_3.375-2.3437 | Label_NY_NP7700_3.375-2.3437 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP260RSL | 8OP260RSL/Label_NY_NP7700_3.375-2.3437 | Label_NY_NP7700_3.375-2.3437 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP260RSL | 8OP260RSL/Label_NY_NP7700_3.375-2.3437 | Label_NY_NP7700_3.375-2.3437 | NY-Analytical | 770 | Label Inspection | Width | Width | 2.2812 |  | in | FDA,4 | 2.2812 | 2.4062 | TM 143 / Visual | Main_Plans
8OP263CK | 8OP263CK/Label_NY_NP7030_8.5-6.5 | Label_NY_NP7030_8.5-6.5 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Label Inspection | Main_Plans
8OP263CK | 8OP263CK/Label_NY_NP7030_8.5-6.5 | Label_NY_NP7030_8.5-6.5 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP263CK | 8OP263CK/Label_NY_NP7030_8.5-6.5 | Label_NY_NP7030_8.5-6.5 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP263CK | 8OP263CK/Label_NY_NP7030_8.5-6.5 | Label_NY_NP7030_8.5-6.5 | NY-Analytical | 770 | Label Inspection | Length | Length | 8.5 |  | in | FDA,4 | 8.4375 | 8.5625 | TM 143 / Visual | Main_Plans
8OP263CK | 8OP263CK/Label_NY_NP7030_8.5-6.5 | Label_NY_NP7030_8.5-6.5 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP263CK | 8OP263CK/Label_NY_NP7030_8.5-6.5 | Label_NY_NP7030_8.5-6.5 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP263CK | 8OP263CK/Label_NY_NP7030_8.5-6.5 | Label_NY_NP7030_8.5-6.5 | NY-Analytical | 770 | Label Inspection | Width | Width | 6.5 |  | in | FDA,4 | 6.4375 | 6.5625 | TM 143 / Visual | Main_Plans
8OP308SIRF | 8OP308SIRF/Label_NY_NP9300_2-1.375 | Label_NY_NP9300_2-1.375 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Label Inspection | Main_Plans
8OP308SIRF | 8OP308SIRF/Label_NY_NP9300_2-1.375 | Label_NY_NP9300_2-1.375 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP308SIRF | 8OP308SIRF/Label_NY_NP9300_2-1.375 | Label_NY_NP9300_2-1.375 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP308SIRF | 8OP308SIRF/Label_NY_NP9300_2-1.375 | Label_NY_NP9300_2-1.375 | NY-Analytical | 770 | Label Inspection | Length | Length | 2 |  | in | FDA,4 | 1.9375 | 2.0625 | TM 143 / Visual | Main_Plans
8OP308SIRF | 8OP308SIRF/Label_NY_NP9300_2-1.375 | Label_NY_NP9300_2-1.375 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP308SIRF | 8OP308SIRF/Label_NY_NP9300_2-1.375 | Label_NY_NP9300_2-1.375 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP308SIRF | 8OP308SIRF/Label_NY_NP9300_2-1.375 | Label_NY_NP9300_2-1.375 | NY-Analytical | 770 | Label Inspection | Width | Width | 1.375 |  | in | FDA,4 | 1.3125 | 1.4375 | TM 143 / Visual | Main_Plans
8OP585SIRF | 8OP585SIRF/Label_NY_NP9300_1.375-2 | Label_NY_NP9300_1.375-2 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Label Inspection | Main_Plans
8OP585SIRF | 8OP585SIRF/Label_NY_NP9300_1.375-2 | Label_NY_NP9300_1.375-2 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP585SIRF | 8OP585SIRF/Label_NY_NP9300_1.375-2 | Label_NY_NP9300_1.375-2 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP585SIRF | 8OP585SIRF/Label_NY_NP9300_1.375-2 | Label_NY_NP9300_1.375-2 | NY-Analytical | 770 | Label Inspection | Length | Length | 1.37 |  | in | FDA,4 | 1.3125 | 1.4375 | TM 143 / Visual | Main_Plans
8OP585SIRF | 8OP585SIRF/Label_NY_NP9300_1.375-2 | Label_NY_NP9300_1.375-2 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP585SIRF | 8OP585SIRF/Label_NY_NP9300_1.375-2 | Label_NY_NP9300_1.375-2 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP585SIRF | 8OP585SIRF/Label_NY_NP9300_1.375-2 | Label_NY_NP9300_1.375-2 | NY-Analytical | 770 | Label Inspection | Width | Width | 1.99 |  | in | FDA,4 | 1.9375 | 2.0625 | TM 143 / Visual | Main_Plans
8OP63000RF | 8OP63000RF/Label_NY_NP9837_85-73 | Label_NY_NP9837_85-73 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Label Inspection | Main_Plans
8OP63000RF | 8OP63000RF/Label_NY_NP9837_85-73 | Label_NY_NP9837_85-73 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP63000RF | 8OP63000RF/Label_NY_NP9837_85-73 | Label_NY_NP9837_85-73 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP63000RF | 8OP63000RF/Label_NY_NP9837_85-73 | Label_NY_NP9837_85-73 | NY-Analytical | 770 | Label Inspection | Length | Length | 84 |  | mm | FDA,4 | 83.4125 | 86.5875 | TM 143 / Visual | Main_Plans
8OP63000RF | 8OP63000RF/Label_NY_NP9837_85-73 | Label_NY_NP9837_85-73 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP63000RF | 8OP63000RF/Label_NY_NP9837_85-73 | Label_NY_NP9837_85-73 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Main_Plans
8OP63000RF | 8OP63000RF/Label_NY_NP9837_85-73 | Label_NY_NP9837_85-73 | NY-Analytical | 770 | Label Inspection | Width | Width | 73 |  | mm | FDA,4 | 71.4125 | 74.5875 | TM 143 / Visual | Main_Plans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
Extra_TestPlans
2AA922SLI | 2AA922SLI/Lid_NY_Skip lot testing | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Print | Print | No smears, streaks, | [5 empty] | Canister & Lid Inspection | Extra_specs
2AA922SLI | 2AA922SLI/Lid_NY_Skip lot testing | Lid_NY_Skip lot testing | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Certificate of Compliance | Extra_specs
2AA922SLI | 2AA922SLI/Lid_NY_Skip lot testing | Lid_NY_Skip lot testing | NY-Analytical | 767 | Certificate of Compliance | Certificate of Compliance | Certificate of Compliance | CofC meets specifications (Done) | [5 empty] | Canister & Lid Inspection | Extra_specs
2AA922SLI | 2AA922SLI/Lid_NY_Skip lot testing | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) | [5 empty] | TM 143 / Visual | Extra_specs
2AA922SLI | 2AA922SLI/Lid_NY_Skip lot testing | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Color/Copy | Color/Copy | Copy/color is in accordance with the item # print specification (Done) | [5 empty] | TM 143 / Visual | Extra_specs
2AA922SLI | 2AA922SLI/Lid_NY_Skip lot testing | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Length | Label Length | 3.6875 |  | in | FDA,4 | 3.6675 | 3.7925 | TM 143 / Visual | Extra_specs
2AA922SLI | 2AA922SLI/Lid_NY_Skip lot testing | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Label Width | Label Width | 1.8438 |  | in | FDA,4 | 1.8125 | 1.9375 | TM 143 / Visual | Extra_specs
2AA922SLI | 2AA922SLI/Lid_NY_Skip lot testing | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | No solvent odor is present (Done) | [5 empty] | TM 143 / Visual | Extra_specs
2AJ123 | 2AJ123/Web_NY_Skip lot testing | Web_NY_Skip lot testing | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Web Measurement | Extra_specs
2AJ123 | 2AJ123/Web_NY_Skip lot testing | Web_NY_Skip lot testing | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 11 |  | in | FDA,4 | 10.9375 | 11.0625 | Measurement | Extra_specs
2AJ123 | 2AJ123/Web_NY_Skip lot testing | Web_NY_Skip lot testing | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream) | Basis Weight | 42.33 |  | #/Ream | FDA,1 | 37.8 | 48.2 | TM 101 | Extra_specs
2AJ123 | 2AJ123/Web_NY_Skip lot testing | Web_NY_Skip lot testing | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@300DegF/30psi/1.5 sec dwell) | Seal Strength (@300°F/30psi/1.5 sec dwell) | 7486 |  | g/in | FDA,0 | 3500 |  | TM 107 | Extra_specs
2AQ471LI | 2AQ471LI/Lid_04B_560_1 | Lid_04B_560_1 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Canister & Lid Inspection | Extra_specs
2AQ471LI | 2AQ471LI/Lid_04B_560_1 | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Appearance Material is clean and free of particulate matter. (Done) | [5 empty] | TM 143 / Visual | Extra_specs
2AQ471LI | 2AQ471LI/Lid_04B_560_1 | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Cap Thickness | Cap Thickness | 0.154 |  | in | FDA,3 | 0.147 | 0.157 | TM 143 / Visual | Extra_specs
2AQ471LI | 2AQ471LI/Lid_04B_560_1 | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Blue (Done) | [5 empty] | TM 143 / Visual | Extra_specs
2AQ471LI | 2AQ471LI/Lid_04B_560_1 | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) | [5 empty] | TM 143 / Visual | Extra_specs
2AQ471LI | 2AQ471LI/Lid_04B_560_1 | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter (as per specification) | Lid Inside Diameter | 3.7695 |  | in | FDA,3 | 3.76 | 3.77 | TM 143 / Visual | Extra_specs
2AQ471LI | 2AQ471LI/Lid_04B_560_1 | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.05 |  | in | FDA,3 | 0.04 | 0.05 | TM 143 / Visual | Extra_specs
2AQ471LI | 2AQ471LI/Lid_04B_560_1 | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) | [5 empty] | TM 143 / Visual | Extra_specs
2AQ471LI | 2AQ471LI/Lid_04B_560_1 | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.5845 |  | in | FDA,3 | 1.58 | 1.6 | TM 143 / Visual | Extra_specs
2AQ471LI | 2AQ471LI/Lid_04B_560_1 | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Outer Diameter (per specification) | Pop-Up Hole Outer Diameter | 0.565 |  | in | FDA,3 | 0.555 | 0.565 | TM 143 / Visual | Extra_specs
2AQ471LI | 2AQ471LI/Lid_04B_560_1 | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Thicknes | Pop-Up Hole Thicknes | 0.02 |  | in | FDA,3 | 0.015 | 0.025 | TM 143 / Visual | Extra_specs
2AQ471LI | 2AQ471LI/Lid_04B_560_1 | Lid_04B_560_1 | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 22.773 |  | g | FDA,1 | 20 | 24 | TM 143 / Visual | Extra_specs
2AU252 | 2AU252/Web_3459_8 | Web_3459_8 | NY-Analytical | 1044 | Basis Weight (lbs/ream) | Basis Weight | Basis Weight |  |  | Lb/Ream | FDA,2 | 57.56 | 70.4 | TM 101 | Extra_specs
2AU252 | 2AU252/Web_3459_8 | Web_3459_8 | NY-Analytical | 761 | Web Measurement | Core ID (Inches) | Core ID |  |  | in | FDA,3 | 2.9375 | 3.0625 | TM 143 | Extra_specs
2AU252 | 2AU252/Web_3459_8 | Web_3459_8 | NY-Analytical | 761 | Web Measurement | Roll OD | Roll OD | 15 |  | in | FDA,3 | 14.9375 | 15.0625 | TM 143 | Extra_specs
2AU252 | 2AU252/Web_3459_8 | Web_3459_8 | NY-Analytical | 761 | Web Measurement | Web Width (Inches) | Web Width | 8 |  | in | FDA,4 | 7.9375 | 8.0625 | TM 143 | Extra_specs
2AU252 | 2AU252/Web_3459_8 | Web_3459_8 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [6 empty] | Visual | Extra_specs
2BQ551LI | 2BQ551LI/Skip Lot Release_NY_2OQ850 | Skip Lot Release_NY_2OQ850 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) | [5 empty] | COA | Extra_specs
2BQ551LI | 2BQ551LI/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [6 empty] | Certificate of Analysis | Extra_specs
2BQ551LI | 2BQ551LI/Lid_NY_Skip lot testing | Lid_NY_Skip lot testing | NY-Analytical | 791 | Flip Lid Force | Flip Lid Force | Flip Lid Force | 1.07 |  | lb | FDA,2 | 0.75 |  | TM 143 | Extra_specs
2BQ551LI | 2BQ551LI/Lid_NY_Skip lot testing | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Bead Diameter | Bead Diameter | 4.85 |  | in | FDA,3 | 4.847 | 4.867 | TM 143 / Visual | Extra_specs
2BQ551LI | 2BQ551LI/Lid_NY_Skip lot testing | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Center to Center Distance | Center to Center Distance | 0.45 |  | in | FDA,3 | 0.445 | 0.455 | TM 143 / Visual | Extra_specs
2BQ551LI | 2BQ551LI/Lid_NY_Skip lot testing | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Translucent Purple (Done) | [5 empty] | TM 143 / Visual | Extra_specs
2BQ551LI | 2BQ551LI/Lid_NY_Skip lot testing | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | End Circle Diameter | End Circle Diameter | 0.13 |  | in | FDA,3 | 0.12 | 0.13 | TM 143 / Visual | Extra_specs
2BQ551LI | 2BQ551LI/Lid_NY_Skip lot testing | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) | [5 empty] | TM 143 / Visual | Extra_specs
2BQ551LI | 2BQ551LI/Lid_NY_Skip lot testing | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Insert Fin Gap (per specification) | Insert Fin Gap | 0.076 |  | in | FDA,3 | 0.07 | 0.08 | TM 143 / Visual | Extra_specs
2BQ551LI | 2BQ551LI/Lid_NY_Skip lot testing | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Lid Inside Diameter (as per specification) | Lid Inside Diameter | 4.92 |  | in | FDA,3 | 4.917 | 4.937 | TM 143 / Visual | Extra_specs
2BQ551LI | 2BQ551LI/Lid_NY_Skip lot testing | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Lid Wall Thickness (per speciffication) | Lid Wall Thickness | 0.042 |  | in | FDA,3 | 0.035 | 0.045 | TM 143 / Visual | Extra_specs
2BQ551LI | 2BQ551LI/Lid_NY_Skip lot testing | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Flap Thickness (per specification) | Orifice Flap Thickness | 0.031 |  | in | FDA,3 | 0.03 | 0.04 | TM 143 / Visual | Extra_specs
2BQ551LI | 2BQ551LI/Lid_NY_Skip lot testing | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Overall Lid Height (per specification) | Overall Lid Height | 1.49 |  | in | FDA,3 | 1.472 | 1.492 | TM 143 / Visual | Extra_specs
2BQ551LI | 2BQ551LI/Lid_NY_Skip lot testing | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Weight (per specification) | Weight | 34.65 |  | g | FDA,1 | 32 | 42 | TM 143 / Visual | Extra_specs
2OD024 | 2OD024/Web_NY_Skip lot testing | Web_NY_Skip lot testing | NY-Analytical | 756 | Basis Weight | Basis_Weight_(#/Ream-Printed/Varnished) | Basis Weight | 59.8 |  | #/Ream | FDA,1 | 55.6 | 68 | TM 101 | Extra_specs
2OD024 | 2OD024/Web_NY_Skip lot testing | Web_NY_Skip lot testing | NY-Analytical | 762 | Seal Strength | Seal_Strength_(@325DegF/40PSI/0.5Sec_Dwell) | Seal Strength (@325°F/40PSI/0.5Sec Dwell) | 3353.04 |  | g/in | FDA,0 | 1500 |  | TM 107 | Extra_specs
2OP067 | 2OP067/Web_NY_Skip lot testing | Web_NY_Skip lot testing | NY-Analytical | 1044 | Basis Weight (lbs/ream) | Basis Weight | Basis Weight | 42.86 |  | Lb/Ream | FDA,2 | 40.1 | 49.02 | TM 101 | Extra_specs
2OP067 | 2OP067/Web_NY_Skip lot testing | Web_NY_Skip lot testing | NY-Analytical | 1049 | Seal strength (NP8536) | Seal Strength (g/inch@320DegF/40psi/0.5 sec dwell) | Seal Strength (@320°F/40psi/0.5 sec dwell) | 9089 |  | g/in | FDA,0 | 2000 |  | TM 107 | Extra_specs
2OP067 | 2OP067/Web_NY_Skip lot testing | Web_NY_Skip lot testing | NY-Analytical | 763 | Web Thickness | Thickness | Thickness | 2.6 |  | mil | FDA,3 | 2.367 | 2.893 | TM 132 | Extra_specs
2OP067 | 2OP067/Web_NY_Skip lot testing | Web_NY_Skip lot testing | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | Visual | Extra_specs
2OP067 | 2OP067/Web_NY_Skip lot testing | Web_NY_Skip lot testing | NY-Analytical | 748 | Visual Inspection | Barcode Legibility/Quality | Barcode Legibility/Quality | Barcode should scan and code should be matched under the barcode (Done) | [5 empty] | Visual | Extra_specs
2OP067 | 2OP067/Web_NY_Skip lot testing | Web_NY_Skip lot testing | NY-Analytical | 748 | Visual Inspection | Color/Copy | Color/Copy | Meets specifications (Done) | [5 empty] | Visual | Extra_specs
2OP093 | 2OP093/Skip Lot Release_NY_2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) | [5 empty] | COA | Extra_specs
2OP093 | 2OP093/Skip Lot Release_NY_2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Canister & Lid Inspection | Extra_specs
2OP093 | 2OP093/Skip Lot Release_NY_2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) | [5 empty] | Canister & Lid Inspection | Extra_specs
2OP093 | 2OP093/Skip Lot Release_NY_2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Matches Specification (Done) | [5 empty] | Canister & Lid Inspection | Extra_specs
2OP093 | 2OP093/Skip Lot Release_NY_2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 765 | Canister & Lid Inspection | Diameter | Diameter | 2.75 |  | in | FDA,3 | 2.709 | 2.803 | Canister & Lid Inspection | Extra_specs
2OP093 | 2OP093/Skip Lot Release_NY_2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 765 | Canister & Lid Inspection | E Diameter | E Diameter | 2.52 |  | in | FDA,3 | 2.509 | 2.539 | Canister & Lid Inspection | Extra_specs
2OP093 | 2OP093/Skip Lot Release_NY_2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) | [5 empty] | Canister & Lid Inspection | Extra_specs
2OP093 | 2OP093/Skip Lot Release_NY_2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 765 | Canister & Lid Inspection | H Dimension | H Dimension | 0.47 |  | in | FDA,3 | 0.457 | 0.487 | Canister & Lid Inspection | Extra_specs
2OP093 | 2OP093/Skip Lot Release_NY_2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 765 | Canister & Lid Inspection | Height | Height | 3.5 |  | in | FDA,3 | 3.45 | 3.544 | Canister & Lid Inspection | Extra_specs
2OP093 | 2OP093/Skip Lot Release_NY_2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) | [5 empty] | Canister & Lid Inspection | Extra_specs
2OP093 | 2OP093/Skip Lot Release_NY_2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 765 | Canister & Lid Inspection | S Dimension | S Dimension | 0.042 |  | in | FDA,3 | 0.035 | 0.065 | Canister & Lid Inspection | Extra_specs
2OP093 | 2OP093/Skip Lot Release_NY_2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 765 | Canister & Lid Inspection | T Diameter | T Diameter | 2.61 |  | in | FDA,3 | 2.593 | 2.623 | Canister & Lid Inspection | Extra_specs
2OP093 | 2OP093/Skip Lot Release_NY_2OP093 | Skip Lot Release_NY_2OP093 | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight | 26.32 |  | g | FDA,2 | 24.5 | 27.5 | Measurement | Extra_specs
2OP093 | 2OP093/ Jar _10 oz Blue HDPE Grade Jar | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Certificate of Analysis | Extra_specs
2OP093 | 2OP093/ Jar _10 oz Blue HDPE Grade Jar | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) | [5 empty] | Canister & Lid Inspection | Extra_specs
2OP093 | 2OP093/ Jar _10 oz Blue HDPE Grade Jar | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Matches Specification (Done) | [5 empty] | Canister & Lid Inspection | Extra_specs
2OP093 | 2OP093/ Jar _10 oz Blue HDPE Grade Jar | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | Diameter | Diameter | 2.746 |  | in | FDA,3 | 2.709 | 2.803 | Canister & Lid Inspection | Extra_specs
2OP093 | 2OP093/ Jar _10 oz Blue HDPE Grade Jar | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | E Diameter | E Diameter | 2.51 |  | in | FDA,3 | 2.509 | 2.539 | Canister & Lid Inspection | Extra_specs
2OP093 | 2OP093/ Jar _10 oz Blue HDPE Grade Jar | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | Functionality | Functionality | Pass (Done) | [5 empty] | Canister & Lid Inspection | Extra_specs
2OP093 | 2OP093/ Jar _10 oz Blue HDPE Grade Jar | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | H Dimension | H Dimension | 0.458 |  | in | FDA,3 | 0.457 | 0.487 | Canister & Lid Inspection | Extra_specs
2OP093 | 2OP093/ Jar _10 oz Blue HDPE Grade Jar | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | Height | Height | 3.52 |  | in | FDA,3 | 3.45 | 3.544 | Canister & Lid Inspection | Extra_specs
2OP093 | 2OP093/ Jar _10 oz Blue HDPE Grade Jar | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) | [5 empty] | Canister & Lid Inspection | Extra_specs
2OP093 | 2OP093/ Jar _10 oz Blue HDPE Grade Jar | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | S Dimension | S Dimension | 0.053 |  | in | FDA,3 | 0.035 | 0.065 | Canister & Lid Inspection | Extra_specs
2OP093 | 2OP093/ Jar _10 oz Blue HDPE Grade Jar | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | T Diameter | T Diameter | 2.6 |  | in | FDA,3 | 2.593 | 2.623 | Canister & Lid Inspection | Extra_specs
2OP093 | 2OP093/ Jar _10 oz Blue HDPE Grade Jar | Jar _10 oz Blue HDPE Grade Jar | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight | 26.19 |  | g | FDA,2 | 24.5 | 27.5 | Measurement | Extra_specs
2OP134 | 2OP134/Skip Lot Release_NY_2OP134 | Skip Lot Release_NY_2OP134 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) | [5 empty] | COA | Extra_specs
2OP138 | 2OP138/Skip Lot Relrase_NY_2OP138C | Skip Lot Relrase_NY_2OP138C | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) | [5 empty] | COA | Extra_specs
2OP138 | 2OP138/Skip Lot Relrase_NY_2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) | [5 empty] | COA | Extra_specs
2OP138 | 2OP138/Skip Lot Relrase_NY_2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | [6 empty] | TM 143 / Visual | Extra_specs
2OP138 | 2OP138/Skip Lot Relrase_NY_2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | White (Done) | [5 empty] | TM 143 / Visual | Extra_specs
2OP138 | 2OP138/Skip Lot Relrase_NY_2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | [6 empty] | TM 143 / Visual | Extra_specs
2OP280LI | 2OP280LI/Lid_NY_Skip lot testing | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Cap Closure | Functionality - Cap Closure | [6 empty] | TM 143 / Visual | Extra_specs
2OP280LI | 2OP280LI/Lid_NY_Skip lot testing | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | [6 empty] | TM 143 / Visual | Extra_specs
2OP280LI | 2OP280LI/Lid_NY_Skip lot testing | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Orifice Diameter | Orifice Diameter |  |  | in | FDA,3 | 1.468 | 1.478 | TM 143 / Visual | Extra_specs
2OP638 | 2OP638/Skip Lot Relrase_NY_2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid has a tight friction fit | [5 empty] | Canister & Lid Inspection | Extra_specs
2OP638 | 2OP638/Skip Lot Relrase_NY_2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Certificate of Analysis | Extra_specs
2OP638 | 2OP638/Skip Lot Relrase_NY_2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Bead OD | Bead OD | 4.9245 |  | in | FDA,3 | 4.922 | 4.958 | Canister & Lid Inspection | Extra_specs
2OP638 | 2OP638/Skip Lot Relrase_NY_2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Canister OD | Canister OD | 5.03 |  | in | FDA,3 | 4.956 | 5.076 | Canister & Lid Inspection | Extra_specs
2OP638 | 2OP638/Skip Lot Relrase_NY_2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Neck OD | Neck OD | 4.8545 |  | in | FDA,3 | 4.824 | 4.86 | Canister & Lid Inspection | Extra_specs
2OP638 | 2OP638/Skip Lot Relrase_NY_2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Total Height OD | Total Height OD | 8.75 |  | in | FDA,3 | 8.688 | 8.808 | Canister & Lid Inspection | Extra_specs
2OP638 | 2OP638/Skip Lot Relrase_NY_2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | HDPE Resin Identification Marker | Meets the Specification (Done) | [5 empty] | Canister & Lid Inspection | Extra_specs
2OP638 | 2OP638/Skip Lot Relrase_NY_2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Lip Width (for reference only) | Lip Width (for reference only) | 0.154 |  | in | FDA,3 | 0.14 |  | Canister & Lid Inspection | Extra_specs
2OP638 | 2OP638/Skip Lot Relrase_NY_2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Wall Thickness (Bottom base area) | Wall Thickness (Bottom base area) | 0.0325 |  | in | FDA,3 | 0.01 |  | Canister & Lid Inspection | Extra_specs
2OP638 | 2OP638/Skip Lot Relrase_NY_2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Wall Thickness (Top area) | Wall Thickness (Top area) | 0.043 |  | in | FDA,3 | 0.02 |  | Canister & Lid Inspection | Extra_specs
2OP638 | 2OP638/Skip Lot Relrase_NY_2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight | 89.75 |  | g | FDA,1 | 85.5 | 94.5 | Measurement | Extra_specs
2OP638 | 2OP638/Skip Lot Relrase_NY_2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 1052 | Top Load Compression Strength | Top load Compression Strength (0.5"/mSR,0.450"TD) | Top load Compression Strength (0.5"/mSR,0.450"TD) | 138.8 |  | lb | FDA,0 | 97 |  | NPTM 1019 | Extra_specs
2OP638 | 2OP638/Skip Lot Relrase_NY_2OP638 | Skip Lot Relrase_NY_2OP638 | NY-Analytical | 1053 | Cavity | cavity Inserts | cavity Inserts | Meets specification (Done) | [5 empty] | Visual | Extra_specs
2OP638 | 2OP638/Bulk_IN_4ABBY12 - 20190218 | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 902 | Bulk Liquid Inspection | Appearance (TM 208) | Appearance | [6 empty] | See Analyte for Method # | Extra_specs
2OP638 | 2OP638/Bulk_IN_4ABBY12 - 20190218 | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 902 | Bulk Liquid Inspection | Color (TM 208) | Color | [6 empty] | See Analyte for Method # | Extra_specs
2OP638 | 2OP638/Bulk_IN_4ABBY12 - 20190218 | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 902 | Bulk Liquid Inspection | Contamination (TM 208) | Contamination | [6 empty] | See Analyte for Method # | Extra_specs
2OP638 | 2OP638/Bulk_IN_4ABBY12 - 20190218 | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 902 | Bulk Liquid Inspection | Odor (TM 208) | Odor | [6 empty] | See Analyte for Method # | Extra_specs
2OP638 | 2OP638/Bulk_IN_4ABBY12 - 20190218 | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 902 | Bulk Liquid Inspection | pH (TM 200) | pH |  |  |  | FDA,2 | 4.5 | 4.9 | See Analyte for Method # | Extra_specs
2OP638 | 2OP638/Bulk_IN_4ABBY12 - 20190218 | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 943 | %Phenoxyethanol & Sodium Benzoate | Phenoxyethanol | Phenoxyethanol |  |  | % | FDA,3 | 0.63 | 0.77 | TM 395 | Extra_specs
2OP638 | 2OP638/Bulk_IN_4ABBY12 - 20190218 | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 943 | %Phenoxyethanol & Sodium Benzoate | Phenoxyethanol_Final | Phenoxyethanol Final |  |  | % | FDA,3 | 0.63 | 0.77 | TM 395 | Extra_specs
2OP638 | 2OP638/Bulk_IN_4ABBY12 - 20190218 | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 943 | %Phenoxyethanol & Sodium Benzoate | Sodium_Benzoate | Sodium Benzoate |  |  | % | FDA,3 | 0.423 | 0.517 | TM 395 | Extra_specs
2OP638 | 2OP638/Bulk_IN_4ABBY12 - 20190218 | Bulk_IN_4ABBY12 - 20190218 | IN-Analytical | 943 | %Phenoxyethanol & Sodium Benzoate | Sodium_Benzoate_Final | Sodium Benzoate Final |  |  | % | FDA,3 | 0.423 | 0.517 | TM 395 | Extra_specs
2OQ753 | 2OQ753/Skip Lot Release_NY_2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid has a tight friction fit | [5 empty] | Canister & Lid Inspection | Extra_specs
2OQ753 | 2OQ753/Skip Lot Release_NY_2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) | [5 empty] | COA | Extra_specs
2OQ753 | 2OQ753/Skip Lot Release_NY_2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Canister & Lid Inspection | Extra_specs
2OQ753 | 2OQ753/Skip Lot Release_NY_2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) | [5 empty] | Canister & Lid Inspection | Extra_specs
2OQ753 | 2OQ753/Skip Lot Release_NY_2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | White (Done) | [5 empty] | Canister & Lid Inspection | Extra_specs
2OQ753 | 2OQ753/Skip Lot Release_NY_2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Bead OD | Bead OD | 4.911 |  | in | FDA,3 | 4.908 | 4.944 | Canister & Lid Inspection | Extra_specs
2OQ753 | 2OQ753/Skip Lot Release_NY_2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Canister OD | Canister OD | 4.998 |  | in | FDA,3 | 4.998 | 5.034 | Canister & Lid Inspection | Extra_specs
2OQ753 | 2OQ753/Skip Lot Release_NY_2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Lip (width) | Lip Width | 0.15 |  | in | FDA,3 | 0.142 | 0.178 | Canister & Lid Inspection | Extra_specs
2OQ753 | 2OQ753/Skip Lot Release_NY_2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Dimensions - Total Height OD | Total Height OD | 6.825 |  | in | FDA,3 | 6.706 | 6.826 | Canister & Lid Inspection | Extra_specs
2OQ753 | 2OQ753/Skip Lot Release_NY_2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | HDPE Resin Identification Marker | Meets the Specification (Done) | [5 empty] | Canister & Lid Inspection | Extra_specs
2OQ753 | 2OQ753/Skip Lot Release_NY_2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) | [5 empty] | Canister & Lid Inspection | Extra_specs
2OQ753 | 2OQ753/Skip Lot Release_NY_2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight | 68.86 |  | g | FDA,1 | 67.5 | 72.5 | Measurement | Extra_specs
2OQ753 | 2OQ753/Skip Lot Release_NY_2OQ753 | Skip Lot Release_NY_2OQ753 | NY-Analytical | 766 | Lid Pull Force | Lid Pull-Off Force | Lid Pull-Off Force | 2979.67 |  | g | FDA,0 | 2000 | 3500 | TM 148 | Extra_specs
2OQ753X | 2OQ753X/Skip Lot Relrase_NY_2OQ753X | Skip Lot Relrase_NY_2OQ753X | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) | [5 empty] | COA | Extra_specs
2OQ753X | 2OQ753X/Skip Lot Release_NY_2OQ850X | Skip Lot Release_NY_2OQ850X | NY-Analytical | 1014 | Canister Release | Certificate of Analysis Review | Certificate of Analysis | C of A Meet Specifications (Done) | [5 empty] | COA | Extra_specs
2OQ850 | 2OQ850/Reduced Testing_NY_Can 2OQ850X | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid has a tight friction fit | [5 empty] | Canister & Lid Inspection | Extra_specs
2OQ850 | 2OQ850/Reduced Testing_NY_Can 2OQ850X | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Canister & Lid Inspection | Extra_specs
2OQ850 | 2OQ850/Reduced Testing_NY_Can 2OQ850X | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) | [5 empty] | Canister & Lid Inspection | Extra_specs
2OQ850 | 2OQ850/Reduced Testing_NY_Can 2OQ850X | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | White (Done) | [5 empty] | Canister & Lid Inspection | Extra_specs
2OQ850 | 2OQ850/Reduced Testing_NY_Can 2OQ850X | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | HDPE Resin Identification Marker | Meets the Specification (Done) | [5 empty] | Canister & Lid Inspection | Extra_specs
2OQ850 | 2OQ850/Reduced Testing_NY_Can 2OQ850X | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) | [5 empty] | Canister & Lid Inspection | Extra_specs
2OQ850 | 2OQ850/Reduced Testing_NY_Can 2OQ850X | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight | 88.41 |  | g | FDA,1 | 85.5 | 94.5 | Measurement | Extra_specs
2OQ850 | 2OQ850/Reduced Testing_NY_Can 2OQ850X | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 1052 | Top Load Compression Strength | Top load Compression Strength (0.5"/mSR,0.450"TD) | Top load Compression Strength (0.5"/mSR,0.450"TD) |  |  | lb | FDA,0 | 97 |  | NPTM 1019 | Extra_specs
2OQ850 | 2OQ850/Reduced Testing_NY_Can 2OQ850X | Reduced Testing_NY_Can 2OQ850X | NY-Analytical | 1053 | Cavity | cavity Inserts | cavity Inserts | [6 empty] | Visual | Extra_specs
2OQ850 | 2OQ850/Skip Lot Release_NY_2OQ850 | Skip Lot Release_NY_2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Functionality - Lid Fit | Functionality - Lid Fit | Lid has a tight friction fit | [5 empty] | Canister & Lid Inspection | Extra_specs
2OQ850 | 2OQ850/Skip Lot Release_NY_2OQ850 | Skip Lot Release_NY_2OQ850 | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Certificate of Analysis | Extra_specs
2OQ850 | 2OQ850/Skip Lot Release_NY_2OQ850 | Skip Lot Release_NY_2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) | [5 empty] | Canister & Lid Inspection | Extra_specs
2OQ850 | 2OQ850/Skip Lot Release_NY_2OQ850 | Skip Lot Release_NY_2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | White (Done) | [5 empty] | Canister & Lid Inspection | Extra_specs
2OQ850 | 2OQ850/Skip Lot Release_NY_2OQ850 | Skip Lot Release_NY_2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | HDPE Resin Identification Marker | HDPE Resin Identification Marker | Meets the Specification (Done) | [5 empty] | Canister & Lid Inspection | Extra_specs
2OQ850 | 2OQ850/Skip Lot Release_NY_2OQ850 | Skip Lot Release_NY_2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Odor | Odor | None detectable (Done) | [5 empty] | Canister & Lid Inspection | Extra_specs
2OQ850 | 2OQ850/Skip Lot Release_NY_2OQ850 | Skip Lot Release_NY_2OQ850 | NY-Analytical | 765 | Canister & Lid Inspection | Weight | Weight | 89.99 |  | g | FDA,1 | 87 | 93 | Measurement | Extra_specs
2OQ850 | 2OQ850/Skip Lot Release_NY_2OQ850 | Skip Lot Release_NY_2OQ850 | NY-Analytical | 766 | Lid Pull Force | Lid Pull-Off Force | Lid Pull-Off Force | 2719.98 |  | g | FDA,0 | 2000 | 3500 | TM 148 | Extra_specs
2OQ850LI | 2OQ850LI/Lid_NY_Skip lot testing | Lid_NY_Skip lot testing | NY-Analytical | 765 | Canister & Lid Inspection | Pop-Up Hole Outer Diameter (per specification) | Pop-Up Hole Outer Diameter | 0.163 |  | in | FDA,3 | 0.16 | 0.17 | TM 143 / Visual | Extra_specs
2OQ850X | 2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | AQL Sample Size | AQL Sample Size |  |  |  | FDA,0 |  |  | TM 141 / Visual | Extra_specs
2OQ850X | 2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | Artwork / Copy | Artwork / Copy | [6 empty] | TM 141 / Visual | Extra_specs
2OQ850X | 2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | Certification Stamp (ECT) | Certification Stamp (ECT) |  |  | Lbs./In. | FDA,0 | 32 |  | TM 141 / Visual | Extra_specs
2OQ850X | 2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | General Inspection | General Inspection | [6 empty] | TM 141 / Visual | Extra_specs
2OQ850X | 2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | Height I.D. | Height |  |  | in | FDA,4 | 4.0625 | 4.3125 | TM 141 / Visual | Extra_specs
2OQ850X | 2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | Length I.D. | Length |  |  | in | FDA,4 | 12.3125 | 12.5625 | TM 141 / Visual | Extra_specs
2OQ850X | 2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | Liner Board Color | Liner Board Color | [6 empty] | TM 141 / Visual | Extra_specs
2OQ850X | 2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | Shipper Style | Shipper Style | [6 empty] | TM 141 / Visual | Extra_specs
2OQ850X | 2OQ850X/SHIP_NY_32K_100_12.4375-5.4375-4.1875 | SHIP_NY_32K_100_12.4375-5.4375-4.1875 | NY-Analytical | 731 | Shipper Inspection | Width I.D. | Width |  |  | in | FDA,4 | 5.3125 | 5.5625 | TM 141 / Visual | Extra_specs
2OX510B | 2OX510B/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Canister & Lid Inspection | Extra_specs
2OX510B | 2OX510B/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 765 | Canister & Lid Inspection | Appearance | Appearance | Meets Specifications (Done) | [5 empty] | Canister & Lid Inspection | Extra_specs
2OX510B | 2OX510B/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 765 | Canister & Lid Inspection | Bottle Internal Diameter at Flange | Bottle Internal Diameter at Flange | 0.4425 |  | in | FDA,3 | 0.442 | 0.452 | Canister & Lid Inspection | Extra_specs
2OX510B | 2OX510B/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 765 | Canister & Lid Inspection | Color | Color | Natural (Done) | [5 empty] | Canister & Lid Inspection | Extra_specs
2OX510B | 2OX510B/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 765 | Canister & Lid Inspection | Overall Length | Overall Length | 2.8805 |  | in | FDA,3 | 2.848 | 2.888 | Canister & Lid Inspection | Extra_specs
2OX510B | 2OX510B/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 765 | Canister & Lid Inspection | Seal Flange Diameter | Seal Flange Diameter | 0.55 |  | in | FDA,3 | 0.544 | 0.554 | Canister & Lid Inspection | Extra_specs
2OX510B | 2OX510B/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight |  |  | g | FDA,2 |  |  | TM 610 | Extra_specs
2OX510B | 2OX510B/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA | [5 empty] | TM 610 | Extra_specs
2OX510B | 2OX510B/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA | [5 empty] | TM 610 | Extra_specs
2OX510B | 2OX510B/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Extra_specs
2OX510B | 2OX510B/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) |  |  | CFU/g | FDA,1 |  | 100 | TM 610 | Extra_specs
2OX510B | 2OX510B/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Extra_specs
2OX510B | 2OX510B/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 1.1 |  | cfu/device | FDA,1 |  |  | TM 610 | Extra_specs
2OX510B | 2OX510B/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0.015027 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Extra_specs
2OX510B | 2OX510B/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Extra_specs
2OX510B | 2OX510B/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Positive Rods (Done) | [5 empty] | WI QAS 076 | Extra_specs
2OX510B | 2OX510B/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Mold (OOS) | [5 empty] | WI QAS 076 | Extra_specs
3BB406 | 3BB406/REDUCED TESTING_NY_3BB406 | REDUCED TESTING_NY_3BB406 | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading | [5 empty] | Color (TM 208) | Extra_specs
3BB406 | 3BB406/REDUCED TESTING_NY_3BB406 | REDUCED TESTING_NY_3BB406 | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | May have residual sulfur odor when wet | [5 empty] | Odor (TM 208) | Extra_specs
3BB406 | 3BB406/REDUCED TESTING_NY_3BB406 | REDUCED TESTING_NY_3BB406 | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Meets specifications (Done) | [5 empty] | See Analyte for Method # | Extra_specs
3BB406 | 3BB406/REDUCED TESTING_NY_3BB406 | REDUCED TESTING_NY_3BB406 | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 200 |  |  | FDA,0 |  |  | See Analyte for Method # | Extra_specs
3BB406 | 3BB406/REDUCED TESTING_NY_3BB406 | REDUCED TESTING_NY_3BB406 | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 52.74 |  | GSM | FDA,1 | 48.9 | 56.3 | See Analyte for Method # | Extra_specs
3BB406 | 3BB406/REDUCED TESTING_NY_3BB406 | REDUCED TESTING_NY_3BB406 | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | See Analyte for Method # | Extra_specs
3BB406 | 3BB406/REDUCED TESTING_NY_3BB406 | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight |  |  | g | FDA,2 |  |  | TM 610 | Extra_specs
3BB406 | 3BB406/REDUCED TESTING_NY_3BB406 | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer | [5 empty] | TM 610 | Extra_specs
3BB406 | 3BB406/REDUCED TESTING_NY_3BB406 | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer | [5 empty] | TM 610 | Extra_specs
3BB406 | 3BB406/REDUCED TESTING_NY_3BB406 | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) |  |  | cfu/device | FDA,1 |  |  | TM 610 | Extra_specs
3BB406 | 3BB406/REDUCED TESTING_NY_3BB406 | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) |  |  | CFU/g | FDA,1 |  | 100 | TM 610 | Extra_specs
3BB406 | 3BB406/REDUCED TESTING_NY_3BB406 | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Extra_specs
3BB406 | 3BB406/REDUCED TESTING_NY_3BB406 | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 1 |  | cfu/device | FDA,1 |  |  | TM 610 | Extra_specs
3BB406 | 3BB406/REDUCED TESTING_NY_3BB406 | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) |  |  | CFU/g | FDA,1 |  | 15 | TM 610 | Extra_specs
3BB406 | 3BB406/REDUCED TESTING_NY_3BB406 | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) |  |  | cfu | FDA,0 |  |  | TM 610 | Extra_specs
3BB406 | 3BB406/REDUCED TESTING_NY_3BB406 | REDUCED TESTING_NY_3BB406 | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Mold (OOS) | [5 empty] | WI QAS 076 | Extra_specs
3OA347 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading | [5 empty] | Color (TM 208) | Extra_specs
3OA347 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | May have residual sulfur odor when wet | [5 empty] | Odor (TM 208) | Extra_specs
3OA347 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Applicator_Inspection | Extra_specs
3OA347 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Smooth free of hard wrinkles, formation uniform matches embossing pattern (Done) | [5 empty] | See Analyte for Method # | Extra_specs
3OA347 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 125 |  |  | FDA,0 |  |  | See Analyte for Method # | Extra_specs
3OA347 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 61.79 |  | GSM | FDA,1 | 60 | 70 | See Analyte for Method # | Extra_specs
3OA347 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (mm) (TM 102) | Caliper | 0.64 |  | mm | FDA,2 | 0.54 | 0.66 | See Analyte for Method # | Extra_specs
3OA347 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Ratio: Wet MD / Wet CD (TM 103) | Ratio: Wet MD / Wet CD | 0 |  |  | FDA,1 |  | 4.5 | See Analyte for Method # | Extra_specs
3OA347 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter | 45 |  | in | FDA,0 | 42 | 48 | See Analyte for Method # | Extra_specs
3OA347 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Roll Width (mm) (Measurement) | Roll Width | 6.062 |  | in | FDA,2 | 6 | 6.125 | See Analyte for Method # | Extra_specs
3OA347 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Wet CD (g/in) (TM 103) | Wet CD | 416 |  | g/in | FDA,0 | 375 | 575 | See Analyte for Method # | Extra_specs
3OA347 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 884 | Applicator_Inspection | Wet MD (g/in) (TM 103) | Wet MD | 692 |  | g/in | FDA,0 | 485 | 915 | See Analyte for Method # | Extra_specs
3OA347 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight |  |  | GSM | FDA,1 | 60 | 70 | TM 101 | Extra_specs
3OA347 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper | 0.63 |  | mm | FDA,2 | 0.54 | 0.66 | TM 102 | Extra_specs
3OA347 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 758 | Tensile Strength | Elongation % Wet CD | Elongation % Wet CD | 25.08 |  | % | FDA,0 | 15 | 65 | TM 103 | Extra_specs
3OA347 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 758 | Tensile Strength | Elongation % Wet MD | Elongation % Wet MD | 17.02 |  | % | FDA,0 | 2 | 50 | TM 103 | Extra_specs
3OA347 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 495 |  | g/in | FDA,0 | 375 | 575 | TM 103 | Extra_specs
3OA347 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 648 |  | g/in | FDA,0 | 485 | 915 | TM 103 | Extra_specs
3OA347 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | TM 208 | Extra_specs
3OA347 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) | [5 empty] | TM 208 | Extra_specs
3OA347 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) | [5 empty] | TM 208 | Extra_specs
3OA347 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 2.7 |  | g | FDA,2 |  |  | TM 610 | Extra_specs
3OA347 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA | [5 empty] | TM 610 | Extra_specs
3OA347 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer | [5 empty] | TM 610 | Extra_specs
3OA347 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Extra_specs
3OA347 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Extra_specs
3OA347 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Extra_specs
3OA347 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 1 |  | cfu/device | FDA,1 |  |  | TM 610 | Extra_specs
3OA347 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0.96774 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Extra_specs
3OA347 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Extra_specs
3OA347 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Positive Rods (Done) | [5 empty] | WI QAS 076 | Extra_specs
3OA347 | 3OA347/Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | Applicator_NY_Nonwoven Rolls Dispersable Unembossed- 6" width - 65 gsm | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Mold (OOS) | [5 empty] | WI QAS 076 | Extra_specs
3OA347 | 3OA347/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading | [5 empty] | Color (TM 208) | Extra_specs
3OA347 | 3OA347/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | May have residual sulfur odor when wet | [5 empty] | Odor (TM 208) | Extra_specs
3OA347 | 3OA347/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Certificate of Analysis | Extra_specs
3OA347 | 3OA347/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Smooth free of hard wrinkles, formation uniform matches embossing pattern (Done) | [5 empty] | See Analyte for Method # | Extra_specs
3OA347 | 3OA347/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 80 |  |  | FDA,0 |  |  | See Analyte for Method # | Extra_specs
3OA347 | 3OA347/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 65.66 |  | GSM | FDA,1 | 60 | 70 | See Analyte for Method # | Extra_specs
3OA347 | 3OA347/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (mm) (TM 102) | Caliper | 0.66 |  | mm | FDA,2 | 0.54 | 0.66 | See Analyte for Method # | Extra_specs
3OA347 | 3OA347/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Ratio: Wet MD / Wet CD (TM 103) | Ratio: Wet MD / Wet CD | 0 |  |  | FDA,1 |  | 4.5 | See Analyte for Method # | Extra_specs
3OA347 | 3OA347/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter | 43 |  | in | FDA,0 | 42 | 48 | See Analyte for Method # | Extra_specs
3OA347 | 3OA347/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Roll Width (mm) (Measurement) | Roll Width | 6 |  | in | FDA,2 | 6 | 6.125 | See Analyte for Method # | Extra_specs
3OA347 | 3OA347/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Wet CD (g/in) (TM 103) | Wet CD | 426 |  | g/in | FDA,0 | 375 | 575 | See Analyte for Method # | Extra_specs
3OA347 | 3OA347/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 884 | Applicator_Inspection | Wet MD (g/in) (TM 103) | Wet MD | 769 |  | g/in | FDA,0 | 485 | 915 | See Analyte for Method # | Extra_specs
3OA347 | 3OA347/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 67.38 |  | GSM | FDA,1 | 60 | 70 | TM 101 | Extra_specs
3OA347 | 3OA347/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_(mm) | Caliper | 0.64 |  | mm | FDA,2 | 0.54 | 0.66 | TM 102 | Extra_specs
3OA347 | 3OA347/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 758 | Tensile Strength | Elongation % Wet CD | Elongation % Wet CD | 43.05 |  | % | FDA,0 | 15 | 65 | TM 103 | Extra_specs
3OA347 | 3OA347/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 758 | Tensile Strength | Elongation % Wet MD | Elongation % Wet MD | 30.3 |  | % | FDA,0 | 2 | 50 | TM 103 | Extra_specs
3OA347 | 3OA347/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 561.7 |  | g/in | FDA,0 | 375 | 575 | TM 103 | Extra_specs
3OA347 | 3OA347/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 694.79 |  | g/in | FDA,0 | 485 | 915 | TM 103 | Extra_specs
3OA347 | 3OA347/Microbiologu - Bioburden_Applicator_NY | Microbiologu - Bioburden_Applicator_NY | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) | [5 empty] | TM 208 | Extra_specs
3OA580 | 3OA580/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 2.05 |  | g | FDA,2 |  |  | TM 610 | Extra_specs
3OA580 | 3OA580/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer | [5 empty] | TM 610 | Extra_specs
3OA580 | 3OA580/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA | [5 empty] | TM 610 | Extra_specs
3OA580 | 3OA580/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Extra_specs
3OA580 | 3OA580/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Extra_specs
3OA580 | 3OA580/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Extra_specs
3OA580 | 3OA580/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Extra_specs
3OA580 | 3OA580/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Extra_specs
3OA580 | 3OA580/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Extra_specs
3OA580 | 3OA580/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Mold (OOS) | [5 empty] | WI QAS 076 | Extra_specs
3OA911 | 3OA911/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Non-Subculturable Growth (Done) | [5 empty] | WI QAS 076 | Extra_specs
3OB108 | 3OB108/REDUCED TESTING_NY_3OB108 | REDUCED TESTING_NY_3OB108 | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading | [5 empty] | Color (TM 208) | Extra_specs
3OB108 | 3OB108/REDUCED TESTING_NY_3OB108 | REDUCED TESTING_NY_3OB108 | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Meets specifications (Done) | [5 empty] | See Analyte for Method # | Extra_specs
3OB108 | 3OB108/REDUCED TESTING_NY_3OB108 | REDUCED TESTING_NY_3OB108 | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 125 |  |  | FDA,0 |  |  | See Analyte for Method # | Extra_specs
3OB108 | 3OB108/REDUCED TESTING_NY_3OB108 | REDUCED TESTING_NY_3OB108 | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 54.9 |  | GSM | FDA,1 | 52.3 | 59.6 | See Analyte for Method # | Extra_specs
3OB108 | 3OB108/REDUCED TESTING_NY_3OB108 | REDUCED TESTING_NY_3OB108 | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | See Analyte for Method # | Extra_specs
3OB108 | 3OB108/REDUCED TESTING_NY_3OB108 | REDUCED TESTING_NY_3OB108 | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | Meets specifications (Done) | [5 empty] | See Analyte for Method # | Extra_specs
3OB108 | 3OB108/REDUCED TESTING_NY_3OB108 | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 1.25 |  | g | FDA,2 |  |  | TM 610 | Extra_specs
3OB108 | 3OB108/REDUCED TESTING_NY_3OB108 | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA | [5 empty] | TM 610 | Extra_specs
3OB108 | 3OB108/REDUCED TESTING_NY_3OB108 | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer | [5 empty] | TM 610 | Extra_specs
3OB108 | 3OB108/REDUCED TESTING_NY_3OB108 | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Extra_specs
3OB108 | 3OB108/REDUCED TESTING_NY_3OB108 | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Extra_specs
3OB108 | 3OB108/REDUCED TESTING_NY_3OB108 | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Extra_specs
3OB108 | 3OB108/REDUCED TESTING_NY_3OB108 | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Extra_specs
3OB108 | 3OB108/REDUCED TESTING_NY_3OB108 | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Extra_specs
3OB108 | 3OB108/REDUCED TESTING_NY_3OB108 | REDUCED TESTING_NY_3OB108 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Extra_specs
3OD438 | 3OD438/REDUCED TESTING_NY_3OD438 | REDUCED TESTING_NY_3OD438 | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | White with no off colors or shading | [5 empty] | Color (TM 208) | Extra_specs
3OD438 | 3OD438/REDUCED TESTING_NY_3OD438 | REDUCED TESTING_NY_3OD438 | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | May have residual sulfur odor when wet | [5 empty] | Odor (TM 208) | Extra_specs
3OD438 | 3OD438/REDUCED TESTING_NY_3OD438 | REDUCED TESTING_NY_3OD438 | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Smooth free of hard wrinkles, formation uniform matches embossing pattern (Done) | [5 empty] | See Analyte for Method # | Extra_specs
3OD438 | 3OD438/REDUCED TESTING_NY_3OD438 | REDUCED TESTING_NY_3OD438 | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 315 |  |  | FDA,0 |  |  | See Analyte for Method # | Extra_specs
3OD438 | 3OD438/REDUCED TESTING_NY_3OD438 | REDUCED TESTING_NY_3OD438 | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 35.09 |  | GSM | FDA,1 | 31.5 | 38.5 | See Analyte for Method # | Extra_specs
3OD438 | 3OD438/REDUCED TESTING_NY_3OD438 | REDUCED TESTING_NY_3OD438 | NY-Analytical | 884 | Applicator_Inspection | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | See Analyte for Method # | Extra_specs
3OD438 | 3OD438/REDUCED TESTING_NY_3OD438 | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 785 | Bioburden_TAMC | Applicator Weight | Applicator Weight | 1.1 |  | g | FDA,2 |  |  | TM 610 | Extra_specs
3OD438 | 3OD438/REDUCED TESTING_NY_3OD438 | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth TSA | [5 empty] | TM 610 | Extra_specs
3OD438 | 3OD438/REDUCED TESTING_NY_3OD438 | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth Phosphate Buffer | [5 empty] | TM 610 | Extra_specs
3OD438 | 3OD438/REDUCED TESTING_NY_3OD438 | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 1 |  | cfu/device | FDA,1 |  |  | TM 610 | Extra_specs
3OD438 | 3OD438/REDUCED TESTING_NY_3OD438 | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0.18182 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Extra_specs
3OD438 | 3OD438/REDUCED TESTING_NY_3OD438 | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Extra_specs
3OD438 | 3OD438/REDUCED TESTING_NY_3OD438 | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Extra_specs
3OD438 | 3OD438/REDUCED TESTING_NY_3OD438 | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Extra_specs
3OD438 | 3OD438/REDUCED TESTING_NY_3OD438 | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Extra_specs
3OD438 | 3OD438/REDUCED TESTING_NY_3OD438 | REDUCED TESTING_NY_3OD438 | NY-Microbiology | 787 | GramStain_Bioburden_TAMC | GramsStain | GramsStain | Gram Positive Rods (Done) | [5 empty] | WI QAS 076 | Extra_specs
3OJ213 | 3OJ213/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width | 12 |  | in | FDA,3 | 11.875 | 12.125 | Applicator Measurements | Extra_specs
3OJ213 | 3OJ213/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight | 40.9 |  | GSM | FDA,0 | 36 | 44 | TM 101 | Extra_specs
3OJ213 | 3OJ213/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 757 | Caliper | Caliper_After_Slitter_CoA_(mm) | Caliper | 0.54 |  | mm | FDA,2 | 0.48 | 0.61 | TM 102 | Extra_specs
3OJ213 | 3OJ213/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 758 | Tensile Strength | Wet_CD_(g/in) | Wet CD | 865 |  | g/in | FDA,0 | 1030 | 1546 | TM 103 | Extra_specs
3OJ213 | 3OJ213/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 758 | Tensile Strength | Wet_MD_(g/in) | Wet MD | 3407 |  | g/in | FDA,0 | 3343 | 4350 | TM 103 | Extra_specs
3OJ213 | 3OJ213/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 748 | Visual Inspection | Appearance | Appearance | Meets specifications (Done) | [5 empty] | TM 208 | Extra_specs
3OJ213 | 3OJ213/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 748 | Visual Inspection | Color | Color | Meets specifications (Done) | [5 empty] | TM 208 | Extra_specs
3OJ213 | 3OJ213/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 748 | Visual Inspection | Odor | Odor | Meets specifications (Done) | [5 empty] | TM 208 | Extra_specs
3OM304 | 3OM304/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Appearance (TM 208) | Appearance | Meets specifications (Done) | [5 empty] | See Analyte for Method # | Extra_specs
3OM304 | 3OM304/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | AQL Sample Size | AQL Sample Size | 315 |  |  | FDA,0 |  |  | See Analyte for Method # | Extra_specs
3OM304 | 3OM304/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Basis Weight (GSM) (TM 101) | Basis Weight | 50.38 |  | GSM | FDA,0 | 45 | 55 | See Analyte for Method # | Extra_specs
3OM304 | 3OM304/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Caliper 1-ply (mm) (TM 102) | Caliper | 0.53 |  | mm | FDA,2 | 0.51 | 0.63 | See Analyte for Method # | Extra_specs
3OM304 | 3OM304/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Color (TM 208) | Color | Meets specifications (Done) | [5 empty] | See Analyte for Method # | Extra_specs
3OM304 | 3OM304/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Dry CD (g/in) (TM103) | Dry CD | 2395.04 |  | g/in | FDA,0 | 1801 | 2821 | See Analyte for Method # | Extra_specs
3OM304 | 3OM304/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Dry MD (g/in) (TM 103) | Dry CD | 8931.7 |  | g/in | FDA,0 | 6675 | 9447 | See Analyte for Method # | Extra_specs
3OM304 | 3OM304/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Odor (TM 208) | Odor | Meets specifications (Done) | [5 empty] | See Analyte for Method # | Extra_specs
3OM304 | 3OM304/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (COA) | Roll Diameter (COA) | 43 |  |  | FDA,0 | 41 | 45 | See Analyte for Method # | Extra_specs
3OM304 | 3OM304/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Roll width (COA) | Roll width (COA) | 9 |  | in | FDA,2 | 8.75 | 9.25 | See Analyte for Method # | Extra_specs
3OP113 | 3OP113/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Caliper after Slitter (mm) (CofA) | Caliper | 0.62 |  | mm | FDA,2 | 0.5 | 0.62 | See Analyte for Method # | Extra_specs
3OP113 | 3OP113/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Roll Diameter (Inches) (Measurement) | Roll Diameter | 40 |  | in | FDA,0 | 40 | 46 | See Analyte for Method # | Extra_specs
3OP113 | 3OP113/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Roll Width (Inches) (Measurement) | Roll Width | 30 |  | in | FDA,3 | 29.875 | 30.125 | See Analyte for Method # | Extra_specs
3OP113 | 3OP113/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Wet CD (g/in) (TM 103) | Wet CD | 1469.66 |  | g/in | FDA,0 | 1142 | 1714 | See Analyte for Method # | Extra_specs
3OP113 | 3OP113/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 884 | Applicator_Inspection | Wet MD (g/in) (TM 103) | Wet MD | 4083.47 |  | g/in | FDA,0 | 3265 | 4897 | See Analyte for Method # | Extra_specs
3OP197A | 3OP197A/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | CofA meets specifications (Done) | [5 empty] | Bioburden_TMYC | Extra_specs
3OP197A | 3OP197A/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Microbiology | 786 | Bioburden_TMYC | Applicator Weight | Applicator Weight | 0.92 |  | g | FDA,2 |  |  | TM 610 | Extra_specs
3OP197A | 3OP197A/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials | Sterility Controls Raw Materials | No Growth Phosphate Buffer | [5 empty] | TM 610 | Extra_specs
3OP197A | 3OP197A/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Microbiology | 997 | Sterility Controls Raw Materials | Sterility Controls Raw Materials SDA | Sterility Controls Raw Materials SDA | No Growth SDA | [5 empty] | TM 610 | Extra_specs
3OP197A | 3OP197A/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Device) | TAMC (Average per Device) | 0 |  | cfu/device | FDA,1 |  |  | TM 610 | Extra_specs
3OP197A | 3OP197A/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Average per Gram Applicator) | TAMC (Average per Gram Applicator) | 0 |  | CFU/g | FDA,1 |  | 100 | TM 610 | Extra_specs
3OP197A | 3OP197A/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Microbiology | 785 | Bioburden_TAMC | TAMC (Individual) | TAMC (Individual) | 0 |  | cfu | FDA,0 |  |  | TM 610 | Extra_specs
3OP197A | 3OP197A/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Device) | TMYC (Average per Device) | 1 |  | cfu/device | FDA,1 |  |  | TM 610 | Extra_specs
3OP197A | 3OP197A/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Average per Gram Applicator) | TMYC (Average per Gram Applicator) | 0.76087 |  | CFU/g | FDA,1 |  | 15 | TM 610 | Extra_specs
3OP197A | 3OP197A/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Microbiology | 786 | Bioburden_TMYC | TMYC (Individual) | TMYC (Individual) | 1 |  | cfu | FDA,0 |  |  | TM 610 | Extra_specs
3OP197A | 3OP197A/Applicator_NY_Skip lot testing | Applicator_NY_Skip lot testing | NY-Microbiology | 788 | GramStain_Bioburden_TMYC | GramStain_Bioburden_TMYC | GramStain Bioburden TMYC | Non-Subculturable Growth (Done) | [5 empty] | WI QAS 076 | Extra_specs
3OP197A-SU | 3OP197A-SU/Label_NY_NP7200_6.25-4.5 | Label_NY_NP7200_6.25-4.5 | NY-Analytical | 771 | Certificate of Conformance | Certificate of Conformance | Certificate of Conformance | CoC meets specifications (Done) | [5 empty] | Label Inspection | Extra_specs
3OP197A-SU | 3OP197A-SU/Label_NY_NP7200_6.25-4.5 | Label_NY_NP7200_6.25-4.5 | NY-Analytical | 770 | Label Inspection | Appearance | Appearance | Material must be clean and free of particulate matter (Done) | [5 empty] | TM 143 / Visual | Extra_specs
3OP197A-SU | 3OP197A-SU/Label_NY_NP7200_6.25-4.5 | Label_NY_NP7200_6.25-4.5 | NY-Analytical | 770 | Label Inspection | Color/Copy | Color/Copy | Copy/color should be in accordance with individual Item # print spec (Done) | [5 empty] | TM 143 / Visual | Extra_specs
3OP197A-SU | 3OP197A-SU/Label_NY_NP7200_6.25-4.5 | Label_NY_NP7200_6.25-4.5 | NY-Analytical | 770 | Label Inspection | Length | Length | 6.25 |  | in | FDA,4 | 6.1875 | 6.3125 | TM 143 / Visual | Extra_specs
3OP197A-SU | 3OP197A-SU/Label_NY_NP7200_6.25-4.5 | Label_NY_NP7200_6.25-4.5 | NY-Analytical | 770 | Label Inspection | Odor | Odor | No solvent odor is to be present (Done) | [5 empty] | TM 143 / Visual | Extra_specs
3OP197A-SU | 3OP197A-SU/Label_NY_NP7200_6.25-4.5 | Label_NY_NP7200_6.25-4.5 | NY-Analytical | 770 | Label Inspection | Print | Print | No smears, streaks, spots, or voids (Done) | [5 empty] | TM 143 / Visual | Extra_specs
3OP197A-SU | 3OP197A-SU/Label_NY_NP7200_6.25-4.5 | Label_NY_NP7200_6.25-4.5 | NY-Analytical | 770 | Label Inspection | Width | Width | 4.5 |  | in | FDA,4 | 4.4375 | 4.5625 | TM 143 / Visual | Extra_specs
3OQ222 | 3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 743 | Certificate of Analysis | Certificate of Analysis | Certificate of Analysis | [6 empty] | Certificate of Analysis | Extra_specs
3OQ222 | 3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 759 | Applicator Measurements | Roll Diameter (Inches) | Roll Diameter |  |  | in | FDA,0 | 42 | 46 | Applicator Measurements | Extra_specs
3OQ222 | 3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 759 | Applicator Measurements | Roll Width (Inches) | Roll Width |  |  | in | FDA,3 | 37.375 | 37.625 | Applicator Measurements | Extra_specs
3OQ222 | 3OQ222/Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | Applicator_NY_ Embossed Spunlaced Nonwoven - 37 1/2" | NY-Analytical | 756 | Basis Weight | Basis_Weight_(GSM) | Basis Weight |  |  | GSM | FDA,1 | 35.5 | 44.5 | TM 101 | Extra_specs